## Sheet1
Unnamed: 0 | WEB_ADDRESS | RECALL_NUMBER | PRODUCT_DESCRIPTION | TRADE_NAME | RECALL_CLASS | CENTER | CENTER_CLASSIFICATION_DT | POSTED_INTERNET_DT | TERMINATION_DT | FIRM_NAME | MANUFACTURER_RECALL_REASON | PAGE_TEXT | ACTION | Recall Type
0 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=212289 | Z-1233-2025 | DxC 500 AU Clinical Chemistry Analyzer, REF: C63519, and C63520 with affected software. | DxC 500 AU Clinical Chemistry Analyzer | 2 | CDRH | 2025-02-25 17:12 | 2025-02-25 20:07 |  | BECKMAN COULTER LABORATORY SYSTEMS (SUZHOU) CO., LTD. | Beckman Coulter has identified an issue which prevents the DxC 500 AU instrument from providing the correct onboard stability (Open Expiration) dates for assays that were loaded prior to a software upgrade (e.g. 1.3 to 1.4, 1.4 to 1.4.1, etc.). As a result, assays that have expired may still be used for sample processing, leading to inaccurate or erroneous patient test results. The issue only happens on assays loaded before instruments doing upgrades and does not impact new installation. It does not impact the assays loaded after upgrade. | Class 2 Device Recall DxC 500 AU Clinical Chemistry Analyzer Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall DxC 500 AU Clinical Chemistry Analyzer FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall DxC 500 AU Clinical Chemistry Analyzer Date Initiated by Firm January 10, 2025 Date Posted February 25, 2025 Recall Status 1 Open 3 , Classified Recall Number Z-1233-2025 Recall Event ID 96193 510(K)Number K220977 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product DxC 500 AU Clinical Chemistry Analyzer, REF: C63519, and C63520 with affected software. Code Information Software Versions: V1.4, V1.4.1, V1.4.2, V1.4.3******REF/UDI-DI/Serial Numbers:***C63519/14987666545058/2024090170;******C63520/14987666545065/2024070116, 2023070009, 2024010038, 2024010035, 2024050072, 2024040056, 2024030049, 2024020046, 2024070111, 2024030051, 2024030052, 2024010033, 2023090015, 2024060103, 2024060104, 2024060101, 2024080145, 2024020045, 2024020042, 2024020043, 2024020047, 2024060096, 2024040059, 2023100019, 2024030053, 2023070010, 2024070129, 2024070130, 2024020041, 2024090183, 2024010036, 2024010037, 2024070132, 2024070131, 2023100020, 2024050064, 2024050070, 2023090016, 2023110022, 2024070110, 2024060075, 2024060106, 2023110023, 2024010024, 2024010025, 2024050066, 2024010029, 2024060107, 2023080012, 2024040060, 2024050074, 2024040062, 2024010031, 2024030055, 2024080144 Recalling Firm/ Manufacturer BECKMAN COULTER LABORATORY SYSTEMS (SUZHOU) CO., LTD. No. 11 Shengpu Street Zhensheng Road Pingjiang Suzhou China For Additional Information Contact Courtney Walton 952-448-4848 Manufacturer Reason for Recall Beckman Coulter has identified an issue which prevents the DxC 500 AU instrument from providing the correct onboard stability (Open Expiration) dates for assays that were loaded prior to a software upgrade (e.g. 1.3 to 1.4, 1.4 to 1.4.1, etc.). As a result, assays that have expired may still be used for sample processing, leading to inaccurate or erroneous patient test results. The issue only happens on assays loaded before instruments doing upgrades and does not impact new installation. It does not impact the assays loaded after upgrade. FDA Determined Cause 2 Software change control Action On 1/10/2025, recall notices were mailed to customers who were asked to do the following:******1) Replace and discard any reagent bottles that were on-board the analyzer at the time of any software upgrade.***2) Review Quality Control data and actions taken, for shifts out of allowed laboratory range which could have indicated potential deterioration in reagent stability.***3) Firm recommends sharing the content of this letter with your laboratory and/or Medical Director to determine if a review of previous patient test results should be conducted.***4) If specific software upgrade date is required, contact firm.***5) Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter.******This issue will not occur with future software updates.******If you have any questions regarding this notice, please contact the firm's Customer Support Center:***Website: http://www.beckmancoulter.com***Hotline: (800) 854-3633 Quantity in Commerce 56 Distribution US: MA, CA, SC, LA, OK, GA, AL, TX, OR, ID, CO, MN, NY, MS, PA, WA, IL, MI, PR.******OUS: Lebanon, Poland, Ghana, Spain, South Africa, Slovakia, Ghana, Croatia, Seychelles, Australia, Republic of Korea Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 1/10/2025, recall notices were mailed to customers who were asked to do the following:******1) Replace and discard any reagent bottles that were on-board the analyzer at the time of any software upgrade.***2) Review Quality Control data and actions taken, for shifts out of allowed laboratory range which could have indicated potential deterioration in reagent stability.***3) Firm recommends sharing the content of this letter with your laboratory and/or Medical Director to determine if a review of previous patient test results should be conducted.***4) If specific software upgrade date is required, contact firm.***5) Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter.******This issue will not occur with future software updates.******If you have any questions regarding this notice, please contact the firm's Customer Support Center:***Website: http://www.beckmancoulter.com***Hotline: (800) 854-3633 | SoftwareChangeControl
1 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211506 | Z-0810-2025 | Xhibit Telemetry Receiver with Software, Model 96280.***Intended to provide monitoring system with patient data. |  | 2 | CDRH | 2025-01-07 13:25 | 2025-01-07 20:07 |  | Spacelabs Healthcare, Inc. | Telemetry receivers with affected software versions may experience unsolicited system shut down and restart (Offline displayed), lasting approximately 1 to 1.5 minutes, which will cause temporary loss of monitoring. If a clinical event occurred during this period, staff may not detect until the system restarts and necessary clinical intervention could be delayed. | Class 2 Device Recall Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Date Initiated by Firm November 25, 2024 Date Posted January 07, 2025 Recall Status 1 Open 3 , Classified Recall Number Z-0810-2025 Recall Event ID 95910 510(K)Number K141156 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product Xhibit Telemetry Receiver with Software, Model 96280.******Intended to provide monitoring system with patient data. Code Information Software Versions: 1.4.0 and 1.4.1. UDI-DI: 10841522107177, 10841522100246 Recalling Firm/ Manufacturer Spacelabs Healthcare, Inc. 35301 Se Center St Snoqualmie WA 98065-9216 For Additional Information Contact Zachary Orlowski 425-396-3300 Manufacturer Reason for Recall Telemetry receivers with affected software versions may experience unsolicited system shut down and restart (Offline displayed), lasting approximately 1 to 1.5 minutes, which will cause temporary loss of monitoring. If a clinical event occurred during this period, staff may not detect until the system restarts and necessary clinical intervention could be delayed. FDA Determined Cause 2 Software change control Action On 11/25/2024, correction notices were emailed to customers who were asked to do the following: ******Firm will contact you at the earliest possible date to schedule a convenient time to update affected devices to software version 1.4.2 for XTR systems which addresses this defect.******You can also contact GTS to have an FSE paged out for discussing upgrade schedule: 1-800-522-7025 and select 2 for Technical Support.******Complete and submit the customer acknowledgement form: https://spacelabshealthcare.com/support/recall-reply-form/recall-reply-form-xtr-15-nov-2024/******For additional information or technical assistance, or to report an adverse event, contact: Technical Support: 1-800-522-7025 and select 2 Quantity in Commerce 434 Distribution US: MN, TX, TN, SC, OK, NM, GA, PA, OH, IA, AL, CO, AR, FL, MO, ME, LA, IL, WA, MA, WI, OR, MS, ID, NY, KS, PR.******OUS: Canada, Taiwan, Czech Republic, France, Greece, Italy, Kuwait, Philippines, Poland. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 11/25/2024, correction notices were emailed to customers who were asked to do the following: ******Firm will contact you at the earliest possible date to schedule a convenient time to update affected devices to software version 1.4.2 for XTR systems which addresses this defect.******You can also contact GTS to have an FSE paged out for discussing upgrade schedule: 1-800-522-7025 and select 2 for Technical Support.******Complete and submit the customer acknowledgement form: https://spacelabshealthcare.com/support/recall-reply-form/recall-reply-form-xtr-15-nov-2024/******For additional information or technical assistance, or to report an adverse event, contact: Technical Support: 1-800-522-7025 and select 2 | SoftwareChangeControl
2 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210913 | Z-0571-2025 | Essenz HLM (Heart-Lung Machine), REF: 49-00-10, with Software | Essenz HLM | 2 | CDRH | 2024-11-27 | 2024-11-27 20:07 |  | LivaNova Deutschland GmbH | Heart-Lung Machine GUI touchscreen may go dark for 10-15 seconds, safety/performance functions remain operational before reset, control maintained through backup control panel, but second reset may occur, may lead to gas blender standby, then gas blender must be reactivated using user interface on blender, sensors deactivated until new case started; GUI reset can lead to Hypoxia and Hypoperfusion. | Class 2 Device Recall Essenz HLM Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Essenz HLM FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Essenz HLM Date Initiated by Firm October 18, 2024 Date Posted November 27, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-0571-2025 Recall Event ID 95665 510(K)Number K232291 Product Classification Console, heart-lung machine, cardiopulmonary bypass - Product Code DTQ Product Essenz HLM (Heart-Lung Machine), REF: 49-00-10, with Software Code Information UDI-DI: 04033817903062. Software Version: 1.5. Serial Numbers: 49PK00436, 49PK00108, 49PK00374, 49PK00479, 49PK00480, 49PK00481, 49PK00093, 49PK00091, 49PK00092, 49PK00163, 49PK00497, 49PK00498, 49PK00219, 49PK00223, 49PK00224, 49PK00225, 49PK00226, 49PK00241, 49PK00255, 49PK00292, 49PK00291, 49PK00293, 49PK00490, 49PK00487, 49PK00488, 49PK00278, 49PK00274, 49PK00277, 49PK00275, 49PK00276, 49PK00273, 49PK00457, 49PK00458, 49PK00459, 49PK00455, 49PK00456, 49PK00290, 49PK00288 Recalling Firm/ Manufacturer LivaNova Deutschland GmbH Lindberghstr. 25 Munich Germany For Additional Information Contact 89-323010 Manufacturer Reason for Recall Heart-Lung Machine GUI touchscreen may go dark for 10-15 seconds, safety/performance functions remain operational before reset, control maintained through backup control panel, but second reset may occur, may lead to gas blender standby, then gas blender must be reactivated using user interface on blender, sensors deactivated until new case started; GUI reset can lead to Hypoxia and Hypoperfusion. FDA Determined Cause 2 Software change control Action On 10/18/24, a correction notice was emailed to customers:******1. Continue using the Heart-Lung Machine with SW HLM.1.5 as intended.***2. In the event of any cockpit reset during a procedure, use the backup control panel***to make any setting adjustments as described in HLM IFU section 3.3.2 "Backup control panel."***A. When prompted, press the "LAST CASE" button to retrieve the previous settings as described in HLM IFU chapter "7.1.3 Returning to the last case." If the gas blender settings were manually changed prior to the reset, adjust the settings as needed.***B. If after selecting the "LAST CASE" button a second reset occurs, the "LAST CASE" button may not be available. Throughout the resetting of the system, the HLM's essential functions, including the pumps, alarms, sensors, and safety features, continue to operate as intended. After the second reset, the Gas Blender may go to Standby. If this occurs, the operator should reactivate the Gas Blender from the user interface on the Gas Blender unit to maintain operation of the Gas Blender for the procedure. After the second reset, a new case should be started from the home screen; while the machine distributes the profile again, the sensors are momentarily deactivated until the user selects the start case button. The user can skip the safety checks since they have already been performed before the reset and enter bypass mode. After starting the case, the user should update the case settings as needed. The start Case function is described in HLM IFU chapter "7.1.1 Starting a case"***3. Complete/return the response form to LivaNova.FSCA@livanova.com***4. Communicate the Correction to all personnel within your organization and any organization where affected devices have been transferred. ******Firm to contact customers to schedule a software update, available no later than 10/31/24.******For questions regarding the correction, contact the firm at: LivaNova.FSCA@livanova.com. Quantity in Commerce 38 Distribution US Nationwide distribution in the states of CA, VA, GA, IN, NJ, FL, CO, MS, NV. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = DTQ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 10/18/24, a correction notice was emailed to customers:******1. Continue using the Heart-Lung Machine with SW HLM.1.5 as intended.***2. In the event of any cockpit reset during a procedure, use the backup control panel***to make any setting adjustments as described in HLM IFU section 3.3.2 "Backup control panel."***A. When prompted, press the "LAST CASE" button to retrieve the previous settings as described in HLM IFU chapter "7.1.3 Returning to the last case." If the gas blender settings were manually changed prior to the reset, adjust the settings as needed.***B. If after selecting the "LAST CASE" button a second reset occurs, the "LAST CASE" button may not be available. Throughout the resetting of the system, the HLM's essential functions, including the pumps, alarms, sensors, and safety features, continue to operate as intended. After the second reset, the Gas Blender may go to Standby. If this occurs, the operator should reactivate the Gas Blender from the user interface on the Gas Blender unit to maintain operation of the Gas Blender for the procedure. After the second reset, a new case should be started from the home screen; while the machine distributes the profile again, the sensors are momentarily deactivated until the user selects the start case button. The user can skip the safety checks since they have already been performed before the reset and enter bypass mode. After starting the case, the user should update the case settings as needed. The start Case function is described in HLM IFU chapter "7.1.1 Starting a case"***3. Complete/return the response form to LivaNova.FSCA@livanova.com***4. Communicate the Correction to all personnel within your organization and any organization where affected devices have been transferred. ******Firm to contact customers to schedule a software update, available no later than 10/31/24.******For questions regarding the correction, contact the firm at: LivaNova.FSCA@livanova.com. | SoftwareChangeControl
3 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210328 | Z-0317-2025 | The FLUOROspeed is intended to be used as a universal diagnostic imaging system for radiographic and fluoroscopic studies. Using a flat detector, it can perform a range of applications including general R/F, angiography and pediatric examinations. FLUOROspeed may be used for emergency treatment on an outpatient basis, as well as for bedside examinations. | FluoroSpeed X1 | 2 | CDRH | 2024-11-07 | 2024-11-07 20:06 |  | SHIMADZU CORPORATION MEDICAL SYSTEMS DIVISION | X-ray R/F system software with certain versions of the control board may occasionally fail to display images when switching from Fluoroscopy mode to Radiography mode, which may lead to the possibility of delay or cancellation of the examination. | Class 2 Device Recall FluoroSpeed X1 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall FluoroSpeed X1 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall FluoroSpeed X1 Date Initiated by Firm September 06, 2024 Date Posted November 07, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-0317-2025 Recall Event ID 95450 510(K)Number K191877 Product Classification System, x-ray, fluoroscopic, image-intensified - Product Code JAA Product The FLUOROspeed is intended to be used as a universal diagnostic imaging system for radiographic and fluoroscopic studies. Using a flat detector, it can perform a range of applications including general R/F, angiography and pediatric examinations. FLUOROspeed may be used for emergency treatment on an outpatient basis, as well as for bedside examinations. Code Information UDI-DI: 04540217065271. Serial Numbers: ***MQ493E3E3001, MQ493E3E3002, MQ493E3D6002, MQ493E3D5001, MQ493E3D8002, MQ493E3E1002, MQ493E3D4002, MQ493E3D6003, MQ493E3DA002, MQ493E3E2001, MQ493E3E2002, MQ493E3E1001, MQ493E3DC001, MQ493E3D9001, MQ493E3D4005, MQ493E3D7001, MQ493E3DB002, MQ493E3D2002, MQ493E3D4001, MQ493E3D7002, MQ493E3D4003, MQ493E3D4004, MQ493E3D2001, MQ493E3D8001, MQ493E3D5002, MQ493E3DB001, MQ493E3D2003, MQ493E3DA001, MQ493E3D6001, MQ493E3E1003 Recalling Firm/ Manufacturer SHIMADZU CORPORATION MEDICAL SYSTEMS DIVISION Nishinokyo Kuwabara -Cho Nakagyo-Ku; 1 Kyoto Japan For Additional Information Contact 8175823-1305 Manufacturer Reason for Recall X-ray R/F system software with certain versions of the control board may occasionally fail to display images when switching from Fluoroscopy mode to Radiography mode, which may lead to the possibility of delay or cancellation of the examination. FDA Determined Cause 2 Software change control Action On 9/13/24, recall notices were emailed to consignees asking them to do the following:******1) If the problem occurs it is possible to solve this problem by turning off the power of the***system and then pushing the reset switch on the Digital radiography unit s control***cabinet.******2) Your firm authorized service representative will be contacting to schedule the implementation of the corrective action but if you are not contacted, contact the firm's National Support. ******3) Complete and return the response form via email to response to drodriguez@shimadzu-usa.com******If you have questions contact the firm's National Technical Support at 844-487-2767, available 8:00AM-5:00PM, Monday - Friday. Quantity in Commerce 30 Distribution US Nationwide distribution in the states of CA, CT, MA, WV, OH, NY, IA, NE, MN, IL, FL, NC, HI, LA, MD, MT, WA, AR, TX, TN, CO. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JAA - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 9/13/24, recall notices were emailed to consignees asking them to do the following:******1) If the problem occurs it is possible to solve this problem by turning off the power of the***system and then pushing the reset switch on the Digital radiography unit s control***cabinet.******2) Your firm authorized service representative will be contacting to schedule the implementation of the corrective action but if you are not contacted, contact the firm's National Support. ******3) Complete and return the response form via email to response to drodriguez@shimadzu-usa.com******If you have questions contact the firm's National Technical Support at 844-487-2767, available 8:00AM-5:00PM, Monday - Friday. | SoftwareChangeControl
4 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210310 | Z-0316-2025 | SONIALVISION G4 is a multi-purpose X-ray R/F system. The equipment is suitable for radiographic and fluoroscopic examinations, including general radiography and pediatric examinations, excluding mammography. | SonialVision G4 | 2 | CDRH | 2024-11-07 | 2024-11-07 20:06 |  | SHIMADZU CORPORATION MEDICAL SYSTEMS DIVISION | X-ray R/F system software with certain versions of the control board may occasionally fail to display images when switching from Fluoroscopy mode to Radiography mode, which may lead to the possibility of delay or cancellation of the examination. | Class 2 Device Recall SonialVision G4 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SonialVision G4 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SonialVision G4 Date Initiated by Firm September 06, 2024 Date Posted November 07, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-0316-2025 Recall Event ID 95450 510(K)Number K190373 Product Classification System, x-ray, fluoroscopic, image-intensified - Product Code JAA Product SONIALVISION G4 is a multi-purpose X-ray R/F system. The equipment is suitable for radiographic and fluoroscopic examinations, including general radiography and pediatric examinations, excluding mammography. Code Information UDI-DI: 04540217064809, 04540217064816. Serial Numbers: ***41C3518D7001, 41C3518E4001, 41C3536D9001 Recalling Firm/ Manufacturer SHIMADZU CORPORATION MEDICAL SYSTEMS DIVISION Nishinokyo Kuwabara -Cho Nakagyo-Ku; 1 Kyoto Japan For Additional Information Contact 8175823-1305 Manufacturer Reason for Recall X-ray R/F system software with certain versions of the control board may occasionally fail to display images when switching from Fluoroscopy mode to Radiography mode, which may lead to the possibility of delay or cancellation of the examination. FDA Determined Cause 2 Software change control Action On 9/13/24, recall notices were emailed to consignees asking them to do the following:******1) If the problem occurs it is possible to solve this problem by turning off the power of the***system and then pushing the reset switch on the Digital radiography unit s control***cabinet.******2) Your firm authorized service representative will be contacting to schedule the implementation of the corrective action but if you are not contacted, contact the firm's National Support. ******3) Complete and return the response form via email to response to drodriguez@shimadzu-usa.com******If you have questions contact the firm's National Technical Support at 844-487-2767, available 8:00AM-5:00PM, Monday - Friday. Quantity in Commerce 3 Distribution US Nationwide distribution in the states of CA, CT, MA, WV, OH, NY, IA, NE, MN, IL, FL, NC, HI, LA, MD, MT, WA, AR, TX, TN, CO. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JAA - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 9/13/24, recall notices were emailed to consignees asking them to do the following:******1) If the problem occurs it is possible to solve this problem by turning off the power of the***system and then pushing the reset switch on the Digital radiography unit s control***cabinet.******2) Your firm authorized service representative will be contacting to schedule the implementation of the corrective action but if you are not contacted, contact the firm's National Support. ******3) Complete and return the response form via email to response to drodriguez@shimadzu-usa.com******If you have questions contact the firm's National Technical Support at 844-487-2767, available 8:00AM-5:00PM, Monday - Friday. | SoftwareChangeControl
5 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209749 | Z-0009-2025 | BD FACSDuet Premium Sample Preparation System, REF: 666340, 666339; and BD FACSDuet Sample Preparation System (Base), REF: 663128, 662588, used with software V 1.4.1.******Intended use: to prepare human specimens to be analyzed.*** |  | 2 | CDRH | 2024-10-03 08:58 | 2024-10-03 21:07 |  | Becton, Dickinson and Company, BD Bio Sciences | Sample preparation system with software may cause: 1) Multidispense feature enabled with 2-tube assay: Specimen volume dispensed into tube1 at least 3% higher than volume in tube2, or 2) Absolute counts calculated using Trucount Tubes: increased volume in tube1 may result in 20% higher absolute counts in tube 1 than tube 2; could result in additional blood collection, delayed results/patient therapy. | Class 2 Device Recall Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Date Initiated by Firm August 12, 2024 Date Posted October 03, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-0009-2025 Recall Event ID 95215 Product Classification Automated pipetting, diluting and specimen processing workstations for flow cytometric analysis - Product Code PER Product BD FACSDuet Premium Sample Preparation System, REF: 666340, 666339; and BD FACSDuet Sample Preparation System (Base), REF: 663128, 662588, used with software V 1.4.1.************Intended use: to prepare human specimens to be analyzed. Code Information REF/UDI-DI/Serial Numbers: ***666340/382906663405/9230010101, 9230010102, 9230010103, 9230010104, 9230010105, 9230010107, 9230010108, 9230010109, 9230010112, 9230010113, 9230010114, 9230010115, 9230010116, 9230010118, 9230010119, 9230010120.******666339/9230000102, 9230000103, 9230000104, 9230000105, 9230000106, 9230000112, 9230000113, 9230000114, 9230000118, 9230000120, 9230000121, 9230000123, 9230000124, 9230000126, 9230000128, 9230000130, 9230000132, 9230000134, 9230000135, 9230000138, 9230000144, 9230000145, 9230000131, 9230000152, 9230000149, 9230000030, 9230000101, 9230000108, 9230000109, 9230000110, 9230000111, 9230000115, 9230000116, 9230000117, 9230000119, 9230000122, 9230000125, 9230000127, 9230000129, 9230000136, 9230000139, 9230000141, 9230000142, 9230000143.******663128/382906631282/9220030004, 9220030005, 9220030008, 9220030016, 9220030017, 9220030027, 9220030029, 9220030031, 9220030043, 9220030053, 9220030057.******662588/9220010119, 9220010179, 9220010182, 9220010232, 9220010236, 9220010250, 9220010265, 9220010275, 9220010293, 9220010311, 9220010313, 9220010316, 9220010324, 9220010327, 9220010335, 9220010338, 9220010341, 9220010344, 9220010348, 9220010349, 9220010292, 9220010112, 9220010120, 9220010125, 9220010127, 9220010128, 9220010130, 9220010133, 9220010138, 9220010154, 9220010161, 9220010183, 9220010188, 9220010197, 9220010215, 9220010225, 9220010231, 9220010238, 9220010240, 9220010242, 9220010244, 9220010256, 9220010257, 9220010260, 9220010263, 9220010268, 9220010269, 9220010279, 9220010280, 9220010289, 9220010290, 9220010294, 9220010296, 9220010329, 9220010332, 9220010345, 9220010354, 9230000031, 99220010292. Recalling Firm/ Manufacturer Becton, Dickinson and Company, BD Bio Sciences 155 N Mccarthy Blvd Milpitas CA 95035-5102 Manufacturer Reason for Recall Sample preparation system with software may cause: 1) Multidispense feature enabled with 2-tube assay: Specimen volume dispensed into tube1 at least 3% higher than volume in tube2, or 2) Absolute counts calculated using Trucount Tubes: increased volume in tube1 may result in 20% higher absolute counts in tube 1 than tube 2; could result in additional blood collection, delayed results/patient therapy. FDA Determined Cause 2 Software change control Action On 8/12/2024, correction notices were sent to customers who were informed of the following; ******1) Cease use of the sample preparation system with software version 1.4.1. until the***Multidispense feature has been disabled. Patient samples can be prepared using single dispense on the sample preparation system or manually.******2) It is recommended that laboratories upgrade affected sample preparation systems to version 1.4.2 to correct the issue as soon as it is available. ******3) Ensure the contents of this notification are read and understood by those within your organization.******4) Complete and return the response form via email to BDRC38@bd.com******If you require further assistance, contact: 1-844-823-5433, Mon-Fri 8:00am and 5:00pm CT or email productcomplaints@bd.com. Recall related questions can be emailed to BDRC38@bd.com Quantity in Commerce 130 Distribution US: NJ, UT, WI, TX, ID, AK, MN, MA, MD, TN, MI, CA, VA******OUS: Ireland, Austria, Italy, Germany, Belgium, Finland, Canada, France, Spain, New Zealand, United Kingdom, Japan, Australia, Norway, Singapore, Sweden, Turkey, Brazil, Colombia, Netherlands, United Arab Emirates, Saudi Arabia, Mexico, Switzerland, Costa Rica, China, Azerbaijan, Algeria, Jordan, India Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 8/12/2024, correction notices were sent to customers who were informed of the following; ******1) Cease use of the sample preparation system with software version 1.4.1. until the***Multidispense feature has been disabled. Patient samples can be prepared using single dispense on the sample preparation system or manually.******2) It is recommended that laboratories upgrade affected sample preparation systems to version 1.4.2 to correct the issue as soon as it is available. ******3) Ensure the contents of this notification are read and understood by those within your organization.******4) Complete and return the response form via email to BDRC38@bd.com******If you require further assistance, contact: 1-844-823-5433, Mon-Fri 8:00am and 5:00pm CT or email productcomplaints@bd.com. Recall related questions can be emailed to BDRC38@bd.com | SoftwareChangeControl
6 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209671 | Z-3262-2024 | PT301US Airvo 3 Respiratory Support Device***with software version 1.2.0 and/or 1.5.1******The Airvo 3 is intended to provide high flow warmed and humidified respiratory gases for administration to spontaneously breathing infant, child, adolescent and adult patients in hospitals and sub-acute facilities. | F&P Airvo 3 | 2 | CDRH | 2024-09-25 | 2024-09-25 21:07 |  | Fisher & Paykel Healthcare, Ltd. | Due to a software issue, affected devices that are set up with High Pressure Oxygen (HPO), if the flow alignment alarm occurs, the device will deliver room air only. If this happens, a patient my experience oxygen desaturation that could lead to hypoxia. | Class 2 Device Recall F&P Airvo 3 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall F&P Airvo 3 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall F&P Airvo 3 Date Initiated by Firm August 07, 2024 Date Posted September 25, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-3262-2024 Recall Event ID 95173 510(K)Number K221338 Product Classification High flow/high velocity humidified oxygen delivery device - Product Code QAV Product PT301US Airvo 3 Respiratory Support Device******with software version 1.2.0 and/or 1.5.1************The Airvo 3 is intended to provide high flow warmed and humidified respiratory gases for administration to spontaneously breathing infant, child, adolescent and adult patients in hospitals and sub-acute facilities. Code Information Model Number: PT301US***UDI-DI code: 09420012466662***Lot/Serial Numbers: Recalling Firm/ Manufacturer Fisher & Paykel Healthcare, Ltd. 15 Maurice Paykel Place East Tamaki Auckland New Zealand For Additional Information Contact 649-574-0100 Manufacturer Reason for Recall Due to a software issue, affected devices that are set up with High Pressure Oxygen (HPO), if the flow alignment alarm occurs, the device will deliver room air only. If this happens, a patient my experience oxygen desaturation that could lead to hypoxia. FDA Determined Cause 2 Software change control Action On 08/13/2024, the firm sent an "Urgent: Medical Device Correction" Letter via email, phone, letter or face to face communication to inform customers that due to a software issue with the Airvo 3 (software version 1.5.1 or earlier) when the flow alignment alarm 3.2.2 occurs, Airvo 3 devices set up with high pressure oxygen the device will deliver room air only. Patients may experience oxygen desaturation that could lead to hypoxia.******For Distributors - Actions required for affected products are:***Step1-Identify affected products using Model Number reference (REF), Serial Number (SN) and Manufacturing Date (YYYY-MM-DD). If the device is still within your inventory, put the product in quarantine to prevent ongoing distribution until Fisher&Paykel (F&P) Healthcare has contacted you with next steps.***Step2-Complete and return the attached response form***Step3-Await instructions from Fisher&Paykel Healthcare******For healthcare facilities:***Until the software is updated, the Airvo 3 device can continue to be used by your customers. When using the device, all instructions, including warnings and cautions in the Airvo 3 User Manual must be followed, particularly those in Sections 1 and 2.***If alarm 3.2.2 occurs, follow the onscreen instructions.******For questions/assistance - contact F&P Field Action Team via email at FieldAction.USA@fphcare.com or directly at +1 (800) 446 3908 ext 5003 or +1 (949) 453 4000 ext 5003. Quantity in Commerce 944 units Distribution U.S.: AZ, CA, CO, CT, DC, FL, GA, IA, IL, KY, MA, MD, MI, MN, MO, MT, ND, NH, NJ, NV, NY, OH, OR, PA, SD, TX, UT, VA and WA.******O.U.S.: Not provided Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = QAV - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 08/13/2024, the firm sent an "Urgent: Medical Device Correction" Letter via email, phone, letter or face to face communication to inform customers that due to a software issue with the Airvo 3 (software version 1.5.1 or earlier) when the flow alignment alarm 3.2.2 occurs, Airvo 3 devices set up with high pressure oxygen the device will deliver room air only. Patients may experience oxygen desaturation that could lead to hypoxia.******For Distributors - Actions required for affected products are:***Step1-Identify affected products using Model Number reference (REF), Serial Number (SN) and Manufacturing Date (YYYY-MM-DD). If the device is still within your inventory, put the product in quarantine to prevent ongoing distribution until Fisher&Paykel (F&P) Healthcare has contacted you with next steps.***Step2-Complete and return the attached response form***Step3-Await instructions from Fisher&Paykel Healthcare******For healthcare facilities:***Until the software is updated, the Airvo 3 device can continue to be used by your customers. When using the device, all instructions, including warnings and cautions in the Airvo 3 User Manual must be followed, particularly those in Sections 1 and 2.***If alarm 3.2.2 occurs, follow the onscreen instructions.******For questions/assistance - contact F&P Field Action Team via email at FieldAction.USA@fphcare.com or directly at +1 (800) 446 3908 ext 5003 or +1 (949) 453 4000 ext 5003. | SoftwareChangeControl
7 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209554 | Z-3073-2024 | Medtronic Handset with Communicator, Model TH90Q01, InterStim Smart Programmer and Communicator Kits, containing the Model HH90 Handset and the TM90 Communicator.*** | Medtronic InterStim | 2 | CDRH | 2024-09-06 | 2024-09-06 21:30 |  | Medtronic Neuromodulation | A subset of the handsets within the kit may not be able to complete the pairing process with the communicator upon initial setup. | Class 2 Device Recall Medtronic InterStim Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Medtronic InterStim FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Medtronic InterStim Date Initiated by Firm June 17, 2024 Date Posted September 06, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-3073-2024 Recall Event ID 95131 PMA Number P970004S340 P080025S235 Product Classification Stimulator, electrical, implantable, for incontinence - Product Code EZW Product Medtronic Handset with Communicator, Model TH90Q01, InterStim Smart Programmer and Communicator Kits, containing the Model HH90 Handset and the TM90 Communicator. Code Information GTIN 0076300092754, serial numbers:******NPV0082351, NPV0085575, NPV0082308, NPV0082309, NPV0082314, NPV0082328, NPV0082329, NPV0082349, NPV0082412, NPV0082483, NPV0082494, NPV0082608, NPV0082631, NPV0082646, NPV0082684, NPV0082689, NPV0082690, NPV0082691, NPV0082756, NPV0082866, NPV0082869, NPV0082899, NPV0082921, NPV0083001, NPV0083071, NPV0083084, NPV0083116, NPV0083125, NPV0083171, NPV0083507,***NPV0083577, NPV0083629, NPV0083632, NPV0083853, NPV0083883, NPV0083917, NPV0083944, NPV0083948, NPV0083954, NPV0083972, NPV0084077, NPV0084084, NPV0084104, NPV0084113, NPV0084119, NPV0084125, NPV0084177, NPV0084211, NPV0084253, NPV0084283, NPV0084293, NPV0084295, NPV0084304, NPV0084305, NPV0084311, NPV0084312, NPV0084313, NPV0084317, NPV0084330, NPV0084340,***NPV0084353, NPV0084355, NPV0084428, NPV0084473, NPV0084523, NPV0084610, NPV0084634, NPV0084679, NPV0084681, NPV0084695, NPV0084697, NPV0084751, NPV0084753, NPV0084754, NPV0084767, NPV0085004, NPV0085034, NPV0085036, NPV0085057, NPV0085059, NPV0085077, NPV0085081, NPV0085473, NPV0085510, NPV0085529, NPV0085580, NPV0085601, NPV0085616, NPV0085626, NPV0085631,***NPV0085647, NPV0085653, NPV0082276, NPV0082278, NPV0082281, NPV0082304, NPV0082313, NPV0082316, NPV0082320, NPV0082340, NPV0082352, NPV0082358, NPV0082376, NPV0082377, NPV0082418, NPV0082421, NPV0082431, NPV0082450, NPV0082451, NPV0082476, NPV0082480, NPV0082508, NPV0082513, NPV0082514, NPV0082520, NPV0082524, NPV0082537, NPV0082539, NPV0082543, NPV0082559,***NPV0082562, NPV0082569, NPV0082571, NPV0082572, NPV0082591, NPV0082635, NPV0082666, NPV0082674, NPV0082695, NPV0082701, NPV0082722, NPV0082723, NPV0082740, NPV0082769, NPV0082779, NPV0082782, NPV0082789, NPV0082806, NPV0082859, NPV0082871, NPV0082879, NPV0082882, NPV0082897, NPV0082913, NPV0082915, NPV0082916, NPV0082923, NPV0082928, NPV0082931, NPV0082944, ***NPV0082945, NPV0082949, NPV0082952, NPV0082960, NPV0082963, NPV0083005, NPV0083027, NPV0083036, NPV0083047, NPV0083052, NPV0083053, NPV0083104, NPV0083129, NPV0083134, NPV0083175, NPV0083178, NPV0083189, NPV0083192, NPV0083200, NPV0083225, NPV0083228, NPV0083229, NPV0083230, NPV0083237, NPV0083277, NPV0083289, NPV0083302, NPV0083312, NPV0083329, NPV0083332,***NPV0083344, NPV0083345, NPV0083349, NPV0083371, NPV0083396, NPV0083484, NPV0083506, NPV0083522, NPV0083526, NPV0083555, NPV0083575, NPV0083596, NPV0083604, NPV0083613, NPV0083618, NPV0083646, NPV0083649, NPV0083688, NPV0083708, NPV0083712, NPV0083716, NPV0083729, NPV0083738, NPV0083827, NPV0083842, NPV0083845, NPV0083846, NPV0083847, NPV0083859, NPV0083937,***NPV0084126. Recalling Firm/ Manufacturer Medtronic Neuromodulation 7000 Central Ave Ne Minneapolis MN 55432-3568 Manufacturer Reason for Recall A subset of the handsets within the kit may not be able to complete the pairing process with the communicator upon initial setup. FDA Determined Cause 2 Software change control Action The recalling firm issued letters dated June 2024 beginning 6/17/2024 to the Medtronic Field Service Representatives to notify consignees and retrieve a subset of the InterStim Smart Programmer and Communicator Kits (TH90Q01). The notification informed the Field Service Representatives of the Field Corrective Action and provided instruction on notifying the consignees and retrieving the device(s). ******Beginning on June 17, 2024, US initiated communications and notifications dated June 2024 were physically hand-delivered to each consignee with the impacted device(s). The letter described the issue and provided Required Customer Actions, which included completing and signing the Customer Confirmation form acknowledging receipt of the notification and sharing the notice with all that needed to be aware. An appendix was enclosed which listed the serial numbers of the affected products. Quantity in Commerce 211 devices Distribution US Nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = EZW - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The recalling firm issued letters dated June 2024 beginning 6/17/2024 to the Medtronic Field Service Representatives to notify consignees and retrieve a subset of the InterStim Smart Programmer and Communicator Kits (TH90Q01). The notification informed the Field Service Representatives of the Field Corrective Action and provided instruction on notifying the consignees and retrieving the device(s). ******Beginning on June 17, 2024, US initiated communications and notifications dated June 2024 were physically hand-delivered to each consignee with the impacted device(s). The letter described the issue and provided Required Customer Actions, which included completing and signing the Customer Confirmation form acknowledging receipt of the notification and sharing the notice with all that needed to be aware. An appendix was enclosed which listed the serial numbers of the affected products. | SoftwareChangeControl
8 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=208032 | Z-2374-2024 | BD Pyxis: MedStation ES (Main), REF: 323;***MedStation ES Tower, REF: 352;***Anesthesia Station ES, REF: 327;***CII Safe ES, REF: 1116-00;***CII Safe ES, Desktop PC, Wired Scanner, REF: 107-254-01;***CII Safe ES, Desktop PC, Wireless Scanner, REF: 107-255-01;***CII Safe ES, Desktop PC, Wired Scanner 1.7.3, REF: 107-256-01;***CII Safe ES, Desktop PC, Wireless Scanner 1.7.3, REF: 107-257-01*************** | BD Pyxis" | 2 | CDRH | 2024-07-17 | 2024-07-17 21:07 |  | CareFusion 303, Inc. | If automated dispensing cabinets have specific software versions, and Component Manager was configured to "installed mode", and software patch KB 5033688 was installed in the production environment, then this will will lead to an error that prevents the user from accessing the dispensing software application, which may delay accessing medication. | Class 2 Device Recall BD Pyxis" Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall BD Pyxis" FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall BD Pyxis" Date Initiated by Firm May 02, 2024 Date Posted July 17, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-2374-2024 Recall Event ID 94699 Product Classification Cabinet, table and tray, anesthesia - Product Code BRY Product BD Pyxis: MedStation ES (Main), REF: 323;******MedStation ES Tower, REF: 352;******Anesthesia Station ES, REF: 327;******CII Safe ES, REF: 1116-00;******CII Safe ES, Desktop PC, Wired Scanner, REF: 107-254-01;******CII Safe ES, Desktop PC, Wireless Scanner, REF: 107-255-01;******CII Safe ES, Desktop PC, Wired Scanner 1.7.3, REF: 107-256-01;******CII Safe ES, Desktop PC, Wireless Scanner 1.7.3, REF: 107-257-01 Code Information REF/UDI-DI: 323/10885403512667,***352/10885403512674,***327/10885403477836,***1116-00/10885403512605******Software Version: ES 1.7.x******Serial Numbers: 13534745, 16220738, 16446299, 16446303, 16446302, 16247443, 16247440, 15962990, 16417275, 16446297, 16446296, 16242783, 16234266, 15425093, 16446295, 16446291, 16242784, 16417277, 15644176, 14019332, 16212329, 16221517, 16395857, 16059358, 16254191, 15319466, 15154945, 15153602, 16215325, 15154942, 14959051, 16395855, 16395853, 16215313, 16361893, 14826969, 14826972, 14826971, 14826973, 14232856, 16241342, 13877154, 15148881, 15148860, 15148880, 15148882, 15854601, 15148883, 15148865, 15148885, 15160417, 16212961, 15854595, 15148859, 15148862, 15148861, 15148863, 15148866, 15148868, 15854596, 15854600, 15148870, 15854597, 15854598, 15854599, 15148873, 15148864, 15148867, 15148884,15148878, 15148879, 16324054, 16223557, 15136666, 13841400, 14138516, 15519749, 16446315, 16141586, 16416617, 16416618, 16416622, 16416619, 16416621, 16416623, 16416616, 16416615, 13781497, 14557551, 16212312, 15500549, 14700155, 16212346, 16214431, 16214432, 16214430, 16214429, 16212344, 16247441, 16247444, 16320157, 16335215, 16204472, 15154938, 14813899, 14697245, 15642731, 16209866, 15135154, 16214435, 16446294, 16227849, 16446311, 16010961, 16375509, 15303370, 16415510, 14668921, 16239128, 16346806, 16242800, 16204479, 16352591, 16324052, 16222213, 16212322, 16271477, 15015227, 16214427, 16373774, 15978927, 16358533, 16215391, 16446313, 16446308, 16247439, 16247445, 16247442, 16347745, 16347746, 14157588, 16271458, 16346807, 15196235, 16322403, 16352593, 14728645, 16212338, 15519703, 16217225, 14806556, 16352598, 16352590, 16352595, 16352596, 15263493, 16216425, 16416639, 16416102, 16416104, 16416632, 16416108, 16416637, 16416110, 16416111, 16416106, 16416638, 14987500, 16234192, 16405456, 16405459, 15136661, 16242578, 16271457, 15911127, 15954728, 16271463, 16271465, 14966741, 16214437, 16324055, 16225327, 16407028, 14613489, 15997584, 14442210, 16357615, 16357616, 16307328, 13839287, 13839299, 15136665, 15045814, 16054085, 15642724, 16244068, 16346335, 16283945, 16231509, 16277495, 16240201, 16211337, 16291566, 16212335, 15933030, 16242781, 16346808, 15135153, 16204480, 16271460, 16212330, 16324056, 16215395, 16214434, 15454500, 16173369, 16423187, 15895362, 16242797, 16204482, 16274624, 16223058, 16223059, 16274625, 15540317, 16223060, 16274626, 16257759, 16352588, 16446301, 12345678, 16271461, 16352594, 16212342, 16120462, 16235542, 16242799, 16212336, 16468210, 16195795, 16211170, 15570657, 16223556, 16412617, 16412618, 16412615, 16412616, 16321316, 16022041, 16242801, 16267770, 15136667, 15136668, 16174009, 14956271, 15236068, 15236067, 16218170, 16061265, 14572826, 16128483, 16128482, 16143915, 15398129, 15449307, 16174023, 16174021, 16143913, 16174036, 16174008, 16174019, 16174020, 15449302, 15691588, 15449303, 16227720, 16446290, 16242780, 16242802, 16405462, 16116814, 16212339, 16008937, 16215393, 16215389, 16215398, 16223558, 15147150, 16342242, 16253003, 16253008, 15495195, 16149377, 16094701, 15560137, 16296089, 16296090, 16296091, 16296092, 13347728, 14826153, 16417216, 16417243, 16417242, 16324044, 16415642, 16419185, 16415641, 16415639, 16415640, 16242803, 16242804, 16212319, 15136664, 16262416, 15136663, 16317160, 16212343, 15136670 Recalling Firm/ Manufacturer CareFusion 303, Inc. 10020 Pacific Mesa Blvd San Diego CA 92121-4386 Manufacturer Reason for Recall If automated dispensing cabinets have specific software versions, and Component Manager was configured to "installed mode", and software patch KB 5033688 was installed in the production environment, then this will will lead to an error that prevents the user from accessing the dispensing software application, which may delay accessing medication. FDA Determined Cause 2 Software change control Action On 6/6/24, correction notices were distributed to customers who were informed that their devices had been remediated. That they had been contacted by phone in May 2024 to schedule device remediation. The firm re-configured the component manager from "installed mode" to "disabled mode" and checked if the Microsoft patch was present on each device. If the patch was identified on a device, it was removed. All devices were then configured to "managed mode" and rebooted to enable the system appropriately.******Customers are asked to complete and return the response form to Email BDRC15@bd.com******For safety information about the system not covered in this document, call the firm's Technical Support Center at 1-800-727-6102, Phone hours: 24 hours, 7 days a week. Additional contacts include: 1) Regional Complaint Center, 1-844-823-5433***Say "Product Complaints" when prompted Mon-Fri 8:00am and 5:00pm CT, Email: productcomplaints@bd.com 2) Post Market Quality Field Action Operations, BDRC15@bd.com Quantity in Commerce 326 Distribution Worldwide - US Nationwide distribution including in the states of OR, KS, NE, LA, ND, OK, AL, AK, WI, WV, WA, NV, PR, TX, CO, OH, CA, MO, IA, AZ, IN, AR, IL, MN, HI, NM, FL, SD, UT, NY, VA, GA, MA, NC and the countries of TW, CA. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 6/6/24, correction notices were distributed to customers who were informed that their devices had been remediated. That they had been contacted by phone in May 2024 to schedule device remediation. The firm re-configured the component manager from "installed mode" to "disabled mode" and checked if the Microsoft patch was present on each device. If the patch was identified on a device, it was removed. All devices were then configured to "managed mode" and rebooted to enable the system appropriately.******Customers are asked to complete and return the response form to Email BDRC15@bd.com******For safety information about the system not covered in this document, call the firm's Technical Support Center at 1-800-727-6102, Phone hours: 24 hours, 7 days a week. Additional contacts include: 1) Regional Complaint Center, 1-844-823-5433***Say "Product Complaints" when prompted Mon-Fri 8:00am and 5:00pm CT, Email: productcomplaints@bd.com 2) Post Market Quality Field Action Operations, BDRC15@bd.com | SoftwareChangeControl
9 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=207225 | Z-1883-2024 | Stryker Blueprint Software, Catalog #BPUE001. | Stryker Blueprint | 2 | CDRH | 2024-05-22 | 2024-05-22 21:06 |  | Tornier S.A.S. | The software bug allows for case planning with anatomic glenoid Perform / Perform Augmented implants and anatomic humeral Tornier Flex implant configurations that are incompatible and do not have regulatory approval. | Class 2 Device Recall Stryker Blueprint Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Stryker Blueprint FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Stryker Blueprint Date Initiated by Firm April 16, 2024 Date Posted May 22, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1883-2024 Recall Event ID 94444 510(K)Number K211359 Product Classification shoulder prosthesis, reverse configuration - Product Code PHX Product Stryker Blueprint Software, Catalog #BPUE001. Code Information Versions 2.1.4 to 4.0.2, UDI-DI numbers 03700434023114, 03700434023107, 03700434015065, and 03700434015058. Recalling Firm/ Manufacturer Tornier S.A.S. 161 Rue Lavoisier Montbonnot St Martin France For Additional Information Contact Meghan Wells 901-201-9298 Manufacturer Reason for Recall The software bug allows for case planning with anatomic glenoid Perform / Perform Augmented implants and anatomic humeral Tornier Flex implant configurations that are incompatible and do not have regulatory approval. FDA Determined Cause 2 Software change control Action The firm issued letters via email on 4/16/2024. The email provided the issue and risks and said that for additional information regarding risks, exact incompatible configuration details, and steps for you to take, please refer to the attached document. The email listed a requested action, which was to review the case information provided on the business acknowledgment form, confirm the final implants used in this surgery, and complete and sign the response form on Page 4 at their earliest convenience.******The attached Urgent Medical Device Recall letter was dated 16-April-2024 and provided the product affected, product description, product issue, potential risks, and actions needed by customers and distributors. Those actions included:******1. Review the case information provided on the business acknowledgment form and confirm the final implants used in this surgery. Return the enclosed business reply form by email to confirm receipt and understanding of this notice.***2. A software patch, Version 4.1.1, is currently in development and is scheduled for release by April 16, 2024. This software patch will resolve this issue and prevent the selection of these incompatible implants.***3. Please ensure that all potential users of these products within your facility are made aware of this notification and instruct all potential users to reference this notice prior to use.***4. Maintain awareness of this communication internally until all required actions have been completed within your facility.***5. Please inform us of any adverse event and/or report them to the Health/Competent Authorities in accordance with current regulations. For questions or concerns, please contact memphis.fieldaction@stryker.com. Distribution US Nationwide distribution in the states of FL, KY, NJ, and WI. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = PHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: |  | SoftwareChangeControl
10 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203285 | Z-0093-2024 | BD Pyxis CII Safe ES, REF: 1116-00; and BD Pyxis : CII Safe ES Desktop PC, Wired SCNR, REF: 107-254-01; CII Safe ES Desktop PC, W-LESS SCNR, BIO, REF: 107-255-01; Enterprise Server, REF: 1115-00, 134157-01, 137670-02, 134148-01, 135292-01, 136449-01, 133760-01 | BD Pyxis CII Safe ES | 2 | CDRH | 2023-10-10 | 2023-10-10 21:06 |  | CareFusion 303, Inc. | When global edit is used to update multiple formulary properties simultaneously, the following properties: 1) Require Lot Number on Recall, 2) Access Destruction Bin (Witness), 3) Add to Destruction Bin (Count/Empty), 4) CII Safe Stock Out Notice are, without user warning, overwritten to the value NO , which could result in diversion of controlled substances from the automated dispensing cabinet. | Class 2 Device Recall BD Pyxis CII Safe ES Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall BD Pyxis CII Safe ES FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall BD Pyxis CII Safe ES Date Initiated by Firm September 06, 2023 Date Posted October 10, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-0093-2024 Recall Event ID 93016 Product Classification Cabinet, table and tray, anesthesia - Product Code BRY Product BD Pyxis CII Safe ES, REF: 1116-00; and BD Pyxis : CII Safe ES Desktop PC, Wired SCNR, REF: 107-254-01; CII Safe ES Desktop PC, W-LESS SCNR, BIO, REF: 107-255-01; Enterprise Server, REF: 1115-00, 134157-01, 137670-02, 134148-01, 135292-01, 136449-01, 133760-01 Code Information REF: 1116-00, UDI-DI: 10885403512605.******Software Versions: 1.7.3 and 1.7.4 (Limited commercial release) Recalling Firm/ Manufacturer CareFusion 303, Inc. 10020 Pacific Mesa Blvd San Diego CA 92121-4386 For Additional Information Contact 888-876-7128 Manufacturer Reason for Recall When global edit is used to update multiple formulary properties simultaneously, the following properties: 1) Require Lot Number on Recall, 2) Access Destruction Bin (Witness), 3) Add to Destruction Bin (Count/Empty), 4) CII Safe Stock Out Notice are, without user warning, overwritten to the value NO , which could result in diversion of controlled substances from the automated dispensing cabinet. FDA Determined Cause 2 Software change control Action On 9/6/23, correction notices were mailed to customers informing them to do the following:******1) Do not use the global edit feature until your software is updated.***2) If a global edit has already been performed, it is recommended to verify that the formulary device settings are accurate for the 4 configuration settings. Controlled substances that have been edited via the global edit feature should be reviewed retrospectively for discrepancies in the destruction bin and a manual witness record may be needed until the settings are adjusted for both "Access Destruction Bin" and "Add to Destruction Bin (count/empty)."***3) Disseminate this notification within your facility network to ensure they are aware of this issue.***4) Complete the attached Customer Response Form and return via email to bdrc44@bd.com******Software patch for v.1.7.3 was released on July 20, 2023, and commercial release software v.1.7.4 was released for on July 31, 2023. Firm will contact customers after the receipt of completed response forms to schedule times to install software update.******The following contacts were provided: Technical Support, 1-800-727-6102, 24 hours, 7 days a week; Customer Self Service Portal, https://service.carefusion.com; North American Regional Complaint Center, 1-844-823-5433, Say "Recall" when prompted, Mon-Fri 8:00am and 5:00pm CT, email: productcomplaints@bd.com Quantity in Commerce 61 Distribution US Nationwide distribution in the states of LA, NC, NY, MI, MA, NV, KS, IA, CA. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 9/6/23, correction notices were mailed to customers informing them to do the following:******1) Do not use the global edit feature until your software is updated.***2) If a global edit has already been performed, it is recommended to verify that the formulary device settings are accurate for the 4 configuration settings. Controlled substances that have been edited via the global edit feature should be reviewed retrospectively for discrepancies in the destruction bin and a manual witness record may be needed until the settings are adjusted for both "Access Destruction Bin" and "Add to Destruction Bin (count/empty)."***3) Disseminate this notification within your facility network to ensure they are aware of this issue.***4) Complete the attached Customer Response Form and return via email to bdrc44@bd.com******Software patch for v.1.7.3 was released on July 20, 2023, and commercial release software v.1.7.4 was released for on July 31, 2023. Firm will contact customers after the receipt of completed response forms to schedule times to install software update.******The following contacts were provided: Technical Support, 1-800-727-6102, 24 hours, 7 days a week; Customer Self Service Portal, https://service.carefusion.com; North American Regional Complaint Center, 1-844-823-5433, Say "Recall" when prompted, Mon-Fri 8:00am and 5:00pm CT, email: productcomplaints@bd.com | SoftwareChangeControl
11 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=200975 | Z-2045-2023 | wDR 2.2 Mobile Digital Diagnostic X-Ray System | Mobile Digital Diagnostic X-Ray System | 2 | CDRH | 2023-06-30 | 2023-06-30 21:07 |  | SEDECAL SA | There is a software login in issue that may prevent the user from logging in. | Class 2 Device Recall Mobile Digital Diagnostic XRay System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Mobile Digital Diagnostic XRay System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Mobile Digital Diagnostic XRay System Date Initiated by Firm May 23, 2023 Date Posted June 30, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-2045-2023 Recall Event ID 92464 510(K)Number K191813 Product Classification System, x-ray, mobile - Product Code IZL Product wDR 2.2 Mobile Digital Diagnostic X-Ray System Code Information Model MobileDiagnost wDR:***1) Model Number 712007, Software versions 2.2.0, 2.2.1, 2.2.2, and 2.2.3***2) Model Number 712006, Software versions 2.2.0, 2.2.1, 2.2.2, and 2.2.3***3) Model Number 712004, Software version 2.1 with Windows 10 Upgrade***4) Model Number 712002, Software version 2.1 with Windows 10 Upgrade Recalling Firm/ Manufacturer SEDECAL SA PELAYA, 9 POL. POL. IND. RIO DE JANEIRO Algete Spain For Additional Information Contact 847394-6960 Manufacturer Reason for Recall There is a software login in issue that may prevent the user from logging in. FDA Determined Cause 2 Software change control Action A recall notification letter was mailed to customers in the United States by the US distributor, Philips, beginning May 23, 2023. Customers are advised to take the following actions:***1) To prevent the system from becoming inoperable, do not reboot or log out of your system until Philips has implemented the software correction.***2) Circulate this notice to all users of the device so that they are aware of the issue.***3) Complete and return the response form to Philips promptly and no later than 30 days from receipt. ***A Philips representative will contact consignees to schedule a time for a Field Service Engineer to visit the site and implement a software solution to resolve the issue. Quantity in Commerce 1061 units Distribution Distribution throughout United States including Puerto Rico and US Virgin Islands******OUS distribution to Albania, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Botswana, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Curacao, Czech Republic, Dominican Republic, Egypt, Finland, France, Germany, Hungary, Indonesia, Iran, Iraq, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kosovo, Kuwait, Latvia,Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Malta, Martinique, Mexico, Myanmar, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, ***Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Suriname, Sweden, Switzerland, Tanzania, Thailand, Trinidad and Tobago, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Vietnam, and Zambia Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = IZL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action A recall notification letter was mailed to customers in the United States by the US distributor, Philips, beginning May 23, 2023. Customers are advised to take the following actions:***1) To prevent the system from becoming inoperable, do not reboot or log out of your system until Philips has implemented the software correction.***2) Circulate this notice to all users of the device so that they are aware of the issue.***3) Complete and return the response form to Philips promptly and no later than 30 days from receipt. ***A Philips representative will contact consignees to schedule a time for a Field Service Engineer to visit the site and implement a software solution to resolve the issue. | SoftwareChangeControl
12 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=200971 | Z-2042-2023 | DigitalDiagnost C90, Model No. 712034 and 712035****** | DigitalDiagnost C90 | 2 | CDRH | 2023-06-29 | 2023-06-30 21:07 |  | Philips North America Llc | There is a software login in issue that may prevent the user from logging in. | Class 2 Device Recall DigitalDiagnost C90 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall DigitalDiagnost C90 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall DigitalDiagnost C90 Date Initiated by Firm May 19, 2023 Date Posted June 30, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-2042-2023 Recall Event ID 92462 510(K)Number K202564 Product Classification Solid state x-ray imager (flat panel/digital imager) - Product Code MQB Product DigitalDiagnost C90, Model No. 712034 and 712035 Code Information 1) UDI-DI (01)00884838090699, Model 712034, serial numbers:***10000074, 10000083, 10000163, 10000188, 10001057, 10001071, 10001072, 10001079, 10001089, 10001106, 10001133, 10001136, 10001137, 10001142, 10001146, 10001169, 10001180, 10001181, 10001186, 10001203, 10001211, 10001229, 10001234, 10001276, 19861179, 19861365, 21860026, 21860084, 21861001, 21861233, 21861386, 21861415, 21861416, 21861493, 21861528, 22860017, 22860114, 22860115, 22860448, 22860454, SN19000006, SN20861092, SN20861107, SN20861108, SN20861115, SN20861125, SN20861126, and SN20861128******2) UDI-DI (01)00884838090705, Model 712035, serial numbers:***10000034, 10000098, 10001051, 10001113, 10001130, 21860040, 22860052, SN20862021, SN20862022, and SN20862024 Recalling Firm/ Manufacturer Philips North America Llc 222 Jacobs St Cambridge MA 02141-2289 For Additional Information Contact Customer Care Solutions Center 800-722-9377 Manufacturer Reason for Recall There is a software login in issue that may prevent the user from logging in. FDA Determined Cause 2 Software change control Action Consignees in the United States were notified via Urgent Medical Device Correction Letter that was distributed on May 26, 2023. US consignees were initially notified via email. Follow-up notifications via FedEx will follow. Philips Market Organizations are responsible for distributing letters outside of the US in accordance with local requirements.******A Philips Field Service Engineer will visit impacted customer sites to perform a software upgrade to prevent the issue from recurring. Philips has requested customers perform the following: a) To prevent the system from becoming inoperable, do not reboot or log out of your system until Philips has implemented the software correction; b) Circulate this notice to all users of the device so that they are aware of the issue; and c) Complete and return the response form to Philips promptly and no later than 30 days from receipt. Quantity in Commerce 58 Distribution Distribution in US: AZ, GA, IN, IA, MN, NE, NH, OH, TX******OUS distribution to Australia, Chile, China, Malta, Namibia, New Zealand, Oman, Panama, Qatar, Suriname, United Kingdom Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = MQB - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Consignees in the United States were notified via Urgent Medical Device Correction Letter that was distributed on May 26, 2023. US consignees were initially notified via email. Follow-up notifications via FedEx will follow. Philips Market Organizations are responsible for distributing letters outside of the US in accordance with local requirements.******A Philips Field Service Engineer will visit impacted customer sites to perform a software upgrade to prevent the issue from recurring. Philips has requested customers perform the following: a) To prevent the system from becoming inoperable, do not reboot or log out of your system until Philips has implemented the software correction; b) Circulate this notice to all users of the device so that they are aware of the issue; and c) Complete and return the response form to Philips promptly and no later than 30 days from receipt. | SoftwareChangeControl
13 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=200448 | Z-1993-2023 | HeartWare Ventricular Assist Device System. ******a) Monitor 1521US Explore Tech, ***b) Monitor 1521GB Explore Tech, ***c) Monitor 1521IL Explore Tech,***d) Monitor 1521DE Explore Tech*** | HeartWare" Ventricular Assist Device (HVAD") System | 2 | CDRH | 2023-06-22 | 2023-06-22 21:07 |  | Heartware, Inc. | Medtronic has identified that the logfiles downloaded from the recently updated Model 1521 monitors (Serial Numbers: MON4xxxxxx for OUS and MON5xxxxxx for US), are unable to be processed by the Autologs web portal. | Class 2 Device Recall HeartWare" Ventricular Assist Device (HVAD") System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall HeartWare" Ventricular Assist Device (HVAD") System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall HeartWare" Ventricular Assist Device (HVAD") System Date Initiated by Firm May 16, 2023 Date Posted June 22, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-1993-2023 Recall Event ID 92350 PMA Number P100047 Product Classification Ventricular (assist) bypass - Product Code DSQ Product HeartWare Ventricular Assist Device System. ******a) Monitor 1521US Explore Tech, ***b) Monitor 1521GB Explore Tech, ***c) Monitor 1521IL Explore Tech,***d) Monitor 1521DE Explore Tech Code Information a) Model Number 1521US, GTIN 00888707009256, Serial/Lot Numbers: MON500010, MON500012, MON500016, MON500028, MON500026, MON500005, MON500006, MON500007.***b) Model Number 1521GB, GTIN 00888707006521, Serial/Lot Numbers: MON400039, MON400050, MON400051, MON400052, MON400053, MON400054, MON400061. ***c) Model Number 1521IL, GTIN 00888707010108, Serial/Lot Number: MON400017. ***d) Model Number 1521DE, GTIN 00888707010085, Serial/Lot Numbers: MON400045, MON400021, MON400013. Recalling Firm/ Manufacturer Heartware, Inc. 14400 Nw 60th Ave Miami Lakes FL 33014-2807 For Additional Information Contact MCS Customer Service 877-367-4823 Manufacturer Reason for Recall Medtronic has identified that the logfiles downloaded from the recently updated Model 1521 monitors (Serial Numbers: MON4xxxxxx for OUS and MON5xxxxxx for US), are unable to be processed by the Autologs web portal. FDA Determined Cause 2 Software change control Action Medtronic Mechanical Circulatory Support issued an URGENT MEDICAL DEVICE COMMUNICATION notice to it consignees on 05/16/2023 by email. The notice explained that logfiles downloaded from the recently updated Model 1521 monitors (Serial Numbers: MON4xxxxxx for OUS and MON5xxxxxx for US), are unable to be processed by the Autologs web portal. ******Medtronic is working on an update to the Autologs web portal to resolve the issue. Until the update is complete, logfiles can be sent to dl.mcsclinicalengineering@medtronic.com for report generation. Please note in the email that you have received a logfile upload error and are requesting a logfile report. ******On about 09/20/2023, a formal Urgent Medical Device Communication notification was sent to Healthcare Professionals (HCPs) informing them of the resolution of the autologs web portal issue previously communicated. Consignees are asked to confirm receipt of the letter and that they have acknowledged the actions to be taken. Quantity in Commerce 19 units Distribution Worldwide - US Nationwide distribution in the states of TX, MO, VA, and the countries of Bahrain, UK. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = DSQ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Medtronic Mechanical Circulatory Support issued an URGENT MEDICAL DEVICE COMMUNICATION notice to it consignees on 05/16/2023 by email. The notice explained that logfiles downloaded from the recently updated Model 1521 monitors (Serial Numbers: MON4xxxxxx for OUS and MON5xxxxxx for US), are unable to be processed by the Autologs web portal. ******Medtronic is working on an update to the Autologs web portal to resolve the issue. Until the update is complete, logfiles can be sent to dl.mcsclinicalengineering@medtronic.com for report generation. Please note in the email that you have received a logfile upload error and are requesting a logfile report. ******On about 09/20/2023, a formal Urgent Medical Device Communication notification was sent to Healthcare Professionals (HCPs) informing them of the resolution of the autologs web portal issue previously communicated. Consignees are asked to confirm receipt of the letter and that they have acknowledged the actions to be taken. | SoftwareChangeControl
14 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197517 | Z-1207-2023 | BD Pyxis Anesthesia ES System (PAS ES), REF: 327; BP Pyxis MedStation ES (Med ES Main), REF: 323; BD Pyxis MedStation ES Tower (Med ES Tower), REF: 352; all with BD Pyxis ES version 1.6.1 drawer firmware (version 1.10.2.5). Pyxis MedStation Enterprise Server (ES) System is an automated medication dispensing system supporting decentralized medication management in a pharmaceutical environment. | BD Pyxis | 2 | CDRH | 2023-03-17 | 2023-03-17 21:07 |  | CareFusion 303, Inc. | Automated dispensing cabinet drawer firmware may have inadvertently been downgraded to an earlier version of drawer firmware during new device installations or device replacement upgrades, which may lead to an increased frequency of drawer failures, which could increase the likelihood of patients experiencing delays in receiving medications. | Class 2 Device Recall BD Pyxis Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall BD Pyxis FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall BD Pyxis Date Initiated by Firm February 14, 2023 Date Posted March 17, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-1207-2023 Recall Event ID 91341 Product Classification Cabinet, table and tray, anesthesia - Product Code BRY Product BD Pyxis Anesthesia ES System (PAS ES), REF: 327; BP Pyxis MedStation ES (Med ES Main), REF: 323; BD Pyxis MedStation ES Tower (Med ES Tower), REF: 352; all with BD Pyxis ES version 1.6.1 drawer firmware (version 1.10.2.5). Pyxis MedStation Enterprise Server (ES) System is an automated medication dispensing system supporting decentralized medication management in a pharmaceutical environment. Code Information REF/UDI-DI: 327/10885403477836, 323/10885403512667, 352/10885403512674 Recalling Firm/ Manufacturer CareFusion 303, Inc. 10020 Pacific Mesa Blvd San Diego CA 92121-4386 For Additional Information Contact 888-876-7128 Manufacturer Reason for Recall Automated dispensing cabinet drawer firmware may have inadvertently been downgraded to an earlier version of drawer firmware during new device installations or device replacement upgrades, which may lead to an increased frequency of drawer failures, which could increase the likelihood of patients experiencing delays in receiving medications. FDA Determined Cause 2 Software change control Action On 01/14/23, correction notices were mailed to directors of pharmacy informing them to take the following actions:******1) Ensure this letter and contents are forwarded and/or discussed with all appropriate personnel to ensure proper awareness of the issue.***2) If you are experiencing any issues, such as drawer failures, call BD technical support and refer to user guides for additional information and troubleshooting procedures.***a. BD Pyxis MedStation ES user guide, refer to sections "Recover storage space" (Chapter 9, pg. 76) and "Emergency medication removal" (Appendix 11, pg. 119). See references 1 and 2 below for additional details.***b. BD Pyxis PAS ES user guide, refer to sections "Emergency access procedures" (Appendix A, pg. 77) and "Storage spaces are not responsive" (Appendix B, pg. 81). See references 3 and 4 below for additional details.***3) Complete and return the Customer Response Form.******The firm will take the following actions:******1) All customers who may have affected devices will be notified.***2) Upon receipt of completed response forms, the firm will coordinate the firmware upgrade with customers.******Technical Support: 1-800-727-6102, Phone hours: 24 hours, 7 days a week Quantity in Commerce 7914 Distribution Worldwide - US Nationwide distribution including in the states of NY, PA, OH, TX, FL, KY, IL, DE, AK, WI, KS, LA, WA, AZ, CO, MI, IN, TN, SD, AR, MO, CA, NV, MA, SC, VA, WV, CT, ME, NE, RI, GA, NC, IA, MN, AL, OK, MS, NM, NH, VT, WY, OR, ID, NJ, UT, DC, HI, MD, MT, ND and the country of Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 01/14/23, correction notices were mailed to directors of pharmacy informing them to take the following actions:******1) Ensure this letter and contents are forwarded and/or discussed with all appropriate personnel to ensure proper awareness of the issue.***2) If you are experiencing any issues, such as drawer failures, call BD technical support and refer to user guides for additional information and troubleshooting procedures.***a. BD Pyxis MedStation ES user guide, refer to sections "Recover storage space" (Chapter 9, pg. 76) and "Emergency medication removal" (Appendix 11, pg. 119). See references 1 and 2 below for additional details.***b. BD Pyxis PAS ES user guide, refer to sections "Emergency access procedures" (Appendix A, pg. 77) and "Storage spaces are not responsive" (Appendix B, pg. 81). See references 3 and 4 below for additional details.***3) Complete and return the Customer Response Form.******The firm will take the following actions:******1) All customers who may have affected devices will be notified.***2) Upon receipt of completed response forms, the firm will coordinate the firmware upgrade with customers.******Technical Support: 1-800-727-6102, Phone hours: 24 hours, 7 days a week | SoftwareChangeControl
15 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197345 | Z-0776-2023 | Guardian iOS app (MMT-8200) and Guardian Android app (MMT-8201), part of the Guardian 4 system | Guardian App for iOS v. 1.3.0; Guardian App for Android v. 1.3.0 | 2 | CDRH | 2022-12-31 | 2023-01-01 20:06 |  | Medtronic MiniMed | An app, part of a continuous glucose monitoring system, for use with smartphone devices may automatically log out from CareLink, then the app is not able to upload data. When logged out, linked care partners will not receive SMS notifications (could result in hypoglycemia or hyperglycemia), and sensor glucose values will also not be sent to the InPen app. | Class 2 Device Recall Guardian App for iOS v. 1.3.0; Guardian App for Android v. 1.3.0 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Guardian App for iOS v. 1.3.0; Guardian App for Android v. 1.3.0 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Guardian App for iOS v. 1.3.0; Guardian App for Android v. 1.3.0 Date Initiated by Firm November 21, 2022 Date Posted January 01, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-0776-2023 Recall Event ID 91260 Product Classification Sensor, glucose, invasive - Product Code MDS Product Guardian iOS app (MMT-8200) and Guardian Android app (MMT-8201), part of the Guardian 4 system Code Information App Version 1.3.0 Recalling Firm/ Manufacturer Medtronic MiniMed 18000 Devonshire St Northridge CA 91325-1219 For Additional Information Contact 818-576-4700 Manufacturer Reason for Recall An app, part of a continuous glucose monitoring system, for use with smartphone devices may automatically log out from CareLink, then the app is not able to upload data. When logged out, linked care partners will not receive SMS notifications (could result in hypoglycemia or hyperglycemia), and sensor glucose values will also not be sent to the InPen app. FDA Determined Cause 2 Software change control Action On 11/21/22, correction notices were sent to patients, care partners, and healthcare providers. The following actions were requested of users:***iOS/Android Users:***1) Verify what version of the app you are running by taking the following steps:***a) On the home screen, navigate to the settings menu on the upper left corner***b) Select "Help"***c) Confirm "Version"***2) If you are not running version 1.3.0, no action is necessary***3) If you have version 1.3.0, update to the latest app version currently available in the App Store (1.3.1) by following these steps:******iOS Users:***a) Open the App Store app.***b) Tap your profile icon at the top of the screen.***c) Scroll to see pending updates.***d) Tap update next to the mobile app icon to update the app.******Android Users:***a) Open the Google Play Store app.***b) Tap Menu and then Manage apps & devices.***c) Apps with an update available are labelled "Update available". You can also search for a specific app.***d) Tap Update.******Once you update to version 1.3.1 and log in to the Guardian App all your past data from version 1.3.0 will remain available and visible.******It is not recommended or necessary to uninstall the current version of the app from your phone before installing the new version 1.3.1. If you choose to do so, your local data will not able available in the app and you will be prompted to pair your transmitter.******Complete and return the Customer Confirmation Form. Form can be returned to rs.safetyinformation@medtronic.com******For questions, contact our 24-Hour Technical Support line at 1-800-646-4633, option 1. Quantity in Commerce 3235 Distribution OUS: Austria, Belgium, Chile, Colombia, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Kazakhstan, Kuwait, Luxembourg, Moldova, Republic of, Netherlands, New Zealand, Norway, Poland, Portugal, Qatar, Romania, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 11/21/22, correction notices were sent to patients, care partners, and healthcare providers. The following actions were requested of users:***iOS/Android Users:***1) Verify what version of the app you are running by taking the following steps:***a) On the home screen, navigate to the settings menu on the upper left corner***b) Select "Help"***c) Confirm "Version"***2) If you are not running version 1.3.0, no action is necessary***3) If you have version 1.3.0, update to the latest app version currently available in the App Store (1.3.1) by following these steps:******iOS Users:***a) Open the App Store app.***b) Tap your profile icon at the top of the screen.***c) Scroll to see pending updates.***d) Tap update next to the mobile app icon to update the app.******Android Users:***a) Open the Google Play Store app.***b) Tap Menu and then Manage apps & devices.***c) Apps with an update available are labelled "Update available". You can also search for a specific app.***d) Tap Update.******Once you update to version 1.3.1 and log in to the Guardian App all your past data from version 1.3.0 will remain available and visible.******It is not recommended or necessary to uninstall the current version of the app from your phone before installing the new version 1.3.1. If you choose to do so, your local data will not able available in the app and you will be prompted to pair your transmitter.******Complete and return the Customer Confirmation Form. Form can be returned to rs.safetyinformation@medtronic.com******For questions, contact our 24-Hour Technical Support line at 1-800-646-4633, option 1. | SoftwareChangeControl
16 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194595 | Z-1403-2022 | RayStation 9B SP1. For radiation treatment planning. | RayStation | 2 | CDRH | 2022-07-19 |  | 2023-10-04 11:09 | RAYSEARCH LABORATORIES AB | An issue with propagation of treatment course information from RayStation with RayTreat to RayCare has been identified. During some workflows, information may not be propagated to RayCare. *** | Class 2 Device Recall RayStation Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall RayStation FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall RayStation Date Initiated by Firm June 01, 2022 Create Date July 19, 2022 Recall Status 1 Terminated 3 on October 04, 2023 Recall Number Z-1403-2022 Recall Event ID 90498 510(K)Number K200569 Product Classification System, planning, radiation therapy treatment - Product Code MUJ Product RayStation 9B SP1. For radiation treatment planning. Code Information UDI-DI: 0735000201029720200310***Build number 9.2.0.483 Recalling Firm/ Manufacturer RAYSEARCH LABORATORIES AB Sveavagen 9 Stockholm Sweden Manufacturer Reason for Recall An issue with propagation of treatment course information from RayStation with RayTreat to RayCare has been identified. During some workflows, information may not be propagated to RayCare. FDA Determined Cause 2 Software change control Action The firm contacted all customers via email on June 1, 2022. Customers were provided with a list of workflows where safety related information may not be propagated.******Customers were instructed to be aware that treatment course information in RayCare may not always be up to date. Customers should verify critical information in RayStation or RayTreat before treatment course information in RayCare is used for critical decisions.******Customers were asked to identify all installed units with the affected software version numbers.******The issue will be resolved in mandatory software updates for RayStation, scheduled for market release in June 2022. Quantity in Commerce 1 Distribution Worldwide distribution - US Nationwide distribution in the state of TN and the countries of Belgium and Switzerland. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MUJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm contacted all customers via email on June 1, 2022. Customers were provided with a list of workflows where safety related information may not be propagated.******Customers were instructed to be aware that treatment course information in RayCare may not always be up to date. Customers should verify critical information in RayStation or RayTreat before treatment course information in RayCare is used for critical decisions.******Customers were asked to identify all installed units with the affected software version numbers.******The issue will be resolved in mandatory software updates for RayStation, scheduled for market release in June 2022. | SoftwareChangeControl
17 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193403 | Z-1075-2022 | Digital Medical X-ray Imaging System, Model: uDR 596i, uDR 592h | uDR 592h, uDR 596i | 2 | CDRH | 2022-05-12 |  | 2022-05-17 11:13 | Shanghai United Imaging Healthcare Co., Ltd. | X-ray imaging system positioning image and protocol label is reversed for Flexion and Extension on C-Spine and L-Spine, this will cause the image to be incorrectly labeled and may cause the patient to have a repeat exposure. | Class 2 Device Recall uDR 592h, uDR 596i Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall uDR 592h, uDR 596i FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall uDR 592h, uDR 596i Date Initiated by Firm July 30, 2019 Create Date May 12, 2022 Recall Status 1 Terminated 3 on May 17, 2022 Recall Number Z-1075-2022 Recall Event ID 90104 510(K)Number K181413 Product Classification System, x-ray, stationary - Product Code KPR Product Digital Medical X-ray Imaging System, Model: uDR 596i, uDR 592h Code Information DI/GTIN: 06971576833023, Serial Numbers: 270012, 270013,270014, 270015 Recalling Firm/ Manufacturer Shanghai United Imaging Healthcare Co., Ltd. No. 2258 Chengbei Rd, Jiading Ind. Shanghai China Manufacturer Reason for Recall X-ray imaging system positioning image and protocol label is reversed for Flexion and Extension on C-Spine and L-Spine, this will cause the image to be incorrectly labeled and may cause the patient to have a repeat exposure. FDA Determined Cause 2 Software change control Action On 07/30/19, Customer Notification Letters started being distributed to customers. Customers were asked to contact the firm if they found an inconsistency between the protocol name and the positioning image. The following phone numbers were provided: (855) 221-1552 (toll free) or (832) 699-2799 (direct). ******The firm issued Field Correction Instructions to Service Engineers to modify the protocol name of Flexion and Extension on C-Spine & L-Spine with the corrected guidelines. In addition, the firm committed to permanently fix impacted systems with a software update. Quantity in Commerce 4 Distribution US Nationwide distribution in the state of TX. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = KPR - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 07/30/19, Customer Notification Letters started being distributed to customers. Customers were asked to contact the firm if they found an inconsistency between the protocol name and the positioning image. The following phone numbers were provided: (855) 221-1552 (toll free) or (832) 699-2799 (direct). ******The firm issued Field Correction Instructions to Service Engineers to modify the protocol name of Flexion and Extension on C-Spine & L-Spine with the corrected guidelines. In addition, the firm committed to permanently fix impacted systems with a software update. | SoftwareChangeControl
18 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192914 | Z-0995-2022 | syngo.via RT Image Suite upgraded to VB60 running on SOMATOM go.Up Model Number 11061620. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade.*** | SOMATOM go.Up | 2 | CDRH | 2022-04-22 |  | 2023-12-19 10:18 | Siemens Medical Solutions USA, Inc | Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step. | Class 2 Device Recall SOMATOM go.Up Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SOMATOM go.Up FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SOMATOM go.Up Date Initiated by Firm March 07, 2022 Create Date April 22, 2022 Recall Status 1 Terminated 3 on December 19, 2023 Recall Number Z-0995-2022 Recall Event ID 89915 510(K)Number K211373 Product Classification System, x-ray, tomography, computed - Product Code JAK Product syngo.via RT Image Suite upgraded to VB60 running on SOMATOM go.Up Model Number 11061620. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade. Code Information Partial UDI: 04056869051314, Serial Number 112261 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact Rebecca Tudor 610-219-4834 Manufacturer Reason for Recall Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step. FDA Determined Cause 2 Software change control Action 4/4/22 update: Urgent Medical Device Correction letter has not been issued. The proposed letter informs customers of software issue and workaround to resolve issue. It also requires electronic signature to acknowledge receipt of the notification. If you have additional questions contact the service organization at 1-800-888-7436. Quantity in Commerce 1 (US) Distribution Worldwide distribution - US Nationwide distribution in the states of AL, AR, AZ, CA, CO, CT, IL, IN, KY, MD, MO, ND, NH, NY, OH, OR, PA, SD, TX, VA and the countries of Argentina, Austria, Bahrain, Belgium, Brazil, Czechia, Denmark, France, Germany, Hungary, India, Italy, Japan, Netherlands, Norway, Poland, Romania, Slovakia, Spain, Switzerland, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action 4/4/22 update: Urgent Medical Device Correction letter has not been issued. The proposed letter informs customers of software issue and workaround to resolve issue. It also requires electronic signature to acknowledge receipt of the notification. If you have additional questions contact the service organization at 1-800-888-7436. | SoftwareChangeControl
19 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192915 | Z-0996-2022 | syngo.via RT Image Suite upgraded to VB60 running on SOMATOM go.Sim Model Number 11061660. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade. | SOMATOM go.Sim | 2 | CDRH | 2022-04-22 |  | 2023-12-19 10:18 | Siemens Medical Solutions USA, Inc | Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step. | Class 2 Device Recall SOMATOM go.Sim Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SOMATOM go.Sim FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SOMATOM go.Sim Date Initiated by Firm March 07, 2022 Create Date April 22, 2022 Recall Status 1 Terminated 3 on December 19, 2023 Recall Number Z-0996-2022 Recall Event ID 89915 510(K)Number K200524 Product Classification System, x-ray, tomography, computed - Product Code JAK Product syngo.via RT Image Suite upgraded to VB60 running on SOMATOM go.Sim Model Number 11061660. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade. Code Information Partial UDI: 04056869249230, Serial Numbers 128123, 128106, 128109, 128144, 128124, 128112, 128132, 128135, 128114, 128125, 128113, 128142, 128107, 128136, 128133 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact Rebecca Tudor 610-219-4834 Manufacturer Reason for Recall Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step. FDA Determined Cause 2 Software change control Action 4/4/22 update: Urgent Medical Device Correction letter has not been issued. The proposed letter informs customers of software issue and workaround to resolve issue. It also requires electronic signature to acknowledge receipt of the notification. If you have additional questions contact the service organization at 1-800-888-7436. Quantity in Commerce 15 (US) Distribution Worldwide distribution - US Nationwide distribution in the states of AL, AR, AZ, CA, CO, CT, IL, IN, KY, MD, MO, ND, NH, NY, OH, OR, PA, SD, TX, VA and the countries of Argentina, Austria, Bahrain, Belgium, Brazil, Czechia, Denmark, France, Germany, Hungary, India, Italy, Japan, Netherlands, Norway, Poland, Romania, Slovakia, Spain, Switzerland, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action 4/4/22 update: Urgent Medical Device Correction letter has not been issued. The proposed letter informs customers of software issue and workaround to resolve issue. It also requires electronic signature to acknowledge receipt of the notification. If you have additional questions contact the service organization at 1-800-888-7436. | SoftwareChangeControl
20 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192916 | Z-0997-2022 | syngo.via RT Image Suite upgraded to VB60 running on SOMATOM go.Open Pro Model Number 11061670. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade. | SOMATOM go.Open Pro | 2 | CDRH | 2022-04-22 |  | 2023-12-19 10:18 | Siemens Medical Solutions USA, Inc | Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step. | Class 2 Device Recall SOMATOM go.Open Pro Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SOMATOM go.Open Pro FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SOMATOM go.Open Pro Date Initiated by Firm March 07, 2022 Create Date April 22, 2022 Recall Status 1 Terminated 3 on December 19, 2023 Recall Number Z-0997-2022 Recall Event ID 89915 510(K)Number K200524 Product Classification System, x-ray, tomography, computed - Product Code JAK Product syngo.via RT Image Suite upgraded to VB60 running on SOMATOM go.Open Pro Model Number 11061670. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade. Code Information Partial UDI: 04056869249247, Serial Numbers 130088, 130081, 130075, 130090, 130087, 130074, 130069, 130091, 130092, 130086, 130079, 130089, 130077, 130071 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact Rebecca Tudor 610-219-4834 Manufacturer Reason for Recall Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step. FDA Determined Cause 2 Software change control Action 4/4/22 update: Urgent Medical Device Correction letter has not been issued. The proposed letter informs customers of software issue and workaround to resolve issue. It also requires electronic signature to acknowledge receipt of the notification. If you have additional questions contact the service organization at 1-800-888-7436. Quantity in Commerce 14 (US) Distribution Worldwide distribution - US Nationwide distribution in the states of AL, AR, AZ, CA, CO, CT, IL, IN, KY, MD, MO, ND, NH, NY, OH, OR, PA, SD, TX, VA and the countries of Argentina, Austria, Bahrain, Belgium, Brazil, Czechia, Denmark, France, Germany, Hungary, India, Italy, Japan, Netherlands, Norway, Poland, Romania, Slovakia, Spain, Switzerland, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action 4/4/22 update: Urgent Medical Device Correction letter has not been issued. The proposed letter informs customers of software issue and workaround to resolve issue. It also requires electronic signature to acknowledge receipt of the notification. If you have additional questions contact the service organization at 1-800-888-7436. | SoftwareChangeControl
21 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192913 | Z-0994-2022 | syngo.via RT Image Suite upgraded to VB60 running on syngo.via Model Number 10496180. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade. | syngo.via RT Image Suite | 2 | CDRH | 2022-04-22 |  | 2023-12-19 10:18 | Siemens Medical Solutions USA, Inc | Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step. | Class 2 Device Recall syngo.via RT Image Suite Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall syngo.via RT Image Suite FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall syngo.via RT Image Suite Date Initiated by Firm March 07, 2022 Create Date April 22, 2022 Recall Status 1 Terminated 3 on December 19, 2023 Recall Number Z-0994-2022 Recall Event ID 89915 510(K)Number K211379 Product Classification System, planning, radiation therapy treatment - Product Code MUJ Product syngo.via RT Image Suite upgraded to VB60 running on syngo.via Model Number 10496180. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade. Code Information Serial Number 133527 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact Rebecca Tudor 610-219-4834 Manufacturer Reason for Recall Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step. FDA Determined Cause 2 Software change control Action 4/4/22 update: Urgent Medical Device Correction letter has not been issued. The proposed letter informs customers of software issue and workaround to resolve issue. It also requires electronic signature to acknowledge receipt of the notification. If you have additional questions contact the service organization at 1-800-888-7436. Quantity in Commerce 1 (US) Distribution Worldwide distribution - US Nationwide distribution in the states of AL, AR, AZ, CA, CO, CT, IL, IN, KY, MD, MO, ND, NH, NY, OH, OR, PA, SD, TX, VA and the countries of Argentina, Austria, Bahrain, Belgium, Brazil, Czechia, Denmark, France, Germany, Hungary, India, Italy, Japan, Netherlands, Norway, Poland, Romania, Slovakia, Spain, Switzerland, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MUJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action 4/4/22 update: Urgent Medical Device Correction letter has not been issued. The proposed letter informs customers of software issue and workaround to resolve issue. It also requires electronic signature to acknowledge receipt of the notification. If you have additional questions contact the service organization at 1-800-888-7436. | SoftwareChangeControl
22 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=191357 | Z-0622-2022 | bellavista 1000 ventilator, Catalog #301.100.030 and bellavista 1000e ventilator, 301.100.130.******The bellavista 1000/1000e ventilator is intended to provide positive pressure ventilatory support to adult and pediatric patients, and optionally, infant and neonatal patients. | bellavista | 1 | CDRH | 2022-02-16 17:27 | 2022-02-16 20:06 | 2024-08-28 09:13 | Vyaire Medical | Potential cessation in ventilation can occur under specific conditions. | Class 1 Device Recall bellavista Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall bellavista FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall bellavista Date Initiated by Firm December 20, 2021 Date Posted February 16, 2022 Recall Status 1 Terminated 3 on August 28, 2024 Recall Number Z-0622-2022 Recall Event ID 89366 510(K)Number K183364 Product Classification Ventilator, continuous, facility use - Product Code CBK Product bellavista 1000 ventilator, Catalog #301.100.030 and bellavista 1000e ventilator, 301.100.130.************The bellavista 1000/1000e ventilator is intended to provide positive pressure ventilatory support to adult and pediatric patients, and optionally, infant and neonatal patients. Code Information Hardware generation G6, software version V6.0.1600.0 or higher installed, all serial numbers. UDI 07640149381115 (bellavista 1000 ventilator) and 07640149388879 (bellavista 1000e ventilator). Recalling Firm/ Manufacturer Vyaire Medical 26125 N Riverwoods Blvd Mettawa IL 60045-3420 For Additional Information Contact Dr. Michael Pedro 339-293-9211 Manufacturer Reason for Recall Potential cessation in ventilation can occur under specific conditions. FDA Determined Cause 2 Software change control Action International consignees were notified of the correction via email notification beginning 12/20/2021 containing a Field Safety Notice, an Appendix to the Field Safety Notice which lists the affected units, an FSCA Distributor Response Form and an FSCA End User Response form, all in English, with the notation that localized versions were currently in translation and will be forwarded as soon as available, if applicable. The email provided information on the affected model numbers and actions to be taken by the distributor. These actions included notifying all affected end users by providing them with the FSCA package containing the notice and the list of affected units, notify the appropriate regulatory authorities as applicable, return the completed and signed FSCA Distributor Response form, and letting them know that if any user facilities have distributed affected products to other persons or facilities, promptly forward a copy of the Field Safety Notice and the list of affected units. The email informs them that imtmedical ag will inform them of the availability of the software patch in a separate communication for installation.******The recalling firm, Vyaire Medical, issued two different notifications titled "URGENT: MEDICAL DEVICE CORRECTION" to distributors and end users on 12/23/2021 via electronic communication. Physical copies of the notification will be issued to each consignee via UPS Next Day Air delivery.******The Distributor letter informed the distributor of the issue and provided a problem description, problem root cause, and said that in order to prevent adverse health consequences associated with this failure mode during use of the bellavista 1000 and bellavista 1000e ventilators, hardware generation G6, end users are encouraged to follow the User Manual and take the immediate mitigation actions listed in the End-User Notification package. The actions to be taken by the distributor included: (1) Immediately notify all affected end users by providi Quantity in Commerce 4189 units Distribution Worldwide distribution - US Nationwide distribution in the states of AL, CA, CO, CT, DC, FL, GA, ID, IL, IN, LA, MA, MD, MI, MO, MS, NC, NJ, NY, OH, OK, PA, SD, TX, VA, WA, WI and WV, including Puerto Rico and the countries of Argentina, Australia, Austria, Bahrain, Belgium, Bosnia and Herzegovina, Brazil, Brunei, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Dominican Republic, Egypt, El Salvador, France, Germany, Guatemala, India, Indonesia, Iran, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Latvia, Lithuania, Malaysia, Mexico, Namibia, Netherlands, Nicaragua, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, South Africa, South Korea, Spain, Switzerland, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam, and Zimbabwe. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = CBK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action International consignees were notified of the correction via email notification beginning 12/20/2021 containing a Field Safety Notice, an Appendix to the Field Safety Notice which lists the affected units, an FSCA Distributor Response Form and an FSCA End User Response form, all in English, with the notation that localized versions were currently in translation and will be forwarded as soon as available, if applicable. The email provided information on the affected model numbers and actions to be taken by the distributor. These actions included notifying all affected end users by providing them with the FSCA package containing the notice and the list of affected units, notify the appropriate regulatory authorities as applicable, return the completed and signed FSCA Distributor Response form, and letting them know that if any user facilities have distributed affected products to other persons or facilities, promptly forward a copy of the Field Safety Notice and the list of affected units. The email informs them that imtmedical ag will inform them of the availability of the software patch in a separate communication for installation.******The recalling firm, Vyaire Medical, issued two different notifications titled "URGENT: MEDICAL DEVICE CORRECTION" to distributors and end users on 12/23/2021 via electronic communication. Physical copies of the notification will be issued to each consignee via UPS Next Day Air delivery.******The Distributor letter informed the distributor of the issue and provided a problem description, problem root cause, and said that in order to prevent adverse health consequences associated with this failure mode during use of the bellavista 1000 and bellavista 1000e ventilators, hardware generation G6, end users are encouraged to follow the User Manual and take the immediate mitigation actions listed in the End-User Notification package. The actions to be taken by the distributor included: (1) Immediately notify all affected end users by providi | SoftwareChangeControl
23 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189350 | Z-0165-2022 | REF/Catalogue Number GSX0020A, GORE CARDIOFORM SEPTAL OCCLUDER 20 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631018 | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. | Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Date Initiated by Firm September 02, 2021 Create Date October 22, 2021 Recall Status 1 Open 3 , Classified Recall Number Z-0165-2022 Recall Event ID 88683 PMA Number P050006S044 Product Classification Transcatheter septal occluder - Product Code MLV Product REF/Catalogue Number GSX0020A, GORE CARDIOFORM SEPTAL OCCLUDER 20 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631018 Code Information serial numbers 22689696 through 23569078 Recalling Firm/ Manufacturer W L Gore & Associates, Inc. 3750 W Kiltie Ln Flagstaff AZ 86005-8712 For Additional Information Contact Gore Customer Service 800-528-8763 Manufacturer Reason for Recall Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. FDA Determined Cause 2 Software change control Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm Quantity in Commerce 138 units Distribution US: AL,AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN,MO,MS,MT,NC, ND, NE, NH, NJ, NM,NV,NY,OH,OK,OR,PA,SC,SD,TN,TX,UT,VA,VT,WA,WI, WV******OUS: AUS, AUT, CAN, CHE, DEU, DNK, ESP, FRA, GBR, GRC, IRL, ITA, MEX, NLD, NOR, PRT, SAU, SWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = MLV - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm | SoftwareChangeControl
24 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189357 | Z-0172-2022 | REF/Catalogue Number GSX0030H, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132648948****Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. | Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Date Initiated by Firm September 02, 2021 Create Date October 22, 2021 Recall Status 1 Open 3 , Classified Recall Number Z-0172-2022 Recall Event ID 88683 PMA Number P050006S044 Product Classification Transcatheter septal occluder - Product Code MLV Product REF/Catalogue Number GSX0030H, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132648948****Not distributed within the US Code Information serial numbers 22689696 through 23569078 Recalling Firm/ Manufacturer W L Gore & Associates, Inc. 3750 W Kiltie Ln Flagstaff AZ 86005-8712 For Additional Information Contact Gore Customer Service 800-528-8763 Manufacturer Reason for Recall Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. FDA Determined Cause 2 Software change control Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm Quantity in Commerce 15 units Distribution US: AL,AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN,MO,MS,MT,NC, ND, NE, NH, NJ, NM,NV,NY,OH,OK,OR,PA,SC,SD,TN,TX,UT,VA,VT,WA,WI, WV******OUS: AUS, AUT, CAN, CHE, DEU, DNK, ESP, FRA, GBR, GRC, IRL, ITA, MEX, NLD, NOR, PRT, SAU, SWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = MLV - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm | SoftwareChangeControl
25 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189355 | Z-0170-2022 | REF/Catalogue Number GSXE0025B, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132620111****Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. | Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Date Initiated by Firm September 02, 2021 Create Date October 22, 2021 Recall Status 1 Open 3 , Classified Recall Number Z-0170-2022 Recall Event ID 88683 PMA Number P050006S044 Product Classification Transcatheter septal occluder - Product Code MLV Product REF/Catalogue Number GSXE0025B, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132620111****Not distributed within the US Code Information serial numbers 22689696 through 23569078 Recalling Firm/ Manufacturer W L Gore & Associates, Inc. 3750 W Kiltie Ln Flagstaff AZ 86005-8712 For Additional Information Contact Gore Customer Service 800-528-8763 Manufacturer Reason for Recall Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. FDA Determined Cause 2 Software change control Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm Quantity in Commerce 187 units Distribution US: AL,AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN,MO,MS,MT,NC, ND, NE, NH, NJ, NM,NV,NY,OH,OK,OR,PA,SC,SD,TN,TX,UT,VA,VT,WA,WI, WV******OUS: AUS, AUT, CAN, CHE, DEU, DNK, ESP, FRA, GBR, GRC, IRL, ITA, MEX, NLD, NOR, PRT, SAU, SWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = MLV - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm | SoftwareChangeControl
26 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189354 | Z-0169-2022 | REF/Catalogue Number GSXE0030, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617661****Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. | Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Date Initiated by Firm September 02, 2021 Create Date October 22, 2021 Recall Status 1 Open 3 , Classified Recall Number Z-0169-2022 Recall Event ID 88683 PMA Number P050006S044 Product Classification Transcatheter septal occluder - Product Code MLV Product REF/Catalogue Number GSXE0030, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617661****Not distributed within the US Code Information serial numbers 22689696 through 23569078 Recalling Firm/ Manufacturer W L Gore & Associates, Inc. 3750 W Kiltie Ln Flagstaff AZ 86005-8712 For Additional Information Contact Gore Customer Service 800-528-8763 Manufacturer Reason for Recall Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. FDA Determined Cause 2 Software change control Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm Quantity in Commerce 846 units Distribution US: AL,AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN,MO,MS,MT,NC, ND, NE, NH, NJ, NM,NV,NY,OH,OK,OR,PA,SC,SD,TN,TX,UT,VA,VT,WA,WI, WV******OUS: AUS, AUT, CAN, CHE, DEU, DNK, ESP, FRA, GBR, GRC, IRL, ITA, MEX, NLD, NOR, PRT, SAU, SWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = MLV - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm | SoftwareChangeControl
27 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189352 | Z-0167-2022 | REF/Catalogue Number GSXE0020, GORE CARDIOFORM SEPTAL OCCLUDER 20 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617630****Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. | Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Date Initiated by Firm September 02, 2021 Create Date October 22, 2021 Recall Status 1 Open 3 , Classified Recall Number Z-0167-2022 Recall Event ID 88683 PMA Number P050006S044 Product Classification Transcatheter septal occluder - Product Code MLV Product REF/Catalogue Number GSXE0020, GORE CARDIOFORM SEPTAL OCCLUDER 20 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617630****Not distributed within the US Code Information serial numbers 22689696 through 23569078 Recalling Firm/ Manufacturer W L Gore & Associates, Inc. 3750 W Kiltie Ln Flagstaff AZ 86005-8712 For Additional Information Contact Gore Customer Service 800-528-8763 Manufacturer Reason for Recall Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. FDA Determined Cause 2 Software change control Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm Quantity in Commerce 95 units Distribution US: AL,AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN,MO,MS,MT,NC, ND, NE, NH, NJ, NM,NV,NY,OH,OK,OR,PA,SC,SD,TN,TX,UT,VA,VT,WA,WI, WV******OUS: AUS, AUT, CAN, CHE, DEU, DNK, ESP, FRA, GBR, GRC, IRL, ITA, MEX, NLD, NOR, PRT, SAU, SWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = MLV - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm | SoftwareChangeControl
28 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189353 | Z-0168-2022 | REF/Catalogue Number GSXE0025, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617647****Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. | Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Date Initiated by Firm September 02, 2021 Create Date October 22, 2021 Recall Status 1 Open 3 , Classified Recall Number Z-0168-2022 Recall Event ID 88683 PMA Number P050006S044 Product Classification Transcatheter septal occluder - Product Code MLV Product REF/Catalogue Number GSXE0025, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617647****Not distributed within the US Code Information serial numbers 22689696 through 23569078 Recalling Firm/ Manufacturer W L Gore & Associates, Inc. 3750 W Kiltie Ln Flagstaff AZ 86005-8712 For Additional Information Contact Gore Customer Service 800-528-8763 Manufacturer Reason for Recall Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. FDA Determined Cause 2 Software change control Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm Quantity in Commerce 1268 unit Distribution US: AL,AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN,MO,MS,MT,NC, ND, NE, NH, NJ, NM,NV,NY,OH,OK,OR,PA,SC,SD,TN,TX,UT,VA,VT,WA,WI, WV******OUS: AUS, AUT, CAN, CHE, DEU, DNK, ESP, FRA, GBR, GRC, IRL, ITA, MEX, NLD, NOR, PRT, SAU, SWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = MLV - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm | SoftwareChangeControl
29 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189356 | Z-0171-2022 | REF/Catalogue Number GSXE0030B, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132620128****Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. | Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Date Initiated by Firm September 02, 2021 Create Date October 22, 2021 Recall Status 1 Open 3 , Classified Recall Number Z-0171-2022 Recall Event ID 88683 PMA Number P050006S044 Product Classification Transcatheter septal occluder - Product Code MLV Product REF/Catalogue Number GSXE0030B, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132620128****Not distributed within the US Code Information serial numbers 22689696 through 23569078 Recalling Firm/ Manufacturer W L Gore & Associates, Inc. 3750 W Kiltie Ln Flagstaff AZ 86005-8712 For Additional Information Contact Gore Customer Service 800-528-8763 Manufacturer Reason for Recall Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. FDA Determined Cause 2 Software change control Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm Quantity in Commerce 222 units Distribution US: AL,AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN,MO,MS,MT,NC, ND, NE, NH, NJ, NM,NV,NY,OH,OK,OR,PA,SC,SD,TN,TX,UT,VA,VT,WA,WI, WV******OUS: AUS, AUT, CAN, CHE, DEU, DNK, ESP, FRA, GBR, GRC, IRL, ITA, MEX, NLD, NOR, PRT, SAU, SWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = MLV - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm | SoftwareChangeControl
30 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189351 | Z-0166-2022 | REF/Catalogue Number GSX0025A, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631025****Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. | Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Date Initiated by Firm September 02, 2021 Create Date October 22, 2021 Recall Status 1 Open 3 , Classified Recall Number Z-0166-2022 Recall Event ID 88683 PMA Number P050006S044 Product Classification Transcatheter septal occluder - Product Code MLV Product REF/Catalogue Number GSX0025A, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631025****Not distributed within the US Code Information serial numbers 22689696 through 23569078 Recalling Firm/ Manufacturer W L Gore & Associates, Inc. 3750 W Kiltie Ln Flagstaff AZ 86005-8712 For Additional Information Contact Gore Customer Service 800-528-8763 Manufacturer Reason for Recall Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. FDA Determined Cause 2 Software change control Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm Distribution US: AL,AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN,MO,MS,MT,NC, ND, NE, NH, NJ, NM,NV,NY,OH,OK,OR,PA,SC,SD,TN,TX,UT,VA,VT,WA,WI, WV******OUS: AUS, AUT, CAN, CHE, DEU, DNK, ESP, FRA, GBR, GRC, IRL, ITA, MEX, NLD, NOR, PRT, SAU, SWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = MLV - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: |  | SoftwareChangeControl
31 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189284 | Z-0164-2022 | REF/Catalogue Number GSX0030A, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631032 | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. | Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GORE CARDIOFORM SEPTAL OCCLUDER Date Initiated by Firm September 02, 2021 Create Date October 22, 2021 Recall Status 1 Open 3 , Classified Recall Number Z-0164-2022 Recall Event ID 88683 PMA Number P050006S044 Product Classification Transcatheter septal occluder - Product Code MLV Product REF/Catalogue Number GSX0030A, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631032 Code Information serial numbers 22689696 through 23569078 Recalling Firm/ Manufacturer W L Gore & Associates, Inc. 3750 W Kiltie Ln Flagstaff AZ 86005-8712 For Additional Information Contact Gore Customer Service 800-528-8763 Manufacturer Reason for Recall Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs. FDA Determined Cause 2 Software change control Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm Quantity in Commerce 2,238 units Distribution US: AL,AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN,MO,MS,MT,NC, ND, NE, NH, NJ, NM,NV,NY,OH,OK,OR,PA,SC,SD,TN,TX,UT,VA,VT,WA,WI, WV******OUS: AUS, AUT, CAN, CHE, DEU, DNK, ESP, FRA, GBR, GRC, IRL, ITA, MEX, NLD, NOR, PRT, SAU, SWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = MLV - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Gore initiated distribution of "Urgent Medical Device Recall" Notification to affected consignees between the dates of September 2, 2021 - September 21, 2021. The customer communication was distributed via Email and UPS. In addition, to informing consignees about the recall, Gore asked consignees to take the following actions:******1. Identify and return any unused devices within the scope of this voluntary recall.***Return Authorization number can be found on the CUSTOMER RESPONSE FORM.******2. Please complete and sign the enclosed CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this notification.******3. Please share this letter with others in your hospital or clinic as appropriate.******4. If a listed device has been used, there is no patient-follow up needed and there are***no further actions needed other than informing Gore the device was used. Please***indicate the used devices on the CUSTOMER RESPONSE FORM and return to***WLGore3120@sedgwick.com within 3 weeks of receipt of this letter. No further***action is needed.******5. Please contact your local Gore field sales associate with any questions regarding this notice, and to coordinate the return and replacement of any unused affected devices. Additionally, you may also contact Gore Customer Service (Email: MPDCustomerCare@wlgore.com; or by phone (U.S. toll free: (800) 528-8763)).******6. In the event that an adverse event occurs:***Any adverse event involving the GORE CARDIOFORM Septal Occluder should be***reported to the manufacturer and the country specific regulatory authorities immediately. To report an event to W. L. Gore & Associates, email: medcomplaints@wlgore.com***or contact: USA: 800 528 1866 Ext. 44922 / 928 864 4922, Fax 928 864 4364******Healthcare professionals and consumers may report adverse events or quality***problems directly to FDA using the FDA MedWatch Website:***https://www.accessdata.fda.gov/scripts/medwatch/index.cfm | SoftwareChangeControl
32 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188585 | Z-2471-2021 | ACUSON Juniper Diagnostic Ultrasound System | ACUSON Juniper ultrasound imaging system | 2 | CDRH | 2021-09-14 |  | 2024-10-16 07:45 | Siemens Medical Solutions USA, Inc. | Due to intermittent failures of the power supply in the ultrasound system which renders it inoperable. | Class 2 Device Recall ACUSON Juniper ultrasound imaging system Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ACUSON Juniper ultrasound imaging system FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ACUSON Juniper ultrasound imaging system Date Initiated by Firm July 12, 2021 Create Date September 14, 2021 Recall Status 1 Terminated 3 on October 16, 2024 Recall Number Z-2471-2021 Recall Event ID 88405 510(K)Number K201130 Product Classification System, imaging, pulsed doppler, ultrasonic - Product Code IYN Product ACUSON Juniper Diagnostic Ultrasound System Code Information Model Number: 11335791***UDI: (01)04056869152059 (21)123456 (240)11335791 (422)410***Firmware version 3.8 for rev. 7******U.S. Serial Number:***555606;******O.U.S. Serial Numbers:***550975;***552770;***553629;***553951;***555060;***555601;***555602;***555604;***555605;***555610;***555611;***555612;***555613;***555626;******O.U.S. Inactive Serial Numbers:***555603;***555607; Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc. 22010 SE 51st St Issaquah WA 98029-7298 For Additional Information Contact Mr. Khalil Thomas 425-392-9180 Manufacturer Reason for Recall Due to intermittent failures of the power supply in the ultrasound system which renders it inoperable. FDA Determined Cause 2 Software change control Action On 07/12/2021, the firm sent a "MEDICAL DEVICE SAFETY CORRECTION" Notification via email to customers informing them that there are intermittent failures of the power supply in the ultrasound system which may render the system inoperable. When an ultrasound system experiences this failure, cycling the power does not recover system functionality. Potential health risks include, lack of ultrasound guidance during interventional procedures (e.g., biopsy, guidance during surgery, patient under sedation) that rely on real-time ultrasound imaging to guide or complete the procedure.******Customer will be contacted by a Customer Service Engineer to schedule a visit to their facility to update the system.******Customers are asked to ensure that all users are aware of this issue and correction within their organization and others who may need to be inform of this situation.******For further questions, contact Siemens Healthineers USD Complaints at usd-complaint.team@siemens-healthineers.com Quantity in Commerce 15 systems Distribution Worldwide distribution - U.S. Nationwide distribution in the Puerto Rico and the countries of Bosnia Herzegovina, Canada, Colombia, Germany, Greece, Italy, North Macedonia, Poland, and Slovenia. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 07/12/2021, the firm sent a "MEDICAL DEVICE SAFETY CORRECTION" Notification via email to customers informing them that there are intermittent failures of the power supply in the ultrasound system which may render the system inoperable. When an ultrasound system experiences this failure, cycling the power does not recover system functionality. Potential health risks include, lack of ultrasound guidance during interventional procedures (e.g., biopsy, guidance during surgery, patient under sedation) that rely on real-time ultrasound imaging to guide or complete the procedure.******Customer will be contacted by a Customer Service Engineer to schedule a visit to their facility to update the system.******Customers are asked to ensure that all users are aware of this issue and correction within their organization and others who may need to be inform of this situation.******For further questions, contact Siemens Healthineers USD Complaints at usd-complaint.team@siemens-healthineers.com | SoftwareChangeControl
33 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187414 | Z-1841-2021 | Medline Tissue Recovery Instrument Set, Reorder #DYNJ64347A, packaged 3/case, Sterile, Rx.****** Convenience kit used to recover tissue, which can be used for medical research, education, or transplantation. | Tissue Recovery Instrument Set | 2 | CDRH | 2021-06-09 |  |  | Medline Industries Inc | The devices may not meet the minimum required sterility assurance level. | Class 2 Device Recall Tissue Recovery Instrument Set Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Tissue Recovery Instrument Set FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Tissue Recovery Instrument Set Date Initiated by Firm April 30, 2021 Create Date June 09, 2021 Recall Status 1 Open 3 , Classified Recall Number Z-1841-2021 Recall Event ID 87910 Product Classification General surgery tray - Product Code LRO Product Medline Tissue Recovery Instrument Set, Reorder #DYNJ64347A, packaged 3/case, Sterile, Rx.************ Convenience kit used to recover tissue, which can be used for medical research, education, or transplantation. Code Information GTIN (01)10193489291186, lot numbers 21CDA959, 21BDB665, 21ADA390, and 20KDC285 Recalling Firm/ Manufacturer Medline Industries Inc 3 Lakes Dr Northfield IL 60093-2753 Manufacturer Reason for Recall The devices may not meet the minimum required sterility assurance level. FDA Determined Cause 2 Software change control Action The recalling firm, Medline Industries, Inc., issued an "IMMEDIATE ACTION REQUIRED RECALL" letter dated and emailed on 4/30/2021. The letter described the product, the reason for recall and actions to be taken. The consignees were instructed to immediately check their inventory for the affected product, complete and submit response form online to:******Website link: https://recalls.medline.com; Recall Reference #: R-21-073; ******Recall Code: R5A2C566 , and quarantine it for subsequent return.************If you are a distributor, or have resold or transferred this product to another company or individual, you are required by FDA regulations to notify them of this recall communication. Have the customers return any affected product to you for collection and return to Medline Industries, Inc.************If you have any questions, please contact the Recall Department at 866-359-1704. Quantity in Commerce 36/3-kit cases (108 kits) Distribution US Distribution to state of: WI. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The recalling firm, Medline Industries, Inc., issued an "IMMEDIATE ACTION REQUIRED RECALL" letter dated and emailed on 4/30/2021. The letter described the product, the reason for recall and actions to be taken. The consignees were instructed to immediately check their inventory for the affected product, complete and submit response form online to:******Website link: https://recalls.medline.com; Recall Reference #: R-21-073; ******Recall Code: R5A2C566 , and quarantine it for subsequent return.************If you are a distributor, or have resold or transferred this product to another company or individual, you are required by FDA regulations to notify them of this recall communication. Have the customers return any affected product to you for collection and return to Medline Industries, Inc.************If you have any questions, please contact the Recall Department at 866-359-1704. | SoftwareChangeControl
34 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=185216 | Z-0918-2021 | FUJIFILM Synapse PACS Software: versions 7.0.0 and 7.0.1 - Product Usage: intended for use as a web based application on an off-the shelf PC which meets or exceeds minimum specifications and is networked with a FUJIFILM Synapse PACS server. | Synapse PACS | 2 | CDRH | 2021-01-22 |  | 2021-11-17 15:24 | Fujifilm Medical Systems U.S.A., Inc. | The software does not update measurements and calculations in the Clinical Reporting Application (CRA) when the ventricular trace is changed in the study by a different user. | Class 2 Device Recall Synapse PACS Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Synapse PACS FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Synapse PACS Date Initiated by Firm December 11, 2020 Create Date January 22, 2021 Recall Status 1 Terminated 3 on November 17, 2021 Recall Number Z-0918-2021 Recall Event ID 87083 510(K)Number K190232 Product Classification System, image processing, radiological - Product Code LLZ Product FUJIFILM Synapse PACS Software: versions 7.0.0 and 7.0.1 - Product Usage: intended for use as a web based application on an off-the shelf PC which meets or exceeds minimum specifications and is networked with a FUJIFILM Synapse PACS server. Code Information Versions 7.0.0 and 7.0.1 Recalling Firm/ Manufacturer Fujifilm Medical Systems U.S.A., Inc. 81 Hartwell Ave Ste 300 Lexington MA 02421-3160 For Additional Information Contact Jeffrey Wan 617-861-5300 Manufacturer Reason for Recall The software does not update measurements and calculations in the Clinical Reporting Application (CRA) when the ventricular trace is changed in the study by a different user. FDA Determined Cause 2 Software change control Action On 12/11/2020 Fujifilm sent consignee notifications via FedEx Priority Overnight stating that there are issues with Synapse PACS Software Version 7.0.000 and 7.0.100. Fujifilm provided a work around until a hotfix is scheduled and completed. Quantity in Commerce 1 unit Distribution US Nationwide distribution in the state of MD. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 12/11/2020 Fujifilm sent consignee notifications via FedEx Priority Overnight stating that there are issues with Synapse PACS Software Version 7.0.000 and 7.0.100. Fujifilm provided a work around until a hotfix is scheduled and completed. | SoftwareChangeControl
35 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184572 | Z-0717-2021 | APEX Compounding System Control Panel Module, REF: 601224, Product Code AX1000 - Product Usage: intended for intravascular administration to a patient. | Apex Compounding System | 2 | CDRH | 2020-12-23 |  | 2023-03-14 14:38 | B Braun Medical, Inc. | There is the potential for the compounding system to not immediately interrupt compounding and alert user to the presence of an air bubble exceeding 3% of the ordered volume for an ingredient. This could result in a a compounded drug that does not meet the specified accuracy range for an individual ingredient. | Class 2 Device Recall Apex Compounding System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Apex Compounding System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Apex Compounding System Date Initiated by Firm November 23, 2020 Create Date December 23, 2020 Recall Status 1 Terminated 3 on March 14, 2023 Recall Number Z-0717-2021 Recall Event ID 86803 Product Classification System/device, pharmacy compounding - Product Code NEP Product APEX Compounding System Control Panel Module, REF: 601224, Product Code AX1000 - Product Usage: intended for intravascular administration to a patient. Code Information All APEX Compounding System with Software Version 1.0.5.57 or earlier Recalling Firm/ Manufacturer B Braun Medical, Inc. 1601 Wallace Dr Ste 150 Carrollton TX 75006-6690 For Additional Information Contact B. Braun Customer Support 800-227-2862 Manufacturer Reason for Recall There is the potential for the compounding system to not immediately interrupt compounding and alert user to the presence of an air bubble exceeding 3% of the ordered volume for an ingredient. This could result in a a compounded drug that does not meet the specified accuracy range for an individual ingredient. FDA Determined Cause 2 Software change control Action On Nov. 23, 2020, the firm issued a "Urgent Medical Device Correction" letter to affected consignees via USPS mail. In addition to informing consignees about the recall device, the firm requested consignees take the following actions:************1.***Review the URGENT MEDICAL DEVICE CORRECTION Notification in its entirety and ensure that all users in your organization and other concerned persons are informed about this voluntary Medical Device Correction and the affected product(s). ************2.***Ensure workflow(s) are in place to review air bubbles and scrap any final containers which exceed your facility***s guidelines until your APEX Compounding System has been upgraded to Software Version 1.0.5.58 or higher. ******"***When compounding is complete, if bubbles were detected, APEX activates the Bubble Events button on the Compounding Complete screen allowing you to review the bubbles detected during compounding. ******"***Until such time as your APEX software has been updated, BBMI recommends having qualified pharmacy personnel review this screen prior to removing the final container to ensure that the size of any air bubbles are within an allowable range per your facility***s guidelines. ******"***Alternatively, your facility can have qualified personnel review each Compounding Activity Report generated by the device which also contains air bubble information. ******"***Instructions for viewing the Bubble Events screen are provided in Attachment 02 of this Notice. Instructions for printing and reviewing the Compounding Activity Report are provided in Attachment 03 of this Notice.******"***Any final containers which exceed your facility***s allowable range should be scrapped. ************3.***Return the completed ***Acknowledgement *** Medical Device Correction*** form to BBMI by utilizing the enclosed self-addressed envelope or via scanned e-mail to BBMUS_Carrollton_QualitySystemManagement@bbraunusa.com within two (2) weeks of receipt, even if the total inventory in your possession is zero (0). It is important this form is Quantity in Commerce 248 units Distribution US Nationwide distribution including in the states of AL, AZ, CA, CO, CT, FL, GA, IA, IL, LA, MA, MD, MI, MN, MO, NC, NH, NJ, NV, NY, OH, OK, PA, TN,TX, WA. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On Nov. 23, 2020, the firm issued a "Urgent Medical Device Correction" letter to affected consignees via USPS mail. In addition to informing consignees about the recall device, the firm requested consignees take the following actions:************1.***Review the URGENT MEDICAL DEVICE CORRECTION Notification in its entirety and ensure that all users in your organization and other concerned persons are informed about this voluntary Medical Device Correction and the affected product(s). ************2.***Ensure workflow(s) are in place to review air bubbles and scrap any final containers which exceed your facility***s guidelines until your APEX Compounding System has been upgraded to Software Version 1.0.5.58 or higher. ******"***When compounding is complete, if bubbles were detected, APEX activates the Bubble Events button on the Compounding Complete screen allowing you to review the bubbles detected during compounding. ******"***Until such time as your APEX software has been updated, BBMI recommends having qualified pharmacy personnel review this screen prior to removing the final container to ensure that the size of any air bubbles are within an allowable range per your facility***s guidelines. ******"***Alternatively, your facility can have qualified personnel review each Compounding Activity Report generated by the device which also contains air bubble information. ******"***Instructions for viewing the Bubble Events screen are provided in Attachment 02 of this Notice. Instructions for printing and reviewing the Compounding Activity Report are provided in Attachment 03 of this Notice.******"***Any final containers which exceed your facility***s allowable range should be scrapped. ************3.***Return the completed ***Acknowledgement *** Medical Device Correction*** form to BBMI by utilizing the enclosed self-addressed envelope or via scanned e-mail to BBMUS_Carrollton_QualitySystemManagement@bbraunusa.com within two (2) weeks of receipt, even if the total inventory in your possession is zero (0). It is important this form is | SoftwareChangeControl
36 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184574 | Z-0719-2021 | APEX Compounding System Control Panel Module, REF: 601238, Product Code AX1000 - Product Usage: intended for intravascular administration to a patient. | Apex Compounding System | 2 | CDRH | 2020-12-23 |  | 2023-03-14 14:38 | B Braun Medical, Inc. | There is the potential for the compounding system to not immediately interrupt compounding and alert user to the presence of an air bubble exceeding 3% of the ordered volume for an ingredient. This could result in a a compounded drug that does not meet the specified accuracy range for an individual ingredient. | Class 2 Device Recall Apex Compounding System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Apex Compounding System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Apex Compounding System Date Initiated by Firm November 23, 2020 Create Date December 23, 2020 Recall Status 1 Terminated 3 on March 14, 2023 Recall Number Z-0719-2021 Recall Event ID 86803 Product Classification System/device, pharmacy compounding - Product Code NEP Product APEX Compounding System Control Panel Module, REF: 601238, Product Code AX1000 - Product Usage: intended for intravascular administration to a patient. Code Information All APEX Compounding System with Software Version 1.0.5.57 or earlier Recalling Firm/ Manufacturer B Braun Medical, Inc. 1601 Wallace Dr Ste 150 Carrollton TX 75006-6690 For Additional Information Contact B. Braun Customer Support 800-227-2862 Manufacturer Reason for Recall There is the potential for the compounding system to not immediately interrupt compounding and alert user to the presence of an air bubble exceeding 3% of the ordered volume for an ingredient. This could result in a a compounded drug that does not meet the specified accuracy range for an individual ingredient. FDA Determined Cause 2 Software change control Action On Nov. 23, 2020, the firm issued a "Urgent Medical Device Correction" letter to affected consignees via USPS mail. In addition to informing consignees about the recall device, the firm requested consignees take the following actions:************1.***Review the URGENT MEDICAL DEVICE CORRECTION Notification in its entirety and ensure that all users in your organization and other concerned persons are informed about this voluntary Medical Device Correction and the affected product(s). ************2.***Ensure workflow(s) are in place to review air bubbles and scrap any final containers which exceed your facility***s guidelines until your APEX Compounding System has been upgraded to Software Version 1.0.5.58 or higher. ******"***When compounding is complete, if bubbles were detected, APEX activates the Bubble Events button on the Compounding Complete screen allowing you to review the bubbles detected during compounding. ******"***Until such time as your APEX software has been updated, BBMI recommends having qualified pharmacy personnel review this screen prior to removing the final container to ensure that the size of any air bubbles are within an allowable range per your facility***s guidelines. ******"***Alternatively, your facility can have qualified personnel review each Compounding Activity Report generated by the device which also contains air bubble information. ******"***Instructions for viewing the Bubble Events screen are provided in Attachment 02 of this Notice. Instructions for printing and reviewing the Compounding Activity Report are provided in Attachment 03 of this Notice.******"***Any final containers which exceed your facility***s allowable range should be scrapped. ************3.***Return the completed ***Acknowledgement *** Medical Device Correction*** form to BBMI by utilizing the enclosed self-addressed envelope or via scanned e-mail to BBMUS_Carrollton_QualitySystemManagement@bbraunusa.com within two (2) weeks of receipt, even if the total inventory in your possession is zero (0). It is important this form is Quantity in Commerce 16 units Distribution US Nationwide distribution including in the states of AL, AZ, CA, CO, CT, FL, GA, IA, IL, LA, MA, MD, MI, MN, MO, NC, NH, NJ, NV, NY, OH, OK, PA, TN,TX, WA. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On Nov. 23, 2020, the firm issued a "Urgent Medical Device Correction" letter to affected consignees via USPS mail. In addition to informing consignees about the recall device, the firm requested consignees take the following actions:************1.***Review the URGENT MEDICAL DEVICE CORRECTION Notification in its entirety and ensure that all users in your organization and other concerned persons are informed about this voluntary Medical Device Correction and the affected product(s). ************2.***Ensure workflow(s) are in place to review air bubbles and scrap any final containers which exceed your facility***s guidelines until your APEX Compounding System has been upgraded to Software Version 1.0.5.58 or higher. ******"***When compounding is complete, if bubbles were detected, APEX activates the Bubble Events button on the Compounding Complete screen allowing you to review the bubbles detected during compounding. ******"***Until such time as your APEX software has been updated, BBMI recommends having qualified pharmacy personnel review this screen prior to removing the final container to ensure that the size of any air bubbles are within an allowable range per your facility***s guidelines. ******"***Alternatively, your facility can have qualified personnel review each Compounding Activity Report generated by the device which also contains air bubble information. ******"***Instructions for viewing the Bubble Events screen are provided in Attachment 02 of this Notice. Instructions for printing and reviewing the Compounding Activity Report are provided in Attachment 03 of this Notice.******"***Any final containers which exceed your facility***s allowable range should be scrapped. ************3.***Return the completed ***Acknowledgement *** Medical Device Correction*** form to BBMI by utilizing the enclosed self-addressed envelope or via scanned e-mail to BBMUS_Carrollton_QualitySystemManagement@bbraunusa.com within two (2) weeks of receipt, even if the total inventory in your possession is zero (0). It is important this form is | SoftwareChangeControl
37 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184573 | Z-0718-2021 | APEX Compounding System Control Panel Module, REF: 601237, Product Code AX1000 - Product Usage: intended for intravascular administration to a patient. | Apex Compounding System | 2 | CDRH | 2020-12-23 |  | 2023-03-14 14:38 | B Braun Medical, Inc. | There is the potential for the compounding system to not immediately interrupt compounding and alert user to the presence of an air bubble exceeding 3% of the ordered volume for an ingredient. This could result in a a compounded drug that does not meet the specified accuracy range for an individual ingredient. | Class 2 Device Recall Apex Compounding System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Apex Compounding System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Apex Compounding System Date Initiated by Firm November 23, 2020 Create Date December 23, 2020 Recall Status 1 Terminated 3 on March 14, 2023 Recall Number Z-0718-2021 Recall Event ID 86803 Product Classification System/device, pharmacy compounding - Product Code NEP Product APEX Compounding System Control Panel Module, REF: 601237, Product Code AX1000 - Product Usage: intended for intravascular administration to a patient. Code Information All APEX Compounding System with Software Version 1.0.5.57 or earlier Recalling Firm/ Manufacturer B Braun Medical, Inc. 1601 Wallace Dr Ste 150 Carrollton TX 75006-6690 For Additional Information Contact B. Braun Customer Support 800-227-2862 Manufacturer Reason for Recall There is the potential for the compounding system to not immediately interrupt compounding and alert user to the presence of an air bubble exceeding 3% of the ordered volume for an ingredient. This could result in a a compounded drug that does not meet the specified accuracy range for an individual ingredient. FDA Determined Cause 2 Software change control Action On Nov. 23, 2020, the firm issued a "Urgent Medical Device Correction" letter to affected consignees via USPS mail. In addition to informing consignees about the recall device, the firm requested consignees take the following actions:************1.***Review the URGENT MEDICAL DEVICE CORRECTION Notification in its entirety and ensure that all users in your organization and other concerned persons are informed about this voluntary Medical Device Correction and the affected product(s). ************2.***Ensure workflow(s) are in place to review air bubbles and scrap any final containers which exceed your facility***s guidelines until your APEX Compounding System has been upgraded to Software Version 1.0.5.58 or higher. ******"***When compounding is complete, if bubbles were detected, APEX activates the Bubble Events button on the Compounding Complete screen allowing you to review the bubbles detected during compounding. ******"***Until such time as your APEX software has been updated, BBMI recommends having qualified pharmacy personnel review this screen prior to removing the final container to ensure that the size of any air bubbles are within an allowable range per your facility***s guidelines. ******"***Alternatively, your facility can have qualified personnel review each Compounding Activity Report generated by the device which also contains air bubble information. ******"***Instructions for viewing the Bubble Events screen are provided in Attachment 02 of this Notice. Instructions for printing and reviewing the Compounding Activity Report are provided in Attachment 03 of this Notice.******"***Any final containers which exceed your facility***s allowable range should be scrapped. ************3.***Return the completed ***Acknowledgement *** Medical Device Correction*** form to BBMI by utilizing the enclosed self-addressed envelope or via scanned e-mail to BBMUS_Carrollton_QualitySystemManagement@bbraunusa.com within two (2) weeks of receipt, even if the total inventory in your possession is zero (0). It is important this form is Quantity in Commerce 85 units Distribution US Nationwide distribution including in the states of AL, AZ, CA, CO, CT, FL, GA, IA, IL, LA, MA, MD, MI, MN, MO, NC, NH, NJ, NV, NY, OH, OK, PA, TN,TX, WA. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On Nov. 23, 2020, the firm issued a "Urgent Medical Device Correction" letter to affected consignees via USPS mail. In addition to informing consignees about the recall device, the firm requested consignees take the following actions:************1.***Review the URGENT MEDICAL DEVICE CORRECTION Notification in its entirety and ensure that all users in your organization and other concerned persons are informed about this voluntary Medical Device Correction and the affected product(s). ************2.***Ensure workflow(s) are in place to review air bubbles and scrap any final containers which exceed your facility***s guidelines until your APEX Compounding System has been upgraded to Software Version 1.0.5.58 or higher. ******"***When compounding is complete, if bubbles were detected, APEX activates the Bubble Events button on the Compounding Complete screen allowing you to review the bubbles detected during compounding. ******"***Until such time as your APEX software has been updated, BBMI recommends having qualified pharmacy personnel review this screen prior to removing the final container to ensure that the size of any air bubbles are within an allowable range per your facility***s guidelines. ******"***Alternatively, your facility can have qualified personnel review each Compounding Activity Report generated by the device which also contains air bubble information. ******"***Instructions for viewing the Bubble Events screen are provided in Attachment 02 of this Notice. Instructions for printing and reviewing the Compounding Activity Report are provided in Attachment 03 of this Notice.******"***Any final containers which exceed your facility***s allowable range should be scrapped. ************3.***Return the completed ***Acknowledgement *** Medical Device Correction*** form to BBMI by utilizing the enclosed self-addressed envelope or via scanned e-mail to BBMUS_Carrollton_QualitySystemManagement@bbraunusa.com within two (2) weeks of receipt, even if the total inventory in your possession is zero (0). It is important this form is | SoftwareChangeControl
38 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=183827 | Z-0282-2021 | Synapse PACS Software Version 5.6.1 - Product Usage: intended for use as a web based application on an off-the shelf PC which meets or exceeds minimum specifications and is networked with a FUJIFILM Synapse PACS server. | Synapse PACS | 2 | CDRH | 2020-10-22 |  | 2021-11-17 15:26 | Fujifilm Medical Systems U.S.A., Inc. | FUJIFILM has become aware of the possibility that certain CT studies may report a much higher 3D Sphere MAX value than expected. This issue is specific to 3D Sphere - other density tools are not affected.****** | Class 2 Device Recall Synapse PACS Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Synapse PACS FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Synapse PACS Date Initiated by Firm September 11, 2020 Create Date October 22, 2020 Recall Status 1 Terminated 3 on November 17, 2021 Recall Number Z-0282-2021 Recall Event ID 86494 510(K)Number K190232 Product Classification System, image processing, radiological - Product Code LLZ Product Synapse PACS Software Version 5.6.1 - Product Usage: intended for use as a web based application on an off-the shelf PC which meets or exceeds minimum specifications and is networked with a FUJIFILM Synapse PACS server. Code Information Synapse Software Version 5.6.1 Recalling Firm/ Manufacturer Fujifilm Medical Systems U.S.A., Inc. 81 Hartwell Ave Ste 300 Lexington MA 02421-3160 For Additional Information Contact Jeffrey Wan 201-675-8947 Manufacturer Reason for Recall FUJIFILM has become aware of the possibility that certain CT studies may report a much higher 3D Sphere MAX value than expected. This issue is specific to 3D Sphere - other density tools are not affected. FDA Determined Cause 2 Software change control Action Letter was sent to all affected customers on 09/11/2020 stating the following: FUJIFILM will offer an upgrade to Synapse PACS 7.1.100 to resolve this issue. Synapse PACS 7.1.100 is estimated to be available by the end of September. ************Until the patch is available, users can continue to safely use Synapse PACS by following these additional instructions:************1. Ensure that the 3D Sphere MAX value is within the sphere***s range. If the 3D Sphere MAX value is observed to be much higher than expected, the value should not be used when rendering a clinical decision until the issue is corrected. ************Customers should complete and return the Field Action Verification Form attached to the recall letter. Quantity in Commerce 1 Distribution US Nationwide distribution including in the state of Minnesota. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Letter was sent to all affected customers on 09/11/2020 stating the following: FUJIFILM will offer an upgrade to Synapse PACS 7.1.100 to resolve this issue. Synapse PACS 7.1.100 is estimated to be available by the end of September. ************Until the patch is available, users can continue to safely use Synapse PACS by following these additional instructions:************1. Ensure that the 3D Sphere MAX value is within the sphere***s range. If the 3D Sphere MAX value is observed to be much higher than expected, the value should not be used when rendering a clinical decision until the issue is corrected. ************Customers should complete and return the Field Action Verification Form attached to the recall letter. | SoftwareChangeControl
39 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=181876 | Z-2551-2020 | MagNA Pure 96 Instrument, clinical sample concentrator, Material number 06541089001 - Product Usage: This protocol is intended for pipetting samples into a MagNA Pure 96 Processing Cartridge (purification rack position 1) before a purification run. | MagNA Pure 96 | 2 | CDRH | 2020-07-10 |  | 2023-10-04 09:25 | Roche Molecular Systems, Inc. | When using Sample Transfer protocol version 3.0, the drop catcher is not activated on the MagNA Pure 96 instrument, which introduces a potential risk for cross-contamination leading to erroneous results. | Class 2 Device Recall MagNA Pure 96 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall MagNA Pure 96 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall MagNA Pure 96 Date Initiated by Firm June 03, 2020 Create Date July 10, 2020 Recall Status 1 Terminated 3 on October 04, 2023 Recall Number Z-2551-2020 Recall Event ID 85838 Product Classification Clinical sample concentrator - Product Code JJH Product MagNA Pure 96 Instrument, clinical sample concentrator, Material number 06541089001 - Product Usage: This protocol is intended for pipetting samples into a MagNA Pure 96 Processing Cartridge (purification rack position 1) before a purification run. Code Information All units when used with Sample Transfer protocol version 3.0 Recalling Firm/ Manufacturer Roche Molecular Systems, Inc. 1080 US Highway 202 S Branchburg NJ 08876-3733 For Additional Information Contact Roche Support Network Customer Support Center 800-526-1247 Manufacturer Reason for Recall When using Sample Transfer protocol version 3.0, the drop catcher is not activated on the MagNA Pure 96 instrument, which introduces a potential risk for cross-contamination leading to erroneous results. FDA Determined Cause 2 Software change control Action On June 5, 2020, the firm distributed Urgent Medical Device Correction letters to consignee accounts. The UMDC instructs the customers to determine whether or not their laboratory is using the affected Sample Transfer protocol version 3.0 with******their MagNA Pure 96 instrument. If the affected protocol is found on the instrument***s Control Unit, it needs to be immediately deleted by following the instructions provided in the UMDC.******A newly updated Sample Transfer protocol version 4.0 is available. Roche Affiliate organizations will provide instructions for obtaining the new version of the protocol.************Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-526-1247 if you have questions about this recall. Quantity in Commerce 917 Distribution Worldwide distribution - US Nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On June 5, 2020, the firm distributed Urgent Medical Device Correction letters to consignee accounts. The UMDC instructs the customers to determine whether or not their laboratory is using the affected Sample Transfer protocol version 3.0 with******their MagNA Pure 96 instrument. If the affected protocol is found on the instrument***s Control Unit, it needs to be immediately deleted by following the instructions provided in the UMDC.******A newly updated Sample Transfer protocol version 4.0 is available. Roche Affiliate organizations will provide instructions for obtaining the new version of the protocol.************Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-526-1247 if you have questions about this recall. | SoftwareChangeControl
40 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180264 | Z-1631-2020 | Pharos Excimer Laser, Model: EX-308 - Product Usage: The Pharos EX-308 Excimer Laser System is designed to treat the skin disorders psoriasis, vitiligo, atopic dermatitis, and leukoderma using phototherapy. The excimer laser system produces collimated 308 nm radiation, within the invisible region of the electromagnetic spectrum. The UV radiation falls within the range of ultraviolet B (UVB), which is used to treat photo-responsive skin disorders. | Pharos Ex-308 Excimer Laser System | 2 | CDRH | 2020-04-01 | 2020-04-01 21:00 | 2021-05-05 15:37 | Ra Medical Systems, Inc. | At doses of less than 100 mJ, in custom and auto repeat modes, a software error in the excimer laser phototherapy system causes the device to fail to cease firing as long as the foot pedal or handpiece switch is depressed. This malfunction can cause overtreatment, which could result in strong erythema, and possibly cause blisters. | Class 2 Device Recall Pharos Ex308 Excimer Laser System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Pharos Ex308 Excimer Laser System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Pharos Ex308 Excimer Laser System Date Initiated by Firm August 23, 2019 Date Posted April 01, 2020 Recall Status 1 Terminated 3 on May 05, 2021 Recall Number Z-1631-2020 Recall Event ID 85200 510(K)Number K062963 Product Classification Powered laser surgical instrument - Product Code GEX Product Pharos Excimer Laser, Model: EX-308 - Product Usage: The Pharos EX-308 Excimer Laser System is designed to treat the skin disorders psoriasis, vitiligo, atopic dermatitis, and leukoderma using phototherapy. The excimer laser system produces collimated 308 nm radiation, within the invisible region of the electromagnetic spectrum. The UV radiation falls within the range of ultraviolet B (UVB), which is used to treat photo-responsive skin disorders. Code Information All serial numbers. Recalling Firm/ Manufacturer Ra Medical Systems, Inc. 2070 Las Palmas Dr Carlsbad CA 92011-1518 For Additional Information Contact Andrew Jackson 877-635-1800 Manufacturer Reason for Recall At doses of less than 100 mJ, in custom and auto repeat modes, a software error in the excimer laser phototherapy system causes the device to fail to cease firing as long as the foot pedal or handpiece switch is depressed. This malfunction can cause overtreatment, which could result in strong erythema, and possibly cause blisters. FDA Determined Cause 2 Software change control Action On 08/23/2019, Medical Device Correction was mailed to customers. Customers were informed that the recalling firm's service department would contact them to replace the software. Unitl the software is replaced, do not use doses less than 100 mJ. During service visits customers were informed that software was updated to 1.06 ZRPL Quantity in Commerce 10 Distribution U.S. Nationwide distribution in the states of CA, FL, OR, MS, NY. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = GEX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 08/23/2019, Medical Device Correction was mailed to customers. Customers were informed that the recalling firm's service department would contact them to replace the software. Unitl the software is replaced, do not use doses less than 100 mJ. During service visits customers were informed that software was updated to 1.06 ZRPL | SoftwareChangeControl
41 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179871 | Z-1476-2020 | QIAGEN Rotor Gene Q Software version 2.3.4. Real-time PCR analyzer designed for rapid thermalcycling and real-time detection of PCR assays. | QIAGEN | 2 | CDRH | 2020-03-12 |  | 2024-09-26 07:32 | Qiagen Sciences LLC | When using the Rotor-Gene Q with Software version 2.3.4 completing a LIMs export, the .csv file reports the***calculated concentration result value as a logarithmic value, and could lead to a false negative result, which could lead to serious medical consequences such as the suspension or non-initiation of treatment | Class 2 Device Recall QIAGEN Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall QIAGEN FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall QIAGEN Date Initiated by Firm January 28, 2020 Create Date March 12, 2020 Recall Status 1 Terminated 3 on September 26, 2024 Recall Number Z-1476-2020 Recall Event ID 85004 510(K)Number K113319 Product Classification real time Nucleic acid amplification system - Product Code OOI Product QIAGEN Rotor Gene Q Software version 2.3.4. Real-time PCR analyzer designed for rapid thermalcycling and real-time detection of PCR assays. Code Information Software version 2.3.4 used in the following: Product No. / Description 9001680 Rotor-Gene Q 2Plex no PC 9001680R Rotor-Gene Q 2Plex no PC refurbished 9001690 Rotor-Gene Q 2Plex HRM no PC 9001690R Rotor-Gene Q 2Plex HRM no PC refurbished 9001700 Rotor-Gene Q 5Plex no PC 9001700R Rotor-Gene Q 5Plex no PC refurbished 9001710 Rotor-Gene Q 5Plex HRM no PC 9001710R Rotor-Gene Q 5Plex HRM no PC refurbished 9001720 Rotor-Gene Q 6Plex no PC 9001720R Rotor-Gene Q 6Plex no PC refurbished 9001960 Rotor-Gene AusDx (Outback Red) 9002000 Rotor-Gene Q MDx 2Plex 9002001 Rotor-Gene Q MDx 2Plex no PC 9002010 Rotor-Gene Q MDx 2Plex HRM 9002011 #Rotor-Gene Q MDx 2Plex HRM no PC 9002020 Rotor-Gene Q MDx 5Plex 9002021 Rotor-Gene Q MDx 5Plex no PC 9002029 Rotor-Gene Q MDx 5plex, PrioPLUS 9002030 Rotor-Gene Q MDx 5Plex HRM 9002030R Rotor-Gene Q MDx 5Plex HRM refurbished 9002031 Rotor-Gene Q MDx 5Plex HRM no PC 9002040 Rotor-Gene Q MDx 6Plex 9002041 Rotor-Gene Q MDx 6Plex no PC90275010 9002204 Rotor-Gene Q DBV 9023241 Rotor-Gene Q Software 9002370 5 plex MDx HRM with laptop (canada) This issue can affect the following QIAGEN IVD kits: Catalogue number Product Name 4514363 artus BK Virus QS-RGQ Kit (24) 4514263 artus BK Virus RG PCR Kit (24) 4514265 artus BK Virus RG PCR Kit (96) 4503363 artus CMV QS-RGQ Kit (24) 4503263 artus CMV RG PCR Kit (24) 4503265 artus CMV RG PCR Kit (96) 4501363 artus EBV QS-RGQ Kit (24) 4501263 artus EBV RG PCR Kit (24) 4501265 artus EBV RG PCR Kit (96) 4506366 artus HBV QS-RGQ (72) 4506356 artus HBV QS-RGQ Kit (24) 4506363 artus HBV QS-RGQ Kit (24) 4506263 artus HBV RG PCR Kit (24) 4506265 artus HBV RG PCR Kit (96) 4518356 artus HCV QS-RGQ Kit (24) 4518363 artus HCV QS-RGQ Kit (24) 4518366 artus HCV QS-RGQ Kit (72) 4538366 artus HCV QS-RGQ Kit V2 4518253 artus HCV RG RT-PCR Kit (24) 4518263 artus HCV RG RT-PCR Kit (24) 4518265 artus HCV RG RT-PCR Kit (96) 4513363 artus HI Virus-1 QS-RGQ Kit (24) 4513366 artus HI Virus-1 QS-RGQ Kit (72) 4513253 artus HI Virus-1 RG RT-PCR Kit (24) 4513263 artus HI Virus-1 RG RT-PCR Kit (24) 4513265 artus HI Virus-1 RG RT-PCR Kit (96) 4502363 artus VZV QS-RGQ Kit (24) 4502263 artus VZV RG PCR Kit (24) 4502265 artus VZV RG PCR Kit (96) 4532265 artus JCV RG PCR Kit (96) 4555263 artus M.tuberculosis RG PCR Kit (24) 4555265 artus M.tuberculosis RG PCR Kit (96) 4504263 artus Parvo B19 RG PCR Kit (24) 4511263 artus SARS RG RT-PCR Kit (24) Recalling Firm/ Manufacturer Qiagen Sciences LLC 19300 Germantown Rd Germantown MD 20874-1415 For Additional Information Contact SAME 240-686-7500 Manufacturer Reason for Recall When using the Rotor-Gene Q with Software version 2.3.4 completing a LIMs export, the .csv file reports the******calculated concentration result value as a logarithmic value, and could lead to a false negative result, which could lead to serious medical consequences such as the suspension or non-initiation of treatment FDA Determined Cause 2 Software change control Action QIAGEN issued Urgent Medical Device Correction via email on 1/28/20 advising reason for recall, health risk and action to take: 1) Discontinue use of the Rotor-Gene Q instrument in the following specific conditions:****** The Rotor-Gene Q instrument is running Software 2.3.4 using the export function to LIMs******2) Upgrade to the latest version of the RGQ software which resolves the issue *** version 2.3.5. This is expected to be available on 1st February 2020. Please visit the QIAGEN website to upgrade: www.qiagen.com/resources/resourcedetail?id=9d8bda8e-1fd7-4519-a1ff-b60bba526b57⟨=en. Apart from correcting this fault, no other changes to the software have been implemented.************Laboratory personnel and clinicians are advised to consider a patient***s previous test results, other diagnostic tests, anamnesis and current clinical condition when interpreting results from this software. If the results do not match the patient***s clinical presentation, or incongruences are found with previous and concomitant tests or the results are otherwise unexpected, the patient sample should be retested using an alternate test method or a reference laboratory.************3) Complete the Acknowledgement of Receipt form and return it to QIAGEN.******If you have further questions, please contact QIAGEN Technical Services Department on 800-362-7737 (US & Canada)******" QIAGEN Subsidiaries: https://www.qiagen.com/about-us/contact/global-contacts/subsidiaries/******" QIAGEN Commercial Partners and Importers: https://www.qiagen.com/about-us/contact/global-contacts/distributors-and-importers/ Quantity in Commerce 11418 instruments Distribution Worldwide distribution - US Nationwide distribution and countries Afghanistan, Algeria, Argentina, Armenia, Australia, Austria, Azerbaijan, Bangladesh, Belarus, Belgium, Benin, Bhutan, Bosnia, Brazil, ******Brunei Darussalam, Bulgaria, Cambodia, Cameroon, Canada, Chile, China, Colombia, Congo, Costa Rica, ******Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Georgia, Germany,******Greece, Hungary, India, Indonesia, Iran, Islamic Republic of, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic of Kuwait, Laos, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, ******Luxembourg, Macao, Macedonia, the former Yugoslav Republic of, Malaysia, Maldives, Malta, Mauritania, Mauritius, Mayotte, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Nigeria, ******Norway, Oman, Pakistan, Palestinian Territory Occupied, Panama, Papua New Guinea, Paraguay, Peru, ******Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syrian Arab, Republic of Taiwan, Tajikistan, Thailand, Timor-Leste (else known as East Timor), Tunisia, Turkey, ******Turkmenistan, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, Bolivarian Republic of, Viet Nam, Yemen. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = OOI - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action QIAGEN issued Urgent Medical Device Correction via email on 1/28/20 advising reason for recall, health risk and action to take: 1) Discontinue use of the Rotor-Gene Q instrument in the following specific conditions:****** The Rotor-Gene Q instrument is running Software 2.3.4 using the export function to LIMs******2) Upgrade to the latest version of the RGQ software which resolves the issue *** version 2.3.5. This is expected to be available on 1st February 2020. Please visit the QIAGEN website to upgrade: www.qiagen.com/resources/resourcedetail?id=9d8bda8e-1fd7-4519-a1ff-b60bba526b57⟨=en. Apart from correcting this fault, no other changes to the software have been implemented.************Laboratory personnel and clinicians are advised to consider a patient***s previous test results, other diagnostic tests, anamnesis and current clinical condition when interpreting results from this software. If the results do not match the patient***s clinical presentation, or incongruences are found with previous and concomitant tests or the results are otherwise unexpected, the patient sample should be retested using an alternate test method or a reference laboratory.************3) Complete the Acknowledgement of Receipt form and return it to QIAGEN.******If you have further questions, please contact QIAGEN Technical Services Department on 800-362-7737 (US & Canada)******" QIAGEN Subsidiaries: https://www.qiagen.com/about-us/contact/global-contacts/subsidiaries/******" QIAGEN Commercial Partners and Importers: https://www.qiagen.com/about-us/contact/global-contacts/distributors-and-importers/ | SoftwareChangeControl
42 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178928 | Z-1124-2020 | Prometra Programmable 20mL Pump. Catalog No. 91827 The device is an implantable infusion pump. GTIN-14 00810335020XXX (XXX = in-country language specific) | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system. | Class 2 Device Recall Prometra Programmable Infusion Pump Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prometra Programmable Infusion Pump FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prometra Programmable Infusion Pump Date Initiated by Firm December 24, 2019 Date Posted February 10, 2020 Recall Status 1 Open 3 , Classified Recall Number Z-1124-2020 Recall Event ID 84653 PMA Number P080012 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product Prometra Programmable 20mL Pump. Catalog No. 91827 The device is an implantable infusion pump. GTIN-14 00810335020XXX (XXX = in-country language specific) Code Information All units with pump firmware version 0.26 manufactured from 2012 until the present time. Recalling Firm/ Manufacturer Flowonix Medical Inc 120 Forbes Blvd Ste 170 Mansfield MA 02048-1150 For Additional Information Contact Technical Solutions Department 855-356-9665 Manufacturer Reason for Recall A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system. FDA Determined Cause 2 Software change control Action ***Updated 5/12/2023******On May 11, 2023, the firm sent an URGENT MEDICAL DEVICE CORRECTION UPDATE letter to all affected customers, which informed them of an update to Flowonix's plan to correct the firmware issue.******Flowonix will NOT be conducting an office-based pump firmware update for affected pumps implanted prior to January 2020.******Healthcare providers were provided ongoing recommendations for Prometra and Prometra II pumps implanted prior to January 2020, which are impacted by this issue.******Healthcare providers should contact their Flowonix Representative or our Technical Solutions Department (1-855-356-9665) in the event cessation of therapy or premature low battery is suspected and/or have general questions regarding this notice.***************On December 24, 2019, the firm notified affected users via an Urgent Medical Device Corrective Action Recall Notice.******The firm advised customers of the product issue and reminded customers that the pumps are currently labeled only for use with Infumorph. Due to the low risk that this failure mode may occur, Flowonix is not recommending that pumps be explanted. Healthcare providers were recommended to inform patients regarding the symptoms and adverse health events associated with an unexpected pump stoppage and the appropriate steps to take in the event they suspect a stoppage.******The firm will be implementing a correction of firmware errors to bring the pumps back into specification. The firm will follow up when the change has been completed. An office-based pump firmware update is being developed for patients currently implanted with the affected pumps. (Note, as of the 5/12/23 update, implanted pumps will not be receiving a firmware update.) Quantity in Commerce 1083 (OUS) Distribution Nationwide domestic distribution. Foreign distribution to United Kingdom, Republic of Ireland, Poland, Greece, Italy, Germany, Netherlands, and Belgium. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = LKK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action ***Updated 5/12/2023******On May 11, 2023, the firm sent an URGENT MEDICAL DEVICE CORRECTION UPDATE letter to all affected customers, which informed them of an update to Flowonix's plan to correct the firmware issue.******Flowonix will NOT be conducting an office-based pump firmware update for affected pumps implanted prior to January 2020.******Healthcare providers were provided ongoing recommendations for Prometra and Prometra II pumps implanted prior to January 2020, which are impacted by this issue.******Healthcare providers should contact their Flowonix Representative or our Technical Solutions Department (1-855-356-9665) in the event cessation of therapy or premature low battery is suspected and/or have general questions regarding this notice.***************On December 24, 2019, the firm notified affected users via an Urgent Medical Device Corrective Action Recall Notice.******The firm advised customers of the product issue and reminded customers that the pumps are currently labeled only for use with Infumorph. Due to the low risk that this failure mode may occur, Flowonix is not recommending that pumps be explanted. Healthcare providers were recommended to inform patients regarding the symptoms and adverse health events associated with an unexpected pump stoppage and the appropriate steps to take in the event they suspect a stoppage.******The firm will be implementing a correction of firmware errors to bring the pumps back into specification. The firm will follow up when the change has been completed. An office-based pump firmware update is being developed for patients currently implanted with the affected pumps. (Note, as of the 5/12/23 update, implanted pumps will not be receiving a firmware update.) | SoftwareChangeControl
43 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178929 | Z-1125-2020 | Prometra II Programmable 20mL Pump. Catalog No. 93827 The device is an implantable infusion pump. GTIN-14 00810335020YYY (YYY= in-country language specific) | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system. | Class 2 Device Recall Prometra Programmable Infusion Pump Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prometra Programmable Infusion Pump FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prometra Programmable Infusion Pump Date Initiated by Firm December 24, 2019 Date Posted February 10, 2020 Recall Status 1 Open 3 , Classified Recall Number Z-1125-2020 Recall Event ID 84653 PMA Number P080012 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product Prometra II Programmable 20mL Pump. Catalog No. 93827 The device is an implantable infusion pump. GTIN-14 00810335020YYY (YYY= in-country language specific) Code Information All units with pump firmware version 0.26 manufactured from 2012 until the present time. Recalling Firm/ Manufacturer Flowonix Medical Inc 120 Forbes Blvd Ste 170 Mansfield MA 02048-1150 For Additional Information Contact Technical Solutions Department 855-356-9665 Manufacturer Reason for Recall A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system. FDA Determined Cause 2 Software change control Action ***Updated 5/12/2023******On May 11, 2023, the firm sent an URGENT MEDICAL DEVICE CORRECTION UPDATE letter to all affected customers, which informed them of an update to Flowonix's plan to correct the firmware issue.******Flowonix will NOT be conducting an office-based pump firmware update for affected pumps implanted prior to January 2020.******Healthcare providers were provided ongoing recommendations for Prometra and Prometra II pumps implanted prior to January 2020, which are impacted by this issue.******Healthcare providers should contact their Flowonix Representative or our Technical Solutions Department (1-855-356-9665) in the event cessation of therapy or premature low battery is suspected and/or have general questions regarding this notice.***************On December 24, 2019, the firm notified affected users via an Urgent Medical Device Corrective Action Recall Notice.******The firm advised customers of the product issue and reminded customers that the pumps are currently labeled only for use with Infumorph. Due to the low risk that this failure mode may occur, Flowonix is not recommending that pumps be explanted. Healthcare providers were recommended to inform patients regarding the symptoms and adverse health events associated with an unexpected pump stoppage and the appropriate steps to take in the event they suspect a stoppage.******The firm will be implementing a correction of firmware errors to bring the pumps back into specification. The firm will follow up when the change has been completed. An office-based pump firmware update is being developed for patients currently implanted with the affected pumps. (Note, as of the 5/12/23 update, implanted pumps will not be receiving a firmware update.) Quantity in Commerce 1083 (OUS) Distribution Nationwide domestic distribution. Foreign distribution to United Kingdom, Republic of Ireland, Poland, Greece, Italy, Germany, Netherlands, and Belgium. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = LKK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action ***Updated 5/12/2023******On May 11, 2023, the firm sent an URGENT MEDICAL DEVICE CORRECTION UPDATE letter to all affected customers, which informed them of an update to Flowonix's plan to correct the firmware issue.******Flowonix will NOT be conducting an office-based pump firmware update for affected pumps implanted prior to January 2020.******Healthcare providers were provided ongoing recommendations for Prometra and Prometra II pumps implanted prior to January 2020, which are impacted by this issue.******Healthcare providers should contact their Flowonix Representative or our Technical Solutions Department (1-855-356-9665) in the event cessation of therapy or premature low battery is suspected and/or have general questions regarding this notice.***************On December 24, 2019, the firm notified affected users via an Urgent Medical Device Corrective Action Recall Notice.******The firm advised customers of the product issue and reminded customers that the pumps are currently labeled only for use with Infumorph. Due to the low risk that this failure mode may occur, Flowonix is not recommending that pumps be explanted. Healthcare providers were recommended to inform patients regarding the symptoms and adverse health events associated with an unexpected pump stoppage and the appropriate steps to take in the event they suspect a stoppage.******The firm will be implementing a correction of firmware errors to bring the pumps back into specification. The firm will follow up when the change has been completed. An office-based pump firmware update is being developed for patients currently implanted with the affected pumps. (Note, as of the 5/12/23 update, implanted pumps will not be receiving a firmware update.) | SoftwareChangeControl
44 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178926 | Z-1122-2020 | Prometra II 40mL Pump. Catalog No. 16827 The device is an implantable infusion pump. GTIN-14 00810335020242 | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system. | Class 2 Device Recall Prometra Programmable Infusion Pump Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prometra Programmable Infusion Pump FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prometra Programmable Infusion Pump Date Initiated by Firm December 24, 2019 Date Posted February 10, 2020 Recall Status 1 Open 3 , Classified Recall Number Z-1122-2020 Recall Event ID 84653 PMA Number P080012 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product Prometra II 40mL Pump. Catalog No. 16827 The device is an implantable infusion pump. GTIN-14 00810335020242 Code Information All units with pump firmware version 0.26 manufactured from 2012 until the present time. Serial Numbers 10GC4A24X 10GC4A02X 10GC4A18X 10GC4A22X 10GC4A14X 4MA00044X 4MA00006X 10GC4A12X 10GC4A09X 10GC4A06X 10GC4A21X 4MA00008X 10GC4A01X 10GC4A20X 10GC4A11X 10GC4A16X 10GC4A17X 4MA00045X 4MA00055X 10GC4A08X 4MA00032X 4MA00031X 4MA00010X 4MA00042X Recalling Firm/ Manufacturer Flowonix Medical Inc 120 Forbes Blvd Ste 170 Mansfield MA 02048-1150 For Additional Information Contact Technical Solutions Department 855-356-9665 Manufacturer Reason for Recall A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system. FDA Determined Cause 2 Software change control Action ***Updated 5/12/2023******On May 11, 2023, the firm sent an URGENT MEDICAL DEVICE CORRECTION UPDATE letter to all affected customers, which informed them of an update to Flowonix's plan to correct the firmware issue.******Flowonix will NOT be conducting an office-based pump firmware update for affected pumps implanted prior to January 2020.******Healthcare providers were provided ongoing recommendations for Prometra and Prometra II pumps implanted prior to January 2020, which are impacted by this issue.******Healthcare providers should contact their Flowonix Representative or our Technical Solutions Department (1-855-356-9665) in the event cessation of therapy or premature low battery is suspected and/or have general questions regarding this notice.***************On December 24, 2019, the firm notified affected users via an Urgent Medical Device Corrective Action Recall Notice.******The firm advised customers of the product issue and reminded customers that the pumps are currently labeled only for use with Infumorph. Due to the low risk that this failure mode may occur, Flowonix is not recommending that pumps be explanted. Healthcare providers were recommended to inform patients regarding the symptoms and adverse health events associated with an unexpected pump stoppage and the appropriate steps to take in the event they suspect a stoppage.******The firm will be implementing a correction of firmware errors to bring the pumps back into specification. The firm will follow up when the change has been completed. An office-based pump firmware update is being developed for patients currently implanted with the affected pumps. (Note, as of the 5/12/23 update, implanted pumps will not be receiving a firmware update.) Quantity in Commerce 6670 (US) Distribution Nationwide domestic distribution. Foreign distribution to United Kingdom, Republic of Ireland, Poland, Greece, Italy, Germany, Netherlands, and Belgium. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = LKK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action ***Updated 5/12/2023******On May 11, 2023, the firm sent an URGENT MEDICAL DEVICE CORRECTION UPDATE letter to all affected customers, which informed them of an update to Flowonix's plan to correct the firmware issue.******Flowonix will NOT be conducting an office-based pump firmware update for affected pumps implanted prior to January 2020.******Healthcare providers were provided ongoing recommendations for Prometra and Prometra II pumps implanted prior to January 2020, which are impacted by this issue.******Healthcare providers should contact their Flowonix Representative or our Technical Solutions Department (1-855-356-9665) in the event cessation of therapy or premature low battery is suspected and/or have general questions regarding this notice.***************On December 24, 2019, the firm notified affected users via an Urgent Medical Device Corrective Action Recall Notice.******The firm advised customers of the product issue and reminded customers that the pumps are currently labeled only for use with Infumorph. Due to the low risk that this failure mode may occur, Flowonix is not recommending that pumps be explanted. Healthcare providers were recommended to inform patients regarding the symptoms and adverse health events associated with an unexpected pump stoppage and the appropriate steps to take in the event they suspect a stoppage.******The firm will be implementing a correction of firmware errors to bring the pumps back into specification. The firm will follow up when the change has been completed. An office-based pump firmware update is being developed for patients currently implanted with the affected pumps. (Note, as of the 5/12/23 update, implanted pumps will not be receiving a firmware update.) | SoftwareChangeControl
45 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178923 | Z-1120-2020 | Prometra Programmable 20mL Pump. Catalog No. 11827 The device is an implantable infusion pump. GTIN-14 00810335020082 | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system. | Class 2 Device Recall Prometra Programmable Infusion Pump Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prometra Programmable Infusion Pump FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prometra Programmable Infusion Pump Date Initiated by Firm December 24, 2019 Date Posted February 10, 2020 Recall Status 1 Open 3 , Classified Recall Number Z-1120-2020 Recall Event ID 84653 PMA Number P080012 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product Prometra Programmable 20mL Pump. Catalog No. 11827 The device is an implantable infusion pump. GTIN-14 00810335020082 Code Information All units with pump firmware version 0.26 manufactured from 2012 until the present time. Serial Numbers 10JW4A32 10HW4A29 10JW4A44 10JW4A35 10HW4A45 10JW4A05 10JW4A08 10JW4A46 10JW4A36 10JW4A30 10JW4A39 10JW4A06 10JW4A21 10JW4A01 10JW4A27 10KW4A47 10JW4A29 10JW4A41 10KW4A122 10JW4A02 10KW4A05 10JW4A25 10KW4A38 10KW4A61 10KW4A11 10KW4A16 10KW4A27 10KW4A08 10KW4A117 10KW4A68 10KW4A99 10KW4A91 10KW4A20 10KW4A53 10KW4A106 10KW4A02 10JW4A20 10KW4A35 10KW4A18 10KW4A101 10AX4A21 10JW4A43 10KW4A73 10KW4A123 10KW4A67 10AX4B27 10KW4A31 10KW4A24 10AX4A37 10KW4A83 10AX4A36 10AX4A22 10AX4A10 10KW4A42 10KW4A119 10AX4B24 10KW4A89 10KW4A71 10KW4A116 10AX4A12 10DX4A79 10DX4A10 10AX4B06 10AX4B09 10DX4A95 10AX4A81 10AX4A98 10KW4A103 10AX4A96 10AX4A95 10AX4B82 10AX4A76 10KW4A48 10HY4C42 10KW4A09 10AX4A88 10DX4A32 10AX4A87 10AX4B92 10GX4B05 10GX4A30 10KW4A104 10AX4A14 10DX4A98 10GX4B21 10AX4A85 10KW4A85 10KW4A06 10AX4A73 10GX4A21 10GX4A09 10DX4A85 10GX4A82 10JW4A31 10GX4A87 10JW4A16 10GX4A85 10KW4A12 10GX4B16 10KW4A43 10GX4A34 10GX4A83 10AX4B78 10GX4B03 10GX4A38 10GX4B08 10KW4A29 10GX4B07 10KW4A30 10GX4A32 10AX4B90 10GX4A70 10AX4B81 10AX4B53 10GX4A92 10KY4A20 10KW4A52 10KW4A39 10GX4A19 10GX4A96 10DX4A21 10KX4A02 10DX4B07 10KX4A12 10KX4A15 10KW4A75 10GX4A99 10KX4A43 10GX4A45 10KX4A40 10KX4A41 10DX4A09 10DX4A74 10DX4A48 10KX4A42 10GX4A97 10GX4A10 10AX4B80 10KX4A59 10JW4A04 10KW4A44 10GX4A73 10KX4A14 10KX4A13 10GX4B15 10KX4A70 10KX4A57 10KX4A74 10KW4A25 10GX4A03 10GX4A63 10GX4A71 10KW4A108 10KX4A85 10KW4A105 10GX4A40 10GX4A14 10DX4A33 10BY4A40 10BY4A51 10BY4A39 10KX4A62 10BY4A25 10BY4A59 10BY4A36 10EY4A53 10KX4A10 10KX4A39 10EY4A27 10KX4A26 10EY4A26 10BY4A60 10EY4A38 10KX4A80 10BY4A57 10EY4A17 10EY4A83 10EY4A03 10BY4A48 10EY4A25 10EY4A29 10EY4A88 10EY4A99 10EY4A59 10EY4A05 10EY4A82 10EY4A19 10EY4A67 10EY4A89 10BY4A52 10KX4A17 10EY4A36 10EY4A58 10EY4A73 10EY4A43 10GX4A12 10KX4A78 10EY4A15 10BY4A34 10GY4A63 10EY4A37 10EY4A98 10KX4A25 10BY4A30 10GY4A16 10EY4A48 10KX4A76 10GY4A11 10GY4A19 10GY4A60 10GY4A55 10GY4A50 10GY4A49 10GY4A54 10BY4A31 10EY4A72 10KX4A11 10GY4A66 10GY4A40 10GY4A18 10EY4A94 10GY4A32 10EY4A13 10EY4A46 10GY4A72 10GY4A20 10GY4A06 10KX4A27 10KX4A08 10GY4A45 10KX4A30 10GY4A25 10KX4A32 10BY4A42 10GY4A15 10HY4B20 10GY4A35 10GY4A12 10GY4A58 10HY4B44 10GY4A23 10HY4B29 10GY4A21 10GY4A31 10GY4A53 10HY4B18 10KX4A29 10GY4A52 10HY4B16 10HY4B42 10HY4A97 10GY4A74 10HY4B48 10GY4A34 10HY4B41 10HY4B03 10EY4A76 10GY4A26 10HY4B23 10HY4B40 10HY4B25 10HY4B37 10HY4B55 10HY4B02 10HY4A85 10HY4B22 10HY4B54 10HY4B07 10HY4B80 10HY4B26 10HY4A78 10HY4C29 10HY4C11 10GY4A52¿ 10GX4B14 10KY4A41 10KX4A71 10GX4B02 10AX4B85 10HY4C03 10HY4B86 10HY4C16 10HY4C46 10HY4C34 10HY4B17 10HY4B65 10HY4B30 10HY4B62 10HY4B56 10HY4C35 10HY4B33 10HY4A81 10HY4C13 10HY4B10 10HY4C07 10HY4C27 10HY4C12 10HY4B96 10HY4B76 10HY4B24 10HY4C08 10HY4C39 10HY4C06 10HY4B83 10HY4C30 10HY4C36 10IY4A46 10HY4C40 10HY4B70 10IY4A29 10HY4C01 10HY4A98 10HY4B08 10IY4A17 10HY4C47 10HY4C32 10HY4C26 10IY4A38 10IY4A25 10HY4A80 10IY4A37 10IY4A43 10HY4C41 10IY4A51 10IY4A19 10IY4A33 10HY4B91 10HY4B51 10JY4A10 10KY4A60 10KY4A93 10KY4A15 10JY4A15 10KY4A21 10IY4A27 10IY4A13 10JY4A18 10JY4A07 10KY4A71 10KY4A39 10KY4A06 10HY4B97 10KY4A72 10KY4A86 10IY4A54 10KY4A87 10KY4A31 10AZ4A51 10KY4A25 10IY4A28 10KY4A50 10AZ4A35 10AZ4A24 10KY4A91 10AZ4A57 10IY4A31 10KY4A78 10KY4A84 10KY4A03 10KY4A82 10KY4A62 10KY4A44 10KY4A13 10HY4A79 10KY4A69 10AZ4A15 10JY4A28 10KY4A51 10JY4A16 10KY4A36 10AZ4A14 10AZ4A65 10KY4A94 10KY4A64 10KY4A47 10KY4A67 10JY4A03 10HY4A91 10AZ4A54 10HY4C23 10AZ4A08 10KY4A40 10AZ4A31 10BZ4A26 10KY4A79 10IY4A04 10EY4A60 10HY4B90 10KY4A90 10EY4A40 10KY4A30 10IY4A02 10AZ4A22 Recalling Firm/ Manufacturer Flowonix Medical Inc 120 Forbes Blvd Ste 170 Mansfield MA 02048-1150 For Additional Information Contact Technical Solutions Department 855-356-9665 Manufacturer Reason for Recall A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system. FDA Determined Cause 2 Software change control Action ***Updated 5/12/2023******On May 11, 2023, the firm sent an URGENT MEDICAL DEVICE CORRECTION UPDATE letter to all affected customers, which informed them of an update to Flowonix's plan to correct the firmware issue.******Flowonix will NOT be conducting an office-based pump firmware update for affected pumps implanted prior to January 2020.******Healthcare providers were provided ongoing recommendations for Prometra and Prometra II pumps implanted prior to January 2020, which are impacted by this issue.******Healthcare providers should contact their Flowonix Representative or our Technical Solutions Department (1-855-356-9665) in the event cessation of therapy or premature low battery is suspected and/or have general questions regarding this notice.***************On December 24, 2019, the firm notified affected users via an Urgent Medical Device Corrective Action Recall Notice.******The firm advised customers of the product issue and reminded customers that the pumps are currently labeled only for use with Infumorph. Due to the low risk that this failure mode may occur, Flowonix is not recommending that pumps be explanted. Healthcare providers were recommended to inform patients regarding the symptoms and adverse health events associated with an unexpected pump stoppage and the appropriate steps to take in the event they suspect a stoppage.******The firm will be implementing a correction of firmware errors to bring the pumps back into specification. The firm will follow up when the change has been completed. An office-based pump firmware update is being developed for patients currently implanted with the affected pumps. (Note, as of the 5/12/23 update, implanted pumps will not be receiving a firmware update.) Quantity in Commerce 6670 (US) Distribution Nationwide domestic distribution. Foreign distribution to United Kingdom, Republic of Ireland, Poland, Greece, Italy, Germany, Netherlands, and Belgium. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = LKK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action ***Updated 5/12/2023******On May 11, 2023, the firm sent an URGENT MEDICAL DEVICE CORRECTION UPDATE letter to all affected customers, which informed them of an update to Flowonix's plan to correct the firmware issue.******Flowonix will NOT be conducting an office-based pump firmware update for affected pumps implanted prior to January 2020.******Healthcare providers were provided ongoing recommendations for Prometra and Prometra II pumps implanted prior to January 2020, which are impacted by this issue.******Healthcare providers should contact their Flowonix Representative or our Technical Solutions Department (1-855-356-9665) in the event cessation of therapy or premature low battery is suspected and/or have general questions regarding this notice.***************On December 24, 2019, the firm notified affected users via an Urgent Medical Device Corrective Action Recall Notice.******The firm advised customers of the product issue and reminded customers that the pumps are currently labeled only for use with Infumorph. Due to the low risk that this failure mode may occur, Flowonix is not recommending that pumps be explanted. Healthcare providers were recommended to inform patients regarding the symptoms and adverse health events associated with an unexpected pump stoppage and the appropriate steps to take in the event they suspect a stoppage.******The firm will be implementing a correction of firmware errors to bring the pumps back into specification. The firm will follow up when the change has been completed. An office-based pump firmware update is being developed for patients currently implanted with the affected pumps. (Note, as of the 5/12/23 update, implanted pumps will not be receiving a firmware update.) | SoftwareChangeControl
46 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178925 | Z-1121-2020 | Prometra II 20mL Pump. Catalog No. 13827 The device is an implantable infusion pump. GTIN-14 00810335020228 | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system. | Class 2 Device Recall Prometra Programmable Infusion Pump Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prometra Programmable Infusion Pump FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prometra Programmable Infusion Pump Date Initiated by Firm December 24, 2019 Date Posted February 10, 2020 Recall Status 1 Open 3 , Classified Recall Number Z-1121-2020 Recall Event ID 84653 PMA Number P080012 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product Prometra II 20mL Pump. Catalog No. 13827 The device is an implantable infusion pump. GTIN-14 00810335020228 Code Information All units with pump firmware version 0.26 manufactured from 2012 until the present time. Serial Numbers 10HW4A31 10AZ4A69X 10AZ4A07X 10BZ4A16X 10BZ4A61X 10BZ4A41X 10AZ4A48X 10AZ4A50X 10AZ4A23X 10AZ4A36X 10AZ4A42X 10BZ4A37X 10BZ4A64X 10BZ4A12X 10CZ4B09X 10CZ4A95X 10CZ4B22X 10BZ4A18X 10CZ4A79X 10CZ4A86X 10CZ4A96X 10CZ4A88X 10CZ4A80X 10CZ4A85X 10CZ4A81X 10CZ4A82X 10CZ4A78X 10CZ4A91X 10CZ4A94X 10CZ4A90X 10CZ4A77X 10CZ4B28X 10CZ4B05X 10CZ4B15X 10CZ4B21X 10CZ4A87X 10CZ4B32X 10AZ4A41X 10CZ4B23X 10CZ4A93X 10CZ4B01X 10CZ4B17X 10CZ4B04X 10CZ4B30X 10BZ4A55X 10CZ4B08X 10CZ4A76X 10CZ4B33X 10CZ4B20X 10DZ4A01X 10DZ4A58X 10DZ4A19X 10CZ4B36X 10CZ4A84X 10DZ4A18X 10CZ4B24X 10DZ4A76X 10BZ4A30X 10CZ4B26X 10BZ4A09X 10DZ4A51X 10DZ4A93X 10DZ4A32X 10CZ4B03X 10DZ4A89X 10BZ4A11X 10DZ4A30X 10CZ4B25X 10DZ4A65X 10DZ4A05X 10DZ4A98X 10DZ4A08X 10DZ4A42X 10DZ4A77X 10CZ4A92X 10DZ4A35X 10DZ4A80X 10DZ4A92X 10DZ4A88X 10DZ4A07X 10BZ4A39X 10DZ4A73X 10DZ4A53X 10DZ4A28X 10DZ4A84X 10DZ4A45X 10CZ4B14X 10DZ4A40X 10DZ4A72X 10DZ4A64X 10DZ4A79X 10DZ4A60X 10DZ4A82X 10DZ4A37X 10DZ4B39X 10DZ4A54X 10DZ4A67X 10CZ4B29X 10BZ4A59X 10DZ4A20X 10DZ4A38X 10DZ4B04X 10DZ4A36X 10DZ4A91X 10DZ4B37X 10DZ4B06X 10DZ4A34X 10DZ4A25X 10DZ4B24X 10DZ4B49X 10DZ4B38X 10CZ4A89X 10DZ4B05X 10DZ4A15X 10DZ4A43X 10DZ4B54X 10BZ4A25X 10DZ4A66X 10DZ4B53X 10DZ4A29X 10DZ4A23X 10DZ4B08X 10DZ4A21X 10DZ4B40X 10DZ4A83X 10DZ4A17X 10DZ4A62X 10DZ4B19X 10DZ4B46X 10EZ4A36X 10DZ4B16X 10DZ4A87X 10DZ4B25X 10DZ4B43X 10DZ4B10X 10DZ4B61X 10DZ4B50X 10DZ4B07X 10DZ4B48X 10DZ4B26X 10DZ4B29X 10DZ4A27X 10CZ4B27X 10EZ4A02X 10DZ4B22X 10DZ4B15X 10EZ4A46X 10EZ4A27X 10EZ4A01X 10DZ4A47X 10DZ4A81X 10DZ4A33X 10EZ4A30X 10EZ4A26X 10EZ4A51X 10DZ4B03X 10EZ4A37X 10EZ4A13X 10EZ4A39X 10EZ4A35X 10DZ4B45X 10EZ4A28X 10DZ4A10X 10FZ4A34X 10CZ4B35X 10EZ4A09X 10EZ4A43X 10EZ4A05X 10EZ4A17X 10EZ4A16X 10DZ4B21X 10DZ4A94X 10EZ4A19X 10EZ4A45X 10DZ4A44X 10DZ4A68X 10EZ4A25X 10FZ4A29X 10EZ4A32X 10DZ4A48X 10FZ4A24X 10EZ4A18X 10DZ4B55X 10DZ4A06X 10FZ4A02X 10EZ4A57X 10EZ4A24X 10EZ4A58X 10EZ4A06X 10FZ4A38X 10FZ4A03X 10EZ4A29X 10EZ4A08X 10EZ4A55X 10FZ4A36X 10EZ4A20X 10EZ4A31X 10DZ4B18X 10GZ4A40X 10FZ4A18X 10FZ4A39X 10DZ4B42X 10FZ4A31X 10EZ4A47X 10GZ4A44X 10GZ4A58X 10DZ4B51X 10GZ4A27X 10DZ4B30X 10FZ4A37X 10GZ4A01X 10GZ4A43X 10GZ4A38X 10GZ4A35X 10EZ4A48X 10GZ4A03X 10GZ4A08X 10GZ4A37X 10DZ4B41X 10GZ4A02X 10DZ4B31X 10DZ4A55X 10GZ4A76X 10GZ4A66X 10DZ4B20X 10FZ4A15X 10FZ4A01X 10GZ4A34X 10EZ4A10X 10FZ4A22X 10GZ4A73X 10EZ4A53X 10FZ4A21X 10GZ4A94X 10GZ4A06X 10GZ4A61X 10GZ4A78X 10EZ4A60X 10FZ4A13X 10GZ4A96X 10GZ4A04X 10FZ4A28X 10GZ4A11X 10EZ4A34X 10GZ4A59X 10EZ4A11X 10GZ4A05X 10FZ4A04X 10GZ4A18X 10EZ4A49X 10GZ4A54X 10GZ4A92X 10GZ4A90X 10GZ4A57X 10EZ4A40X 10FZ4A40X 10GZ4A63X 10DZ4A49X 10GZ4A41X 10GZ4A29X 10GZ4A83X 10GZ4B54X 10GZ4A70X 10GZ4A19X 10EZ4A23X 10FZ4A06X 10GZ4A95X 10FZ4A20X 10GZ4A30X 10GZ4A89X 10GZ4A72X 10GZ4A71X 10GZ4A12X 10GZ4B34X 10GZ4B18X 10GZ4A09X 10GZ4B22X 10GZ4A99X 10GZ4A93X 10GZ4A64X 10GZ4A77X 10FZ4A05X 10DZ4A59X 10GZ4A75X 10GZ4A98X 10GZ4B26X 10GZ4A33X 10GZ4B08X 10FZ4A32X 10GZ4B59X 10GZ4B38X 10GZ4B07X 10GZ4B44X 10GZ4B41X 10GZ4A69X 10GZ4A87X 10HZ4A02X 10GZ4B55X 10GZ4B28X 10HZ4A26X 10GZ4A91X 10HZ4A41X 10GZ4B52X 10HZ4A46X 10FZ4A35X 10GZ4B16X 10FZ4A09X 10HZ4A70X 10DZ4A12X 10GZ4B37X 10GZ4A42X 10DZ4B35X 10GZ4A20X 10HZ4A82X 10HZ4A49X 10GZ4A97X 10EZ4A59X 10HZ4A16X 10GZ4A21X 10GZ4A53X 10HZ4A07X 10HZ4A33X 10HZ4A81X 10GZ4B06X 10HZ4A19X 10HZ4A22X 10HZ4A83X 10HZ4A20X 10HZ4A21X 10HZ4A12X 10HZ4A04X 10DZ4B12X 10GZ4A07X 10HZ4A15X 10GZ4B01X 10HZ4A80X 10GZ4A46X 10HZ4A35X 10HZ4A39X 10GZ4B51X 10GZ4B12X 10HZ4A09X 10HZ4A57X 10GZ4B32X 10GZ4B57X 10HZ4A71X 10GZ4B10X 10HZ4A25X 10FZ4A07X 10GZ4B11X 10HZ4A64X 10IZ4B09X 10GZ4A81X 10GZ4B21X******10GZ4B53X 10IZ4A84X 10HZ4A11X 10JZ4A67X 10HZ4A05X 10HZ4A37X 10IZ4A64X 10GZ4B36X 10IZ4A38X 10GZ4B04X 10GZ4B35X 10HZ4A61X 10HZ4A28X 10HZ4A44X 10HZ4A23X 10IZ4A78X 10GZ4B61X 10GZ4A39X 10IZ4A91X 10HZ4A40X 10IZ4B05X 10IZ4A60X 10IZ4B14X 10IZ4A61X 10GZ4B25X 10GZ4B47X 10IZ4A92X 10IZ4B08X 10IZ4B07X 10HZ4A55X 10GZ4B45X 10GZ4B42X 10HZ4A69X 10IZ4A50X 10IZ4A85X 10GZ4A15X 10IZ4B17X 10HZ4A42X 10CZ4B06X 10HZ4A53X 10IZ4A79X 10HZ4A24X 10IZ4A69X 10IZ4A42X 10IZ4A96X 10IZ4A72X 10HZ4A17X 10IZ4B21X 10IZ4A73X 10IZ4A74X 10IZ4A31X 10JZ4A18X 10IZ4A77X 10IZ4A49X 10IZ4A57X 10IZ4A35X 10IZ4B16X 10IZ4B15X 10IZ4A46X 10GZ4A10X 10GZ4B03X 10IZ4A71X 10IZ4A43X 10JZ4A05X 10IZ4B06X 10JZ4A60X 10HZ4A45X 10JZ4A35X 10IZ4A40X 10IZ4A34X 10IZ4A93X 10HZ4A85X 10JZ4A11X 10IZ4B04X 10JZ4A54X 10JZ4A56X 10IZ4A67X 10GZ4A36X 10IZ4B01X 10HZ4A08X 10JZ4A12X 10JZ4A27X 10HZ4A56X 10JZ4A46X 10JZ4A78X 10JZ4A23X 10HZ4A79X 10JZ4A02X 10JZ4A89X 10JZ4A19X 10JZ4A82X 10JZ4A44X 10IZ4A76X 10IZ4A53X 10JZ4A90X 10JZ4A81X 10IZ4B12X 10GZ4B15X 10JZ4A80X 10JZ4A06X 10JZ4A53X 10JZ4A87X 10GZ4B60X 10JZ4A88X 10JZ4A17X 10HZ4A31X 10JZ4A72X 10JZ4A01X 10JZ4A16X 10IZ4A75X 10JZ4A07X 10HZ4A03X 10HZ4A54X 10JZ4A39X 10HZ4A13X 10JZ4A29X 10JZ4A65X 10JZ4A63X 10JZ4A69X 10JZ4A32X 10EZ4A22X 10JZ4A37X 10IZ4A66X 10JZ4A62X 10JZ4A34X 10JZ4A43X 10JZ4A42X 10IZ4A83X 10IZ4A59X 10JZ4B16X 10HZ4A34X 10JZ4A08X 10JZ4B17X 10JZ4B25X 10JZ4B10X 10JZ4B08X 10JZ4A13X 10JZ4B02X 10JZ4B24X 10JZ4A95X 10JZ4A25X 10JZ4A64X 10IZ4B13X 10JZ4B12X 10JZ4B31X 10JZ4B29X 10JZ4A74X 10JZ4A76X 10JZ4B26X 10JZ4B30X 10JZ4A97X 10JZ4B40X 10JZ4A98X 10JZ4B36X 10JZ4A40X 10JZ4B06X 10GZ4A50X 10JZ4B23X 10JZ4C20X 10JZ4C09X 10JZ4B39X 10JZ4B73X 10HZ4A50X 10JZ4C26X 10JZ4B13X 10JZ4B42X 10JZ4B78X 10HZ4A65X 10JZ4C17X 10JZ4B72X 10JZ4B68X 10JZ4C24X 10GZ4B05X 10JZ4B11X 10JZ4B65X 10JZ4B66X 10JZ4C32X 10IZ4A44X 10JZ4C18X 10JZ4B34X 10JZ4B53X 10JZ4B22X 10JZ4B43X 10JZ4C13X 10JZ4B52X 10GZ4A79X 10JZ4B38X 10JZ4B09X 10JZ4A58X 10JZ4C05X 10JZ4C02X 10JZ4B03X 10JZ4B37X 10JZ4B88X 10JZ4B90X 10JZ4C01X 10JZ4B07X 10JZ4B71X 10JZ4C25X 10JZ4A26X 10HZ4A66X 10JZ4B67X 10JZ4C21X 10JZ4B99X 10JZ4C23X 10JZ4A75X 10JZ4C11X 10JZ4A83X 10JZ4C27X 10KZ4A01X 10JZ4B76X 10KZ4A39X 10JZ4B58X 10FA4A69X 10JZ4B74X 10KZ4A27X 10JZ4C30X 10HZ4A63X 10JZ4B59X 10GZ4B13X 10JZ4A20X 10JZ4A52X 10KZ4A08X 10JZ4B50X 10JZ4B70X 10JZ4B98X 10JZ4B04X 10IZ4A82X 10JZ4B80X 10KZ4A09X 10JZ4B45X 10GZ4B24X 10KZ4A06X 10KZ4A18X 10KZ4A23X 10JZ4A99X 10JZ4C14X 10KZ4A20X 10JZ4B89X 10KZ4A17X 10KZ4A40X 10KZ4A05X 10KZ4A26X 10KZ4A19X 10KZ4A30X 10JZ4C28X 10JZ4B79X 10JZ4A85X 10JZ4B96X 10KZ4A41X 10KZ4A22X 10JZ4C31X 10IZ4A65X 10JZ4A84X 10KZ4A29X 10LZ4A87X 10KZ4A25X 10KZ4A45X 10JZ4B44X 10JZ4B87X 10KZ4A33X 10KZ4A42X 10KZ4A31X 10JZ4B63X 10JZ4A91X 10LZ4A48X 10KZ4A14X 10JZ4B32X 10LZ4A38X 10LZ4A74X 10LZ4A43X 10JZ4B57X 10JZ4A47X 10JZ4B94X 10LZ4A58X 10LZ4A46X 10LZ4A75X 10KZ4A04X 10LZ4A64X 10JZ4B61X 10LZ4A84X 10LZ4A67X 10JZ4C04X 10LZ4A45X 10LZ4A39X 10AA4A44X 10LZ4A55X 10LZ4A40X 10LZ4A61X 10LZ4A60X 10KZ4A10X 10KZ4A11X 10LZ4A66X 10LZ4A52X 10JZ4A33X 10LZ4A69X 10JZ4A36X 10LZ4A90X 10LZ4A82X 10AA4A38X 10JZ4B62X 10LZ4A68X 10LZ4A41X 10LZ4A81X 10KZ4A02X 10AA4A06X 10LZ4A37X 10LZ4A49X 10KZ4A32X 10LZ4A80X 10LZ4A85X 10AA4A34X 10LZ4A53X 10KZ4A38X 10LZ4A59X 10LZ4A76X 10JZ4A96X 10LZ4A77X 10JZ4C15X 10LZ4A50X 10LZ4A78X 10AA4A58X 10AA4A79X 10LZ4A94X 10AA4A27X 10AA4A32X 10HZ4A58X 10AA4A71X 10JZ4C29X 10AA4A24X 10LZ4A51X 10AA4A85X 10KZ4A12X 10LZ4A88X 10AA4A12X 10AA4A65X 10JZ4B54X 10AA4A40X 10AA4A46X 10AA4A19X 10AA4A55X 10AA4A28X 10AA4A14X 10AA4A15X 10AA4A21X 10LZ4A42X 10AA4A41X 10LZ4A44X 10AA4A05X 10AA4A77X 10AA4A49X 10IZ4A55X 10JZ4B64X 10JZ4B60X 10AA4A09X 10AA4A07X 10AA4A78X 10JZ4B49X 10KZ4A03X 10AA4A17X 10AA4A47X 10KZ4A28X 10AA4A01X 10AA4A59X 10LZ4A93X 10AA4A67X 10AA4A86X 10AA4A30X 10KZ4A36X 10LZ4A62X 10LZ4A95X 10AA4A03X 10AA4A11X 10AA4A72X 10AA4A90X 10AA4A52X 10AA4A83X 10AA4A31X 10AA4A10X 10KZ4A24X 10AA4A76X 10AA4A51X 10IZ4B11X 10AA4A63X 10IZ4A88X 10AA4A60X 10LZ4A89X 10BA4A54X 10AA4A04X 10AA4A20X 10AA4A13X 10AA4A69X 10BA4A30X 10BA4A40X 10BA4A36X 10LZ4A72X 10AA4A08X 10LZ4A91X 10IZ4A33X 10BA4A79X 10BA4A74X 10AA4A57X 10AA4A02X 10BA4A62X 10AA4A50X 10KZ4A16X 10BA4A38X 10BA4A76X 10BA4A28X 10AA4A61X 10BA4A64X 10BA4A75X 10BA4A50X 10AA4A39X 10BA4A34X 10AA4A81X 10AA4A54X 10AA4A80X 10KZ4A13X 10LZ4A54X 10BA4A68X 10AA4A23X 10BA4A57X 10LZ4A92X 10BA4A49X 10BA4A60X 10AA4A26X 10BA4A47X 10AA4A36X 10BA4A72X 10JZ4C22X 10JZ4B47X 10AA4A87X 10AA4A29X 10BA4A65X 10AA4A45X 10AA4A88X 10BA4A70X 10BA4A39X 10BA4A71X 10AA4A42X 10AA4A56X 10JZ4A59X 10JZ4A15X 10EZ4A42X 10FZ4A08X 10CA4A43X 10CA4A44X 10CA4A23X 10BA4A84X 10CA4A05X 10BA4A41X 10BA4A82X 10CA4A11X 10CA4A08X 10BA4A51X 10IZ4A68X 10BA4A55X 10CA4A30X 10BA4A33X 10BA4A66X 10CA4A01X 10BA4A52X 10CA4A10X 10BA4A29X 10BA4A78X 10CA4A07X 10CA4A45X 10JZ4A14X 10CA4A38X 10CA4A17X 10CA4A53X 10CA4A16X 10CA4A94X 10DA4A48X 10CA4A04X 10CA4A42X 10CA4A39X 10BA4A35X 10CA4A82X 10BA4A53X 10BA4A58X 10CA4A13X 10CA4A18X 10CA4A19X 10CA4A02X 10CA4A83X 10CA4A46X 10CA4A09X 10CA4A26X 10CA4A80X 10CA4A63X 10BA4A31X 10CA4A57X 10CA4A49X 10CA4A62X 10CA4A61X 10BA4A81X 10CA4A87X 10CA4A89X 10CA4A25X 10BA4A85X 10CA4A56X 10CA4A88X 10CA4A60X 10CA4A85X 10DA4A01X 10CA4A97X 10DA4A21X 10CA4A95X 10DA4A31X 10CA4A93X 10CA4A74X 10DA4A53X 10DA4A49X 10CA4A73X 10CA4A64X 10CA4A84X 10CA4A67X 10DA4A52X 10DA4A09X 10CA4A69X 10DA4A69X 10BA4A27X 10DA4A02X 10BA4A83X 10DA4A62X 10DA4A07X 10DA4A70X 10DA4A39X 10DA4A41X 10DA4A67X 10DA4A66X 10DA4A58X 10DA4A54X 10DA4A35X 10DA4A43X 10CA4A66X 10CA4A34X 10DA4A42X 10JZ4B84X 10CA4A58X 10DA4A71X 10DA4A44X 10DA4A30X 10CA4A14X 10CA4A70X 10EA4A65X 10EA4A71X 10EA4A26X 10EA4A58X 10EA4A73X 10EA4A47X 10DA4A68X 10EA4A78X 10EA4A11X 10DA4A14X 10DA4A57X 10DA4A17X 10DA4A74X 10DA4A34X 10DA4A05X 10EA4A25X 10EA4A33X 10DA4A72X 10EA4A12X 10DA4A37X 10EA4A01X 10CA4A90X 10EA4A21X 10EA4A24X 10EA4A77X 10EA4A30X 10CA4A50X 10DA4A73X 10EA4A62X 10EA4A53X 10DA4A12X 10CA4A68X 10DA4A47X 10EA4A07X 10EA4A15X 10DA4A15X 10EA4A64X 10EA4A63X 10CA4A03X 10EA4A35X 10EA4A27X 10AA4A37X 10CA4A12X 10CA4A78X 10EA4A32X 10EA4A56X 10EA4A03X 10DA4A27X 10EA4A18X 10EA4A44X 10EA4A42X 10DA4A60X 10EA4A28X 10DA4A63X 10DA4A18X 10EA4A06X 10EA4A68X 10EA4A38X 10FA4A90X 10FA4A81X 10BA4A43X 10EA4A08X 10EA4A34X 10EA4A59X 10DA4A32X 10FA4A39X 10FA4A94X 10EA4A13X 10EA4A49X 10FA4A79X 10DA4A59X 10FA4B02X 10FA4A59X 10FA4A35X 10CA4A92X 10EA4A69X 10FA4A53X 10FA4B10X 10FA4A70X 10FA4A46X 10FA4A83X 10EA4A04X 10EA4A46X 10DA4A26X 10FA4A49X 10FA4A40X 10FA4A76X 10CA4A35X 10FA4A97X 10CA4A71X 10FA4B06X 10FA4A43X 10FA4A37X 10FA4B01X 10FA4A34X 10EA4A20X 10FA4A32X 10EA4A70X 10GA4A46X 10FA4A44X 10GA4A50X 10EA4A60X 10FA4A82X 10GA4A65X 10EA4A45X 10EA4A23X 10FA4A51X 10GA4A44X 10FA4A85X 10DA4A64X 10GA4A34X 10GA4A52X 10DA4A51X 10GA4A32X 10FA4B05X 10CA4A52X 10DA4A61X 10EA4A09X 10CA4A48X 10FA4A42X 10FA4A99X 10EA4A50X 10FA4A36X 10GA4A07X 10EA4A31X 10FA4A96X 10EA4A74X 10FA4A55X 10FA4A62X 10FA4A56X 10CA4A51X 10GA4A11X 10GA4A23X 10FA4A31X 10GA4A67X 10FA4A74X 10EA4A22X 10FA4A66X 10GA4A68X 10FA4A67X 10GA4A10X 10FA4A88X 10GA4A40X 10FA4A72X 10GA4A61X 10GA4A64X 10FA4A57X 10EA4A66X 10FA4A65X 10FA4A64X 10HA4A09X 10HA4A25X 10EA4A75X 10FA4A75X 10FA4A89X 10DA4A06X 10FA4A77X 10FA4A91X 10FA4A33X 10GA4A28X 10GA4A73X 10GA4A43X 10HA4A54X 10DA4A16X 10HA4A38X 10GA4A71X 10FA4B07X 10GA4A53X 10HA4A36X 10HA4A61X 10FA4A80X 10HA4A40X 10HA4A20X 10GA4A01X 10HA4A12X 10GA4A63X 10HA4A66X 10HA4A32X 10GA4A09X 10HA4A18X 10HA4A75X 10HA4A63X 10HA4A72X 10FA4A47X 10HA4A16X 10HA4A46X 10HA4A07X 10HA4A98X 10HA4A11X 10HA4A14X 10GA4A14X 10GA4A05X 10FA4A86X 10FA4B03X 10HA4A19X 10HA4A05X 10HA4A74X 10FA4A93X 10GA4A74X 10FA4A87X 10HA4A82X 10HA4B26X 10HA4A59X 10HA4A65X 10HA4A06X 10GA4A18X 10HA4A77X 10HA4A76X 10HA4A57X 10GA4A51X 10HA4A79X 10HA4A89X 10HA4B21X 10HA4B27X 10CA4A96X 10HA4B04X 10HA4B54X 10HA4B11X 10HA4A87X 10HA4B02X 10FA4A78X 10HA4B03X 10HA4A50X 10HA4A93X 10HA4A68X 10HA4B15X 10GA4A47X 10HA4A02X 10HA4A80X 10EA4A19X 10HA4B82X 10HA4B83X 10HA4B05X 10HA4A78X 10GA4A75X 10HA4B47X 10HA4A45X 10HA4B32X 10HA4A99X 10HA4A91X 10HA4B39X 10HA4B23X 10CA4A79X 10HA4A58X 10HA4A30X 10HA4A28X 10HA4A81X 10HA4B44X 10GA4A24X 10EA4A76X 10HA4A53X 10HA4B25X 10HA4A96X 10HA4A86X 10HA4B72X 10HA4B29X 10HA4B52X 10HA4C17X 10HA4B17X 10HA4A90X 10HA4A42X 10HA4D20X 10HA4B49X 10HA4B80X 10EA4A51X 10GA4A72X 10HA4A15X 10HA4C85X 10GA4A08X 10HA4A73X 10HA4C30X 10HA4A35X 10HA4B30X 10GA4A21X 10HA4D11X 10HA4C23X 10HA4B68X 10HA4B85X 10HA4B65X 10HA4B86X 10HA4B64X 10HA4B77X 10HA4C28X 10HA4B76X 10HA4D15X 10HA4C21X 10HA4C26X 10HA4B20X 10HA4B59X 10HA4A04X 10HA4B12X 10HA4B51X 10HA4B90X 10HA4B63X 10HA4C91X 10HA4B06X 10HA4A37X 10HA4D09X 10HA4D33X 10HA4D17X 10HA4D13X 10HA4D19X 10HA4B98X 10HA4B57X 10HA4C04X 10HA4D29X 10DA4A29X 10HA4C15X 10HA4D87X 10HA4B94X 10HA4D22X 10GA4A56X 10HA4D04X 10HA4C87X 10FA4B08X 10HA4B33X 10HA4C94X 10HA4C16X 10GA4A12X 10HA4C19X 10HA4D86X 10HA4C92X 10HA4D30X 10HA4A85X 10HA4B38X 10HA4C84X 10HA4D85X 10GA4A54X 10HA4C27X 10HA4D88X 10HA4B46X 10HA4E28X 10HA4D99X 10HA4D08X 10GA4A41X 10HA4A23X 10HA4D45X 10HA4E29X 10HA4B92X 10HA4A21X 10HA4A43X 10HA4D21X 10HA4D32X 10HA4B40X 10HA4D60X 10HA4C99X 10HA4D05X 10HA4D25X 10HA4B41X 10HA4C18X 10HA4D58X 10HA4D46X 10HA4D35X 10HA4D68X 10HA4B18X 10HA4D93X 10HA4B66X 10HA4E03X 10HA4D59X 10HA4D14X 10HA4A10X 10HA4C08X 10HA4D56X 10HA4D79X 10FA4A98X 10HA4D34X 10HA4D53X 10HA4B89X 10FA4A48X 10HA4B75X 10HA4A49X 10HA4C13X 10HA4A51X 10HA4D26X 10HA4D27X 10HA4B31X 10HA4B45X 10HA4B55X 10HA4B73X 10HA4A97X 10HA4A88X 10HA4D44X 10HA4B48X 10HA4B16X 10HA4D01X 10HA4A48X 10HA4D62X 10HA4E19X 10HA4C14X 10IA4A96X 10HA4D41X 10HA4C25X 10HA4B69X 10HA4D70X 10HA4B37X 10HA4D18X 10HA4D48X 10HA4D69X 10HA4D39X 10HA4C20X 10GA4A39X 10HA4D36X 10HA4D57X 10HA4C10X 10HA4C90X 10HA4D54X 10HA4D65X 10HA4D23X 10HA4B95X 10HA4A84X 10HA4E26X 10HA4C09X 10HA4C05X 10HA4D12X 10HA4D98X 10HA4D38X 10IA4A19X 10HA4B81X 10HA4D82X 10HA4B13X 10IA4A44X 10HA4E06X 10IA4A24X 10HA4D78X 10IA4A02X 10IA4A15X 10IA4B51X 10IA4A90X 10IA4B24X 10IA4A37X 10HA4D55X 10IA4A31X 10HA4E11X 10IA4A41X 10HA4E17X 10HA4E30X 10HA4E05X 10HA4E10X 10IA4A78X 10FA4A95X 10HA4D43X 10HA4B28X 10IA4B38X 10IA4A06X 10IA4A93X 10HA4D95X 10IA4B53X 10IA4A69X 10IA4B71X 10IA4A32X 10IA4A51X 10IA4A72X 10IA4A98X 10HA4D63X 10HA4D73X 10IA4A25X 10HA4B08X 10HA4E04X 10IA4B30X 10HA4E27X 10IA4B27X 10HA4E23X 10HA4D37X 10IA4B45X 10IA4B73X 10IA4B58X 10HA4D96X 10JA4A14X 10IA4B52X 10IA4B68X 10IA4B36X 10IA4A59X 10HA4D31X 10GA4A35X 10IA4B02X 10IA4B06X 10HA4E14X 10IA4A53X 10IA4B42X 10HA4C11X 10IA4B10X 10IA4B41X 10IA4B34X 10IA4A34X 10HA4E15X 10HA4D81X 10HA4A62X 10IA4A94X 10HA4D67X 10JA4A16X 10IA4B62X 10IA4A88X 10HA4A70X 10HA4D24X 10IA4B66X 10IA4B08X 10IA4B18X 10IA4B09X 10IA4A60X 10IA4A38X 10IA4B57X 10IA4A30X 10IA4B76X 10HA4D51X 10IA4A17X 10FA4A61X 10JA4A65X 10IA4B20X 10JA4A57X 10HA4D90X 10IA4B72X 10HA4D06X 10HA4D91X 10IA4B39X 10JA4A32X 10JA4A39X 10IA4B61X 10IA4A04X 10JA4A98X 10JA4A43X 10FA4A41X 10FA4A73X 10JA4A42X 10IA4A95X 10IA4A29X 10AA4A16X 10GA4A37X 10IA4B07X 10IA4A36X 10HA4E12X 10HA4B50X 10HA4D52X 10JA4A35X 10IA4B47X 10JA4B42X 10JA4B28X 10HA4A44X 10IA4B56X 10JA4A55X 10JA4B09X 10JA4B39X 10JA4A37X 10IA4A55X 10JA4B12X 10JA4A49X 10IA4A56X 10HA4D71X 10HA4B70X 10JA4A29X 10IA4B75X 10JA4A17X 10JA4A75X 10IA4B49X 10JA4A44X 10JA4A06X 10IA4B63X 10JA4B31X 10JA4A08X 10JA4A51X 10HA4E07X 10JA4B32X 10JA4B24X 10JA4A05X 10HA4D40X 10JA4B23X 10JA4A80X 10IZ4A97X 10JA4B58X 10JA4A86X 10HA4B43X 10JA4A30X 10HA4D97X 10JA4A47X 10JA4B30X 10JA4B16X 10HA4D02X 10JA4B13X 10JA4B46X 10JA4A93X 10IA4A87X 10JA4B29X 10IA4A77X 10HA4D72X 10JA4A64X 10IA4A05X 10JA4B10X 10IA4A14X 10JA4B22X 10JA4A78X 10JA4A91X 10IA4A16X 10KA4A15X 10IA4B67X 10JA4B50X 10JA4A90X 10IA4A23X 10JA4A97X 10KA4A19X 10JA4A72X 10KA4A61X 10JA4A07X 10JA4A74X 10KA4A40X 10IA4A35X 10IA4A80X 10JA4B45X 10JA4A31X 10JA4A03X 10JA4B18X 10JA4A24X 10JA4A79X 10JA4A45X 10KA4A13X 10JA4A83X 10JA4B11X 10JA4A61X 10KA4A17X 10KA4A29X 10JA4B52X 10JA4B61X 10HA4B42X 10JA4A77X 10JA4A89X 10JA4B17X 10JA4A23X 10JA4A58X 10KA4A50X 10JA4A94X 10LA4A88X 10JA4B14X 10JA4B53X 10LA4A94X 10KA4A33X 10KA4A59X 10LA4B22X 10KA4A01X 10JA4A99X 10IA4A70X 10KA4A63X 10JA4B47X 10IA4B44X 10JA4B41X 10IA4B28X 10JA4A50X 10JA4A48X 10JA4A15X 10KA4A34X 10LA4A76X 10JA4A85X 10KA4A05X 10JA4A41X 10HA4B10X 10JA4B43X 10JA4B54X 10LA4A81X 10JA4B55X 10KA4A56X 10KA4A49X 10KA4A27X 10KA4A43X 10KA4A07X 10LA4A78X 10IA4B11X 10JA4B38X 10KA4A22X 10IA4B16X 10JA4A69X 10LA4B10X 10JA4A59X 10KA4A24X 10LA4B09X 10JA4A28X 10JA4A88X 10JA4B20X 10LA4B12X 10LA4A71X 10LA4A87X 10KA4A14X 10KA4A52X 10LA4A99X 10JA4A95X 10IA4B40X 10KA4A21X 10KA4A42X 10KA4A11X 10JA4A71X 10KA4A08X 10KA4A55X 10KA4A35X 10LA4A74X 10LA4B21X 10JA4B06X 10LA4A96X 10LA4B17X 10IA4A07X 10LA4B23X 10JA4B26X 10LA4A84X 10IA4B65X 10KA4A36X 10KA4A28X 10KA4A51X 10LA4A91X 10AB4A77X 10AB4A87X 10KA4A44X 10LA4B16X 10HA4D61X 10AB4A33X 10AB4A39X 10AB4A50X 10LA4A79X 10HA4D89X 10KA4A38X 10HA4D49X 10AB4A46X 10AB4A02X 10AB4A80X 10AB4A09X 10AB4A32X 10KA4A03X 10AB4A30X 10LA4A80X 10AB4A22X 10IA4A85X 10AB4A35X 10AB4A03X 10JA4A26X 10AB4A48X 10AB4A59X 10JA4A62X 10AB4A74X 10AB4A51X 10AB4A17X 10KA4A57X 10AB4A96X 10AB4A07X 10KA4A41X 10AB4A27X 10AB4A81X 10AB4A20X 10KA4A32X 10AB4A86X 10JA4B44X 10AB4A01X 10AB4A10X 10JA4A63X 10KA4A45X 10EA4A40X 10AB4A29X 10JA4B33X 10LA4B07X 10KA4A48X 10JA4B37X 10AB4A83X 10AB4A69X 10LA4B13X 10AB4A45X 10AB4A60X 10LA4A89X 10AB4B46X 10LA4A97X 10AB4A38X 10AB4A92X 10KA4A25X 10AB4A21X 10AB4A58X 10AB4A05X 10HA4D47X 10AB4A44X 10AB4A24X 10AB4A71X 10JA4A60X 10AB4A31X 10AB4A19X 10JA4B15X 10JA4A87X 10JA4A04X 10LA4A98X 10AB4B30X 10EA4A05X 10AB4B21X 10KA4A37X 10IA4A62X 10IA4A81X 10IA4A58X 10HA4E18X 10IA4A75X 10AB4B13X 10BB4A46X 10AB4A04X 10KA4A46X 10AB4B35X 10AB4A79X 10BB4A19X 10BB4A22X 10BB4A34X 10AB4A40X 10AB4A95X 10LA4B24X 10AB4B04X 10AB4A28X 10AB4B31X 10LA4B15X 10AB4A57X 10LA4A72X 10BB4A49X 10BB4A14X 10BB4A11X 10JA4B05X 10BB4A13X 10CB4A01X 10AB4B41X 10AB4A98X 10JA4B36X 10JA4B04X 10LA4B08X 10AB4B10X 10JA4A76X 10IA4B35X 10CB4A46X 10KA4A09X 10JA4B02X 10JA4A12X 10BB4A73X 10CB4A81X 10AB4B33X 10JA4B03X 10IA4B54X 10AB4A89X 10AB4A65X 10JA4B27X 10AB4A53X 10CB4A28X 10AB4A25X 10JA4B07X 10BB4A02X 10AB4B40X 10AB4A14X 10KA4A65X 10AB4B15X 10KA4A47X 10BB4A58X 10CB4A08X 10AB4B45X 10CB4A41X 10BB4A20X 10AB4A13X 10BB4A23X 10BB4A69X 10AB4A88X 10EA4A37X 10AB4A70X 10AB4A55X 10JA4B19X 10AB4A63X 10BB4A38X 10AB4B34X 10BB4A28X 10KA4A10X 10BB4A44X 10KA4A39X 10CB4A64X 10AB4A15X 10AB4A67X 10AB4B43X 10KA4A64X 10AB4B36X 10BB4A36X 10KA4A04X 10AB4B24X 10AB4B47X 10HZ4A27X 10AB4A37X 10AB4A42X 10CB4A58X 10AB4A16X 10BB4A65X 10AB4B07X 10CB4A19X 10CB4A37X 10BB4A05X 10CB4A84X 10AB4B14X 10AB4A97X 10BB4A43X 10AB4B03X 10LA4B19X 10CB4A61X 10EA4A41X 10CB4A39X 10AB4B48X 10CB4A23X 10CB4A35X 10CB4A45X 10CB4A12X 10CB4A07X 10JA4B49X 10CB4A74X 10BB4A53X 10AB4A49X 10CB4A14X 10CB4A47X 10CB4A76X 10IA4B43X 10CB4A44X 10CB4B72X 10CB4A67X 10CB4A59X 10IA4B46X 10JA4B08X 10AB4A54X 10CB4A29X 10CB4A43X 10CB4A24X 10AB4B29X 10CB4A42X 10AB4B37X 10CB4A77X 10AB4B11X 10CB4B64X 10CB4B62X 10CB4B74X 10CB4B50X 10AB4A41X 10CB4B68X 10AB4A08X 10CB4B47X 10FA4A45X 10CB4B35X 10CB4A17X 10JA4B34X 10CB4A71X 10BB4A01X 10BB4A15X 10LA4A86X 10BB4A07X 10CB4A79X 10BB4A29X 10CB4A85X 10CB4B49X 10BB4A64X 10AB4B49X 10CB4B76X 10DB4A25X 10CB4B59X 10KA4A18X 10CB4B37X 10CB4A52X 10HA4D77X 10JA4B01X 10CB4A06X 10DB4A45X 10CB4A69X 10CB4B69X 10AB4B27X 10BB4A12X 10CB4B75X 10CB4B70X 10DB4A24X 10AB4B42X 10CB4B71X 10CB4A40X 10CB4A48X 10BB4A61X 10DZ4B34X 10DZ4A69X 10DZ4A96X 10CB4B73X 10DB4A19X 10CB4A70X 10DB4A49X 10DB4A55X 10DB4A44X 10DB4A42X 10CB4A72X 10CB4B56X 10CB4A82X 10BB4A72X 10CB4A33X 10CB4A78X 10CB4A31X 10AB4B26X 10CB4B42X 10BB4A26X 10CB4A26X 10CB4A18X 10CB4A51X 10LA4B01X 10JA4B25X 10CB4B66X 10DB4A27X 10CB4B41X 10BB4A08X 10BB4A03X 10AB4B17X 10BB4A04X 10BB4A30X 10CB4A21X 10CB4B28X 10CB4B38X 10DB4A53X 10AB4B19X 10CB4B44X 10CB4B32X 10JA4A38X 10AB4B01X 10CB4A83X 10CB4B60X 10LA4B06X 10DB4A36X 10BB4A09X 10AB4B02X 10BB4A39X 10JA4A33X 10CB4B61X 10DB4A48X 10DB4A43X 10CB4B46X 10DB4A32X 10AB4A34X 10CB4A49X 10BB4A52X 10AB4B50X 10EB4A66X 10DB4A50X 10EB4A72X 10CB4A80X 10EB4A80X 10CB4B36X 10CB4A68X 10CB4B27X 10EB4A22X 10EB4A47X 10EB4A03X 10DB4A35X 10EB4A67X 10LA4A82X 10CB4A55X 10DB4A52X 10EB4A74X 10CB4B40X 10EB4A70X 10CB4A50X 10EB4A85X 10CB4A03X 10CB4B33X 10BB4A54X 10EB4A41X 10AB4A66X 10EB4A04X 10EB4A17X 10CB4B58X 10EB4A24X 10EB4A95X 10EB4A88X 10EB4A61X 10AB4B25X 10EB4A76X 10DB4A54X 10EB4A48X 10EB4A42X 10LA4A93X 10CB4B52X 10EA4A14X 10DB4A23X 10EB4A96X 10EB4A94X 10CB4B43X 10CB4B29X 10CB4B54X 10EB4A12X 10CB4A57X 10CB4B55X 10EB4A75X 10EB4A91X 10EB4A49X 10CB4A63X 10DB4A20X 10CB4B31X 10EB4A35X 10CB4A02X 10EB4A52X 10EB4A19X 10EB4B34X 10EB4B04X 10EB4A36X 10EB4A21X 10BB4A06X 10BB4A37X 10EB4A63X 10CB4A25X 10EB4A23X 10EB4B06X 10EB4A55X 10EB4A05X 10CB4B67X 10DB4A47X 10LA4A90X 10EB4A92X 10EB4B17X 10EB4B15X 10BB4A17X 10HA4B88X 10EB4B47X 10CB4A09X 10FB4A42X 10DB4A26X 10FB4A45X 10EB4A06X 10EB4A09X 10EB4B27X 10AB4B22X 10CB4A16X 10CB4B53X 10EB4B21X 10EB4B43X 10FB4A82X 10AB4B05X 10EB4A57X 10EB4A84X 10EB4A31X 10EB4A27X 10EB4A77X 10EB4B25X 10EB4B39X 10CB4B65X 10EB4B11X 10EB4B50X 10EB4B19X 10FB4A81X 10HA4E16X 10HA4E22X 10EB4B32X 10FB4A88X 10EB4A08X 10EB4B09X 10EB4B18X 10DB4A14X 10EB4A79X 10EB4A78X 10FB4A19X 10FB4A41X 10FB4A25X 10FB4A06X 10FB4A80X 10FB4A04X 10FB4A07X 10FB4A89X 10FB4A91X 10FB4A74X 10FB4A48X 10EB4A07X 10FB4A52X 10FB4A86X 10EB4A73X 10FB4A20X 10DA4A45X 10DB4A18X 10FB4A43X 10FB4A23X 10CB4A53X 10FB4A03X 10EB4B08X 10EB4B02X 10DB4A21X 10FB4A87X 10EB4A45X 10EB4A29X 10FB4A08X 10FB4A13X 10FB4A98X 10EB4A50X 10EB4A01X 10BB4A56X 10BB4A59X 10EB4A33X 10HA4D74X 10HA4D07X 10HA4B35X 10FB4A33X 10FB4A65X 10CB4A75X 10CB4B34X 10CB4A13X 10EB4B01X 10BB4A21X 10EB4B33X 10FB4A66X 10EB4A14X 10BB4A55X 10EB4B23X 10FB4A12X 10CB4A30X 10FB4A32X 10FB4A17X 10BB4A60X 10AB4A06X 10BB4A71X 10JA4B56X 10BB4A31X 10CB4A56X 10CB4B51X 10DB4A41X 10EB4B40X 10FB4A71X 10EB4B24X 10EB4A44X 10FB4A78X 10EB4B30X 10FB4A01X 10BB4A40X 10EB4B10X 10BB4A25X 10EB4B26X 10FB4A18X 10FB4A94X 10BB4A32X 10FB4A60X 10FB4A09X 10EB4A25X 10FB4A76X 10FB4A79X 10EB4A54X 10FB4A75X 10EB4B38X 10DB4A17X 10FB4A62X 10EB4B28X 10DB4A04X 10BB4A63X 10FB4A55X 10FB4A29X 10FB4A93X 10FB4A95X 10FB4B04X 10FB4B31X 10FB4B27X 10HA4E21X 10CB4A05X 10CB4A20X 10AB4B20X 10FB4A53X 10CB4B39X 10DB4A13X 10EB4B22X 10FB4A83X 10FB4A49X 10DB4A02X 10EB4A89X 10FB4A59X 10EB4A13X 10FB4B28X 10FB4B06X 10FB4B35X 10FB4B46X 10DB4A05X 10FB4A54X 10HA4E31X 10FB4B10X 10DB4A03X 10FB4B23X 10FB4B33X 10FB4B51X 10FB4B47X 10FB4B44X 10FB4A35X 10FB4B86X 10EB4B35X 10GB4B11X 10AB4A61X 10FB4B14X 10FB4A21X 10LA4A92X 10FB4B98X 10FB4B39X 10FB4B38X 10FB4B89X 10FB4B92X 10BB4A35X 10FB4B24X 10FB4B07X 10FB4B01X 10FB4B02X 10EB4A82X 10FB4A44X 10FB4B26X 10FB4B05X 10FB4B19X 10HA4A55X 10FB4B62X 10FB4B41X 10DB4A12X 10FB4B60X 10FB4B52X 10FB4B88X 10FB4B57X 10GB4A03X 10EB4B48X 10FB4B43X 10FB4B54X 10FB4A56X 10FB4B64X 10EB4B12X 10FB4A28X 10FB4B48X 10FA4B09X 10FB4B91X 10FB4B85X 10FB4B71X 10FB4B55X 10AB4B09X 10KA4A30X 10JA4B21X 10FB4A96X 10DA4A40X 10FB4B76X 10FB4B09X 10FB4B70X 10FB4B17X 10FB4B63X 10FB4B65X 10FB4B56X 10FB4B79X 10FB4C02X 10FB4B58X 10CB4A60X 10EB4A71X 10FB4B99X 10FB4B68X 10GB4B40X 10BB4A57X 10FB4B49X 10FB4B13X 10FB4B15X 10FB4B96X 10EB4A99X 10FB4B16X 10FB4B73X 10GB4A94X 10GB4A99X 10GB4A28X 10GB4B15X 10BB4A24X 10FB4B67X 10FB4B22X 10FB4B12X 10FB4B69X 10FB4B78X 10GB4B49X 10FB4A27X 10AB4B39X 10GB4A71X 10GB4A37X 10FB4B08X 10GB4B48X 10GB4A45X 10FB4C01X 10FB4B97X 10FB4B20X 10GB4A40X 10FB4B74X 10GB4A57X 10GB4B16X 10GB4B35X 10GB4A10X 10GB4A22X 10GB4A95X 10GB4A86X 10EB4A81X 10GB4A35X 10HA4E25X 10AB4A91X 10EB4A02X 10GB4A02X 10GB4A27X 10FB4B32X 10FB4B66X 10GB4A78X 10GB4A49X 10GB4B06X 10GB4A42X 10GB4A52X 10FB4B87X 10GB4A60X 10GB4A11X 10GB4A20X 10GB4A74X 10FB4B42X 10GB4A21X 10FB4A46X 10GB4B09X 10GB4B05X 10GB4A41X 10GB4A33X 10GB4A19X 10GB4A01X 10GB4A63X 10GB4A24X 10GB4A83X 10GB4A88X 10FB4A22X 10FB4B82X 10GB4A97X 10GB4A39X 10GB4A91X 10GB4A70X 10EB4B13X 10GB4B17X 10HB4B34X 10GB4B41X 10HB4A70X 10GB4B18X 10GB4B20X 10GB4B21X 10HB4A69X 10HB4B23X 10GB4A62X 10GB4A14X 10HB4A46X 10FB4B72X 10FB4B75X 10FB4B30X 10GB4B30X 10HB4A13X 10GB4A68X 10HB4A92X 10HB4A62X 10GB4A89X 10GB4A07X 10GB4B23X 10GB4B33X 10HB4A45X 10GB4A61X 10GB4A31X 10GB4A64X 10GB4A80X 10HB4B07X 10HB4A77X 10GB4B43X 10GB4B25X 10GB4A08X 10CB4A66X 10FB4A84X 10HB4A20X 10GB4A04X 10HA4C98X 10GB4A46X 10GB4B32X 10HB4A04X 10GB4A53X 10GB4A90X 10HB4B60X 10HB4A52X 10HB4B01X 10HB4A88X 10GB4A23X 10HB4B13X 10HB4A32X 10GB4A29X 10GB4B50X 10HB4B73X 10GB4A05X 10HB4B42X 10GB4A92X 10FB4B59X 10HB4A09X 10HB4B55X 10GB4B34X 10HB4A29X 10GB4B07X 10FB4B45X 10GB4B22X 10GB4B42X 10CB4A65X 10FB4B77X 10LA4B05X 10AB4A99X 10GB4A48X 10GB4A96X 10GB4B46X 10GB4A56X 10GB4B28X 10HB4A59X 10GB4B31X 10GB4B44X 10HB4A75X 10EB4B42X 10GB4A82X 10EB4A26X 10GB4A69X 10HB4B72X 10GB4B45X 10FB4A90X 10EB4A64X 10GB4A32X 10FB4B61X 10GB4A54X 10FB4A26X 10HB4A53X 10HB4A41X 10HB4A48X 10FB4A64X 10GB4A75X 10HB4B26X 10EB4A69X 10GB4B02X 10HB4A27X 10GB4B37X 10EB4A56X 10FB4B25X 10HB4B97X 10GB4A73X 10HB4B68X 10HB4B71X 10GB4A47X 10HB4A49X 10FB4B93X 10HB4A02X 10GB4A51X 10HB4B36X 10HB4B59X 10HB4B70X 10HB4A97X 10HB4B99X 10FB4A63X 10GB4B29X 10HA4A71X 10HB4B75X 10HB4A37X 10HB4B50X 10HB4B84X 10HB4B74X 10HB4B10X 10HB4A10X 10HB4A34X 10HB4B21X 10HB4B54X 10EB4A62X 10HB4A40X 10HB4B14X 10HB4B15X 10HB4A24X 10HB4A22X 10HB4B32X 10GB4A13X 10HB4A14X 10HB4B37X 10EB4B44X 10HB4C28X 10HB4C42X 10HB4B33X 10HB4A57X 10HB4A06X 10HB4A35X 10HB4A12X 10HB4A11X 10HB4B62X 10HB4A07X 10HB4B76X 10HB4B67X 10HB4B47X 10HB4A60X 10HB4B93X 10HB4A89X 10HB4C26X 10HB4A82X 10HB4A54X 10HB4B16X 10HB4B45X 10HB4B56X 10HB4B61X 10HA4A29X 10HB4B57X 10HB4B96X 10HB4A43X 10HB4A28X 10HA4C96X 10HB4B53X 10HB4A61X 10HB4C07X 10HB4B22X 10HB4A19X 10HB4A79X 10GB4A18X 10HB4B64X 10HB4B86X 10GB4B12X 10GB4B08X 10HB4A23X 10HB4C22X 10HB4B77X 10HB4A38X 10HB4B69X 10HB4A36X 10HB4B87X 10HB4C33X 10HB4A05X 10HB4C36X 10HB4B85X 10HB4A91X 10FB4B94X 10GB4A15X 10HB4C44X 10HB4C43X 10HB4B88X 10HB4C23X 10HB4C40X 10HB4C50X 10GB4B26X 10EB4A34X 10HB4B78X 10HB4C51X 10HB4B80X 10HB4C47X 10HB4C25X 10HB4C38X 10HB4C30X 10HB4A21X 10HB4A33X 10HA4A27X 10HB4A16X 10HB4C20X 10HB4C11X 10HB4C39X 10HB4A31X 10IB4B84X 10IB4B30X 10IB4B95X 10HB4A93X 10HB4A25X 10HB4C29X 10HB4C49X 10IB4C25X 10HB4C16X 10IB4A69X 10HB4A68X 10HB4A67X 10HB4B81X 10HB4A08X 10IB4C36X 10IB4B66X 10HB4B91X 10HB4B43X 10HB4A01X 10HB4C46X 10IB4C09X 10HB4C48X 10IB4A61X 10HB4A85X 10HB4B08X 10IB4C52X 10HB4C08X 10HB4C35X 10HB4C27X 10HB4A03X 10IB4A68X 10IB4A86X 10IB4B16X 10IB4A40X 10IB4B18X 10IB4C27X 10IB4A56X 10IB4A96X 10IB4A74X 10IB4B44X 10IB4A70X 10HB4C34X 10IB4C02X 10HB4B49X 10HB4B51X 10HB4B66X 10HB4C09X 10IB4A48X 10IB4B70X 10IB4B40X 10IB4B35X 10IB4C31X 10HB4A15X 10IB4A82X 10EB4A68X 10IB4A64X 10HB4C02X 10IB4C13X 10IB4B46X 10IB4B12X 10IB4B38X 10IB4A38X 10IB4A36X 10IB4A31X 10IB4B05X 10IB4B06X 10IB4A94X 10IB4C14X 10IB4B25X 10HB4B38X 10HB4B98X 10IB4C07X 10IB4A47X 10HB4A76X 10IB4B59X 10HB4C05X 10HB4B79X 10HB4B11X 10HB4C41X 10IB4C17X 10IB4B14X 10IB4C06X 10IB4A95X 10IB4A73X 10IB4B37X 10HB4C17X 10IB4A92X 10IB4B21X 10IB4B39X 10IB4C48X 10HB4A87X 10IB4C51X 10IB4A88X 10HB4B94X 10HB4C24X 10HB4C18X 10HB4A81X 10IB4C28X 10IB4A76X 10IB4B89X 10IB4C47X 10IB4C34X 10IB4C43X 10IB4C55X 10IB4A65X 10IB4C30X 10HB4C52X 10HB4A84X 10IB4B82X 10IB4B24X 10IB4A46X 10IB4A79X 10IB4A97X 10IB4A91X 10IB4C32X 10IB4B20X 10IB4A42X 10IB4C20X 10HB4B02X 10IB4B98X 10IB4A41X 10IB4A44X 10IB4C57X 10HB4B17X 10IB4B28X 10IB4B74X 10IB4B65X 10HB4C13X 10IB4B68X 10IB4C15X 10HB4B03X 10HB4C06X 10IB4B79X 10IB4C40X 10JB4A80X 10IB4C24X 10IB4C38X 10IB4C42X 10IB4C35X 10IB4B08X 10IB4B58X 10IB4B22X 10IB4B81X 10IB4B15X 10IB4C16X 10HB4B04X 10HB4B58X 10EB4A60X 10IB4B85X 10IB4B93X 10HB4B05X 10IB4C19X 10IB4A83X 10IB4A52X 10IB4B97X 10IB4B01X 10JB4A83X 10IB4A59X 10IB4A39X 10IB4C46X 10IB4A60X 10JB4B50X 10JB4B24X 10IB4B54X 10IB4B31X 10IB4A87X 10HB4B25X 10JB4B47X 10JB4A94X 10HB4B82X 10JB4A42X 10IB4A93X 10JB4A79X 10JB4B02X 10JB4B05X 10IB4C04X 10IB4B17X 10IB4A57X 10HB4C12X 10IB4C49X 10IB4B23X 10IB4B03X 10IB4B76X 10IB4C05X 10JB4B40X 10IB4B09X 10IB4B77X 10IB4B56X 10IB4C11X 10IB4A33X 10IB4A66X 10IB4C39X 10JB4A54X 10IB4B32X 10JB4A16X 10JB4A07X 10JB4A89X 10IB4A32X 10IB4B43X 10JB4A05X 10HB4A95X 10JB4B34X 10JB4B27X 10JB4A88X 10JB4A44X 10JB4A35X 10JB4B45X 10JB4A06X 10JB4A02X 10JB4B28X 10JB4C82X 10JB4C73X 10IB4B02X 10HB4B27X 10HB4A18X 10IB4B88X 10IB4C10X 10JB4B10X 10JB4C26X 10HB4B83X 10IB4C18X 10IB4B72X 10JB4A75X 10JB4A76X 10JB4A56X 10IB4C12X 10JB4A03X 10JB4C29X 10JB4B51X 10IB4B04X 10JB4A74X 10HB4A47X 10HB4B06X 10JB4A78X 10JB4B29X 10JB4A60X 10JB4A72X 10JB4C46X 10JB4A25X 10JB4B35X 10JB4A96X 10IB4C08X 10JB4B20X 10IB4B90X 10JB4C91X 10HB4C01X 10JB4B32X 10JB4A23X 10IB4B91X 10JB4C64X 10HB4A56X 10JB4A41X 10JB4C20X 10HB4A66X 10JB4A37X 10JB4A48X 10JB4C11X 10JB4B12X 10JB4B18X 10JB4C12X 10JB4B46X 10IB4C54X 10HB4C32X 10JB4A13X 10JB4C51X 10IB4B61X 10JB4A64X 10GB4A38X 10JB4C27X 10IB4C26X 10IB4C50X 10JB4B67X 10IB4A81X 10IB4B80X 10JB4B36X 10JB4B80X 10JB4A43X 10IB4A34X 10JB4B25X 10JB4C06X 10JB4C08X 10JB4B38X 10JB4C43X 10JB4A69X 10JB4B71X 10JB4B13X 10IB4B45X 10JB4A47X 10JB4A81X 10HB4C19X 10JB4A49X 10JB4A99X 10JB4C54X 10JB4C62X 10IB4A51X 10JB4B16X 10JB4B23X 10JB4B03X 10JB4C25X 10JB4A32X 10IB4A89X 10JB4B21X 10JB4C63X 10JB4C10X 10JB4C37X 10JB4B64X 10JB4B65X 10IB4A67X 10JB4B15X 10JB4A87X 10HB4A42X 10IB4A84X 10JB4A36X 10JB4B41X 10JB4B19X 10JB4C33X 10JB4A46X 10JB4C86X 10JB4B43X 10JB4B52X 10JB4C30X 10JB4B93X 10IB4B42X 10JB4A09X 10JB4A61X 10JB4C18X 10JB4A26X 10JB4B11X 10JB4B99X 10JB4A86X 10JB4C81X 10JB4C76X 10JB4A73X 10JB4C96X 10JB4C31X 10JB4B57X 10JB4D21X 10JB4C32X 10JB4A51X 10JB4B96X 10JB4C87X 10JB4A24X 10JB4A31X 10JB4D16X 10JB4D38X 10JB4D43X 10JB4A39X 10JB4C04X 10JB4C66X 10JB4D12X 10JB4D71X 10JB4B01X 10JB4A84X 10IB4B55X 10JB4A14X 10JB4B31X 10JB4C16X 10JB4B58X 10JB4B17X 10JB4C56X 10JB4D66X 10JB4C21X 10JB4D42X 10IB4B69X 10JB4D24X 10JA4A54X 10JB4B61X 10JB4B07X 10JA4A25X 10JB4C84X 10JB4D18X 10JB4D68X 10JB4D23X 10KB4A35X 10JB4C39X 10JB4D11X 10JB4D29X 10JB4A58X 10JB4D39X 10JB4B79X 10IB4A62X 10IB4A80X 10JB4D08X 10HB4C15X 10IB4A63X 10KB4A03X 10JB4A50X 10JB4D25X 10JB4A66X 10KB4A62X 10KB4A46X 10JB4A12X 10JB4A38X 10JB4D49X 10KB4A18X 10JB4C13X 10JB4B66X 10KB4A06X 10IB4A77X 10JB4B56X 10JB4D06X 10JB4B92X 10KB4A19X 10JB4D22X 10KB4A40X 10JB4D60X 10JB4D55X 10KB4A78X 10KB4A13X 10JB4D46X 10JB4D77X 10JB4D05X 10JB4D76X 10JB4A20X 10JB4C38X 10KB4A74X 10KB4A38X 10KB4A57X 10KB4A39X 10JB4C98X 10KB4A36X 10JB4B54X 10KB4A07X 10KB4A12X 10JB4D37X 10KB4A47X 10JB4D62X 10KB4A22X 10JB4D75X 10JB4D73X 10KB4A16X 10JB4A18X 10JB4D32X 10KB4A10X 10IB4A54X 10KB4A61X 10KB4A04X 10KB4B25X 10KB4A28X 10JB4A15X 10JB4D34X 10JB4D54X 10KB4B09X 10KB4A17X 10KB4B13X 10KB4A09X 10JB4C42X 10KB4B34X 10KB4B14X 10JB4C50X 10KB4B01X 10JB4A19X 10KB4A63X 10KB4B33X 10JB4B84X 10JB4B90X 10KB4B21X 10JB4D45X 10KB4A73X 10JB4C57X 10JB4D51X 10KB4B42X 10JB4B91X 10EB4B46X 10KB4B19X 10KB4B61X 10IB4B27X 10IB4B13X 10JB4A55X 10KB4A58X 10JB4D52X 10KB4A94X 10JB4D50X 10KB4B72X 10JB4B30X 10KB4A76X 10JB4D65X 10JB4A29X 10JB4D59X 10KB4B17X 10JB4D07X 10JB4C17X 10JB4B62X 10JB4A52X 10IB4A35X 10KB4A82X 10JB4B68X 10JB4C36X 10KB4B66X 10KB4A49X 10KB4A91X 10KB4B29X 10JB4B63X 10KB4B56X 10KB4B58X 10JB4C23X 10IB4B67X 10JB4C44X 10JB4D41X 10KB4A27X 10KB4B39X 10KB4A15X 10KB4B51X 10KB4B16X 10HB4C21X 10HB4A51X 10JB4D48X 10JB4C40X 10IB4A75X 10KB4B40X 10KB4A66X 10KB4A95X 10JB4A95X 10JB4B26X 10JB4C49X 10KB4A01X 10JB4C24X 10KB4B41X 10JB4D31X 10KB4A87X 10KB4A77X 10KB4B06X 10JB4A93X 10KB4B55X 10HB4A72X 10KB4B73X 10HB4A64X 10KB4B67X 10LB4A04X 10KB4A79X 10KB4A84X 10LB4A29X 10LB4A15X 10LB4A64X 10KB4B48X 10JB4B08X 10JB4D04X 10IB4C45X 10JB4C59X 10JB4D58X 10HB4C03X 10KB4B76X 10KB4B43X 10IB4C21X 10KB4A21X 10KB4A89X 10JB4A53X 10JB4B77X 10JB4C71X 10JB4D57X 10JB4C58X 10JB4D63X 10LB4A82X 10LB4A39X 10LB4A24X 10LB4A14X 10JB4A65X 10KB4A51X 10JB4C94X 10JB4C19X 10KB4A97X 10LB4A69X 10KB4A42X 10LB4A80X 10LB4A05X 10JB4C03X 10KB4B24X 10AC4A03X 10JB4D78X 10AC4A09X 10LB4A11X 10KB4B27X 10LB4A01X 10KB4A56X 10AC4B03X 10KB4B64X 10LB4A85X 10KB4B07X 10LB4A32X 10KB4A44X 10LB4A13X 10JB4C48X 10LB4A34X 10JB4A04X 10LB4A54X 10KB4A02X 10IB4B86X 10KB4B37X 10LB4A98X 10LB4A51X 10KB4B32X 10JB4D30X 10KB4B12X 10AC4A32X 10LB4A40X 10IB4B71X 10IB4A55X 10LB4A86X 10LB4A16X 10LB4A48X 10LB4A63X 10KB4B44X 10AC4A98X 10KB4A32X 10JB4C65X 10LB4A38X 10LB4A55X 10AC4B05X 10LB4A37X 10LB4A18X 10JB4B85X 10KB4B69X 10KB4A67X 10HB4A78X 10IB4A58X 10LB4A91X 10LB4A23X 10KB4B23X 10AC4A82X 10KB4B53X 10LB4A19X 10KB4A71X 10JB4C34X 10LB4A02X 10AC4A48X 10HB4 |  | SoftwareChangeControl
47 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178927 | Z-1123-2020 | US Clinical 20mL Programmable Pump. Catalog No. 01827 The device is an implantable infusion pump. GTIN-14 00810335020082 | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system. | Class 2 Device Recall Prometra Programmable Infusion Pump Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prometra Programmable Infusion Pump FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prometra Programmable Infusion Pump Date Initiated by Firm December 24, 2019 Date Posted February 10, 2020 Recall Status 1 Open 3 , Classified Recall Number Z-1123-2020 Recall Event ID 84653 PMA Number P080012 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product US Clinical 20mL Programmable Pump. Catalog No. 01827 The device is an implantable infusion pump. GTIN-14 00810335020082 Code Information All units with pump firmware version 0.26 manufactured from 2012 until the present time. Serial Numbers 10HW4A57 10AX4A06 36ES4A15 36AR4A04 36AR4A06 36AR4A13 36HS4A10 36HS4A54 36JN4A04 36JN4A06 36JN4A07 36JN4A09 36JN4A13 36JN4A14 36JN4A15 36JN4A18 36JN4A23 36JN4A28 36JN4A35 36KN4A11 36KN4A16 36KN4A23 36KN4A27 36KN4A33 36KN4A34 36KN4A41 36HS4A01 10IT4A21 36HS4A14 36HS4A26 36BS4A17 36ES4A28 36ES4A33 36ES4A08 10IT4A16 36FR4A13 36HS4A61 10ET4A07 36HS4A46 10ET4A20 36HS4A25 36HS4A43 10ET4A02 Recalling Firm/ Manufacturer Flowonix Medical Inc 120 Forbes Blvd Ste 170 Mansfield MA 02048-1150 For Additional Information Contact Technical Solutions Department 855-356-9665 Manufacturer Reason for Recall A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system. FDA Determined Cause 2 Software change control Action ***Updated 5/12/2023******On May 11, 2023, the firm sent an URGENT MEDICAL DEVICE CORRECTION UPDATE letter to all affected customers, which informed them of an update to Flowonix's plan to correct the firmware issue.******Flowonix will NOT be conducting an office-based pump firmware update for affected pumps implanted prior to January 2020.******Healthcare providers were provided ongoing recommendations for Prometra and Prometra II pumps implanted prior to January 2020, which are impacted by this issue.******Healthcare providers should contact their Flowonix Representative or our Technical Solutions Department (1-855-356-9665) in the event cessation of therapy or premature low battery is suspected and/or have general questions regarding this notice.***************On December 24, 2019, the firm notified affected users via an Urgent Medical Device Corrective Action Recall Notice.******The firm advised customers of the product issue and reminded customers that the pumps are currently labeled only for use with Infumorph. Due to the low risk that this failure mode may occur, Flowonix is not recommending that pumps be explanted. Healthcare providers were recommended to inform patients regarding the symptoms and adverse health events associated with an unexpected pump stoppage and the appropriate steps to take in the event they suspect a stoppage.******The firm will be implementing a correction of firmware errors to bring the pumps back into specification. The firm will follow up when the change has been completed. An office-based pump firmware update is being developed for patients currently implanted with the affected pumps. (Note, as of the 5/12/23 update, implanted pumps will not be receiving a firmware update.) Quantity in Commerce 6670 (US) Distribution Nationwide domestic distribution. Foreign distribution to United Kingdom, Republic of Ireland, Poland, Greece, Italy, Germany, Netherlands, and Belgium. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = LKK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action ***Updated 5/12/2023******On May 11, 2023, the firm sent an URGENT MEDICAL DEVICE CORRECTION UPDATE letter to all affected customers, which informed them of an update to Flowonix's plan to correct the firmware issue.******Flowonix will NOT be conducting an office-based pump firmware update for affected pumps implanted prior to January 2020.******Healthcare providers were provided ongoing recommendations for Prometra and Prometra II pumps implanted prior to January 2020, which are impacted by this issue.******Healthcare providers should contact their Flowonix Representative or our Technical Solutions Department (1-855-356-9665) in the event cessation of therapy or premature low battery is suspected and/or have general questions regarding this notice.***************On December 24, 2019, the firm notified affected users via an Urgent Medical Device Corrective Action Recall Notice.******The firm advised customers of the product issue and reminded customers that the pumps are currently labeled only for use with Infumorph. Due to the low risk that this failure mode may occur, Flowonix is not recommending that pumps be explanted. Healthcare providers were recommended to inform patients regarding the symptoms and adverse health events associated with an unexpected pump stoppage and the appropriate steps to take in the event they suspect a stoppage.******The firm will be implementing a correction of firmware errors to bring the pumps back into specification. The firm will follow up when the change has been completed. An office-based pump firmware update is being developed for patients currently implanted with the affected pumps. (Note, as of the 5/12/23 update, implanted pumps will not be receiving a firmware update.) | SoftwareChangeControl
48 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=176760 | Z-0354-2020 | EndoTool SubQ (Versions 1.7.1, 1.7.4, 1.7.5) | EndoTool SubQ | 2 | CDRH | 2019-11-08 |  |  | Monarch Medical Technologies | Product was distributed prior to approval or clearance from FDA. | Class 2 Device Recall EndoTool SubQ Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall EndoTool SubQ FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall EndoTool SubQ Date Initiated by Firm October 01, 2019 Create Date November 08, 2019 Recall Status 1 Open 3 , Classified Recall Number Z-0354-2020 Recall Event ID 83926 Product Classification Calculator, drug dose - Product Code NDC Product EndoTool SubQ (Versions 1.7.1, 1.7.4, 1.7.5) Code Information versions: v 1.7.1, 1.7.4, 1.7.5 Recalling Firm/ Manufacturer Monarch Medical Technologies 2137 South Blvd Ste 300 Charlotte NC 28203-5189 For Additional Information Contact Bruce Lisante 855-363-7475 Manufacturer Reason for Recall Product was distributed prior to approval or clearance from FDA. FDA Determined Cause 2 Software change control Action The firm held meetings and emailed the consignees beginning on 09/27/2019 and followed with a letter. The notice stated that no medical decision should be based solely on the recommended guidance provided by the software program and requested the facility take additional precautions when following EndoTool SubQ dose recommendations. Quantity in Commerce 6 units Distribution SC, IL IN, CA Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm held meetings and emailed the consignees beginning on 09/27/2019 and followed with a letter. The notice stated that no medical decision should be based solely on the recommended guidance provided by the software program and requested the facility take additional precautions when following EndoTool SubQ dose recommendations. | SoftwareChangeControl
49 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=176305 | Z-0260-2020 | Dario Blood Glucose Monitoring System, Model No. 1015-05, running Android software versions 4.3.0-4.3.2 | Dario Blood Glucose Monitoring System | 2 | CDRH | 2019-10-31 |  | 2020-05-05 13:58 | LABSTYLE INNOVATIONS | The Dario-Blood Glucose Tracker Android App versions 4.3.0-4.3.2 may experience duplicate logging of a blood glucose level reading. | Class 2 Device Recall Dario Blood Glucose Monitoring System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Dario Blood Glucose Monitoring System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Dario Blood Glucose Monitoring System Date Initiated by Firm August 29, 2019 Create Date October 31, 2019 Recall Status 1 Terminated 3 on May 05, 2020 Recall Number Z-0260-2020 Recall Event ID 83761 510(K)Number K150817 Product Classification System, test, blood glucose, over the counter - Product Code NBW Product Dario Blood Glucose Monitoring System, Model No. 1015-05, running Android software versions 4.3.0-4.3.2 Code Information Android App Software versions 4.3.0-4.3.2 Recalling Firm/ Manufacturer LABSTYLE INNOVATIONS 7, Derekh Menakhem Begin Ramat Gan Israel For Additional Information Contact Customer Service 800-8955921 Manufacturer Reason for Recall The Dario-Blood Glucose Tracker Android App versions 4.3.0-4.3.2 may experience duplicate logging of a blood glucose level reading. FDA Determined Cause 2 Software change control Action The firm released Android App version 4.3.3 on August 29, 2019. Users were informed about the issue via push notification on 8/30/19, asking to update to the new version. An additional push notification was sent on 8/31/19, along with a Field Safety Notice letter sent to their "inbox". On 9/9/19, remaining potential relevant users received a third notification in which they were asked to upgrade their version through in-app and push notifications.************Customers were instructed to do the following:******Please update the latest Dario - Blood Glucose Tracker Android App version as follows******o Open the Google Play store******o Search for Dario Blood Glucose Tracker App******o Press Uninstall button******o Press OK to confirm the question "Do you want to uninstall this app?"******o Once completed, press Install and verify you installed 4.3.3 version (or above) by pressing the "About this App" link.************Customers with questions should contact the firm's Customer Service Center at 1-800-895-5921, Monday *** Friday, 9AM *** 5PM Eastern. Quantity in Commerce 126,271 (83,141 US; 43,130 OUS) Distribution Nationwide domestic distribution. Foreign distribution to Australia, Canada, UK, and Italy. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = NBW - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm released Android App version 4.3.3 on August 29, 2019. Users were informed about the issue via push notification on 8/30/19, asking to update to the new version. An additional push notification was sent on 8/31/19, along with a Field Safety Notice letter sent to their "inbox". On 9/9/19, remaining potential relevant users received a third notification in which they were asked to upgrade their version through in-app and push notifications.************Customers were instructed to do the following:******Please update the latest Dario - Blood Glucose Tracker Android App version as follows******o Open the Google Play store******o Search for Dario Blood Glucose Tracker App******o Press Uninstall button******o Press OK to confirm the question "Do you want to uninstall this app?"******o Once completed, press Install and verify you installed 4.3.3 version (or above) by pressing the "About this App" link.************Customers with questions should contact the firm's Customer Service Center at 1-800-895-5921, Monday *** Friday, 9AM *** 5PM Eastern. | SoftwareChangeControl
50 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175374 | Z-0171-2020 | Plum 360 Infusion System With MedNet/ Smart Card Plug And Play Module, Model # 30010. Infusion pump. | Plum 360 Infusion System with MedNet/ Smart Card Plug n Play Module | 2 | CDRH | 2019-10-22 | 2019-10-22 21:00 | 2021-04-27 16:12 | ICU Medical Inc | A software malfunction of the infuser when used with a MedNet Custom Drug Library (CDL). When a new CDL is downloaded and installed on the infuser the infuser can become inoperable and show a blank screen on the User Interface. | Class 2 Device Recall Plum 360 Infusion System with MedNet/ Smart Card Plug n Play Module Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Plum 360 Infusion System with MedNet/ Smart Card Plug n Play Module FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Plum 360 Infusion System with MedNet/ Smart Card Plug n Play Module Date Initiated by Firm August 01, 2018 Date Posted October 22, 2019 Recall Status 1 Terminated 3 on April 27, 2021 Recall Number Z-0171-2020 Recall Event ID 83480 510(K)Number K161469 Product Classification Pump, infusion - Product Code FRN Product Plum 360 Infusion System With MedNet/ Smart Card Plug And Play Module, Model # 30010. Infusion pump. Code Information All Plum 360 Infusers with software versions 15.10.00.010/.011/.012 and 15.11.00.017 when used with an ICU Medical MedNet CDL Recalling Firm/ Manufacturer ICU Medical Inc 600 N Field Dr Lake Forest IL 60045-4835 Manufacturer Reason for Recall A software malfunction of the infuser when used with a MedNet Custom Drug Library (CDL). When a new CDL is downloaded and installed on the infuser the infuser can become inoperable and show a blank screen on the User Interface. FDA Determined Cause 2 Software change control Action On 8/1/2018 a "Urgent Medical Device Correction" letter was mailed to consignees via traceable mailing. The medical device correction letter informed consignees of the following:************1. To reduce the likelihood of the issue occurring at the bedside, you can consider not utilizing the option to "Force Pump to Accept Drug Library During Power Down" when finalizing the CDL with MedNet. In doing so, when powering down the infuser the clinician will have the option to install the CDL or not, based on the clinical situation. ************2. ICU Medical will release a software update to correct the issue and will contact you when the update is available. ************3. . Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular************4.Please inform potential users in your organization of this notification, complete the attached response form and return it to the fax number (1-888-628-6849 or e-mail address (ICUMedical3618@stericycle.com) on the form, even if you do not have the affected product. ************5. For further inquiries, please contact ICU Medical using the information provided below:************Global Complaint Management - 1-844-654-7780 or productcomplaintsPP@icumed.com******Technical Support Center - 1-800-241-4002, option 4 or TSC.support@icumed.com Quantity in Commerce 47,152 devices Distribution US nationwide distribution, and Canada Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = FRN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 8/1/2018 a "Urgent Medical Device Correction" letter was mailed to consignees via traceable mailing. The medical device correction letter informed consignees of the following:************1. To reduce the likelihood of the issue occurring at the bedside, you can consider not utilizing the option to "Force Pump to Accept Drug Library During Power Down" when finalizing the CDL with MedNet. In doing so, when powering down the infuser the clinician will have the option to install the CDL or not, based on the clinical situation. ************2. ICU Medical will release a software update to correct the issue and will contact you when the update is available. ************3. . Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular************4.Please inform potential users in your organization of this notification, complete the attached response form and return it to the fax number (1-888-628-6849 or e-mail address (ICUMedical3618@stericycle.com) on the form, even if you do not have the affected product. ************5. For further inquiries, please contact ICU Medical using the information provided below:************Global Complaint Management - 1-844-654-7780 or productcomplaintsPP@icumed.com******Technical Support Center - 1-800-241-4002, option 4 or TSC.support@icumed.com | SoftwareChangeControl
51 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172822 | Z-1786-2019 | Monarch Medical Technologies EndoTool IV, Versions 1.8, Version 1.8.5 and higher, and Version 1.9 and Version 1.10. | Monarch Medical Technologies EndoTool IV | 2 | CDRH | 2019-06-15 |  |  | Monarch Medical Technologies | Insulin dosing calculations were erroneously high. | Class 2 Device Recall Monarch Medical Technologies EndoTool IV Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Monarch Medical Technologies EndoTool IV FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Monarch Medical Technologies EndoTool IV Date Initiated by Firm May 03, 2019 Create Date June 15, 2019 Recall Status 1 Open 3 , Classified Recall Number Z-1786-2019 Recall Event ID 82828 510(K)Number K132547 Product Classification Calculator, drug dose - Product Code NDC Product Monarch Medical Technologies EndoTool IV, Versions 1.8, Version 1.8.5 and higher, and Version 1.9 and Version 1.10. Code Information Versions 1.8, Version 1.8.5 and higher, and Version 1.9 and Version 1.10 Recalling Firm/ Manufacturer Monarch Medical Technologies 2137 South Blvd Ste 300 Charlotte NC 28203-5189 For Additional Information Contact Bruce Lisanti 855-363-7475 Manufacturer Reason for Recall Insulin dosing calculations were erroneously high. FDA Determined Cause 2 Software change control Action The firm notified the consignees by email beginning on 05/03/2019. The notice alerted the consignee of the problem and provided an educational brief and a tip card to provide to the staff to remind them to always use clinical judgement in following EndoTool IV dose recommendations. Quantity in Commerce 83 downloads Distribution Nationwide and Dubai Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = NDC - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm notified the consignees by email beginning on 05/03/2019. The notice alerted the consignee of the problem and provided an educational brief and a tip card to provide to the staff to remind them to always use clinical judgement in following EndoTool IV dose recommendations. | SoftwareChangeControl
52 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172979 | Z-1715-2019 | Prometra II Programmable Pump, Cat. No. 13827 For intrathecal infusion. | Prometra Programmable Pump System | 2 | CDRH | 2019-06-03 |  | 2020-06-05 09:23 | Flowonix Medical Inc | The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. Error 115 can also occur if dendrites form or if the pump is dropped. | Class 2 Device Recall Prometra Programmable Pump System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prometra Programmable Pump System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prometra Programmable Pump System Date Initiated by Firm March 21, 2019 Create Date June 03, 2019 Recall Status 1 Terminated 3 on June 05, 2020 Recall Number Z-1715-2019 Recall Event ID 82586 PMA Number P080012 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product Prometra II Programmable Pump, Cat. No. 13827 For intrathecal infusion. Code Information Serial Numbers: 10AA4A01X 10BC4A84X 10CC4C06X 10EB4A27X 10GB4A05X 10HA4E28X 10IA4B69X 10JA4B52X 10KB4A76X 10AA4A02X 10BC4A86X 10CC4C07X 10EB4A28X 10GB4A07X 10HA4E29X 10IA4B70X 10JA4B53X 10KB4A77X 10AA4A03X 10BC4A87X 10CC4C08X 10EB4A29X 10GB4A08X 10HA4E30X 10IA4B71X 10JA4B54X 10KB4A78X 10AA4A04X 10BC4A88X 10CC4C09X 10EB4A30X 10GB4A10X 10HA4E31X 10IA4B72X 10JA4B55X 10KB4A79X 10AA4A05X 10BC4A89X 10CC4C10X 10EB4A31X 10GB4A11X 10HB4A01X 10IA4B73X 10JA4B56X 10KB4A80X 10AA4A06X 10BC4A90X 10CC4C11X 10EB4A33X 10GB4A13X 10HB4A02X 10IA4B74X 10JA4B57X 10KB4A81X 10AA4A07X 10BC4A92X 10CC4C12X 10EB4A34X 10GB4A14X 10HB4A03X 10IA4B75X 10JA4B58X 10KB4A82X 10AA4A08X 10BC4A94X 10CC4C13X 10EB4A35X 10GB4A15X 10HB4A04X 10IA4B76X 10JA4B59X 10KB4A84X 10AA4A09X 10BC4A96X 10CC4C14X 10EB4A36X 10GB4A17X 10HB4A05X 10IB4A31X 10JA4B60X 10KB4A86X 10AA4A10X 10BC4A97X 10CC4C15X 10EB4A39X 10GB4A18X 10HB4A06X 10IB4A32X 10JA4B61X 10KB4A87X 10AA4A11X 10BC4A98X 10CC4C16X 10EB4A40X 10GB4A19X 10HB4A07X 10IB4A33X 10JB4A01X 10KB4A88X 10AA4A12X 10BC4A99X 10CC4C17X 10EB4A41X 10GB4A20X 10HB4A08X 10IB4A34X 10JB4A02X 10KB4A89X 10AA4A13X 10BC4B01X 10CC4C18X 10EB4A42X 10GB4A21X 10HB4A09X 10IB4A35X 10JB4A03X 10KB4A90X 10AA4A14X 10BC4B02X 10CC4C19X 10EB4A43X 10GB4A22X 10HB4A10X 10IB4A36X 10JB4A04X 10KB4A91X 10AA4A15X 10BC4B04X 10CC4C20X 10EB4A44X 10GB4A23X 10HB4A11X 10IB4A38X 10JB4A05X 10KB4A93X 10AA4A16X 10BC4B05X 10CC4C21X 10EB4A45X 10GB4A24X 10HB4A12X 10IB4A39X 10JB4A06X 10KB4A94X 10AA4A17X 10BC4B06X 10CC4C22X 10EB4A46X 10GB4A27X 10HB4A13X 10IB4A40X 10JB4A07X 10KB4A95X 10AA4A18X 10BC4B07X 10CC4C23X 10EB4A47X 10GB4A28X 10HB4A14X 10IB4A41X 10JB4A08X 10KB4A96X 10AA4A19X 10BC4B08X 10CC4C24X 10EB4A48X 10GB4A29X 10HB4A15X 10IB4A42X 10JB4A09X 10KB4A97X 10AA4A20X 10BC4B09X 10CC4C25X 10EB4A49X 10GB4A30X 10HB4A16X 10IB4A43X 10JB4A10X 10KB4A99X 10AA4A32X 10BC4B22X 10CC4C37X 10EB4A63X 10GB4A44X 10HB4A28X 10IB4A58X 10JB4A23X 10KB4B18X 10AA4A33X 10BC4B24X 10CC4C38X 10EB4A64X 10GB4A45X 10HB4A29X 10IB4A59X 10JB4A24X 10KB4B19X 10AA4A34X 10BC4B25X 10CC4C39X 10EB4A65X 10GB4A46X 10HB4A30X 10IB4A60X 10JB4A25X 10KB4B21X 10AA4A35X 10BC4B26X 10CC4C40X 10EB4A66X 10GB4A47X 10HB4A31X 10IB4A61X 10JB4A26X 10KB4B22X 10AA4A36X 10BC4B27X 10CC4C41X 10EB4A67X 10GB4A48X 10HB4A32X 10IB4A62X 10JB4A27X 10KB4B23X 10AA4A37X 10BC4B28X 10CC4C42X 10EB4A68X 10GB4A49X 10HB4A33X 10IB4A63X 10JB4A29X 10KB4B24X 10AA4A38X 10BC4B29X 10CC4C43X 10EB4A69X 10GB4A51X 10HB4A34X 10IB4A64X 10JB4A30X 10KB4B25X 10AA4A39X 10BC4B30X 10CC4C44X 10EB4A70X 10GB4A52X 10HB4A35X 10IB4A65X 10JB4A31X 10KB4B26X 10AA4A40X 10BC4B31X 10CC4C45X 10EB4A71X 10GB4A53X 10HB4A36X 10IB4A66X 10JB4A32X 10KB4B27X 10AA4A41X 10BC4B32X 10CC4C46X 10EB4A72X 10GB4A54X 10HB4A37X 10IB4A67X 10JB4A33X 10KB4B28X 10AA4A42X 10BC4B33X 10CC4C47X 10EB4A73X 10GB4A55X 10HB4A38X 10IB4A68X 10JB4A34X 10KB4B29X 10AA4A43X 10BC4B34X 10CC4C48X 10EB4A74X 10GB4A56X 10HB4A39X 10IB4A69X 10JB4A35X 10KB4B30X 10AA4A44X 10BC4B35X 10CC4C49X 10EB4A75X 10GB4A57X 10HB4A40X 10IB4A70X 10JB4A36X 10KB4B31X 10AA4A45X 10BC4B36X 10CC4C50X 10EB4A76X 10GB4A60X 10HB4A41X 10IB4A71X 10JB4A37X 10KB4B32X 10AA4A46X 10BC4B37X 10CC4C51X 10EB4A77X 10GB4A61X 10HB4A42X 10IB4A72X 10JB4A38X 10KB4B33X 10AA4A47X 10BC4B38X 10CC4C52X 10EB4A78X 10GB4A62X 10HB4A43X 10IB4A73X 10JB4A39X 10KB4B34X 10AA4A48X 10BC4B39X 10CZ4A76X 10EB4A79X 10GB4A63X 10HB4A44X 10IB4A74X 10JB4A40X 10KB4B35X 10AA4A49X 10BC4B40X 10CZ4A77X 10EB4A80X 10GB4A64X 10HB4A45X 10IB4A75X 10JB4A41X 10KB4B36X 10AA4A50X 10BC4B41X 10CZ4A78X 10EB4A81X 10GB4A65X 10HB4A46X 10IB4A76X 10JB4A42X 10KB4B37X 10AA4A51X 10BC4B42X 10CZ4A79X 10EB4A82X 10GB4A66X 10HB4A47X 10IB4A77X 10JB4A43X 10KB4B38X 10AA4A52X 10BC4B43X******10CZ4A80X 10EB4A84X 10GB4A67X 10HB4A48X 10IB4A79X 10JB4A44X 10KB4B39X 10AA4A53X 10BC4B44X 10CZ4A81X 10EB4A85X 10GB4A68X 10HB4A49X 10IB4A80X 10JB4A45X 10KB4B40X 10AA4A54X 10BC4B46X 10CZ4A82X 10EB4A86X 10GB4A69X 10HB4A51X 10IB4A81X 10JB4A46X 10KB4B41X 10AA4A55X 10BC4B52X 10CZ4A83X 10EB4A88X 10GB4A70X 10HB4A52X 10IB4A82X 10JB4A47X 10KB4B42X 10AA4A56X 10BC4B53X 10CZ4A84X 10EB4A89X 10GB4A71X 10HB4A53X 10IB4A83X 10JB4A48X 10KB4B43X 10AA4A57X 10BC4B55X 10CZ4A85X 10EB4A91X 10GB4A72X 10HB4A54X 10IB4A84X 10JB4A49X 10KB4B44X 10AA4A58X 10BC4B56X 10CZ4A86X 10EB4A92X 10GB4A73X 10HB4A55X 10IB4A86X 10JB4A50X 10KB4B45X 10AA4A59X 10BC4B57X 10CZ4A87X 10EB4A93X 10GB4A74X 10HB4A56X 10IB4A87X 10JB4A51X 10KB4B46X 10AA4A60X 10BC4B58X 10CZ4A88X 10EB4A94X 10GB4A75X 10HB4A57X 10IB4A88X 10JB4A52X 10KB4B47X 10AA4A61X 10BC4B59X 10CZ4A89X 10EB4A95X 10GB4A76X 10HB4A59X 10IB4A89X 10JB4A53X 10KB4B48X 10AA4A62X 10BC4B60X 10CZ4A90X 10EB4A96X 10GB4A78X 10HB4A60X 10IB4A90X 10JB4A54X 10KB4B49X 10AA4A63X 10BC4B61X 10CZ4A91X 10EB4A99X 10GB4A79X 10HB4A61X 10IB4A91X 10JB4A55X 10KB4B50X 10AA4A64X 10BC4B62X 10CZ4A92X 10EB4B01X 10GB4A80X 10HB4A62X 10IB4A92X 10JB4A56X 10KB4B51X 10AA4A65X 10BC4B63X 10CZ4A93X 10EB4B02X 10GB4A81X 10HB4A63X 10IB4A93X 10JB4A58X 10KB4B52X 10AA4A66X 10BC4B64X 10CZ4A94X 10EB4B04X 10GB4A82X 10HB4A64X 10IB4A94X 10JB4A59X 10KB4B53X 10AA4A67X 10BC4B65X 10CZ4A95X 10EB4B05X 10GB4A83X 10HB4A65X 10IB4A95X 10JB4A60X 10KB4B54X 10AA4A68X 10BC4B66X 10CZ4A96X 10EB4B06X 10GB4A84X 10HB4A66X 10IB4A96X 10JB4A61X 10KB4B55X 10AA4A69X 10BC4B67X 10CZ4A97X 10EB4B07X 10GB4A85X 10HB4A67X 10IB4A97X 10JB4A62X 10KB4B56X 10AA4A70X 10BC4B68X 10CZ4A98X 10EB4B08X 10GB4A86X 10HB4A68X 10IB4A98X 10JB4A63X 10KB4B57X 10AA4A71X 10BC4B69X 10CZ4A99X 10EB4B09X 10GB4A87X 10HB4A69X 10IB4A99X 10JB4A64X 10KB4B58X 10AA4A72X 10BC4B70X 10CZ4B01X 10EB4B10X 10GB4A88X 10HB4A70X 10IB4B01X 10JB4A65X 10KB4B59X 10AA4A73X 10BC4B71X 10CZ4B02X 10EB4B11X 10GB4A89X 10HB4A71X 10IB4B02X 10JB4A66X 10KB4B60X 10AA4A74X 10BC4B72X 10CZ4B03X 10EB4B12X 10GB4A90X 10HB4A72X 10IB4B03X 10JB4A67X 10KB4B61X 10AA4A75X 10BC4B73X 10CZ4B04X 10EB4B13X 10GB4A91X 10HB4A73X 10IB4B04X 10JB4A68X 10KB4B62X 10AA4A76X 10BC4B74X 10CZ4B05X 10EB4B14X 10GB4A92X 10HB4A75X 10IB4B05X 10JB4A69X 10KB4B63X 10AA4A77X 10BC4B75X 10CZ4B06X 10EB4B15X 10GB4A93X 10HB4A76X 10IB4B06X 10JB4A72X 10KB4B64X 10AA4A78X 10BC4B76X 10CZ4B07X 10EB4B17X 10GB4A94X 10HB4A77X 10IB4B07X 10JB4A73X 10KB4B66X 10AA4A79X 10BC4B77X 10CZ4B08X 10EB4B18X 10GB4A95X 10HB4A78X 10IB4B08X 10JB4A74X 10KB4B67X 10AA4A80X 10BC4B78X 10CZ4B09X 10EB4B19X 10GB4A96X 10HB4A79X 10IB4B09X 10JB4A75X 10KB4B68X 10AA4A81X 10BC4B79X 10CZ4B10X 10EB4B20X 10GB4A97X 10HB4A80X 10IB4B11X 10JB4A76X 10KB4B69X 10AA4A82X 10BC4B80X 10CZ4B11X 10EB4B21X 10GB4A99X 10HB4A81X 10IB4B12X 10JB4A77X 10KB4B70X 10AA4A83X 10BC4B81X 10CZ4B12X 10EB4B22X 10GB4B01X 10HB4A82X 10IB4B13X 10JB4A78X 10KB4B71X 10AA4A84X 10BC4B82X 10CZ4B13X 10EB4B23X 10GB4B02X 10HB4A84X 10IB4B14X 10JB4A79X 10KB4B72X 10AA4A85X 10BC4B83X 10CZ4B14X 10EB4B24X 10GB4B05X 10HB4A85X 10IB4B15X 10JB4A80X 10KB4B73X 10AA4A86X 10BC4B84X 10CZ4B15X 10EB4B25X 10GB4B06X 10HB4A86X 10IB4B16X 10JB4A81X 10KB4B74X 10AA4A87X 10BC4B85X 10CZ4B16X 10EB4B26X 10GB4B07X 10HB4A87X 10IB4B17X 10JB4A82X 10KB4B75X 10AA4A88X 10BC4B86X 10CZ4B17X 10EB4B27X 10GB4B08X 10HB4A88X 10IB4B18X 10JB4A83X 10KB4B76X 10AA4A90X 10BC4B87X 10CZ4B18X 10EB4B28X 10GB4B09X 10HB4A89X 10IB4B19X 10JB4A84X 10KC4A00X 10AB4A01X 10BC4B88X 10CZ4B19X 10EB4B30X 10GB4B10X 10HB4A91X 10IB4B20X 10JB4A85X 10KC4A01X 10AB4A02X 10BC4B89X 10CZ4B20X 10EB4B32X 10GB4B11X 10HB4A92X 10IB4B21X 10JB4A86X 10KC4A02X 10AB4A03X 10BC4B90X 10CZ4B21X 10EB4B33X 10GB4B12X 10HB4A93X 10IB4B22X 10JB4A87X 10KC4A03X 10AB4A04X 10BC4B91X 10CZ4B22X 10EB4B34X 10GB4B14X 10HB4A95X 10IB4B23X 10JB4A88X 10KC4A04X 10AB4A05X 10BC4B92X 10CZ4B23X 10EB4B35X 10GB4B15X 10HB4A96X 10IB4B24X 10JB4A89X 10KC4A06X 10AB4A06X 10BC4B93X 10CZ4B24X 10EB4B36X 10GB4B16X 10HB4A97X 10IB4B25X 10JB4A91X 10KC4A07X 10AB4A07X 10BC4B94X 10CZ4B25X 10EB4B37X 10GB4B17X 10HB4A98X 10IB4B27X 10JB4A92X 10KC4A08X 10AB4A08X 10BC4B95X 10CZ4B26X 10EB4B38X 10GB4B18X 10HB4A99X 10IB4B28X 10JB4A93X 10KC4A09X 10AB4A09X 10BC4B96X 10CZ4B27X 10EB4B39X 10GB4B20X 10HB4B01X 10IB4B29X 10JB4A94X 10KC4A10X 10AB4A10X 10BC4B97X 10CZ4B28X 10EB4B40X 10GB4B21X 10HB4B02X 10IB4B30X 10JB4A95X 10KC4A11X 10AB4A11X 10BC4B98X 10CZ4B29X 10EB4B42X 10GB4B22X 10HB4B03X 10IB4B31X 10JB4A96X 10KC4A12X 10AB4A12X 10BC4B99X 10CZ4B30X 10EB4B43X 10GB4B23X 10HB4B04X 10IB4B32X 10JB4A97X 10KC4A13X 10AB4A13X 10BC4C01X 10CZ4B31X 10EB4B44X 10GB4B24X 10HB4B05X 10IB4B33X 10JB4A98X 10KC4A14X 10AB4A14X 10BC4C02X 10CZ4B32X 10EB4B45X 10GB4B25X 10HB4B06X 10IB4B34X 10JB4A99X 10KC4A15X 10AB4A15X 10BC4C03X 10CZ4B33X 10EB4B46X 10GB4B26X 10HB4B07X 10IB4B35X 10JB4B01X 10KC4A16X 10AB4A16X 10BC4C04X 10CZ4B34X 10EB4B47X 10GB4B28X 10HB4B08X 10IB4B36X 10JB4B02X 10KC4A17X 10AB4A17X 10BC4C05X 10CZ4B35X 10EB4B48X 10GB4B29X 10HB4B09X 10IB4B37X 10JB4B03X 10KC4A18X 10AB4A19X 10BC4C06X 10CZ4B36X 10EB4B50X 10GB4B30X 10HB4B10X 10IB4B38X 10JB4B04X 10KC4A19X 10AB4A20X 10BC4C07X 10CZ4B57X 10EC4A01X 10GB4B31X 10HB4B11X 10IB4B39X 10JB4B05X 10KC4A20X 10AB4A21X 10BC4C08X 10DA4A01X 10EC4A02X 10GB4B32X 10HB4B13X 10IB4B40X 10JB4B06X 10KC4A21X 10AB4A22X 10BC4C09X 10DA4A02X 10EC4A03X 10GB4B33X 10HB4B14X 10IB4B41X 10JB4B07X 10KC4A22X 10AB4A24X 10BC4C10X 10DA4A03X 10EC4A04X 10GB4B34X 10HB4B15X 10IB4B42X 10JB4B08X 10KC4A23X 10AB4A25X 10BC4C11X 10DA4A04X 10EC4A05X 10GB4B35X 10HB4B16X 10IB4B43X 10JB4B09X 10KC4A24X 10AB4A26X 10BC4C12X 10DA4A05X 10EC4A06X 10GB4B36X 10HB4B17X 10IB4B44X 10JB4B10X 10KC4A25X 10AB4A27X 10BC4C13X 10DA4A06X 10EC4A07X 10GB4B37X 10HB4B18X 10IB4B45X 10JB4B11X 10KC4A26X 10AB4A28X 10BC4C14X 10DA4A07X 10EC4A08X 10GB4B38X 10HB4B19X 10IB4B46X 10JB4B12X 10KC4A27X 10AB4A29X 10BC4C15X 10DA4A08X 10EC4A09X 10GB4B39X 10HB4B21X 10IB4B47X 10JB4B13X 10KC4A28X 10AB4A30X 10BC4C16X 10DA4A09X 10EC4A10X 10GB4B40X 10HB4B22X 10IB4B48X 10JB4B14X 10KC4A29X 10AB4A31X 10BC4C17X 10DA4A11X 10EC4A11X 10GB4B41X 10HB4B23X 10IB4B50X 10JB4B15X 10KC4A30X 10AB4A32X 10BC4C18X 10DA4A12X 10EC4A12X 10GB4B42X 10HB4B24X 10IB4B51X 10JB4B16X 10KC4A31X 10AB4A33X 10BC4C19X 10DA4A13X 10EC4A13X 10GB4B43X 10HB4B25X 10IB4B52X 10JB4B17X 10KC4A32X 10AB4A34X 10BC4C20X 10DA4A14X 10EC4A14X 10GB4B44X 10HB4B26X 10IB4B54X 10JB4B18X 10KC4A33X 10AB4A35X 10BC4C21X 10DA4A15X 10EC4A15X 10GB4B45X 10HB4B27X 10IB4B55X 10JB4B19X 10KC4A34X 10AB4A36X 10BC4C22X 10DA4A16X 10EC4A16X 10GB4B46X 10HB4B28X 10IB4B56X 10JB4B20X 10KC4A35X 10AB4A37X 10BC4C23X 10DA4A17X 10EC4A17X 10GB4B47X 10HB4B30X 10IB4B57X 10JB4B21X 10KC4A36X 10AB4A38X 10BC4C25X 10DA4A18X 10EC4A18X 10GB4B48X 10HB4B32X 10IB4B58X 10JB4B22X 10KC4A37X 10AB4A39X 10BC4C26X 10DA4A19X 10EC4A19X 10GB4B49X 10HB4B33X 10IB4B59X 10JB4B23X 10KC4A38X 10AB4A40X 10BC4C27X 10DA4A20X 10EC4A20X 10GB4B50X 10HB4B34X 10IB4B60X 10JB4B24X 10KC4A39X 10AB4A41X 10BC4C28X 10DA4A21X 10EC4A21X 10GC4A26X 10HB4B36X 10IB4B61X 10JB4B25X 10KC4A40X 10AB4A42X 10BC4C29X 10DA4A22X 10EC4A22X 10GC4A27X 10HB4B37X 10IB4B62X 10JB4B26X 10KC4A41X 10AB4A43X 10BC4C30X 10DA4A23X 10EC4A23X 10GC4A28X 10HB4B38X 10IB4B64X 10JB4B27X 10KC4A42X 10AB4A44X 10BC4C31X 10DA4A24X 10EC4A24X 10GC4A29X 10HB4B42X 10IB4B65X 10JB4B28X 10KC4A43X 10AB4A45X 10BC4C32X 10DA4A25X 10EC4A25X 10GC4A30X 10HB4B43X 10IB4B66X 10JB4B29X 10KC4A44X 10AB4A46X 10BC4C33X 10DA4A26X 10EC4A26X 10GC4A31X 10HB4B45X 10IB4B67X 10JB4B30X 10KC4A45X 10AB4A47X 10BC4C34X 10DA4A27X 10EC4A27X 10GC4A32X 10HB4B46X 10IB4B68X 10JB4B31X 10KC4A46X 10AB4A48X 10BC4C35X 10DA4A28X 10EC4A29X 10GC4A33X 10HB4B47X 10IB4B69X 10JB4B32X 10KC4A47X 10AB4A49X 10BC4C36X 10DA4A29X 10EC4A30X 10GC4A34X 10HB4B49X 10IB4B70X 10JB4B33X 10KC4A48X 10AB4A50X 10BC4C38X 10DA4A30X 10EC4A31X 10GC4A35X 10HB4B50X 10IB4B71X 10JB4B34X 10KC4A49X 10AB4A51X 10BC4C39X 10DA4A31X 10EC4A32X 10GC4A36X 10HB4B51X 10IB4B72X 10JB4B35X 10KC4A50X 10AB4A53X 10BC4C40X 10DA4A32X 10EC4A33X 10GC4A37X 10HB4B53X 10IB4B73X 10JB4B36X 10KC4A51X 10AB4A54X 10BC4C41X 10DA4A33X 10EC4A34X 10GC4A38X 10HB4B54X 10IB4B74X 10JB4B37X 10KC4A52X 10AB4A55X 10BC4C42X 10DA4A34X 10EC4A35X 10GC4A39X 10HB4B55X 10IB4B75X 10JB4B38X 10KC4A53X 10AB4A57X 10BC4C43X 10DA4A35X 10EC4A36X 10GC4A40X 10HB4B56X 10IB4B76X 10JB4B39X 10KC4A54X 10AB4A58X 10BC4C44X 10DA4A36X 10EC4A37X 10GC4A41X 10HB4B57X 10IB4B77X 10JB4B40X 10KC4A55X 10AB4A59X 10BC4C46X 10DA4A37X 10EC4A38X 10GC4A42X 10HB4B58X 10IB4B78X 10JB4B41X 10KC4A56X 10AB4A60X 10BC4C47X 10DA4A38X 10EC4A39X 10GC4A43X 10HB4B59X 10IB4B79X 10JB4B42X 10KC4A57X 10AB4A61X 10BC4C48X 10DA4A39X 10EC4A40X 10GC4A44X 10HB4B60X 10IB4B80X 10JB4B43X 10KC4A58X 10AB4A63X 10BC4C49X 10DA4A40X 10EC4A41X 10GC4A45X 10HB4B61X 10IB4B81X 10JB4B44X 10KC4A59X 10AB4A65X 10BC4C50X 10DA4A41X 10EC4A42X 10GC4A46X 10HB4B62X 10IB4B82X 10JB4B45X 10KC4A60X 10AB4A66X 10BC4C51X 10DA4A42X 10EC4A43X 10GC4A47X 10HB4B63X 10IB4B84X 10JB4B46X 10KC4A61X 10AB4A67X 10BC4C52X 10DA4A43X 10EC4A44X 10GC4A48X 10HB4B64X 10IB4B85X 10JB4B47X 10KC4A62X 10AB4A69X 10BC4C53X 10DA4A44X 10EC4A45X 10GC4A49X 10HB4B65X 10IB4B86X 10JB4B48X 10KC4A63X 10AB4A70X 10BC4C54X 10DA4A45X 10EC4A46X 10GC4A50X 10HB4B66X 10IB4B88X 10JB4B49X 10KC4A64X 10AB4A71X 10BC4C55X 10DA4A46X 10EC4A47X 10GC4A51X 10HB4B67X 10IB4B89X 10JB4B50X 10KC4A65X 10AB4A73X 10BC4C57X 10DA4A47X 10EC4A48X 10GC4A52X 10HB4B68X 10IB4B90X 10JB4B51X 10KC4A66X 10AB4A74X 10BC4C58X 10DA4A48X 10EC4A49X 10GC4A53X 10HB4B69X 10IB4B91X 10JB4B52X 10KC4A67X 10AB4A77X 10BC4C59X 10DA4A49X 10EC4A50X 10GC4A54X 10HB4B70X 10IB4B92X 10JB4B54X 10KC4A68X 10AB4A78X 10BC4C60X 10DA4A50X 10EC4A51X 10GC4A55X 10HB4B71X 10IB4B93X 10JB4B55X 10KC4A70X 10AB4A79X 10BC4C61X 10DA4A51X 10EC4A52X 10GC4A57X 10HB4B72X 10IB4B94X 10JB4B56X 10KC4A71X 10AB4A80X 10BC4C62X 10DA4A52X 10EC4A53X 10GC4A58X 10HB4B73X 10IB4B95X 10JB4B57X 10KC4A72X 10AB4A81X 10BC4C63X 10DA4A53X 10EC4A54X 10GC4A59X 10HB4B74X 10IB4B96X 10JB4B58X 10KC4A73X 10AB4A83X 10BC4C64X 10DA4A54X 10EC4A55X 10GC4A60X 10HB4B75X 10IB4B97X 10JB4B59X 10KC4A74X 10AB4A86X 10BC4C65X 10DA4A55X 10EC4A56X 10GC4A61X 10HB4B76X 10IB4B98X 10JB4B60X 10KC4A75X 10AB4A87X 10BC4C66X 10DA4A56X 10EC4A57X 10GC4A62X 10HB4B77X 10IB4B99X 10JB4B61X 10KC4A76X 10AB4A88X 10BC4C67X 10DA4A57X 10EC4A58X 10GC4A63X 10HB4B78X 10IB4C01X 10JB4B62X 10KC4A77X 10AB4A89X 10BC4C68X 10DA4A58X 10EC4A59X 10GC4A64X 10HB4B79X 10IB4C02X 10JB4B63X 10KC4A78X 10AB4A91X 10BC4C69X 10DA4A59X 10EC4A60X 10GC4A65X 10HB4B80X 10IB4C03X 10JB4B64X 10KC4A79X 10AB4A92X 10BC4C70X 10DA4A60X 10EC4A61X 10GC4A67X 10HB4B81X 10IB4C04X 10JB4B65X 10KC4A80X 10AB4A95X 10BC4C71X 10DA4A61X 10EC4A62X 10GC4A68X 10HB4B82X 10IB4C05X 10JB4B66X 10KC4A81X 10AB4A96X 10BC4C72X 10DA4A62X 10EC4A63X 10GC4A69X 10HB4B83X 10IB4C06X 10JB4B67X 10KC4A82X 10AB4A97X 10BC4C73X 10DA4A63X 10EC4A64X 10GC4A70X 10HB4B84X 10IB4C07X 10JB4B68X 10KC4A83X 10AB4A98X 10BC4C74X 10DA4A64X 10EC4A65X 10GC4A71X 10HB4B85X 10IB4C08X 10JB4B69X 10KC4A84X 10AB4A99X 10BC4C75X 10DA4A65X 10EC4A67X 10GC4A72X 10HB4B86X 10IB4C09X 10JB4B70X 10KC4A85X 10AB4B01X 10BC4C76X 10DA4A66X 10EC4A68X 10GC4A73X 10HB4B87X 10IB4C10X 10JB4B71X 10KC4A86X 10AB4B02X 10BC4C77X 10DA4A67X 10EC4A69X 10GC4A74X 10HB4B88X 10IB4C11X 10JB4B72X 10KC4A87X 10AB4B03X 10BC4C78X 10DA4A68X 10EC4A70X 10GC4A75X 10HB4B89X 10IB4C12X 10JB4B73X 10KC4A88X 10AB4B04X 10BC4C79X 10DA4A69X 10EC4A71X 10GC4A76X 10HB4B91X 10IB4C13X 10JB4B74X 10KC4A89X 10AB4B05X 10BC4C80X 10DA4A70X 10EC4A72X 10GC4A77X 10HB4B93X 10IB4C14X 10JB4B76X 10KC4A90X 10AB4B06X 10BC4C81X 10DA4A71X 10EC4A73X 10GC4A78X 10HB4B94X 10IB4C15X 10JB4B77X 10KC4A91X 10AB4B07X 10BC4C82X 10DA4A72X 10EC4A74X 10GC4A79X 10HB4B95X 10IB4C16X 10JB4B78X 10KC4A92X 10AB4B08X 10BC4C83X 10DA4A73X 10EC4A75X 10GC4A80X 10HB4B96X 10IB4C17X 10JB4B79X 10KC4A93X 10AB4B09X 10BC4C84X 10DA4A74X 10EC4A76X 10GC4A81X 10HB4B97X 10IB4C18X 10JB4B80X 10KC4A94X 10AB4B10X 10BC4C85X 10DA4A75X 10EC4A77X 10GC4A82X 10HB4B98X 10IB4C19X 10JB4B81X 10KC4A95X 10AB4B11X 10BC4C86X 10DB4A02X 10EC4A78X 10GC4A83X 10HB4B99X 10IB4C20X 10JB4B82X 10KC4A96X 10AB4B12X 10BC4C87X 10DB4A03X 10EC4A79X 10GC4A84X 10HB4C01X 10IB4C21X 10JB4B83X 10KC4A97X 10AB4B13X 10BC4C88X 10DB4A04X 10EC4A80X 10GC4A85X 10HB4C02X 10IB4C22X 10JB4B84X 10KC4A98X 10AB4B14X 10BC4C89X 10DB4A05X 10EC4A81X 10GC4A86X 10HB4C03X 10IB4C23X 10JB4B85X 10KC4A99X 10AB4B15X 10BC4C90X 10DB4A06X 10EC4A82X 10GC4A87X 10HB4C04X 10IB4C24X 10JB4B86X 10KC4B00X 10AB4B16X 10BC4C91X 10DB4A12X 10EC4A83X 10GC4A88X 10HB4C05X 10IB4C25X 10JB4B87X 10KC4B01X 10AB4B17X 10BC4C92X 10DB4A13X 10EC4A84X 10GC4A89X 10HB4C06X 10IB4C26X 10JB4B88X 10KC4B02X 10AB4B19X 10BC4C93X 10DB4A14X 10EC4A85X 10GC4A90X 10HB4C07X 10IB4C27X 10JB4B90X 10KC4B03X 10AB4B20X 10BC4C94X 10DB4A17X 10EC4A86X 10GC4A91X 10HB4C08X 10IB4C28X 10JB4B91X 10KC4B04X 10AB4B21X 10BC4C95X 10DB4A18X 10EC4A87X 10GC4A92X 10HB4C09X 10IB4C29X 10JB4B92X 10KC4B05X 10AB4B22X 10BC4C96X 10DB4A19X 10EC4A88X 10GC4A93X 10HB4C10X 10IB4C30X 10JB4B93X 10KC4B06X 10AB4B23X 10BC4C97X 10DB4A20X 10EC4A89X 10GC4A94X 10HB4C11X 10IB4C31X 10JB4B94X 10KC4B07X 10AB4B24X 10BD4A00X 10DB4A21X 10EC4A90X 10GC4A95X 10HB4C12X 10IB4C32X 10JB4B95X 10KC4B08X 10AB4B25X 10BD4A01X 10DB4A23X 10EC4A91X 10GC4A96X 10HB4C13X 10IB4C33X 10JB4B96X 10KC4B09X 10AB4B26X 10BD4A02X 10DB4A24X 10EC4A92X 10GC4A97X 10HB4C15X 10IB4C34X 10JB4B97X 10KC4B10X 10AB4B27X 10BD4A03X 10DB4A25X 10EC4A93X 10GC4A98X 10HB4C16X 10IB4C35X 10JB4B98X 10KC4B12X 10AB4B29X 10BD4A04X 10DB4A26X 10EC4A94X 10GC4A99X 10HB4C17X 10IB4C36X 10JB4B99X 10KC4B13X 10AB4B30X 10BD4A05X 10DB4A27X 10EC4A95X 10GC4B01X 10HB4C18X 10IB4C37X 10JB4C03X 10KC4B14X 10AB4B31X 10BD4A06X 10DB4A28X 10EC4A96X 10GC4B02X 10HB4C19X 10IB4C38X 10JB4C04X 10KC4B15X 10AB4B33X 10BD4A07X 10DB4A32X 10EC4A97X 10GC4B03X 10HB4C20X 10IB4C39X 10JB4C06X 10KC4B16X 10AB4B34X 10BD4A08X 10DB4A35X 10EC4A98X 10GC4B04X 10HB4C21X 10IB4C40X 10JB4C07X 10KC4B17X 10AB4B35X 10BD4A09X 10DB4A36X 10EC4A99X 10GC4B05X 10HB4C22X 10IB4C41X 10JB4C08X 10KC4B18X 10AB4B36X 10BD4A10X 10DB4A37X 10EC4B01X 10GC4B06X 10HB4C23X 10IB4C42X 10JB4C09X 10KC4B19X 10AB4B37X 10BD4A11X 10DB4A38X 10EZ4A01X 10GC4B07X 10HB4C24X 10IB4C43X 10JB4C10X 10KC4B20X 10AB4B38X 10BD4A12X 10DB4A41X 10EZ4A02X 10GC4B08X 10HB4C25X 10IB4C44X 10JB4C11X 10KC4B21X 10AB4B39X 10BD4A13X 10DB4A42X 10EZ4A03X 10GC4B09X 10HB4C26X 10IB4C45X 10JB4C12X 10KC4B22X 10AB4B40X 10BD4A14X 10DB4A43X 10EZ4A04X 10GC4B10X 10HB4C27X 10IB4C46X 10JB4C13X 10KC4B23X 10AB4B41X 10BD4A15X 10DB4A44X 10EZ4A05X 10GC4B11X 10HB4C28X 10IB4C47X 10JB4C14X 10KC4B24X 10AB4B42X 10BD4A16X 10DB4A45X 10EZ4A06X 10GC4B13X 10HB4C29X 10IB4C48X 10JB4C15X 10KC4B25X 10AB4B43X 10BD4A17X 10DB4A46X 10EZ4A07X 10GC4B14X 10HB4C30X 10IB4C49X 10JB4C16X 10KC4B26X 10AB4B44X 10BD4A18X 10DB4A47X 10EZ4A08X 10GC4B15X 10HB4C31X 10IB4C50X 10JB4C17X 10KC4B27X 10AB4B45X 10BD4A19X 10DB4A48X 10EZ4A09X 10GC4B16X 10HB4C32X 10IB4C51X 10JB4C18X 10KC4B28X 10AB4B46X 10BD4A20X 10DB4A49X 10EZ4A10X 10GC4B17X 10HB4C33X 10IB4C52X 10JB4C19X 10KC4B29X 10AB4B47X 10BD4A21X 10DB4A50X 10EZ4A11X 10GC4B18X 10HB4C34X 10IB4C53X 10JB4C20X 10KC4B31X 10AB4B48X 10BD4A22X 10DB4A51X 10EZ4A12X 10GC4B19X 10HB4C35X 10IB4C54X 10JB4C21X 10KC4B32X 10AB4B49X 10BD4A23X 10DB4A52X 10EZ4A13X 10GC4B20X 10HB4C36X 10IB4C55X 10JB4C22X 10KC4B33X 10AB4B50X 10BD4A24X 10DB4A53X 10EZ4A16X 10GC4B21X 10HB4C37X 10IB4C56X 10JB4C23X 10KC4B35X 10AB4B51X 10BD4A25X 10DB4A54X 10EZ4A17X 10GC4B22X 10HB4C38X 10IB4C57X 10JB4C24X 10KC4B36X 10AC4A01X 10BD4A26X 10DB4A55X 10EZ4A18X 10GC4B23X 10HB4C39X 10IC4A00X 10JB4C25X 10KC4B37X 10AC4A02X 10BD4A27X 10DC4A01X 10EZ4A19X 10GC4B24X 10HB4C40X 10IC4A01X 10JB4C26X 10KC4B39X 10AC4A03X 10BD4A28X 10DC4A02X 10EZ4A20X 10GC4B25X 10HB4C41X 10IC4A02X 10JB4C27X 10KC4B40X 10AC4A04X 10BD4A29X 10DC4A03X 10EZ4A21X 10GC4B26X 10HB4C42X 10IC4A03X 10JB4C28X 10KC4B41X 10AC4A05X 10BD4A30X 10DC4A04X 10EZ4A22X 10GZ4A01X 10HB4C43X 10IC4A04X 10JB4C29X 10KC4B42X 10AC4A09X 10BD4A31X 10DC4A05X 10EZ4A23X 10GZ4A02X 10HB4C44X 10IC4A05X 10JB4C30X 10KC4B43X 10AC4A13X 10BD4A32X 10DC4A06X 10EZ4A24X 10GZ4A03X 10HB4C45X 10IC4A06X 10JB4C31X 10KC4B44X 10AC4A15X 10BD4A33X 10DC4A07X 10EZ4A25X 10GZ4A04X 10HB4C46X 10IC4A07X 10JB4C32X 10KC4B45X 10AC4A16X 10BD4A34X 10DC4A08X 10EZ4A26X 10GZ4A05X 10HB4C47X 10IC4A08X 10JB4C33X 10KC4B46X 10AC4A19X 10BD4A35X 10DC4A09X 10EZ4A27X 10GZ4A06X 10HB4C48X 10IC4A09X 10JB4C34X 10KC4B47X 10AC4A20X 10BD4A36X 10DC4A10X 10EZ4A28X 10GZ4A07X 10HB4C49X 10IC4A10X 10JB4C35X 10KC4B48X 10AC4A23X 10BD4A37X 10DC4A11X 10EZ4A29X 10GZ4A08X 10HB4C50X 10IC4A11X 10JB4C36X 10KC4B49X 10AC4A24X 10BD4A38X 10DC4A12X 10EZ4A30X 10GZ4A09X 10HB4C51X 10IC4A12X 10JB4C37X 10KC4B50X 10AC4A25X 10BD4A39X 10DC4A13X 10EZ4A31X 10GZ4A10X 10HB4C52X 10IC4A13X 10JB4C38X 10KC4B51X 10AC4A26X 10BD4A40X 10DC4A14X 10EZ4A32X 10GZ4A11X 10HC4A00X 10IC4A14X 10JB4C39X 10KC4B52X 10AC4A28X 10BD4A41X 10DC4A15X 10EZ4A33X 10GZ4A12X 10HC4A01X 10IC4A15X 10JB4C40X 10KC4B53X 10AC4A29X 10BD4A42X 10DC4A16X 10EZ4A34X 10GZ4A13X 10HC4A02X 10IC4A16X 10JB4C41X 10KC4B54X 10AC4A30X 10BD4A43X 10DC4A17X 10EZ4A35X 10GZ4A15X 10HC4A03X 10IC4A17X 10JB4C42X 10KC4B55X 10AC4A31X 10BD4A45X 10DC4A18X 10EZ4A36X 10GZ4A16X 10HC4A04X 10IC4A18X 10JB4C43X 10KC4B56X 10AC4A32X 10BD4A46X 10DC4A19X 10EZ4A37X 10GZ4A17X 10HC4A05X 10IC4A19X 10JB4C44X 10KC4B57X 10AC4A35X 10BD4A47X 10DC4A20X 10EZ4A38X 10GZ4A18X 10HC4A06X 10IC4A20X 10JB4C45X 10KC4B58X 10AC4A36X 10BD4A48X 10DC4A21X 10EZ4A39X 10GZ4A19X 10HC4A07X 10IC4A21X 10JB4C46X 10KC4B59X 10AC4A37X 10BD4A49X 10DC4A22X 10EZ4A40X 10GZ4A20X 10HC4A08X 10IC4A22X 10JB4C47X 10KC4B60X 10AC4A38X 10BD4A51X 10DC4A23X 10EZ4A41X 10GZ4A21X 10HC4A09X 10IC4A23X 10JB4C48X 10KC4B61X 10AC4A39X 10BD4A52X 10DC4A24X 10EZ4A42X 10GZ4A22X 10HC4A10X 10IC4A25X 10JB4C49X 10KC4B62X 10AC4A40X 10BD4A53X 10DC4A25X 10EZ4A43X 10GZ4A24X 10HC4A12X 10IC4A26X 10JB4C50X 10KC4B63X 10AC4A41X 10BD4A54X 10DC4A26X 10EZ4A44X 10GZ4A25X 10HC4A13X 10IC4A27X 10JB4C51X 10KC4B64X 10AC4A43X 10BD4A55X 10DC4A27X 10EZ4A45X 10GZ4A26X 10HC4A14X 10IC4A28X 10JB4C52X 10KC4B65X 10AC4A44X 10BD4A56X 10DC4A28X 10EZ4A46X 10GZ4A27X 10HC4A15X 10IC4A29X 10JB4C53X 10KC4B66X 10AC4A45X 10BD4A57X 10DC4A29X 10EZ4A47X 10GZ4A28X 10HC4A16X 10IC4A30X 10JB4C54X 10KC4B67X 10AC4A46X 10BD4A58X 10DC4A30X 10EZ4A48X 10GZ4A29X 10HC4A17X 10IC4A31X 10JB4C55X 10KC4B68X 10AC4A48X 10BD4A59X 10DC4A31X 10EZ4A49X 10GZ4A30X 10HC4A18X 10IC4A32X 10JB4C56X 10KC4B69X 10AC4A50X 10BD4A60X 10DC4A32X 10EZ4A50X 10GZ4A31X 10HC4A19X 10IC4A33X 10JB4C57X 10KC4B70X 10AC4A51X 10BD4A61X 10DC4A33X 10EZ4A51X 10GZ4A32X 10HC4A20X 10IC4A34X 10JB4C58X 10KC4B71X 10AC4A52X 10BD4A62X 10DC4A34X 10EZ4A52X 10GZ4A33X 10HC4A21X 10IC4A35X 10JB4C59X 10KC4B72X 10AC4A53X 10BD4A63X 10DC4A35X 10EZ4A53X 10GZ4A34X 10HC4A22X 10IC4A36X 10JB4C61X 10KC4B73X 10AC4A56X 10BD4A73X 10DC4A36X 10EZ4A55X 10GZ4A35X 10HC4A23X 10IC4A37X 10JB4C62X 10KC4B74X 10AC4A57X 10BD4A75X 10DC4A37X 10EZ4A56X 10GZ4A36X 10HC4A24X 10IC4A38X 10JB4C63X 10KC4B75X 10AC4A58X 10BD4A78X 10DC4A38X 10EZ4A57X 10GZ4A37X 10HC4A25X 10IC4A39X 10JB4C64X 10KC4B76X 10AC4A61X 10BD4A84X 10DC4A39X 10EZ4A58X 10GZ4A38X 10HC4A26X 10IC4A40X 10JB4C65X 10KC4B77X 10AC4A62X 10BD4A87X 10DC4A40X 10EZ4A59X 10GZ4A39X 10HC4A27X 10IC4A41X 10JB4C66X 10KC4B78X 10AC4A65X 10BD4A88X 10DC4A41X 10EZ4A60X 10GZ4A40X 10HC4A28X 10IC4A42X 10JB4C67X 10KC4B79X 10AC4A66X 10BD4A89X 10DC4A42X 10EZ4Z74X 10GZ4A41X 10HC4A29X 10IC4A43X 10JB4C68X 10KC4B80X 10AC4A70X 10BD4A91X 10DC4A43X 10FA4A02 10GZ4A42X 10HC4A30X 10IC4A44X 10JB4C69X 10KC4B81X 10AC4A71X 10BD4A92X 10DC4A44X 10FA4A31X 10GZ4A43X 10HC4A31X 10IC4A46X 10JB4C70X 10KC4B82X 10AC4A72X 10BD4A93X 10DC4A45X 10FA4A32X 10GZ4A44X 10HC4A32X 10IC4A47X 10JB4C71X 10KC4B83X 10AC4A78X 10BD4A94X 10DC4A46X 10FA4A33X 10GZ4A45X 10HC4A33X 10IC4A49X 10JB4C72X 10KC4B84X 10AC4A79X 10BD4A96X 10DC4A47X 10FA4A34X 10GZ4A46X 10HC4A34X 10IC4A50X 10JB4C73X 10KC4B85X 10AC4A80X 10BD4A98X 10DC4A48X 10FA4A35X 10GZ4A47X 10HC4A35X 10IC4A51X 10JB4C75X 10KC4B86X 10AC4A81X 10BD4A99X 10DC4A49X 10FA4A36X 10GZ4A48X 10HC4A36X 10IC4A52X 10JB4C76X 10KC4B87X 10AC4A82X 10BD4B00X 10DC4A50X 10FA4A37X 10GZ4A49X 10HC4A37X 10IC4A53X 10JB4C77X 10KC4B88X 10AC4A83X 10BD4B01X 10DC4A51X 10FA4A38X 10GZ4A50X 10HC4A38X 10IC4A54X 10JB4C78X 10KC4B89X 10AC4A85X 10BD4B02X 10DC4A52X 10FA4A39X 10GZ4A51X 10HC4A39X 10IC4A55X 10JB4C79X 10KC4B90X 10AC4A86X 10BD4B03X 10DC4A53X 10FA4A40X 10GZ4A52X 10HC4A40X 10IC4A56X 10JB4C80X 10KC4B91X 10AC4A88X 10BD4B05X 10DC4A54X 10FA4A41X 10GZ4A53X 10HC4A41X 10IC4A57X 10JB4C81X 10KC4B92X 10AC4A89X 10BD4B06X 10DC4A55X 10FA4A42X 10GZ4A54X 10HC4A42X 10IC4A58X 10JB4C82X 10KC4B93X 10AC4A90X 10BD4B07X 10DC4A56X 10FA4A43X 10GZ4A55X 10HC4A43X 10IC4A59X 10JB4C84X 10KC4B94X 10AC4A94X 10BZ4A09X 10DC4A57X 10FA4A44X 10GZ4A56X 10HC4A44X 10IC4A60X 10JB4C85X 10KC4B95X 10AC4A97X 10BZ4A11X 10DC4A58X 10FA4A45X 10GZ4A57X 10HC4A45X 10IC4A61X 10JB4C86X 10KC4B96X 10AC4A98X 10BZ4A12X 10DC4A59X 10FA4A46X 10GZ4A58X 10HC4A46X 10IC4A62X 10JB4C87X 10KC4B97X 10AC4A99X 10BZ4A14X 10DC4A60X 10FA4A47X 10GZ4A59X 10HC4A47X 10IC4A63X 10JB4C88X 10KC4B98X 10AC4B02X 10BZ4A16X 10DC4A61X 10FA4A48X 10GZ4A61X 10HC4A48X 10IC4A64X 10JB4C89X 10KC4B99X 10AC4B03X 10BZ4A173X 10DC4A62X 10FA4A49X 10GZ4A62X 10HC4A49X 10IC4A65X 10JB4C90X 10KC4C00X 10AC4B05X 10BZ4A17X 10DC4A63X 10FA4A50X 10GZ4A63X 10HC4A50X 10IC4A66X 10JB4C91X 10KC4C01X 10AC4B07X 10BZ4A18X 10DC4A64X 10FA4A51X 10GZ4A64X 10HC4A51X 10IC4A67X 10JB4C92X 10KC4C02X 10AC4B08X 10BZ4A25X 10DC4A65X 10FA4A52X 10GZ4A66X 10HC4A52X 10IC4A68X 10JB4C93X 10KC4C03X 10AC4B11X 10BZ4A30X 10DC4A66X 10FA4A53X 10GZ4A68X 10HC4A53X 10IC4A69X 10JB4C94X 10KC4C05X 10AC4B17X 10BZ4A36X 10DC4A67X 10FA4A54X 10GZ4A69X 10HC4A54X 10IC4A70X 10JB4C95X 10KC4C06X 10AC4B18X 10BZ4A37X 10DC4A68X 10FA4A55X 10GZ4A70X 10HC4A55X 10IC4A71X 10JB4C96X 10KC4C07X 10AC4B19X 10BZ4A39X 10DC4A69X 10FA4A56X 10GZ4A71X 10HC4A56X 10IC4A72X 10JB4C97X 10KC4C08X 10AC4B22X 10BZ4A41X 10DC4A70X 10FA4A57X 10GZ4A72X 10HC4A57X 10IC4A73X 10JB4C98X 10KC4C09X 10AC4B23X 10BZ4A44X 10DC4A71X 10FA4A58X 10GZ4A73X 10HC4A58X 10IC4A74X 10JB4D01X 10KC4C11X 10AC4B24X 10BZ4A53X 10DC4A72X 10FA4A59X 10GZ4A74X 10HC4A59X 10IC4A75X 10JB4D02X 10KC4C12X 10AC4B28X 10BZ4A55X 10DC4A73X 10FA4A61X 10GZ4A75X 10HC4A60X 10IC4A76X 10JB4D04X 10KC4C13X 10AC4B29X 10BZ4A59X 10DC4A74X 10FA4A62X 10GZ4A76X 10HC4A61X 10IC4A77X 10JB4D05X 10KC4C14X 10AC4B31X 10BZ4A60X 10DC4A75X 10FA4A63X 10GZ4A77X 10HC4A62X 10IC4A78X 10JB4D06X 10KC4C16X 10AC4B32X 10BZ4A61X 10DC4A76X 10FA4A64X 10GZ4A78X 10HC4A63X 10IC4A79X 10JB4D07X 10KC4C17X 10AC4B34X 10BZ4A64X 10DC4A77X 10FA4A65X 10GZ4A79X 10HC4A64X 10IC4A80X 10JB4D08X 10KC4C18X 10AC4B35X 10BZ4A65X 10DC4A78X 10FA4A66X 10GZ4A80X 10HC4A65X 10IC4A81X 10JB4D09X 10KC4C19X 10AC4B37X 10BZ4A66X 10DC4A79X 10FA4A67X 10GZ4A81X 10HC4A66X 10IC4A82X 10JB4D10X 10KC4C21X 10AC4B38X 10BZ4A73X 10DC4A80X 10FA4A68X 10GZ4A82X 10HC4A67X 10IC4A83X 10JB4D11X 10KC4C22X 10AC4B39X 10BZ4A74X 10DC4A81X 10FA4A69X 10GZ4A83X 10HC4A68X 10IC4A84X 10JB4D12X 10KC4C23X 10AC4B40X 10BZ4A75X 10DC4A82X 10FA4A70X 10GZ4A84X 10HC4A69X 10IC4A85X 10JB4D13X 10KC4C24X 10AC4B42X 10CA4A01X 10DC4A83X 10FA4A71X 10GZ4A85X 10HC4A70X 10IC4A86X 10JB4D14X 10KZ4A01X 10AC4B44X 10CA4A02X 10DC4A84X 10FA4A72X 10GZ4A86X 10HC4A71X 10IC4A87X 10JB4D15X 10KZ4A02X 10AC4B45X 10CA4A03X 10DC4A85X 10FA4A73X 10GZ4A87X 10HC4A72X 10IC4A88X 10JB4D16X 10KZ4A03X 10AC4B48X 10CA4A04X 10DC4A86X 10FA4A74X 10GZ4A88X 10HC4A73X 10IC4A89X 10JB4D17X 10KZ4A04X 10AC4B49X 10CA4A05X 10DC4A87X 10FA4A75X 10GZ4A89X 10HC4A74X 10IC4A90X 10JB4D18X 10KZ4A05X 10AC4B50X 10CA4A07X 10DC4A88X 10FA4A76X 10GZ4A90X 10HC4A75X 10IC4A91X 10JB4D19X 10KZ4A06X 10AC4B51X 10CA4A08X 10DC4A89X 10FA4A77X 10GZ4A91X 10HC4A76X 10IC4A92X 10JB4D20X 10KZ4A07X 10AC4B52X 10CA4A09X 10DC4A90X 10FA4A78X 10GZ4A92X 10HC4A77X 10IC4A93X 10JB4D21X 10KZ4A08X 10AC4B53X 10CA4A10X 10DC4A91X 10FA4A79X 10GZ4A93X 10HC4A78X 10IC4A94X 10JB4D22X 10KZ4A09X 10AC4B54X 10CA4A11X 10DC4A92X 10FA4A80X 10GZ4A94X 10HC4A79X 10IC4A95X 10JB4D23X 10KZ4A10X 10AC4B55X 10CA4A12X 10DC4A93X 10FA4A81X 10GZ4A95X 10HC4A80X 10IC4A96X 10JB4D24X 10KZ4A11X 10AC4B56X 10CA4A13X 10DC4A94X 10FA4A82X 10GZ4A96X 10HC4A81X 10IC4A97X 10JB4D25X 10KZ4A12X 10AC4B57X 10CA4A14X 10DC4A95X 10FA4A83X 10GZ4A97X 10HC4A82X 10IC4A98X 10JB4D26X 10KZ4A13X 10AC4B59X 10CA4A16X 10DC4A96X 10FA4A84X 10GZ4A98X 10HC4A83X 10IC4A99X 10JB4D27X 10KZ4A14X 10AC4B60X 10CA4A17X 10DC4A97X 10FA4A85X 10GZ4A99X 10HC4A84X 10IC4B00X 10JB4D28X 10KZ4A15X 10AC4B62X 10CA4A18X 10DC4A98X 10FA4A86X 10GZ4B01X 10HC4A85X 10IC4B01X 10JB4D29X 10KZ4A16X 10AC4B63X 10CA4A19X 10DC4A99X 10FA4A87X 10GZ4B03X 10HC4A86X 10IC4B03X 10JB4D30X 10KZ4A17X 10AC4B64X 10CA4A21X 10DC4B01X 10FA4A88X 10GZ4B04X 10HC4A87X 10IC4B04X 10JB4D31X 10KZ4A18X 10AC4B66X 10CA4A23X 10DC4B02X 10FA4A89X 10GZ4B05X 10HC4A88X 10IC4B05X 10JB4D32X 10KZ4A19X 10AC4B68X 10CA4A25X 10DC4B03X 10FA4A90X 10GZ4B06X 10HC4A89X 10IC4B06X 10JB4D33X 10KZ4A20X 10AC4B70X 10CA4A26X 10DC4B04X 10FA4A91X 10GZ4B07X 10HC4A90X 10IC4B07X 10JB4D34X 10KZ4A21X 10AC4B71X 10CA4A27X 10DC4B05X 10FA4A92X 10GZ4B08X 10HC4A91X 10IC4B08X 10JB4D35X 10KZ4A22X 10AC4B73X 10CA4A28X 10DC4B06X 10FA4A93X 10GZ4B09X 10HC4A92X 10IC4B09X 10JB4D36X 10KZ4A23X 10AC4B74X 10CA4A30X 10DC4B07X 10FA4A94X 10GZ4B10X 10HC4A93X 10IC4B10X 10JB4D37X 10KZ4A24X 10AC4B75X 10CA4A32X 10DC4B08X 10FA4A95X 10GZ4B11X 10HC4A94X 10IC4B11X 10JB4D38X 10KZ4A25X 10AC4B76X 10CA4A34X 10DC4B09X 10FA4A96X 10GZ4B12X 10HC4A97X 10IC4B12X 10JB4D39X 10KZ4A26X 10AC4B77X 10CA4A35X 10DC4B10X 10FA4A97X 10GZ4B13X 10HC4A98X 10IC4B13X 10JB4D40X 10KZ4A27X 10AC4B78X 10CA4A36X 10DC4B11X 10FA4A98X 10GZ4B15X 10HC4A99X 10IC4B14X 10JB4D41X 10KZ4A28X 10AC4B79X 10CA4A38X 10DC4B12X 10FA4A99X 10GZ4B16X 10HC4B02X 10IC4B15X 10JB4D42X 10KZ4A29X 10AC4B80X 10CA4A39X 10DC4B13X 10FA4B01X 10GZ4B18X 10HC4B03X 10IC4B16X 10JB4D43X 10KZ4A30X 10AC4B81X 10CA4A42X 10DC4B14X 10FA4B02X 10GZ4B21X 10HC4B05X 10IC4B17X 10JB4D44X 10KZ4A31X 10AC4B82X 10CA4A43X 10DC4B15X 10FA4B03X 10GZ4B22X 10HC4B06X 10IC4B18X 10JB4D45X 10KZ4A32X 10AC4B84X 10CA4A44X 10DC4B16X 10FA4B05X 10GZ4B24X 10HC4B07X 10IC4B19X 10JB4D46X 10KZ4A33X 10AC4B85X 10CA4A45X 10DC4B17X 10FA4B06X 10GZ4B25X 10HC4B08X 10IC4B20X 10JB4D47X 10KZ4A34X 10AC4B86X 10CA4A46X 10DC4B18X 10FA4B07X 10GZ4B26X 10HC4B09X 10IC4B21X 10JB4D48X 10KZ4A35X 10AC4B87X 10CA4A48X 10DC4B19X 10FA4B08X 10GZ4B27X 10HC4B10X 10IC4B22X 10JB4D49X 10KZ4A36X 10AC4B88X 10CA4A49X 10DC4B20X 10FA4B09X 10GZ4B28X 10HC4B11X 10IC4B23X 10JB4D50X 10KZ4A38X 10AC4B89X 10CA4A50X 10DC4B21X 10FA4B10X 10GZ4B30X 10HC4B12X 10IC4B24X 10JB4D51X 10KZ4A39X 10AC4B90X 10CA4A51X 10DC4B22X 10FA4B11X 10GZ4B31X 10HC4B13X 10IC4B25X 10JB4D52X 10KZ4A40X 10AC4B91X 10CA4A52X 10DC4B23X 10FB4A01X 10GZ4B32X 10HC4B14X 10IC4B26X 10JB4D53X 10KZ4A41X 10AC4B92X 10CA4A53X 10DC4B24X 10FB4A02X 10GZ4B33X 10HC4B15X 10IC4B27X 10JB4D54X 10KZ4A42X 10AC4B93X 10CA4A54X 10DC4B25X 10FB4A03X 10GZ4B34X 10HC4B16X 10IC4B28X 10JB4D55X 10KZ4A43X 10AC4B94X 10CA4A55X 10DC4B26X 10FB4A04X 10GZ4B35X 10HC4B17X 10IC4B29X 10JB4D56X 10KZ4A45X 10AC4B95X 10CA4A56X 10DC4B27X 10FB4A06X 10GZ4B36X 10HC4B18X 10IC4B30X 10JB4D57X 10LA4A71X 10AC4B96X 10CA4A57X 10DC4B28X 10FB4A07X 10GZ4B37X 10HC4B19X 10IC4B31X 10JB4D58X 10LA4A72X 10AC4B97X 10CA4A58X 10DC4B29X 10FB4A08X 10GZ4B38X 10HC4B21X 10IC4B32X 10JB4D59X 10LA4A73X 10AC4B98X 10CA4A59X 10DC4B30X 10FB4A09X 10GZ4B39X 10HC4B22X 10IC4B33X 10JB4D60X 10LA4A74X 10AC4B99X 10CA4A60X 10DC4B31X 10FB4A12X 10GZ4B40X 10HC4B23X 10IC4B34X 10JB4D61X 10LA4A75X 10AC4C01X 10CA4A61X 10DC4B32X 10FB4A13X 10GZ4B41X 10HC4B24X 10IC4B35X 10JB4D62X 10LA4A76X 10AC4C02X 10CA4A62X 10DC4B33X 10FB4A14X 10GZ4B42X 10HC4B25X 10IC4B36X 10JB4D63X 10LA4A78X 10AC4C03X 10CA4A63X 10DC4B34X 10FB4A16X 10GZ4B43X 10HC4B26X 10IC4B37X 10JB4D64X 10LA4A79X 10AC4C06X 10CA4A64X 10DC4B35X 10FB4A17X 10GZ4B44X 10HC4B27X 10IC4B38X 10JB4D65X 10LA4A80X 10AC4C07X 10CA4A65X 10DC4B36X 10FB4A18X 10GZ4B45X 10HC4B28X 10IC4B39X 10JB4D66X 10LA4A81X 10AC4C08X 10CA4A66X 10DC4B37X 10FB4A19X 10GZ4B47X 10HC4B29X 10IC4B40X 10JB4D67X 10LA4A82X 10AC4C10X 10CA4A67X 10DC4B38X 10FB4A20X 10GZ4B50X 10HC4B30X 10IC4B41X 10JB4D68X 10LA4A83X 10AC4C11X 10CA4A68X 10DC4B39X 10FB4A21X 10GZ4B51X 10HC4B31X 10IC4B42X 10JB4D69X 10LA4A84X 10AC4C12X 10CA4A69X 10DC4B40X 10FB4A22X 10GZ4B52X 10HC4B33X 10IC4B43X 10JB4D70X 10LA4A85X 10AC4C13X 10CA4A70X 10DC4B41X 10FB4A23X 10GZ4B53X 10HC4B34X 10IC4B44X 10JB4D71X 10LA4A86X 10AC4C14X 10CA4A71X 10DC4B42X 10FB4A25X 10GZ4B54X 10HC4B35X 10IC4B45X 10JB4D72X 10LA4A87X 10AC4C15X 10CA4A72X 10DC4B43X 10FB4A26X 10GZ4B55X 10HC4B36X |  | SoftwareChangeControl
53 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171923 | Z-1712-2019 | Prometra Clinician Programmer (Touchscreen) Software version 2.00.29, Cat. No. 12828, used with Prometra and Prometra II Programmable Pumps. For intrathecal infusion. | Prometra Programmable Infusion Pump System, Prometra Clinician Programmer | 2 | CDRH | 2019-06-03 |  | 2020-06-05 09:23 | Flowonix Medical Inc | The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. | Class 2 Device Recall Prometra Programmable Infusion Pump System, Prometra Clinician Programmer Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prometra Programmable Infusion Pump System, Prometra Clinician Programmer FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prometra Programmable Infusion Pump System, Prometra Clinician Programmer Date Initiated by Firm March 21, 2019 Create Date June 03, 2019 Recall Status 1 Terminated 3 on June 05, 2020 Recall Number Z-1712-2019 Recall Event ID 82586 PMA Number P080012 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product Prometra Clinician Programmer (Touchscreen) Software version 2.00.29, Cat. No. 12828, used with Prometra and Prometra II Programmable Pumps. For intrathecal infusion. Code Information Serial Numbers: A15E0017 A15J0019 A16D0004 RA17B0003 RA17I0001 RA18B0008 RA18D0046 RA18G0074 RA18J0074 A15E0023 A15J0035 A16D0013 RA17B0004 RA17I0002 RA18B0009 RA18D0047 RA18G0079 RA18J0077 A15E0032 A15J0038 A16D0015 RA17B0007 RA17I0003 RA18B0012 RA18D0048 RA18G0080 RA18J0078 A15H0003 A15J0047 A16D0018 RA17C0001 RA17I0004 RA18B0014 RA18D0052 RA18J0001 RA18J0081 A15H0007 A15J0050 A16D0026 RA17C0005 RA17I0005 RA18B0016 RA18D0053 RA18J0002 RA18J050 A15H0031 A15J0051 A16D0032 RA17C0006 RA17I0007 RA18B0018 RA18G0004 RA18J0003 RA18K0003 A15H0038 A15K0018 A16D0033 RA17C0007 RA17I0008 RA18B0020 RA18G0005 RA18J0004 RA18K0004 A15H0039 A15K0036 A16D0036 RA17C0025 RA17I0010 RA18B0023 RA18G0006 RA18J0009 RA18K0007 A15H0054 A15K0042 A16D0037 RA17C0028 RA17I0012 RA18B0026 RA18G0007 RA18J0011 RA18K0008 A15H0067 A15K0045 A16D0039 RA17C0038 RA17I0021 RA18B0027 RA18G0010 RA18J0012 RA18K001 A15I0005 A15K0046 A16D0040 RA17C0039 RA17I0024 RA18B0030 RA18G0011 RA18J0013 RA18K0010 A15I0015 A15K0053 A16D0042 RA17C0041 RA17J0003 RA18B0033 RA18G0012 RA18J0014 RA18L0001 A15I0019 A15K0055 A16D0046 RA17F0006 RA17J0004 RA18B0036 RA18G0017 RA18J0015 RA18L0002 A15I0020 A15L0003 A16D0047 RA17F0007 RA17J0005 RA18B0037 RA18G0021 RA18J0016 RA18L0004 A15I0039 A15L0029 A16D0049 RA17F0008 RA17J0008 RA18B0040 RA18G0022 RA18J0019 RA18L0007 A15I0043 A15L0067 A16D0051 RA17F0009 RA17J0011 RA18B0042 RA18G0023 RA18J0021 RA18L0010 A15I0046 A15L0068 A16D0058 RA17F0011 RA17J0016 RA18B0047 RA18G0025 RA18J0022 RA18L019 A15I0049 A15L0069 A16D0073 RA17F0012 RA17K0002 RA18B0053 RA18G0026 RA18J0023 RA18L020 A15I0058 A15L0070 A16D0074 RA17F0018 RA17K0008 RA18B0055 RA18G0027 RA18J0025 RA18L021 A15I0062 A16B0002 A17C0011 RA17F0024 RA17K0009 RA18C0003 RA18G0031 RA18J0030 RA18L022 A15I0067 A16B0003 A17D0009 RA17F0025 RA17L0001 RA18D0002 RA18G0032 RA18J0031 RA18L023 A15I0068 A16B0008 A17D0010 RA17F0028 RA17L0002 RA18D0003 RA18G0033 RA18J0034 RA18L024 A15I0070 A16C0006 A17E0003 RA17F0030 RA17L0003 RA18D0006 RA18G0040 RA18J0038 RA18L027 A15I0072 A16C0011 A17E0008 RA17F0032 RA17L0004 RA18D0007 RA18G0042 RA18J0039 RA18L028 A15I0073 A16C0012 A17E0009 RA17F0033 RA17L0008 RA18D0010 RA18G0043 RA18J0040 RA18L029 A15I0075 A16C0017 A17F0001 RA17F0037 RA17L0009 RA18D0022 RA18G0044 RA18J0041 RA18L030 A15I0089 A16C0024 A17F0003 RA17F0038 RA17L0010 RA18D0024 RA18G0045 RA18J0046 RA18L031 A15I0139 A16C0028 A17F0004 RA17F0039 RA17L0012 RA18D0026 RA18G0050 RA18J0048 RA18L032 A15I0140 A16C0030 A17F0005 RA17F0041 RA17L0014 RA18D0028 RA18G0055 RA18J0053 RA18L033 A15J0002 A16C0038 RA17A0002 RA17F0042 RA17L0022 RA18D0033 RA18G0056 RA18J0063 RA18L034 A15J0004 A16C0041 RA17A0004 RA17F0043 RA18B0001 RA18D0039 RA18G0058 RA18J0064 RA18L036 A15J0006 A16C0044 RA17A0005 RA17G0005 RA18B0002 RA18D0040 RA18G0062 RA18J0068 RA18L051 A15J0010 A16C0060 RA17A0011 RA17G0006 RA18B0004 RA18D0041 RA18G0064 RA18J0070 RA18L052 A15J0011 A16C0063 RA17A0035 RA17G0016 RA18B0006 RA18D0043 RA18G0065 RA18J0071 RA18L053 A15J0017 A16C0072 RA17A0041 RA17G0017 RA18B0007 RA18D0045 RA18G0068 RA18J0073 RA18L054 RA18L057 RA18L055 RA18L056 Recalling Firm/ Manufacturer Flowonix Medical Inc 120 Forbes Blvd Ste 170 Mansfield MA 02048-1150 For Additional Information Contact Technical Solutions Department 855-356-9665 Manufacturer Reason for Recall The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. FDA Determined Cause 2 Software change control Action On March 21, 2019, the firm issued Urgent Recall Notices to customers to inform them of the product issue. ************In May 2019, the firm sent updated communications to customers:************Healthcare providers who have Clinician Programmers with SW 2.00.29 were sent a follow-up recall letter. This letter included the information that the pump stop following "Error Code 115" will only occur after the Bridge or Demand Bolus is complete, which could take up to seventy-two (72 hours). ************In addition, in the event of any cause of an "Error Code 115", the pump sounds a one-beep alarm and then shuts down. This behavior differs from the Instructions for Use, which state that the pump will sound three long (half second) beeps every 30 minutes for all critical errors. As it is unlikely that the patient will notice the single beep, there may be a significant delay in the patient becoming aware that the pump has stopped delivering therapy which may in turn delay resumption of drug therapy.************Healthcare providers who have Clinician Programmers with any other version of the software and the implant pumps received a Corrective Action Notice. This notice included the information from the preceding paragraph, noting the discrepant alarm behavior for "Error Code 115" could be missed by the patient.************Customers with the new Prometra Clinician Programmer software (v. 2.00.29) were asked to do the following:************1. STOP using the Periodic Flow or Multiple Rates flow mode functions in combination with a Bridge or******Demand Bolus using the Clinician Programmer with Software Version 2.00.29. The error only occurs when******these functions are used together.************a. In the event that a Bridge or Demand Bolus is needed, convert the patient to Constant Flow Mode,******then program the Bridge or Demand Bolus as needed (instructions for use PL-11803-04, page 30).******Once the bolus is complete, revisit the patient or have him/her return to your clinic for programming******back into their original flow mode.************2. IDENTIFY Quantity in Commerce 318 Distribution US nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LKK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On March 21, 2019, the firm issued Urgent Recall Notices to customers to inform them of the product issue. ************In May 2019, the firm sent updated communications to customers:************Healthcare providers who have Clinician Programmers with SW 2.00.29 were sent a follow-up recall letter. This letter included the information that the pump stop following "Error Code 115" will only occur after the Bridge or Demand Bolus is complete, which could take up to seventy-two (72 hours). ************In addition, in the event of any cause of an "Error Code 115", the pump sounds a one-beep alarm and then shuts down. This behavior differs from the Instructions for Use, which state that the pump will sound three long (half second) beeps every 30 minutes for all critical errors. As it is unlikely that the patient will notice the single beep, there may be a significant delay in the patient becoming aware that the pump has stopped delivering therapy which may in turn delay resumption of drug therapy.************Healthcare providers who have Clinician Programmers with any other version of the software and the implant pumps received a Corrective Action Notice. This notice included the information from the preceding paragraph, noting the discrepant alarm behavior for "Error Code 115" could be missed by the patient.************Customers with the new Prometra Clinician Programmer software (v. 2.00.29) were asked to do the following:************1. STOP using the Periodic Flow or Multiple Rates flow mode functions in combination with a Bridge or******Demand Bolus using the Clinician Programmer with Software Version 2.00.29. The error only occurs when******these functions are used together.************a. In the event that a Bridge or Demand Bolus is needed, convert the patient to Constant Flow Mode,******then program the Bridge or Demand Bolus as needed (instructions for use PL-11803-04, page 30).******Once the bolus is complete, revisit the patient or have him/her return to your clinic for programming******back into their original flow mode.************2. IDENTIFY | SoftwareChangeControl
54 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172978 | Z-1714-2019 | Prometra Programmable Pump, Cat. No. 11827 For intrathecal infusion. | Prometra Programmable Pump System | 2 | CDRH | 2019-06-03 |  | 2020-06-05 09:23 | Flowonix Medical Inc | The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. Error 115 can also occur if dendrites form or if the pump is dropped. | Class 2 Device Recall Prometra Programmable Pump System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prometra Programmable Pump System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prometra Programmable Pump System Date Initiated by Firm March 21, 2019 Create Date June 03, 2019 Recall Status 1 Terminated 3 on June 05, 2020 Recall Number Z-1714-2019 Recall Event ID 82586 PMA Number P080012 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product Prometra Programmable Pump, Cat. No. 11827 For intrathecal infusion. Code Information Serial Numbers: 10AX4A03 10BY4A40 10EY4A35 10GX4A83 10HW4A31 10HY4C11 10JW4A28 10KW4A42 10KX4A48 10AX4A04 10BY4A41 10EY4A36 10GX4A84 10HW4A45 10HY4C12 10JW4A29 10KW4A43 10KX4A49 10AX4A06 10BY4A42 10EY4A37 10GX4A85 10HW4A57 10HY4C13 10JW4A30 10KW4A44 10KX4A50 10AX4A10 10BY4A44 10EY4A38 10GX4A86 10HY4A77 10HY4C15 10JW4A31 10KW4A45 10KX4A52 10AX4A12 10BY4A45 10EY4A40 10GX4A87 10HY4A78 10HY4C16 10JW4A32 10KW4A46 10KX4A56 10AX4A14 10BY4A46 10EY4A41 10GX4A88 10HY4A79 10HY4C17 10JW4A33 10KW4A47 10KX4A57 10AX4A21 10BY4A47 10EY4A43 10GX4A89 10HY4A80 10HY4C18 10JW4A34 10KW4A48 10KX4A58 10AX4A22 10BY4A48 10EY4A44 10GX4A92 10HY4A81 10HY4C19 10JW4A35 10KW4A49 10KX4A59 10AX4A23 10BY4A49 10EY4A45 10GX4A94 10HY4A82 10HY4C20 10JW4A36 10KW4A50 10KX4A60 10AX4A28 10BY4A50 10EY4A46 10GX4A96 10HY4A83 10HY4C22 10JW4A37 10KW4A51 10KX4A62 10AX4A30 10BY4A51 10EY4A47 10GX4A97 10HY4A84 10HY4C23 10JW4A38 10KW4A52 10KX4A63 10AX4A36 10BY4A52 10EY4A48 10GX4A98 10HY4A85 10HY4C26 10JW4A39 10KW4A53 10KX4A64 10AX4A37 10BY4A53 10EY4A51 10GX4A99 10HY4A87 10HY4C27 10JW4A40 10KW4A54 10KX4A65 10AX4A39 10BY4A54 10EY4A52 10GX4B02 10HY4A90 10HY4C28 10JW4A41 10KW4A55 10KX4A66 10AX4A42 10BY4A55 10EY4A53 10GX4B03 10HY4A91 10HY4C29 10JW4A42 10KW4A56 10KX4A67 10AX4A48 10BY4A56 10EY4A54 10GX4B04 10HY4A92 10HY4C30 10JW4A43 10KW4A57 10KX4A68 10AX4A51 10BY4A57 10EY4A55 10GX4B05 10HY4A95 10HY4C32 10JW4A44 10KW4A58 10KX4A69 10AX4A73 10BY4A58 10EY4A56 10GX4B06 10HY4A97 10HY4C34 10JW4A45 10KW4A59 10KX4A70 10AX4A74 10BY4A59 10EY4A58 10GX4B07 10HY4A98 10HY4C35 10JW4A46 10KW4A60 10KX4A71 10AX4A76 10BY4A60 10EY4A59 10GX4B08 10HY4B01 10HY4C36 10JW4A47 10KW4A61 10KX4A72 10AX4A81 10BZ4A24 10EY4A60 10GX4B09 10HY4B02 10HY4C37 10JW4A48 10KW4A63 10KX4A73 10AX4A83 10BZ4A26 10EY4A61 10GX4B10 10HY4B03 10HY4C39 10JY4A01 10KW4A64 10KX4A74 10AX4A85 10BZ4A32 10EY4A62 10GX4B11 10HY4B05 10HY4C40 10JY4A03 10KW4A65 10KX4A76 10AX4A87 10BZ4A46 10EY4A63 10GX4B12 10HY4B07 10HY4C41 10JY4A06 10KW4A66 10KX4A77 10AX4A88 10BZ4A70 10EY4A64 10GX4B13 10HY4B08 10HY4C42 10JY4A07 10KW4A67 10KX4A78 10AX4A89 10BZ4A71 10EY4A65 10GX4B14 10HY4B10 10HY4C46 10JY4A09 10KW4A68 10KX4A79 10AX4A93 10CZ4A02 10EY4A66 10GX4B15 10HY4B12 10HY4C47 10JY4A10 10KW4A70 10KX4A80 10AX4A95 10CZ4A03 10EY4A67 10GX4B16 10HY4B16 10IY2A05 10JY4A15 10KW4A71 10KX4A81 10AX4A96 10CZ4A04 10EY4A69 10GX4B21 10HY4B17 10IY2A26 10JY4A16 10KW4A72 10KX4A82 10AX4A98 10DX4A03 10EY4A70 10GX4B22 10HY4B18 10IY4A01 10JY4A18 10KW4A73 10KX4A85 10AX4B06 10DX4A07 10EY4A72 10GX4B25 10HY4B19 10IY4A02 10JY4A20 10KW4A74 10KY4A02 10AX4B08 10DX4A09 10EY4A73 10GY4A02 10HY4B20 10IY4A03 10JY4A21 10KW4A75 10KY4A03 10AX4B09 10DX4A10 10EY4A74 10GY4A03 10HY4B22 10IY4A04 10JY4A25 10KW4A76 10KY4A04 10AX4B16 10DX4A13 10EY4A75 10GY4A05 10HY4B23 10IY4A05 10JY4A28 10KW4A77 10KY4A06 10AX4B17 10DX4A21 10EY4A76 10GY4A06 10HY4B24 10IY4A09 10JY4A29 10KW4A78 10KY4A07 10DX4A24 10EY4A81 10GY4A11 10HY4B26 10IY4A11 10KV4A01 10KW4A80 10KY4A10 10AX4B26 10DX4A32 10EY4A82 10GY4A12 10HY4B27 10IY4A13 10KW4A01 10KW4A81 10KY4A11 10AX4B27 10DX4A33 10EY4A83 10GY4A13 10HY4B28 10IY4A15 10KW4A02 10KW4A83 10KY4A12 10AX4B29 10DX4A34 10EY4A87 10GY4A14 10HY4B29 10IY4A17 10KW4A03 10KW4A84 10KY4A13 10AX4B53 10DX4A41 10EY4A88 10GY4A15 10HY4B30 10IY4A19 10KW4A04 10KW4A85 10KY4A15 10AX4B78 10DX4A44 10EY4A89 10GY4A16 10HY4B33 10IY4A25 10KW4A05 10KW4A86 10KY4A17 10AX4B79 10DX4A48 10EY4A90 10GY4A18 10HY4B34 10IY4A26 10KW4A06 10KW4A87 10KY4A18 10AX4B80 10DX4A52 10EY4A92 10GY4A19 10HY4B35 10IY4A27 10KW4A07 10KW4A88 10KY4A20 10AX4B81 10DX4A60 10EY4A93 10GY4A20 10HY4B36 10IY4A28 10KW4A08 10KW4A89 10KY4A21 10AX4B82 10DX4A62 10EY4A94 10GY4A21 10HY4B37 10IY4A29 10KW4A09 10KW4A90 10KY4A25 10AX4B84 10DX4A65 10EY4A95 10GY4A22 10HY4B40 10IY4A31 10KW4A10 10KW4A91 10KY4A26 10AX4B85 10DX4A67 10EY4A96 10GY4A23 10HY4B41 10IY4A33 10KW4A101 10KW4A92 10KY4A27 10AX4B86 10DX4A71 10EY4A97 10GY4A25 10HY4B42 10IY4A37 10KW4A103 10KW4A93 10KY4A30 10AX4B88 10DX4A74 10EY4A98 10GY4A26 10HY4B43 10IY4A38 10KW4A104 10KW4A94 10KY4A31 10AX4B90 10DX4A79 10EY4A99 10GY4A27 10HY4B44 10IY4A41 10KW4A105 10KW4A95 10KY4A33 10AX4B91 10DX4A82 10GX4A03 10GY4A28 10HY4B45 10IY4A43 10KW4A106 10KW4A96 10KY4A35 10AX4B92 10DX4A85 10GX4A04 10GY4A30 10HY4B48 10IY4A44 10KW4A108 10KW4A98 10KY4A36 10AX4B94 10DX4A89 10GX4A08 10GY4A31 10HY4B49 10IY4A46 10KW4A109 10KW4A99 10KY4A37 10AX4B96 10DX4A95 10GX4A09 10GY4A32 10HY4B50 10IY4A50 10KW4A11 10KX4A01 10KY4A39 10AX4B97 10DX4A96 10GX4A10 10GY4A34 10HY4B51 10IY4A51 10KW4A110 10KX4A02 10KY4A40 10AZ4A08 10DX4A98 10GX4A12 10GY4A35 10HY4B53 10IY4A52 10KW4A111 10KX4A04 10KY4A41 10AZ4A10 10DX4A99 10GX4A14 10GY4A36 10HY4B54 10IY4A54 10KW4A113 10KX4A05 10KY4A43 10AZ4A12 10DX4B02 10GX4A15 10GY4A39 10HY4B55 10IY4A56 10KW4A116 10KX4A06 10KY4A44 10AZ4A14 10DX4B03 10GX4A16 10GY4A40 10HY4B56 10IY4A58 10KW4A117 10KX4A07 10KY4A45 10AZ4A15 10DX4B07 10GX4A19 10GY4A41 10HY4B62 10IY4A59 10KW4A119 10KX4A08 10KY4A47 10AZ4A18 10DX4B15 10GX4A21 10GY4A42 10HY4B63 10IY4A60 10KW4A12 10KX4A09 10KY4A50 10AZ4A19 10DX4B17 10GX4A30 10GY4A44 10HY4B65 10IY4A61 10KW4A120 10KX4A10 10KY4A51 10AZ4A22 10DX4B21 10GX4A31 10GY4A45 10HY4B66 10IY4A62 10KW4A121 10KX4A11 10KY4A60 10AZ4A24 10DX4B25 10GX4A32 10GY4A47 10HY4B69 10JW4A01 10KW4A122 10KX4A12 10KY4A61 10AZ4A31 10DX4B26 10GX4A33 10GY4A49 10HY4B70 10JW4A02 10KW4A123 10KX4A13 10KY4A62 10AZ4A33 10EY4A01 10GX4A34 10GY4A50 10HY4B74 10JW4A03 10KW4A125 10KX4A14 10KY4A63 10AZ4A43 10EY4A02 10GX4A37 10GY4A51 10HY4B76 10JW4A04 10KW4A13 10KX4A15 10KY4A64 10AZ4A54 10EY4A03 10GX4A38 10GY4A52 10HY4B78 10JW4A05 10KW4A14 10KX4A16 10KY4A65 10AZ4A57 10EY4A04 10GX4A39 10GY4A53 10HY4B79 10JW4A06 10KW4A15 10KX4A17 10KY4A67 10AZ4A59 10EY4A05 10GX4A40 10GY4A54 10HY4B80 10JW4A07 10KW4A16 10KX4A18 10KY4A68 10AZ4A61 10EY4A08 10GX4A44 10GY4A55 10HY4B81 10JW4A08 10KW4A17 10KX4A19 10KY4A69 10AZ4A65 10EY4A11 10GX4A45 10GY4A56 10HY4B83 10JW4A09 10KW4A18 10KX4A20 10KY4A71 10AZ4A70 10EY4A13 10GX4A46 10GY4A57 10HY4B84 10JW4A10 10KW4A19 10KX4A21 10KY4A72 10AZ4A73 10EY4A14 10GX4A47 10GY4A58 10HY4B85 10JW4A11 10KW4A20 10KX4A22 10KY4A75 10BY4A25 10EY4A15 10GX4A48 10GY4A59 10HY4B86 10JW4A12 10KW4A21 10KX4A25 10KY4A78 10BY4A26 10EY4A16 10GX4A57 10GY4A60 10HY4B90 10JW4A13 10KW4A23 10KX4A26 10KY4A79 10BY4A27 10EY4A17 10GX4A62 10GY4A61 10HY4B91 10JW4A14 10KW4A24 10KX4A27 10KY4A82 10GX4A65 10GY4A63 10HY4B96 10JW4A16 10KW4A26 10KX4A29 10KY4A84 10BY4A30 10EY4A21 10GX4A67 10GY4A64 10HY4B97 10JW4A17 10KW4A27 10KX4A30 10KY4A86 10BY4A31 10EY4A22 10GX4A69 10GY4A65 10HY4B99 10JW4A18 10KW4A29 10KX4A32 10KY4A87 10BY4A32 10EY4A24 10GX4A70 10GY4A66 10HY4C01 10JW4A19 10KW4A30 10KX4A39 10KY4A90 10BY4A33 10EY4A25 10GX4A71 10GY4A67 10HY4C02 10JW4A20 10KW4A31 10KX4A40 10KY4A91 10BY4A34 10EY4A26 10GX4A72 10GY4A70 10HY4C03 10JW4A21 10KW4A32 10KX4A41 10KY4A92 10BY4A35 10EY4A27 10GX4A73 10GY4A71 10HY4C04 10JW4A22 10KW4A35 10KX4A42 10KY4A93 10BY4A36 10EY4A28 10GX4A74 10GY4A72 10HY4C06 10JW4A24 10KW4A38 10KX4A43 10KY4A94 10BY4A37 10EY4A29 10GX4A76 10GY4A73 10HY4C07 10JW4A25 10KW4A39 10KX4A44 10KY4A97 10BY4A38 10EY4A31 10GX4A81 10GY4A74 10HY4C08 10JW4A26 10KW4A40 10KX4A46 10KY4A99 10BY4A39 10EY4A33 10GX4A82 10HW4A29 10HY4C09 10JW4A27 10KW4A41 10KX4A47 Recalling Firm/ Manufacturer Flowonix Medical Inc 120 Forbes Blvd Ste 170 Mansfield MA 02048-1150 For Additional Information Contact Technical Solutions Department 855-356-9665 Manufacturer Reason for Recall The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. Error 115 can also occur if dendrites form or if the pump is dropped. FDA Determined Cause 2 Software change control Action On March 21, 2019, the firm issued Urgent Recall Notices to customers to inform them of the product issue. ************In May 2019, the firm sent updated communications to customers:************Healthcare providers who have Clinician Programmers with SW 2.00.29 were sent a follow-up recall letter. This letter included the information that the pump stop following "Error Code 115" will only occur after the Bridge or Demand Bolus is complete, which could take up to seventy-two (72 hours). ************In addition, in the event of any cause of an "Error Code 115", the pump sounds a one-beep alarm and then shuts down. This behavior differs from the Instructions for Use, which state that the pump will sound three long (half second) beeps every 30 minutes for all critical errors. As it is unlikely that the patient will notice the single beep, there may be a significant delay in the patient becoming aware that the pump has stopped delivering therapy which may in turn delay resumption of drug therapy.************Healthcare providers who have Clinician Programmers with any other version of the software and the implant pumps received a Corrective Action Notice. This notice included the information from the preceding paragraph, noting the discrepant alarm behavior for "Error Code 115" could be missed by the patient.************Customers with the new Prometra Clinician Programmer software (v. 2.00.29) were asked to do the following:************1. STOP using the Periodic Flow or Multiple Rates flow mode functions in combination with a Bridge or******Demand Bolus using the Clinician Programmer with Software Version 2.00.29. The error only occurs when******these functions are used together.************a. In the event that a Bridge or Demand Bolus is needed, convert the patient to Constant Flow Mode,******then program the Bridge or Demand Bolus as needed (instructions for use PL-11803-04, page 30).******Once the bolus is complete, revisit the patient or have him/her return to your clinic for programming******back into their original flow mode.************2. IDENTIFY Quantity in Commerce 809 Distribution US nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LKK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On March 21, 2019, the firm issued Urgent Recall Notices to customers to inform them of the product issue. ************In May 2019, the firm sent updated communications to customers:************Healthcare providers who have Clinician Programmers with SW 2.00.29 were sent a follow-up recall letter. This letter included the information that the pump stop following "Error Code 115" will only occur after the Bridge or Demand Bolus is complete, which could take up to seventy-two (72 hours). ************In addition, in the event of any cause of an "Error Code 115", the pump sounds a one-beep alarm and then shuts down. This behavior differs from the Instructions for Use, which state that the pump will sound three long (half second) beeps every 30 minutes for all critical errors. As it is unlikely that the patient will notice the single beep, there may be a significant delay in the patient becoming aware that the pump has stopped delivering therapy which may in turn delay resumption of drug therapy.************Healthcare providers who have Clinician Programmers with any other version of the software and the implant pumps received a Corrective Action Notice. This notice included the information from the preceding paragraph, noting the discrepant alarm behavior for "Error Code 115" could be missed by the patient.************Customers with the new Prometra Clinician Programmer software (v. 2.00.29) were asked to do the following:************1. STOP using the Periodic Flow or Multiple Rates flow mode functions in combination with a Bridge or******Demand Bolus using the Clinician Programmer with Software Version 2.00.29. The error only occurs when******these functions are used together.************a. In the event that a Bridge or Demand Bolus is needed, convert the patient to Constant Flow Mode,******then program the Bridge or Demand Bolus as needed (instructions for use PL-11803-04, page 30).******Once the bolus is complete, revisit the patient or have him/her return to your clinic for programming******back into their original flow mode.************2. IDENTIFY | SoftwareChangeControl
55 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172980 | Z-1716-2019 | Prometra II Programmable Pump (Clinical), Cat. No. 01827 For intrathecal infusion. | Prometra Programmable Pump System | 2 | CDRH | 2019-06-03 |  | 2020-06-05 09:23 | Flowonix Medical Inc | The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. Error 115 can also occur if dendrites form or if the pump is dropped. | Class 2 Device Recall Prometra Programmable Pump System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prometra Programmable Pump System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prometra Programmable Pump System Date Initiated by Firm March 21, 2019 Create Date June 03, 2019 Recall Status 1 Terminated 3 on June 05, 2020 Recall Number Z-1716-2019 Recall Event ID 82586 PMA Number P080012 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product Prometra II Programmable Pump (Clinical), Cat. No. 01827 For intrathecal infusion. Code Information Serial Numbers: 10ET4A02 36AR4A11 36AR4A33 36FR4A01 36HS4A25 36JN4A09 36JN4A30 36KN4A13 36KN4A33 10ET4A06 36AR4A12 36AR4A34 36FR4A08 36HS4A26 36JN4A10 36JN4A31 36KN4A15 36KN4A34 10ET4A07 36AR4A13 36AR4A35 36FR4A09 36HS4A27 36JN4A11 36JN4A32 36KN4A16 36KN4A35 10ET4A20 36AR4A14 36AR4A37 36FR4A12 36HS4A29 36JN4A12 36JN4A33 36KN4A17 36KN4A36 10ET4A21 36AR4A15 36BS4A08 36FR4A13 36HS4A34 36JN4A13 36JN4A35 36KN4A18 36KN4A38 10ET4A46 36AR4A16 36BS4A17 36FR4A16 36HS4A39 36JN4A14 36JN4A37 36KN4A19 36KN4A40 10IT4A16 36AR4A17 36BS4A18 36FR4A19 36HS4A42 36JN4A15 36JN4A39 36KN4A20 36KN4A41 10IT4A21 36AR4A18 36ES4A03 36FR4A22 36HS4A43 36JN4A16 36KN4A01 36KN4A21 36KN4A42 10IT4A22 36AR4A19 36ES4A07 36FR4A24 36HS4A46 36JN4A17 36KN4A02 36KN4A22 36KN4A43 36AR4A01 36AR4A20 36ES4A08 36HS4A01 36HS4A51 36JN4A18 36KN4A03 36KN4A23 36KN4A44 36AR4A02 36AR4A21 36ES4A14 36HS4A02 36HS4A53 36JN4A19 36KN4A04 36KN4A24 36KN4A45 36AR4A03 36AR4A22 36ES4A15 36HS4A04 36HS4A54 36JN4A21 36KN4A05 36KN4A25 36AR4A04 36AR4A23 36ES4A17 36HS4A07 36HS4A55 36JN4A22 36KN4A06 36KN4A26 36AR4A05 36AR4A24 36ES4A18 36HS4A10 36HS4A61 36JN4A23 36KN4A07 36KN4A27 36AR4A06 36AR4A27 36ES4A19 36HS4A14 36JN4A03 36JN4A25 36KN4A08 36KN4A28 36AR4A07 36AR4A29 36ES4A20 36HS4A20 36JN4A04 36JN4A26 36KN4A09 36KN4A29 36AR4A08 36AR4A30 36ES4A22 36HS4A21 36JN4A06 36JN4A27 36KN4A10 36KN4A30 36AR4A09 36AR4A31 36ES4A28 36HS4A22 36JN4A07 36JN4A28 36KN4A11 36KN4A31 36AR4A10 36AR4A32 36ES4A33 36HS4A23 36JN4A08 36JN4A29 36KN4A12 36KN4A32 Recalling Firm/ Manufacturer Flowonix Medical Inc 120 Forbes Blvd Ste 170 Mansfield MA 02048-1150 For Additional Information Contact Technical Solutions Department 855-356-9665 Manufacturer Reason for Recall The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. Error 115 can also occur if dendrites form or if the pump is dropped. FDA Determined Cause 2 Software change control Action On March 21, 2019, the firm issued Urgent Recall Notices to customers to inform them of the product issue. ************In May 2019, the firm sent updated communications to customers:************Healthcare providers who have Clinician Programmers with SW 2.00.29 were sent a follow-up recall letter. This letter included the information that the pump stop following "Error Code 115" will only occur after the Bridge or Demand Bolus is complete, which could take up to seventy-two (72 hours). ************In addition, in the event of any cause of an "Error Code 115", the pump sounds a one-beep alarm and then shuts down. This behavior differs from the Instructions for Use, which state that the pump will sound three long (half second) beeps every 30 minutes for all critical errors. As it is unlikely that the patient will notice the single beep, there may be a significant delay in the patient becoming aware that the pump has stopped delivering therapy which may in turn delay resumption of drug therapy.************Healthcare providers who have Clinician Programmers with any other version of the software and the implant pumps received a Corrective Action Notice. This notice included the information from the preceding paragraph, noting the discrepant alarm behavior for "Error Code 115" could be missed by the patient.************Customers with the new Prometra Clinician Programmer software (v. 2.00.29) were asked to do the following:************1. STOP using the Periodic Flow or Multiple Rates flow mode functions in combination with a Bridge or******Demand Bolus using the Clinician Programmer with Software Version 2.00.29. The error only occurs when******these functions are used together.************a. In the event that a Bridge or Demand Bolus is needed, convert the patient to Constant Flow Mode,******then program the Bridge or Demand Bolus as needed (instructions for use PL-11803-04, page 30).******Once the bolus is complete, revisit the patient or have him/her return to your clinic for programming******back into their original flow mode.************2. IDENTIFY Quantity in Commerce 163 Distribution US nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LKK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On March 21, 2019, the firm issued Urgent Recall Notices to customers to inform them of the product issue. ************In May 2019, the firm sent updated communications to customers:************Healthcare providers who have Clinician Programmers with SW 2.00.29 were sent a follow-up recall letter. This letter included the information that the pump stop following "Error Code 115" will only occur after the Bridge or Demand Bolus is complete, which could take up to seventy-two (72 hours). ************In addition, in the event of any cause of an "Error Code 115", the pump sounds a one-beep alarm and then shuts down. This behavior differs from the Instructions for Use, which state that the pump will sound three long (half second) beeps every 30 minutes for all critical errors. As it is unlikely that the patient will notice the single beep, there may be a significant delay in the patient becoming aware that the pump has stopped delivering therapy which may in turn delay resumption of drug therapy.************Healthcare providers who have Clinician Programmers with any other version of the software and the implant pumps received a Corrective Action Notice. This notice included the information from the preceding paragraph, noting the discrepant alarm behavior for "Error Code 115" could be missed by the patient.************Customers with the new Prometra Clinician Programmer software (v. 2.00.29) were asked to do the following:************1. STOP using the Periodic Flow or Multiple Rates flow mode functions in combination with a Bridge or******Demand Bolus using the Clinician Programmer with Software Version 2.00.29. The error only occurs when******these functions are used together.************a. In the event that a Bridge or Demand Bolus is needed, convert the patient to Constant Flow Mode,******then program the Bridge or Demand Bolus as needed (instructions for use PL-11803-04, page 30).******Once the bolus is complete, revisit the patient or have him/her return to your clinic for programming******back into their original flow mode.************2. IDENTIFY | SoftwareChangeControl
56 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171924 | Z-1713-2019 | Prometra Programmer (Grand Prime) Software version 2.00.29, Cat. No. 13828, used with Prometra and Prometra II Programmable Pumps. For intrathecal infusion. | Prometra Clinician Programmer | 2 | CDRH | 2019-06-03 |  | 2020-06-05 09:23 | Flowonix Medical Inc | The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. | Class 2 Device Recall Prometra Clinician Programmer Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prometra Clinician Programmer FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prometra Clinician Programmer Date Initiated by Firm March 21, 2019 Create Date June 03, 2019 Recall Status 1 Terminated 3 on June 05, 2020 Recall Number Z-1713-2019 Recall Event ID 82586 PMA Number P080012 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product Prometra Programmer (Grand Prime) Software version 2.00.29, Cat. No. 13828, used with Prometra and Prometra II Programmable Pumps. For intrathecal infusion. Code Information Software v.2.00.29 Serial numbers: D18L0003 D18L0022 D18L0013 D18L0035 D18L0004 D18L0025 D18L0014 D18L0036 D18L0006 D18L0026 D18L0015 D18L0037 D18L0010 D18L0027 D18L0016 D18L0038 D18L0009 D18L0028 D18L0017 D18L0034 D18L0008 D18L0029 D18L0018 D18L0039 D18L0007 D18L0030 D18L0019 D18L0040 D18L0011 D18L0031 D18L0021 D18L0041 D18L0005 D18L0032 D18L0023 D18L0042 D18L0012 D18L0033 D18L0024 D18L0020 Recalling Firm/ Manufacturer Flowonix Medical Inc 120 Forbes Blvd Ste 170 Mansfield MA 02048-1150 For Additional Information Contact Technical Solutions Department 855-356-9665 Manufacturer Reason for Recall The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. FDA Determined Cause 2 Software change control Action On March 21, 2019, the firm issued Urgent Recall Notices to customers to inform them of the product issue. ************In May 2019, the firm sent updated communications to customers:************Healthcare providers who have Clinician Programmers with SW 2.00.29 were sent a follow-up recall letter. This letter included the information that the pump stop following "Error Code 115" will only occur after the Bridge or Demand Bolus is complete, which could take up to seventy-two (72 hours). ************In addition, in the event of any cause of an "Error Code 115", the pump sounds a one-beep alarm and then shuts down. This behavior differs from the Instructions for Use, which state that the pump will sound three long (half second) beeps every 30 minutes for all critical errors. As it is unlikely that the patient will notice the single beep, there may be a significant delay in the patient becoming aware that the pump has stopped delivering therapy which may in turn delay resumption of drug therapy.************Healthcare providers who have Clinician Programmers with any other version of the software and the implant pumps received a Corrective Action Notice. This notice included the information from the preceding paragraph, noting the discrepant alarm behavior for "Error Code 115" could be missed by the patient.************Customers with the new Prometra Clinician Programmer software (v. 2.00.29) were asked to do the following:************1. STOP using the Periodic Flow or Multiple Rates flow mode functions in combination with a Bridge or******Demand Bolus using the Clinician Programmer with Software Version 2.00.29. The error only occurs when******these functions are used together.************a. In the event that a Bridge or Demand Bolus is needed, convert the patient to Constant Flow Mode,******then program the Bridge or Demand Bolus as needed (instructions for use PL-11803-04, page 30).******Once the bolus is complete, revisit the patient or have him/her return to your clinic for programming******back into their original flow mode.************2. IDENTIFY Quantity in Commerce 40 Distribution US nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LKK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On March 21, 2019, the firm issued Urgent Recall Notices to customers to inform them of the product issue. ************In May 2019, the firm sent updated communications to customers:************Healthcare providers who have Clinician Programmers with SW 2.00.29 were sent a follow-up recall letter. This letter included the information that the pump stop following "Error Code 115" will only occur after the Bridge or Demand Bolus is complete, which could take up to seventy-two (72 hours). ************In addition, in the event of any cause of an "Error Code 115", the pump sounds a one-beep alarm and then shuts down. This behavior differs from the Instructions for Use, which state that the pump will sound three long (half second) beeps every 30 minutes for all critical errors. As it is unlikely that the patient will notice the single beep, there may be a significant delay in the patient becoming aware that the pump has stopped delivering therapy which may in turn delay resumption of drug therapy.************Healthcare providers who have Clinician Programmers with any other version of the software and the implant pumps received a Corrective Action Notice. This notice included the information from the preceding paragraph, noting the discrepant alarm behavior for "Error Code 115" could be missed by the patient.************Customers with the new Prometra Clinician Programmer software (v. 2.00.29) were asked to do the following:************1. STOP using the Periodic Flow or Multiple Rates flow mode functions in combination with a Bridge or******Demand Bolus using the Clinician Programmer with Software Version 2.00.29. The error only occurs when******these functions are used together.************a. In the event that a Bridge or Demand Bolus is needed, convert the patient to Constant Flow Mode,******then program the Bridge or Demand Bolus as needed (instructions for use PL-11803-04, page 30).******Once the bolus is complete, revisit the patient or have him/her return to your clinic for programming******back into their original flow mode.************2. IDENTIFY | SoftwareChangeControl
57 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171710 | Z-1137-2019 | Biomerieux Vitek 2 Test kit ******VITEK 2 is an automated system consisting of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. *** | Biomerieux Vitek 2 Test kit | 2 | CDRH | 2019-04-12 |  | 2021-08-17 17:54 | bioMerieux, Inc. | Unexpected ESBL (Extended Spectrum Beta Lactamase) phenotype had been proposed for some Escherichia coli strains in conjunction with the VITEK 2 AST-N254, AST-N330, AST-N371 (ref.413723,418674, 422024) test kits. | Class 2 Device Recall Biomerieux Vitek 2 Test kit Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Biomerieux Vitek 2 Test kit FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Biomerieux Vitek 2 Test kit Date Initiated by Firm March 15, 2019 Create Date April 12, 2019 Recall Status 1 Terminated 3 on August 17, 2021 Recall Number Z-1137-2019 Recall Event ID 82501 Product Classification System, test, automated, antimicrobial susceptibility, short incubation - Product Code LON Product Biomerieux Vitek 2 Test kit ************VITEK 2 is an automated system consisting of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information Ref: 413723 (AST-N254), 418674 (AST-N330) and 422024 (AST-N371 Card) with 8.01/9.01 SW Unique Device Identifier: 03573026404550 (AST-N254) / 03573026474836 (AST-N330) / 03573026604363 (AST-N371) Recalling Firm/ Manufacturer bioMerieux, Inc. 100 Rodolphe St Durham NC 27712-9402 For Additional Information Contact 919-620-2000 Manufacturer Reason for Recall Unexpected ESBL (Extended Spectrum Beta Lactamase) phenotype had been proposed for some Escherichia coli strains in conjunction with the VITEK 2 AST-N254, AST-N330, AST-N371 (ref.413723,418674, 422024) test kits. FDA Determined Cause 2 Software change control Action On 03/15/2019, the firm, BioMerieux sent notifications to all their consignees with the following instructions:******"Please take the following actions at this time:****** Confirm this letter has been distributed to, and reviewed by, all appropriate personnel within your organization.****** Apply the VITEK 2 Systems bioART rule criteria as previously mentioned (i.e., using the criteria listed below for ESBL / E. coli isolates). Refer to Appendix A for detailed instructions.******o If the card is AST-NXXX (where XXX indicates the card type designation, 254 / 330 / 371)******o The organism is Escherichia coli******o And the phenotype is ESBL******o And the drug cefuroxime is <=8******o And the drug ceftazidime is <=0.5******o Then stop for consultation and add a comment: ***There is no evidence that the isolate is an ESBL producer. The category results for cephalosporins should be interpreted as******tested, before the AES expertise.********* Please store this letter with your bioMrieux instrument documentation.****** Complete the Acknowledgement Form and return it to your local bioMrieux representative.************bioMrieux, Inc. is committed to providing our customers with the highest quality products, and we apologize for any inconvenience this may have caused in your laboratory. If you have any questions or concerns, please contact your local bioMrieux representative." Quantity in Commerce 10784 cartons Distribution Internationally to: United Kingdom, Latvia, Poland, Ukraine, Germany, and Scotland. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 03/15/2019, the firm, BioMerieux sent notifications to all their consignees with the following instructions:******"Please take the following actions at this time:****** Confirm this letter has been distributed to, and reviewed by, all appropriate personnel within your organization.****** Apply the VITEK 2 Systems bioART rule criteria as previously mentioned (i.e., using the criteria listed below for ESBL / E. coli isolates). Refer to Appendix A for detailed instructions.******o If the card is AST-NXXX (where XXX indicates the card type designation, 254 / 330 / 371)******o The organism is Escherichia coli******o And the phenotype is ESBL******o And the drug cefuroxime is <=8******o And the drug ceftazidime is <=0.5******o Then stop for consultation and add a comment: ***There is no evidence that the isolate is an ESBL producer. The category results for cephalosporins should be interpreted as******tested, before the AES expertise.********* Please store this letter with your bioMrieux instrument documentation.****** Complete the Acknowledgement Form and return it to your local bioMrieux representative.************bioMrieux, Inc. is committed to providing our customers with the highest quality products, and we apologize for any inconvenience this may have caused in your laboratory. If you have any questions or concerns, please contact your local bioMrieux representative." | SoftwareChangeControl
58 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=169845 | Z-0821-2019 | MOSAIQ¿ Oncology Information System and Sequencer******MOSAIQ is an oncology information system used to manage workflows for treatment planning and delivery. It supports information flow among healthcare facility personnel and can be used wherever radiotherapy and/or chemotherapy are prescribed. | MOSAIQ Oncology Information System and Sequencer | 2 | CDRH | 2019-02-08 |  |  | Elekta, Inc. | There is a potential that Wedge IDs were not included in the DICOM RT PLAN sent from the TPS, which will cause the field in MOSAIQ to be created with no wedge. Treatment of a field without the planned wedge will receive more dose than the treatment plan indicates. | Class 2 Device Recall MOSAIQ Oncology Information System and Sequencer Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall MOSAIQ Oncology Information System and Sequencer FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall MOSAIQ Oncology Information System and Sequencer Date Initiated by Firm December 13, 2018 Create Date February 08, 2019 Recall Status 1 Open 3 , Classified Recall Number Z-0821-2019 Recall Event ID 81785 510(K)Number K120067 K123230 K141572 K981313 Product Classification Accelerator, linear, medical - Product Code IYE Product MOSAIQ Oncology Information System and Sequencer************MOSAIQ is an oncology information system used to manage workflows for treatment planning and delivery. It supports information flow among healthcare facility personnel and can be used wherever radiotherapy and/or chemotherapy are prescribed. Code Information MOSAIQ Radiation Oncology version 1.0 and higher. Multi-ACCESS version 6.1 and higher. UDI 858164002084 858164002091 858164002107 858164002220 858164002237 Recalling Firm/ Manufacturer Elekta, Inc. 400 Perimeter Center Ter NE Ste 50 Atlanta GA 30346-1227 For Additional Information Contact Alejandra Carrillo 404-202-4333 Manufacturer Reason for Recall There is a potential that Wedge IDs were not included in the DICOM RT PLAN sent from the TPS, which will cause the field in MOSAIQ to be created with no wedge. Treatment of a field without the planned wedge will receive more dose than the treatment plan indicates. FDA Determined Cause 2 Software change control Action Elekta sent a letter to affected customers titled, Urgent Field Safety Notification, on about 12/13/2018. The letter contained recommended user action to review the TPS system to verify that Wedge IDs are defined for all wedges and check to make sure that the exported RT Plan information contains the necessary Wedge IDs. Instructions also included to post the recall notice in a place accessible to all users until the correction is closed, and to advise the appropriate personnel working with the affected product of the recall. ************The letter stated that the issue will be resolved in MOSAIQ Release 2.80 Beta 01. The update will be provided in Service Packs to MOSAIQ and Multi-ACCESS who will be notified with a Product Bulletin when the Service Packs are available. An Acknowledgement Form was sent with the letter that should be completed and returned to Elekta within 30 days. ************The firm can be contacted at 408-830- 8000 or 855-693-5358. Quantity in Commerce 3,950 units total (1,747 US) Distribution U.S. Nationwide distribution to the following states AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and the District of Columbia and Puerto Rico. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Elekta sent a letter to affected customers titled, Urgent Field Safety Notification, on about 12/13/2018. The letter contained recommended user action to review the TPS system to verify that Wedge IDs are defined for all wedges and check to make sure that the exported RT Plan information contains the necessary Wedge IDs. Instructions also included to post the recall notice in a place accessible to all users until the correction is closed, and to advise the appropriate personnel working with the affected product of the recall. ************The letter stated that the issue will be resolved in MOSAIQ Release 2.80 Beta 01. The update will be provided in Service Packs to MOSAIQ and Multi-ACCESS who will be notified with a Product Bulletin when the Service Packs are available. An Acknowledgement Form was sent with the letter that should be completed and returned to Elekta within 30 days. ************The firm can be contacted at 408-830- 8000 or 855-693-5358. | SoftwareChangeControl
59 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167480 | Z-0442-2019 | 3.2mm Proximal Reamer/Cannulated Drill provided as part of the Fibulock Fibular Nail System Pack. FibuLock Nail Procedure Pack: Contains instrumentation for the implantation of the FibuLock implant. Contents include: Actuation Driver, 6.2mm Reamer, 3.2mm Reamer, 2 mm Drill, Spade Tip Guide Wire, and 1.6mm x 12 K-Wire. | 3.2mm Proximal Reamer/Cannulated Drill | 2 | CDRH | 2018-11-09 |  |  | Arthrex, Inc. | There is potential to break during use. | Class 2 Device Recall 3.2mm Proximal Reamer/Cannulated Drill Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall 3.2mm Proximal Reamer/Cannulated Drill FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall 3.2mm Proximal Reamer/Cannulated Drill Date Initiated by Firm September 13, 2018 Create Date November 09, 2018 Recall Status 1 Open 3 , Classified Recall Number Z-0442-2019 Recall Event ID 81119 Product Classification Reamer - Product Code HTO Product 3.2mm Proximal Reamer/Cannulated Drill provided as part of the Fibulock Fibular Nail System Pack. FibuLock Nail Procedure Pack: Contains instrumentation for the implantation of the FibuLock implant. Contents include: Actuation Driver, 6.2mm Reamer, 3.2mm Reamer, 2 mm Drill, Spade Tip Guide Wire, and 1.6mm x 12 K-Wire. Code Information Part #ST6100, Lot/Serial #'s: A251217-04, A011217-01, A011217-03, A091017-05, A101017-01, A111017-01, A121017-01, A121017-02, A131017-01, A140917-01, A161017-01, A171017-01, A181017-07, A221117-01, A271117-01, A281117-02, A291117-01, A291117-03, A301117-01, A281117-01. Recalling Firm/ Manufacturer Arthrex, Inc. 1370 Creekside Blvd Naples FL 34108-1945 Manufacturer Reason for Recall There is potential to break during use. FDA Determined Cause 2 Software change control Action Arthrex sent letters to their consignees with the following instructions.******1. Immediately discontinue use, sale, and distribution of the affected product.******2. Please contact Arthrex Product Surveillance at 866-267-9138 or******complaints@arthrex.com as soon as possible to arrange for return of the******affected product. Our Product Surveillance Specialists can provide******assistance regarding alternative solutions and are available to answer******questions regarding credit for affected devices in your possession.******3. If you are an Arthrex agency, you do not need to notify any customers as******Arthrex will be notifying customers directly.******4. If you are an Arthrex customer, you do not need to notify any patients of******this product recall.******5. If you have any questions about this product recall, please contact Arthrex******Product Surveillance at 866-267-9138 or complaints@arthrex.com. Quantity in Commerce 832 Distribution Distributed throughout the U.S. to the following states: NC, CA, NJ, CO, NV, UT, LA, KY, OH, AZ, MO, MN, TN, NY, WI, MA, PA, TX, IL, PA, AR, FL, IA, MI, AL, SD, OR, WA, MD, KS, and GA. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Arthrex sent letters to their consignees with the following instructions.******1. Immediately discontinue use, sale, and distribution of the affected product.******2. Please contact Arthrex Product Surveillance at 866-267-9138 or******complaints@arthrex.com as soon as possible to arrange for return of the******affected product. Our Product Surveillance Specialists can provide******assistance regarding alternative solutions and are available to answer******questions regarding credit for affected devices in your possession.******3. If you are an Arthrex agency, you do not need to notify any customers as******Arthrex will be notifying customers directly.******4. If you are an Arthrex customer, you do not need to notify any patients of******this product recall.******5. If you have any questions about this product recall, please contact Arthrex******Product Surveillance at 866-267-9138 or complaints@arthrex.com. | SoftwareChangeControl
60 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=165490 | Z-0353-2019 | Intelligent Peri-Analytical Workstation (HW+SW), Catalog #: A89492 (iPAW Software v4.2.1), Beckman Coulter, Inc.****** The iPAW ((intelligent Peri-Analytical Workstation) can act as: a software client for specimen tracking and workload statistics or a physical workstation for manual sorting and label printing. | Intelligent Peri-Analytical Workstation | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes. Sorting Drive software version 4.2 was replaced by 4.2.1 (both versions contain the same option for manual barcode printing).******When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. | Class 3 Device Recall Intelligent PeriAnalytical Workstation Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall Intelligent PeriAnalytical Workstation FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall Intelligent PeriAnalytical Workstation Date Initiated by Firm May 24, 2018 Create Date October 31, 2018 Recall Status 1 Terminated 3 on December 08, 2022 Recall Number Z-0353-2019 Recall Event ID 80434 Product Classification Calculator/data processing module, for clinical use - Product Code JQP Product Intelligent Peri-Analytical Workstation (HW+SW), Catalog #: A89492 (iPAW Software v4.2.1), Beckman Coulter, Inc.************ The iPAW ((intelligent Peri-Analytical Workstation) can act as: a software client for specimen tracking and workload statistics or a physical workstation for manual sorting and label printing. Code Information v4.2.1 (and 4.2, discontinued) Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact 469-961-4941 Manufacturer Reason for Recall Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes. Sorting Drive software version 4.2 was replaced by 4.2.1 (both versions contain the same option for manual barcode printing).************When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. FDA Determined Cause 2 Software change control Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. Quantity in Commerce 283 Distribution Worldwide Distribution: US (nationwide) and countries of: Algeria, Argentina, Austria, Belgium, Brazil, Chile, China, Czech Republic, Denmark ,France, Germany, Greece ,Hong Kong, India, Ireland, Israel, Italy ,Republic of Korea, Lebanon, Lithuania ,Malaysia, Netherlands, Poland, Portugal , Qatar, Romania, Russian Federation, Saudi Arabia ,South Africa , Spain, Sweden, Taiwan, Turkey , and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. | SoftwareChangeControl
61 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167213 | Z-0357-2019 | Software Package, iPAW 4.2.1 User Kit for Windows XP and Windows 10, Part No. B95343, Manufactured by: Beckman Coulter Biomedical GmbH******Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes. | Software Package, Software Package, Sorting-Drive 4.2.1 User Kit for Windows XP and Windows 10 | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes.******When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. | Class 3 Device Recall Software Package, Software Package, SortingDrive 4.2.1 User Kit for Windows XP and Windows 10 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall Software Package, Software Package, SortingDrive 4.2.1 User Kit for Windows XP and Windows 10 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall Software Package, Software Package, SortingDrive 4.2.1 User Kit for Windows XP and Windows 10 Date Initiated by Firm May 24, 2018 Create Date October 31, 2018 Recall Status 1 Terminated 3 on December 08, 2022 Recall Number Z-0357-2019 Recall Event ID 80434 Product Classification Calculator/data processing module, for clinical use - Product Code JQP Product Software Package, iPAW 4.2.1 User Kit for Windows XP and Windows 10, Part No. B95343, Manufactured by: Beckman Coulter Biomedical GmbH************Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes. Code Information v4.2.1 (and 4.2, discontinued) Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact 469-961-4941 Manufacturer Reason for Recall Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes.************When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. FDA Determined Cause 2 Software change control Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. Quantity in Commerce 29 Distribution Worldwide Distribution: US (nationwide) and countries of: Algeria, Argentina, Austria, Belgium, Brazil, Chile, China, Czech Republic, Denmark ,France, Germany, Greece ,Hong Kong, India, Ireland, Israel, Italy ,Republic of Korea, Lebanon, Lithuania ,Malaysia, Netherlands, Poland, Portugal , Qatar, Romania, Russian Federation, Saudi Arabia ,South Africa , Spain, Sweden, Taiwan, Turkey , and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. | SoftwareChangeControl
62 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167211 | Z-0356-2019 | Software Package, iPAW 4.2 User Kit, Part No. B73489, Manufactured by: Beckman Coulter Biomedical GmbH******The iPAW 4.2 User Kit (PN B73489) includes the iPAW software application CD (version 4.2). This function is primarily used to copy an existing barcode ID to replace the label of an individual sample tube because the original label is deteriorated. | Software Package, iPAW 4.2 User Kit | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes.******When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. | Class 3 Device Recall Software Package, iPAW 4.2 User Kit Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall Software Package, iPAW 4.2 User Kit FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall Software Package, iPAW 4.2 User Kit Date Initiated by Firm May 24, 2018 Create Date October 31, 2018 Recall Status 1 Terminated 3 on December 08, 2022 Recall Number Z-0356-2019 Recall Event ID 80434 Product Classification Calculator/data processing module, for clinical use - Product Code JQP Product Software Package, iPAW 4.2 User Kit, Part No. B73489, Manufactured by: Beckman Coulter Biomedical GmbH************The iPAW 4.2 User Kit (PN B73489) includes the iPAW software application CD (version 4.2). This function is primarily used to copy an existing barcode ID to replace the label of an individual sample tube because the original label is deteriorated. Code Information v4.2.1 (and 4.2, discontinued) Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact 469-961-4941 Manufacturer Reason for Recall Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes.************When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. FDA Determined Cause 2 Software change control Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. Quantity in Commerce 14 Distribution Worldwide Distribution: US (nationwide) and countries of: Algeria, Argentina, Austria, Belgium, Brazil, Chile, China, Czech Republic, Denmark ,France, Germany, Greece ,Hong Kong, India, Ireland, Israel, Italy ,Republic of Korea, Lebanon, Lithuania ,Malaysia, Netherlands, Poland, Portugal , Qatar, Romania, Russian Federation, Saudi Arabia ,South Africa , Spain, Sweden, Taiwan, Turkey , and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. | SoftwareChangeControl
63 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167207 | Z-0354-2019 | Software Package, Sorting-Drive 4.2 User Kit, Part No. B73488, Manufactured by: Beckman Coulter Biomedical GmbH****** The Sorting-Drive itself is an optional software to be used with the Automate 2500 Family of products. The system automates the process of sorting, decapping, and archiving samples. The Sorting-Drive 4.2. User Kit includes the Sorting-Drive Client software application that includes an option for manual printing of barcodes for sample tubes. | Software Package, Sorting-Drive 4.2 User Kit | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option which included the Sorting-Drive Client that is used for manual printing of barcodes.******When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. | Class 3 Device Recall Software Package, SortingDrive 4.2 User Kit Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall Software Package, SortingDrive 4.2 User Kit FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall Software Package, SortingDrive 4.2 User Kit Date Initiated by Firm May 24, 2018 Create Date October 31, 2018 Recall Status 1 Terminated 3 on December 08, 2022 Recall Number Z-0354-2019 Recall Event ID 80434 Product Classification Calculator/data processing module, for clinical use - Product Code JQP Product Software Package, Sorting-Drive 4.2 User Kit, Part No. B73488, Manufactured by: Beckman Coulter Biomedical GmbH************ The Sorting-Drive itself is an optional software to be used with the Automate 2500 Family of products. The system automates the process of sorting, decapping, and archiving samples. The Sorting-Drive 4.2. User Kit includes the Sorting-Drive Client software application that includes an option for manual printing of barcodes for sample tubes. Code Information v4.2.1 (and 4.2, discontinued) Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact 469-961-4941 Manufacturer Reason for Recall Sorting-Drive software version 4.2 (released in October 2015) introduced a new option which included the Sorting-Drive Client that is used for manual printing of barcodes.************When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. FDA Determined Cause 2 Software change control Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. Quantity in Commerce 25 Distribution Worldwide Distribution: US (nationwide) and countries of: Algeria, Argentina, Austria, Belgium, Brazil, Chile, China, Czech Republic, Denmark ,France, Germany, Greece ,Hong Kong, India, Ireland, Israel, Italy ,Republic of Korea, Lebanon, Lithuania ,Malaysia, Netherlands, Poland, Portugal , Qatar, Romania, Russian Federation, Saudi Arabia ,South Africa , Spain, Sweden, Taiwan, Turkey , and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. | SoftwareChangeControl
64 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167209 | Z-0355-2019 | Software Package, Sorting-Drive 4.2.1 User Kit for Windows XP and Windows 10, Part No. B95347, Manufactured by: Beckman Coulter Biomedical GmbH******The Sorting-Drive itself is an optional software to be used with the Automate 2500 Family of products. . The Sorting-Drive 4.2.1 User Kit includes the Sorting-Drive Client software application that includes an option for manual printing of barcodes for sample tubes; version 4.2 was replaced by version 4.2.1 (they both contain the same option for manual barcode printing). This function is primarily used to copy an existing barcode ID to replace the label of an individual sample tube because the original label is deteriorated. | Software Package, Sorting-Drive 4.2.1 User Kit for Windows XP and Windows 10 | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option which included the Sorting-Drive Client that is used for manual printing of barcodes; version 4.2 was replaced by 4.2.1 (they both contain the same option for manual barcode printing).******When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs.The firm reports this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. | Class 3 Device Recall Software Package, SortingDrive 4.2.1 User Kit for Windows XP and Windows 10 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall Software Package, SortingDrive 4.2.1 User Kit for Windows XP and Windows 10 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall Software Package, SortingDrive 4.2.1 User Kit for Windows XP and Windows 10 Date Initiated by Firm May 24, 2018 Create Date October 31, 2018 Recall Status 1 Terminated 3 on December 08, 2022 Recall Number Z-0355-2019 Recall Event ID 80434 Product Classification Calculator/data processing module, for clinical use - Product Code JQP Product Software Package, Sorting-Drive 4.2.1 User Kit for Windows XP and Windows 10, Part No. B95347, Manufactured by: Beckman Coulter Biomedical GmbH************The Sorting-Drive itself is an optional software to be used with the Automate 2500 Family of products. . The Sorting-Drive 4.2.1 User Kit includes the Sorting-Drive Client software application that includes an option for manual printing of barcodes for sample tubes; version 4.2 was replaced by version 4.2.1 (they both contain the same option for manual barcode printing). This function is primarily used to copy an existing barcode ID to replace the label of an individual sample tube because the original label is deteriorated. Code Information v4.2.1 (and 4.2, discontinued) Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact 469-961-4941 Manufacturer Reason for Recall Sorting-Drive software version 4.2 (released in October 2015) introduced a new option which included the Sorting-Drive Client that is used for manual printing of barcodes; version 4.2 was replaced by 4.2.1 (they both contain the same option for manual barcode printing).************When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs.The firm reports this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. FDA Determined Cause 2 Software change control Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. Quantity in Commerce 54 Distribution Worldwide Distribution: US (nationwide) and countries of: Algeria, Argentina, Austria, Belgium, Brazil, Chile, China, Czech Republic, Denmark ,France, Germany, Greece ,Hong Kong, India, Ireland, Israel, Italy ,Republic of Korea, Lebanon, Lithuania ,Malaysia, Netherlands, Poland, Portugal , Qatar, Romania, Russian Federation, Saudi Arabia ,South Africa , Spain, Sweden, Taiwan, Turkey , and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. | SoftwareChangeControl
65 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167218 | Z-0359-2019 | Sorting-Drive Professional Part, Sorting-Drive Professional Kit, (Sorting Drive Software v4.2.1), Part No. A94144, Beckman Coulter, Inc.****** The Sorting-Drive Professional Kit (PN A94144) includes the Sorting-Drive 4.2.1 User Kit (PN B95347) which includes the Sorting-Drive 4.2.1 Software DVD (PN B95346). This software version includes an option for manual barcode printing for sample tubes. Sorting-Drive software facilitates the Laboratory Information System (LIS) connection to AutoMate 2500 Family systems. | Sorting-Drive Professional Part, Sorting-Drive Professional Kit, (Sorting Drive Software v4.2.1) | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes.Sorting Drive software version 4.2 was replaced by 4.2.1 (both versions contain the same option for manual barcode printing).******When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. | Class 3 Device Recall SortingDrive Professional Part, SortingDrive Professional Kit, (Sorting Drive Software v4.2.1) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall SortingDrive Professional Part, SortingDrive Professional Kit, (Sorting Drive Software v4.2.1) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall SortingDrive Professional Part, SortingDrive Professional Kit, (Sorting Drive Software v4.2.1) Date Initiated by Firm May 24, 2018 Create Date October 31, 2018 Recall Status 1 Terminated 3 on December 08, 2022 Recall Number Z-0359-2019 Recall Event ID 80434 Product Classification Calculator/data processing module, for clinical use - Product Code JQP Product Sorting-Drive Professional Part, Sorting-Drive Professional Kit, (Sorting Drive Software v4.2.1), Part No. A94144, Beckman Coulter, Inc.************ The Sorting-Drive Professional Kit (PN A94144) includes the Sorting-Drive 4.2.1 User Kit (PN B95347) which includes the Sorting-Drive 4.2.1 Software DVD (PN B95346). This software version includes an option for manual barcode printing for sample tubes. Sorting-Drive software facilitates the Laboratory Information System (LIS) connection to AutoMate 2500 Family systems. Code Information v4.2.1 (and 4.2, discontinued) Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact 469-961-4941 Manufacturer Reason for Recall Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes.Sorting Drive software version 4.2 was replaced by 4.2.1 (both versions contain the same option for manual barcode printing).************When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. FDA Determined Cause 2 Software change control Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. Quantity in Commerce 58 Distribution Worldwide Distribution: US (nationwide) and countries of: Algeria, Argentina, Austria, Belgium, Brazil, Chile, China, Czech Republic, Denmark ,France, Germany, Greece ,Hong Kong, India, Ireland, Israel, Italy ,Republic of Korea, Lebanon, Lithuania ,Malaysia, Netherlands, Poland, Portugal , Qatar, Romania, Russian Federation, Saudi Arabia ,South Africa , Spain, Sweden, Taiwan, Turkey , and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. | SoftwareChangeControl
66 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167217 | Z-0358-2019 | Sorting-Drive Product Part, Sorting-Drive Plus Kit (Sorting Drive Software v4.2.1), Part No. A89483, Beckman Coulter, Inc.******The Sorting-Drive Plus Kit includes the Sorting-Drive 4.2.1 User Kit which includes the Sorting-Drive 4.2.1 Software DVD. . The Sorting-Drive software facilitates the Laboratory Information System (LIS) connection to AutoMate 2500 Family systems. . | Sorting-Drive Product Part, Sorting-Drive Plus Kit (Sorting Drive Software v4.2.1) | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes. Sorting Drive software version 4.2 was replaced by 4.2.1 (both versions contain the same option for manual barcode printing).******When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. | Class 3 Device Recall SortingDrive Product Part, SortingDrive Plus Kit (Sorting Drive Software v4.2.1) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall SortingDrive Product Part, SortingDrive Plus Kit (Sorting Drive Software v4.2.1) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall SortingDrive Product Part, SortingDrive Plus Kit (Sorting Drive Software v4.2.1) Date Initiated by Firm May 24, 2018 Create Date October 31, 2018 Recall Status 1 Terminated 3 on December 08, 2022 Recall Number Z-0358-2019 Recall Event ID 80434 Product Classification Calculator/data processing module, for clinical use - Product Code JQP Product Sorting-Drive Product Part, Sorting-Drive Plus Kit (Sorting Drive Software v4.2.1), Part No. A89483, Beckman Coulter, Inc.************The Sorting-Drive Plus Kit includes the Sorting-Drive 4.2.1 User Kit which includes the Sorting-Drive 4.2.1 Software DVD. . The Sorting-Drive software facilitates the Laboratory Information System (LIS) connection to AutoMate 2500 Family systems. . Code Information v4.2.1 (and 4.2, discontinued) Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact 469-961-4941 Manufacturer Reason for Recall Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes. Sorting Drive software version 4.2 was replaced by 4.2.1 (both versions contain the same option for manual barcode printing).************When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected. FDA Determined Cause 2 Software change control Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. Quantity in Commerce 495 Distribution Worldwide Distribution: US (nationwide) and countries of: Algeria, Argentina, Austria, Belgium, Brazil, Chile, China, Czech Republic, Denmark ,France, Germany, Greece ,Hong Kong, India, Ireland, Israel, Italy ,Republic of Korea, Lebanon, Lithuania ,Malaysia, Netherlands, Poland, Portugal , Qatar, Romania, Russian Federation, Saudi Arabia ,South Africa , Spain, Sweden, Taiwan, Turkey , and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 6/4/2018 Beckman Coulter (BEC) reported that the issue will be corrected via a field modification performed by Field Service. BEC reported that they do not intend to send a customer letter for this issue.US customers can contact the Technical Support Troubleshooting Hotline at (800) 854-3633. | SoftwareChangeControl
67 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=165764 | Z-0111-2019 | GE Healthcare MRI Systems, Software Version 12.x, sold under the following product names: ******SIGNA MR/I ECHOSPD W/ EXCITE, 1.5 HISPEED HD 8-CH 400-AP, MR GOLDSEAL 1.5T HD 8CH USED, UPGRADE 1.5 EXC-II TO EXC-III HD, SIGNA MR/I HISPD W/ EXCITE, 1.5 ECHOSPEED HD 16-CH 800-AP, 1.5 HISPEED HD 4-CH 200-AP, 1.5 TWINSPEED HD 8CH-400-AP, MR UPGRADE LX/HZ TO HD FOR NON LCC, 1.5 TWINSPEED HD 16CH-800-AP, ***1-5T HDXT TWINSPEED 16 CHANNEL, SIGNA MR/I TWISP TR W/EXCITE, 3T HD 16CH-800-AP, 1.5 ECHOSPEED HD 4-CH 200-AP, UPGRADE PREEXCITE TO 1.5TS HD 8CH, MR GOLDSEAL 1.5T HDXT 16CH NEW, MR GOLDSEAL 1.5T HD ES 8CH MOB, 1-5T NON LCC LX/5_5 TO HDIUPG8CH, UPGRADE 1.5 EXC-II TO EXC-III 16HD, UPGRADE 1.5 HORIZON TO HS/ES EXCITE, UPGRADE FROM 1.5 4X TO HISPEED LX, 1.5 LX TO HD 8CH FORKLIFT UPG., MR GOLDSEAL 1.5 LX HISP PL NEW, MR UPGRADE 4X/5X TO HD FOR NON LCC, MR GOLDSEAL 1.5 ES+ LX, SIGNA HISPEED LX 1.5T FIXED, 3T HD T/R W/EXCITE, 1.5 ECHOSPEED HD 8-CH 400-A, SIGNA 1.5 ECHOSPEED HD 8-CH 400-A, SIGNA EXCITE HD 1.5T, MR EXCITE12X, MR GOLDSEAL 1.5T HD 8CH USE, 1.5T BRM ES 16CH 2 VRE FIXED F, UPG 1.5 EXC-II TO EXC-III HD, EXCITE HD - UPG 3.5 EXC, 1.5 ECHOSPEED HD 4-CH 200, SIGNA HD ADVANCED, EXCITE HD, 1.5 ECHOSPEED HD 16-CH, MR 1.5T HD, 1.5 ECHOSPEED HD 8-CH, MR 1.5 EXCITE HD, MR GOLDSEAL 1.5T HD ES 8CH NEW, 1.5T EXCITE HD, ***SIGNA ECHOSPD LX K4 1.5T FIXED, 1.5T LX TO 8CH HD 23 UPG, 1.5 HISPEED HD 8-C, SIGNA EXCITE 1.5T HD, MR 1.5T HD ECHOSPEED WITH EXCITE, MR 1.5T HD TWINSPE, MR 1.5T HD ECHOSPEED, MR SIGNA VH/I 3.0T, MR 3.0T HD 16CH, MR 1.5T HD TWINSPEED, MR 3.0T HD 8CH, MR UPG TO 1.5 HD, MR UPG TO 1.5T HD, MR UPG 1.5ES HD 8CH, SIGNA EXCITE HD 1.5 T, SIGNA EXCITE HD, SIGNA MR/I ECHOSPD 1.5T FIX(+75798), SIGNA ES EXCITE, SIGNA 1.5T ECHO HD, MR HD (VON FA. LINK), SIGNA 1.5T ECHOSPEED HD, SIGNA MR 1.5T ECHO HD, SIGNA EXITE-III HD, SIGNA 1.5T HD, ECHOSPEED HD 8-CH, 1.5 ECHOSPEED HD 8CH, MR 1.5T EXCITE HD, MR SIGNA EXCITE 3T HD, MR 1.5 ECHOSPEED HD 8-CH 400-AP, 1.5 ECHOSPEED HD, 1.5T HD, 1.5T HD PREOWN, 1.5T ECHO SPEED, 1.5 ECHOSPEED HD , SIGNA MR 1.5 ECHOSPEED 8CH, ***SIGNA HD 1.5 ES, MR ECHOSPEED 1.5 HD, SIGNA 1.5TECHOSPEED, MR 1.5 ECHOSPEED HD 8-CH, 1.5T HD EXCITE, SIGNA 1.5 T ECHOSPEED, SIGNA VH/I 3.0T, MR 1.5T HD 16CH, MR 1.5T HD 8CH USED, SIGNA EXCITE3 HD, MR SIGNA EXCITE HD (UPGR.), 1.5T EXCITE-HD, 1.5T EX-HD TWIN, SIGNA 3T, MR 1.5 ECHOSPEED HD 8-CH 400-A, 1.5T EXCITE HD GS, 3T E2, 1.5 ECHOSPEED HD 8-CH 400-, SIGNA EXCITE III HD 1.5T TWIN SPEED, MR GOLDSEAL 1.5T HD 16CH USED, EXCITE II UPG, SIGNA EXCITE III HD 3.0T, MR SIGNA EXCITE, 1.5T SIGNA EXCITE III HD, MR SIGNA HD, SIGNA EXCITE HD 1.5, 3T EXCITE HD, UPG PREEXCITE TO 1.5ES HD 8CH, MR SIGNA SMARTSPEED 1.5, 1.5T HD ECHOSPEED 8 CHANNEL, SIGNA 1.5 MOBILE EXCITE HD, SIGNA 3T HD 16CH-800-AP, SIGNA 3T HD 8CH-400-AP, 1.5T MRI EXCITE 3 UPGRADE, MR 1.5 ECHOSPEED HD 4-CH 200-AP, SIGNA 1.5 ECHOSPEED HD 8-CH 400-AP, SIGNA EXCITE 1.5T ECHOSPEED HD 4CH, MR 3T HD 8CH-400-AP, SIGNA 1.5T ECHOSPEED HD 4-CH 200-AP, 1.5 ECHOSPEED HD 16CH 800AP-2669585, 1.5 ECHOSPEED HD 8-CH 400-AP2669575, 1.5T ECHOSPEED HD16-CH800-AP2669583, UPG 1.5 EXC-II TO EXC-III HD-561936, 1.5 TWINSPEED HD 8CH-400-AP-2583394, 1.5 ECHOSPEED HD 8-CH 400-AP2700482, 1.5 ECHOSPEED HDMR, 1.5 ECHOSPEED HD 8-CH 400-AP EXCITE, MR 1.5 ECHOSPEED HD 8-C, MR SIGNA 1.5T HD 8CH, MR SIGNA 15T HD UPGRADED, MR SIGNA 3T HD 8CH-400-AP, MR SIGNA 1.5T HD UPGRADED, MR SIGNA 1.5T HD 8CH MOBIL, EXCITE HD 1.5T UPGRADES, MR SIGNA EXCITE HD 15T, MR SIGNA EXCITE HD 1.5T, MR 1.5 ECHOSPEED HD 16-CH, MR PROFILE HD PHASE ARRAY, SIGNA 3T EX HD, MR SIGNA VH/I 3.0, MR SIGNA MR/I HISPEED 1.5T, 1.5T ECHOSPEED HD, 1.5 ECHOSPEED HD 8, 1.5T ECHOSPEED HD 8, 1.5 T ECHOSPEED HD 4-CH 200-AP, SIGNA EXCITE 3T HD 8CH-400-AP, MR 1.5T ECHOSPEED HD 8-CH 400-AP, MR 1.5 ECHOSPEED HD 4-CH 200-A, MR/I 1.5 ECHOSPEED HD 4CH 200-, MR 15T ECHOSPEED HD 8-CH 400-, MR 15 ECHOSPEED HD 4-CH 200-A, MR 1.5 ECHOSPEED HD 8-CH 400-AP//, SIGNA 1.5T, SIGNA MR/I ECHOSPD W/ EXCIT, 1.5T EXCITE HD 16 CH, 1.5T HD E | GE Healthcare MRI Systems, Software Version 12.x | 2 | CDRH | 2018-10-15 |  | 2020-06-18 09:37 | GE Medical Systems, LLC | The latest on-site software version may not have been reinstalled at some sites after service activities were performed resulting in images possibly flipped left/right and/or there can be patient data mismatch. | Class 2 Device Recall GE Healthcare MRI Systems, Software Version 12.x Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GE Healthcare MRI Systems, Software Version 12.x FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GE Healthcare MRI Systems, Software Version 12.x Date Initiated by Firm June 07, 2018 Create Date October 15, 2018 Recall Status 1 Terminated 3 on June 18, 2020 Recall Number Z-0111-2019 Recall Event ID 80323 510(K)Number K012200 K013636 K030874 K040444 K041476 K052293 Product Classification System, nuclear magnetic resonance imaging - Product Code LNH Product GE Healthcare MRI Systems, Software Version 12.x, sold under the following product names: ************SIGNA MR/I ECHOSPD W/ EXCITE, 1.5 HISPEED HD 8-CH 400-AP, MR GOLDSEAL 1.5T HD 8CH USED, UPGRADE 1.5 EXC-II TO EXC-III HD, SIGNA MR/I HISPD W/ EXCITE, 1.5 ECHOSPEED HD 16-CH 800-AP, 1.5 HISPEED HD 4-CH 200-AP, 1.5 TWINSPEED HD 8CH-400-AP, MR UPGRADE LX/HZ TO HD FOR NON LCC, 1.5 TWINSPEED HD 16CH-800-AP, ******1-5T HDXT TWINSPEED 16 CHANNEL, SIGNA MR/I TWISP TR W/EXCITE, 3T HD 16CH-800-AP, 1.5 ECHOSPEED HD 4-CH 200-AP, UPGRADE PREEXCITE TO 1.5TS HD 8CH, MR GOLDSEAL 1.5T HDXT 16CH NEW, MR GOLDSEAL 1.5T HD ES 8CH MOB, 1-5T NON LCC LX/5_5 TO HDIUPG8CH, UPGRADE 1.5 EXC-II TO EXC-III 16HD, UPGRADE 1.5 HORIZON TO HS/ES EXCITE, UPGRADE FROM 1.5 4X TO HISPEED LX, 1.5 LX TO HD 8CH FORKLIFT UPG., MR GOLDSEAL 1.5 LX HISP PL NEW, MR UPGRADE 4X/5X TO HD FOR NON LCC, MR GOLDSEAL 1.5 ES+ LX, SIGNA HISPEED LX 1.5T FIXED, 3T HD T/R W/EXCITE, 1.5 ECHOSPEED HD 8-CH 400-A, SIGNA 1.5 ECHOSPEED HD 8-CH 400-A, SIGNA EXCITE HD 1.5T, MR EXCITE12X, MR GOLDSEAL 1.5T HD 8CH USE, 1.5T BRM ES 16CH 2 VRE FIXED F, UPG 1.5 EXC-II TO EXC-III HD, EXCITE HD - UPG 3.5 EXC, 1.5 ECHOSPEED HD 4-CH 200, SIGNA HD ADVANCED, EXCITE HD, 1.5 ECHOSPEED HD 16-CH, MR 1.5T HD, 1.5 ECHOSPEED HD 8-CH, MR 1.5 EXCITE HD, MR GOLDSEAL 1.5T HD ES 8CH NEW, 1.5T EXCITE HD, ******SIGNA ECHOSPD LX K4 1.5T FIXED, 1.5T LX TO 8CH HD 23 UPG, 1.5 HISPEED HD 8-C, SIGNA EXCITE 1.5T HD, MR 1.5T HD ECHOSPEED WITH EXCITE, MR 1.5T HD TWINSPE, MR 1.5T HD ECHOSPEED, MR SIGNA VH/I 3.0T, MR 3.0T HD 16CH, MR 1.5T HD TWINSPEED, MR 3.0T HD 8CH, MR UPG TO 1.5 HD, MR UPG TO 1.5T HD, MR UPG 1.5ES HD 8CH, SIGNA EXCITE HD 1.5 T, SIGNA EXCITE HD, SIGNA MR/I ECHOSPD 1.5T FIX(+75798), SIGNA ES EXCITE, SIGNA 1.5T ECHO HD, MR HD (VON FA. LINK), SIGNA 1.5T ECHOSPEED HD, SIGNA MR 1.5T ECHO HD, SIGNA EXITE-III HD, SIGNA 1.5T HD, ECHOSPEED HD 8-CH, 1.5 ECHOSPEED HD 8CH, MR 1.5T EXCITE HD, MR SIGNA EXCITE 3T HD, MR 1.5 ECHOSPEED HD 8-CH 400-AP, 1.5 ECHOSPEED HD, 1.5T HD, 1.5T HD PREOWN, 1.5T ECHO SPEED, 1.5 ECHOSPEED HD , SIGNA MR 1.5 ECHOSPEED 8CH, ******SIGNA HD 1.5 ES, MR ECHOSPEED 1.5 HD, SIGNA 1.5TECHOSPEED, MR 1.5 ECHOSPEED HD 8-CH, 1.5T HD EXCITE, SIGNA 1.5 T ECHOSPEED, SIGNA VH/I 3.0T, MR 1.5T HD 16CH, MR 1.5T HD 8CH USED, SIGNA EXCITE3 HD, MR SIGNA EXCITE HD (UPGR.), 1.5T EXCITE-HD, 1.5T EX-HD TWIN, SIGNA 3T, MR 1.5 ECHOSPEED HD 8-CH 400-A, 1.5T EXCITE HD GS, 3T E2, 1.5 ECHOSPEED HD 8-CH 400-, SIGNA EXCITE III HD 1.5T TWIN SPEED, MR GOLDSEAL 1.5T HD 16CH USED, EXCITE II UPG, SIGNA EXCITE III HD 3.0T, MR SIGNA EXCITE, 1.5T SIGNA EXCITE III HD, MR SIGNA HD, SIGNA EXCITE HD 1.5, 3T EXCITE HD, UPG PREEXCITE TO 1.5ES HD 8CH, MR SIGNA SMARTSPEED 1.5, 1.5T HD ECHOSPEED 8 CHANNEL, SIGNA 1.5 MOBILE EXCITE HD, SIGNA 3T HD 16CH-800-AP, SIGNA 3T HD 8CH-400-AP, 1.5T MRI EXCITE 3 UPGRADE, MR 1.5 ECHOSPEED HD 4-CH 200-AP, SIGNA 1.5 ECHOSPEED HD 8-CH 400-AP, SIGNA EXCITE 1.5T ECHOSPEED HD 4CH, MR 3T HD 8CH-400-AP, SIGNA 1.5T ECHOSPEED HD 4-CH 200-AP, 1.5 ECHOSPEED HD 16CH 800AP-2669585, 1.5 ECHOSPEED HD 8-CH 400-AP2669575, 1.5T ECHOSPEED HD16-CH800-AP2669583, UPG 1.5 EXC-II TO EXC-III HD-561936, 1.5 TWINSPEED HD 8CH-400-AP-2583394, 1.5 ECHOSPEED HD 8-CH 400-AP2700482, 1.5 ECHOSPEED HDMR, 1.5 ECHOSPEED HD 8-CH 400-AP EXCITE, MR 1.5 ECHOSPEED HD 8-C, MR SIGNA 1.5T HD 8CH, MR SIGNA 15T HD UPGRADED, MR SIGNA 3T HD 8CH-400-AP, MR SIGNA 1.5T HD UPGRADED, MR SIGNA 1.5T HD 8CH MOBIL, EXCITE HD 1.5T UPGRADES, MR SIGNA EXCITE HD 15T, MR SIGNA EXCITE HD 1.5T, MR 1.5 ECHOSPEED HD 16-CH, MR PROFILE HD PHASE ARRAY, SIGNA 3T EX HD, MR SIGNA VH/I 3.0, MR SIGNA MR/I HISPEED 1.5T, 1.5T ECHOSPEED HD, 1.5 ECHOSPEED HD 8, 1.5T ECHOSPEED HD 8, 1.5 T ECHOSPEED HD 4-CH 200-AP, SIGNA EXCITE 3T HD 8CH-400-AP, MR 1.5T ECHOSPEED HD 8-CH 400-AP, MR 1.5 ECHOSPEED HD 4-CH 200-A, MR/I 1.5 ECHOSPEED HD 4CH 200-, MR 15T ECHOSPEED HD 8-CH 400-, MR 15 ECHOSPEED HD 4-CH 200-A, MR 1.5 ECHOSPEED HD 8-CH 400-AP//, SIGNA 1.5T, SIGNA MR/I ECHOSPD W/ EXCIT, 1.5T EXCITE HD 16 CH, 1.5T HD E Code Information Model Number, Serial Number (System ID) : 2226300, 000000000R4938 (256329MR2); 2226300, 000000000R5598 (256ALOPMR); N/A, N/A (334699MR); 2377062-2, 00000002547YR8 (256265HMR); 5107849, 00000274691MR5 (205487MR2); 46-294231G1, 0002005-0411-5 (870972E16); 46-294231G1, 002005-1201-35 (623583MR1); 2377062-2, 00000963660YM6 (480323MR1); 5107849-2, 00000278561MR6 (818445MR); 46-294231G1, 002004-1028-04 (3825); 2377062, 00000941225YM5 (626335MR); 46-294231G1, 002004-0913-01 (3591); 46-294231G1, 002005-1018-15 (352333MR3T); N/A, N/A (954587MR1); 46-294231G1, 002002-0605-11 (781340MR); 5107849, 266201MR3 (RADNETMR1585); 46-294231G1, 002005-0110-28 (313593MR02); 5308573, 00000292781MR2 (609890MR); 2214149-2, 00000010224GQ4 (732745MR); 2377062-2, 00000963098YM9 (607798MREXCT); N/A, N/A (516825MR); N/A, N/A (405307NMR); N/A, N/A (814535MR1); 5107849-2, 00000283624MR5 (615588MRI); N/A, N/A (615473MR1); N/A, N/A (817922MRLX); 5107849-2, 00000278623MR4 (972378NSC3TMR1); 46-294231G1, 0002005-0823-7 (DZ1057MR01); 46-294231G1, 002006-0127-10 (DZ1193MR01); 2226300, 00000000R5456A (2722SIGNA15); 2226300, 000000000R6230 (4291EXC15T); 2226300, 000000000R4601 (MRR4601); 2226300, 000000000R4973 (0910272001); 2226300, 000000000R5899 (0910274027); M3000PT, 00000280570MR3 (0910273001); 46-294231G1, 002005-0511-13 (AZ1003MR01); 46-294231G1, 0000099-0111-2 (110037MR01); 46-294231G1, 002005-0301-06 (2560124); 2226300, 000000000R5054 (BA4073MR01); N/A, 000000000R6295 (2712857); 5308573, 00000290734MR3 (2771768); 2226300, 000000000R6853 (2776713); 5107849-2, 00000284596MR4 (1236631); 2226300, 000000000R5190 (1191369); N/A, N/A (270004); 2189500-3, 00000201911MR5 (22946-1); 46-294231G1, 002005-0301-18 (CATMR01); 46-294231G1, 000005-1212-11 (CRMR01); 2226300, 000000000R5504 (082427310017); 2215261, 00000215559MR6 (082427040057); 2226300, 000000000R5271 (082427040072); N/A, 00000243366MR2 (082427040040); 46-294231G1, 0000003-0315-2 (082427040061); 2226300, 000000000R6344 (082427040074); 5107849-2, 00000289798MR1 (082427100093); 5107849-2, 00000283174MR1 (082427140047); 5107849-2, 00000282831MR7 (082427140046); 5107849-2, 00000283662MR5 (082427210041); 5107849-2, 00000282959MR6 (2697548EXCITE); 2226300, 000000000R5692 (2694100HD); N/A, N/A (CY1028MR01); 2226300, 000000000R4124 (PC0114MR03); 46-294231G1, 000006-0425-19 (HC4719MR02); 5107849, 00000272346MR8 (CC5340MR02); N/A, 0002005-1020-4 (126597MR01); 46-294231G1, 0002005-0511-6 (125390MR01); 46-294231G1, 000091-0919-12 (PSTU02MR01); 2226300, 000000000R5851 (2704833SIGNA); N/A, 000000000R5205 (GF1002MR01); 2226300, 000000000R4942 (24527633818MRI); 46-294231G1, 002005-0519-10 (0828270022IB01); 46-294231G1, 002005-0216-12 (0828270021IB01); N/A, 0000006-0323-2 (HU1099MR01); 5107849-2, 00000284297MR9 (083027062101067); 5107849, 00000267657MR5 (083027826593614); 46-294231G1, 000007-0530-11 (083027592002490); N/A, N/A (083027062101391); 2226300, 000000000R5902 (083027602101557); N/A, N/A (0835270007); 2108800-2, 535779 (B5174306); 46-294231G1, 0002005-0714-6 (A126322002); N/A, N/A (B5183714); 2226300, 000000000R5542 (B5367312); 2377062-2, 00000963588YM9 (A6118004); N/A, N/A (A148761302); N/A, 00000189273MR6 (A5155133); 2226300, 000000000R5739 (A377686701); 2226300, 000000000R5390 (A6134706); 46-294231G1, 002001-0702-56 (A5104397); 2357500-3, 00000300710MR1 (A5739706); 46-294231G1, 002005-0805-20 (YM1503); 2226300, 000000000R4952 (YM1516); 2198999-2, 00000158807MR8 (YM0451); N/A, 000000000R4216 (JO1029MR01); 2226300, 000000000R1382 (0850270404); N/A, 00000271322MR0 (0850270094); 5107849, 00000282503MR2 (LV4038MR01); 46-294231G1, 0002005-1130-3 (1084MREINP); 5107849, 00000263976MR3 (MR262474MR0); 2377062-2, 00000963560YM8 (5606SIGNA); 46-294231G1, 002005-0411-13 (56MRE3TH2565370); 2226300, 000000000R5911 (813EXCITEHD); 2226300, 000000000R5037 (2640781MR-1); 2226300, 000000000R5966 (99251SIGNAHD); 46-294231G1, 000006-0403-14 (99201MR15); 2395001-2, 00000280669MR3 (0920279025); 46-294231G1, 0000099-0722-5 (0920279011); N/A, N/A (25565916MR); 2226300, 00000000R4926A (PL1530MR02); 46-294231G1, 000006-0308-13 (787740MOBILE); 5107849, 00000276591MR5 (787653MR3T); 46-294231G1, 0000006-0315-5 (787840CIRT3T); 46-294231G1, 002001-0530-30 (787840EXCITE); N/A, 000000000R5552 (RO1115MR01); 2226300, 000000000R5057 (RU1195MR01); 46-294231G1, 002005-1212-26 (600130MR02); 5107849, 00000275845MR6 (RU1283MR01); 46-294231G1, 000006-0412-16 (RU1296MR01); 2226300, 000000000R5862 (SA1014MR02); 2226300, 000000000R5545 (SA2051MR01); 2226300, 000000000R5699 (SA1041MR02); 46-294231G1, 002005-0328-20 (SA1004MR02); N/A, 002005-1018-11 (SA1006MR02); 2226300, 000000000R5764 (SA1021MR01); 2226300, 000000000R4655 (0847270017); 46-294231G1, 002005-1201-14 (ZA1180MR01); 2226300, 000000000R5102 (NPV31703); 2226300, 000000000R4415 (MPX33111); 2357500-2, 00000265818MR5 (MPX49302); 5107849, 00000275642MR7 (GEE34040); 2357500-2, 00000265818MR5 (MPX94700); 5107849, 00000278210MR0 (UPW21613); 2377062-2, 00000963343YM9 (DPI53205); 5107849, 00000267036MR2 (DPN62300); 46-294231G1, 0000094-0131-5 (DPM19607); 46-294231G1, 0000005-0209-1 (FPG33602); 5127452, 00000VALUE1177 (NPV77700); N/A, 002005-0627-13 (387040MR06); 2357500, 00000259106MR3 (0853270052IB1); N/A, N/A (0853270070); 46-294231G1, 0002005-0519-8 (0856270035); 2226300, 000000000R4976 (0856270037); N/A, N/A (0856270042); 5107849-2, 00000279892MR4 (06108MRS01); 46-294231G1, 002005-1201-26 (34006MRS03); 46-294231G1, 000006-0502-12 (35045MRS01); 2226300, 00000000R4866A (78004MRS01); N/A, N/A (45007MRS01); 5107849, 00000266746MR7 (54014MRS01); 2226300, 000000000R5885 63025MRS01); 2226300, 000000000R5892 (UK1015MR02); N/A, N/A (00133MRS02); N/A, N/A (05270MRS01); 46-294231G1, 002002-0221-13 (70065MRS02); 5107849-2, 00000280006MR8 (10079MRM73); 46-294231G1, 0000006-0511-7 (3882630908MRMET); Recalling Firm/ Manufacturer GE Medical Systems, LLC 3200 N Grandview Blvd Waukesha WI 53188-1693 For Additional Information Contact 800-437-1171 Manufacturer Reason for Recall The latest on-site software version may not have been reinstalled at some sites after service activities were performed resulting in images possibly flipped left/right and/or there can be patient data mismatch. FDA Determined Cause 2 Software change control Action The firm, GE Healthcare, notified its customers on or before 06/07/2018 via a letter entitled "URGENT MEDICAL DEVICE CORRECTION". The letter described the product, problem and actions to be taken. The letter informed customers that GE Healthcare will inspect all affected systems and ensure the proper software version is installed. Instructions for customers to confirm the current software included the following steps: 1. Click on tools icon located on the top left of the screen; 2. Click on Service Browser Button; 3. The MR Service Desktop browser will launch (may take 1 minute to launch); 4. Look at the Release Information on the screen and compare both MrpApps Build Number and Service Pack Build Number on the screen to two tables provided in the safety issue section of the customer letter. ************If you have identified that your system has invalid software version, please contact GE Healthcare service representative immediately.****** ******GE Healthcare will advise you after your software version has been inspected and inform you if any correction is needed. If a correction is needed, a service representative will contact you to arrange for this correction.************Questions or concerns regarding this notification can be directed to GE Healthcare Service at 1-800-437-1171 or your local Service Representative. Quantity in Commerce 5,200 total units Distribution Worldwide Distribution: US (nationwide) including states of:: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV; and countries of: Algeria, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, CHINA, Colombia, Costa Rica, Croatia, CYPRUS, Czech Republic, DENMARK, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, Georgia, Greece, Guadeloupe, Guatemala, Guyana, Honduras, Hong Kong, Iceland, India, Indonesia, Iran, IRAQ, Ireland, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea (Republic of), Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Macedonia, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Pakistan, Paraguay, Philippines, Portugal, Puerto Rico, Russia, Saudi Arabia, SINGAPORE, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, United Arab Emirates, Uruguay, Uzbekistan, Venezuela, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LNH 510(K)s with Product Code = LNH 510(K)s with Product Code = LNH - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, GE Healthcare, notified its customers on or before 06/07/2018 via a letter entitled "URGENT MEDICAL DEVICE CORRECTION". The letter described the product, problem and actions to be taken. The letter informed customers that GE Healthcare will inspect all affected systems and ensure the proper software version is installed. Instructions for customers to confirm the current software included the following steps: 1. Click on tools icon located on the top left of the screen; 2. Click on Service Browser Button; 3. The MR Service Desktop browser will launch (may take 1 minute to launch); 4. Look at the Release Information on the screen and compare both MrpApps Build Number and Service Pack Build Number on the screen to two tables provided in the safety issue section of the customer letter. ************If you have identified that your system has invalid software version, please contact GE Healthcare service representative immediately.****** ******GE Healthcare will advise you after your software version has been inspected and inform you if any correction is needed. If a correction is needed, a service representative will contact you to arrange for this correction.************Questions or concerns regarding this notification can be directed to GE Healthcare Service at 1-800-437-1171 or your local Service Representative. | SoftwareChangeControl
68 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=165729 | Z-0110-2019 | GE Healthcare MRI Systems, Software Version 11.x, sold under the following product names: ******1.5T CXK4 LCC MAGNET,1.5T EXCITE, 1.5T EXCITE E/S, 1.5T EXCITE GS, ***1.5T EXCITE HYBRID, 1.5T EXCITE TWIN, 1.5T EXCITE UPGRAD, 1.5T HD ES/HS TR W/EXCITE, 1.5T MR MOBILE (SERIAL # R350), 1.5T SIGNA ECHOSPEED, 1.5T SIGNA EXCITE, 1.5T SIGNA EXCITE MRI, 1.5T SIGNA VIBRANT 16CH FIXED, 3.0T MRI TWINSPEED 8CH, 3.0T SIGNA EXCITE FIXED, 3.0T SIGNA EXCITE UPGRADE, 3T EXCITE, EX ES+8CH USED***EXCITE, EXCITE 1.5T, EXCITE 3.0T, EXCITE TWIN, GOLDSEAL 1.5T EXCITE 4CH, GOLDSEAL 1.5T EXCITE 8CH, GOLDSEAL 1.5T EXCITE REL***GOLDSEAL EXCITE 1.5T, GOLDSEAL EXCITE 1.5T 8CH, GS 1.5 EX ES+ 4CH NEW, GS 1.5 EX ES+4CH USED, GS 1.5 EX ES+8CH NEW, GS 1.5T EXCITE MRI, GS CLASS 1.5T ES+ EX 4CHN, GS SIGNA EXCITE II 1.5T, GS SIGNA MR/I 1.5T, LCC EXCITE 1.5T, MR 1.5 EX ES 8CH NEW, MR 1.5 EX ES+ 4CH, MR 1.5 EX ES+4CH MOB, MR 1.5 EX ES+8CH, MR 1.5EX8CH, MR 1.5EX8CH USE -HUNG, MR 1.5T EX8CH, MR 1.5T EXCITE 4CH, MR 1.5T EXCITE EXPERT MR, MR 1.5T EXCITE GS, MR 1.5T EXCITE GS MR, MR 1.5T GS ES+ EXCITE 4 CHANNE, MR 1.5T SIGNA EXCITE 8 CH GOLD SEAL, MR 1.5T SIGNA EXCITE II 4CH, MR 3.0T TWINSPEED, MR 3.0T SIGNA EXCI, MR 3.0T SIGNA EXCITE, MR 3.0T SIGNA EXCITE FIXED, MR EXCITE, MR EXCITE 1.5T, MR EXCITE 3, MR EXCITE ECHOSPEE, MR EXCITE ECHOSPEE UPG, MR EXCITE ECHOSPEED, MR EXCITE HISPEED***MR EXCITE TWINSPEE, MR EXCITE TWINSPEED, MR GOLD SEAL 1.5 EX ES+8CH NEW, MR GOLDSEAL 1.5 ES PL EX 4 CH, MR GOLDSEAL 1.5 EX,*** MR GOLDSEAL 1.5 EX, MR GOLDSEAL 1.5 EX 4CH, MR GOLDSEAL 1.5 EX 8CH, MR GOLDSEAL 1.5 EX ES+ 4CH, MR GOLDSEAL 1.5 EX ES+ 4CH NEW, MR GOLDSEAL 1.5 EX ES+4CH, MR GOLDSEAL 1.5 EX ES+4CH USD, MR GOLDSEAL 1.5 EX ES+4CH USE, MR GOLDSEAL 1.5 EX ES+4CH USED, R GOLDSEAL 1.5 EX ES+8CH, MR GOLDSEAL 1.5 EX ES+8CH NEW, MR GOLDSEAL 1.5 EX ES+8CH USED, MR GOLDSEAL 1.5 EXCITE S+ 4CH, MR GOLDSEAL 1.5 FIXED ACTIVE, MR GOLDSEAL 1.5EX4CH NEW -HUNG, MR GOLDSEAL 3.0T EXCITE HD FIX, MR GS 1.5 EX, MR GS 1.5 EX ES+8C, MR GS 1.5 EX ES+8CH, MR SIGNA 1.5 EXCIT, MR SIGNA 1.5 EXCITE EXPERT, MR SIGNA 1.5T EXCITE, MR SIGNA 1.5T EXCITE 4CH***MR SIGNA 1.5T EXCITE 8CH, MR SIGNA 1.5T EXCITE 8CH MOBIL, MR SIGNA 1.5T EXCITE EXPERT, MR SIGNA 1.5T EXCITE II, MR SIGNA 15T EXCITE 8CH, MR SIGNA 15T EXCITE 8CH MOBIL, MR SIGNA EXCITE 1.5 T GS (4 CHANNEL, MR SIGNA MR/I ECHOS W/EXCITE, MR SIGNA MR/I ECHOSPD W/ EXCITE, MR SIGNA MR/I ECHOSPD W/ EXCITE , MR SIGNA MR/I ECHOSPEED W/ EXCITE, MR SIGNA MR/I HISPEED W/ EXCITE***MR SIGNA MR/I TWISP TR W/EXCIT, MR TWINSPEED 1.5T W/ EXCITE, MR UPG 1.5 SIGNA LX TO EXCITE, MR UPG 3.0T94 ECLIPSE TO EXCIT, MR/I ECHOSPD W/ EXCITE, MR/I ECHOSPEED W/ EXCITE, MR/I SIGNA ECHOSPD W/ EXCITE, MRI EXCITE, MRI EXCITE 1.5T FIXED, MRI SIGNA EXCITE 1.5T, PREOWN 1.5T EXITE, SIGNA 1.5 EXCITE, SIGNA 1.5T , CHOSPEED W EXCITE, SIGNA 1.5T ECHOSPEED W/EXCITE, SIGNA 1.5T ES EXCITE, SIGNA 1.5T EXCITE, SIGNA 1.5T EXCITE 2, SIGNA 1.5T EXCITE EXPERT, SIGNA 1.5T EXCITE II, SIGNA 1.5T SMART SPEED, /EXCITE, SIGNA 1.5T W/ EXCITE, SIGNA 3.0T E, SIGNA 3.0T SIGNA EXCITE FIXED, SIGNA ECHOSPEED 1.5T FIXED, SIGNA ECHOSPEED EXCITE 1.5T, SIGNA ECHOSPEED LX 1.5T FIXED***SIGNA ECHOSPEED PLUS + EXCITE, SIGNA EXCITE, SIGNA EXCITE 1.5***SIGNA EXCITE 1.5 T HISPEED, SIGNA EXCITE 1.5T, SIGNA EXCITE 1.5T RELEASE 12.X/MR, SIGNA EXCITE 1.5T SMART SPEED, SIGNA EXCITE 11X, SIGNA EXCITE 3.0T, SIGNA EXCITE ECHOSPEED 1,5 T, SIGNA EXCITE ECHOSPEEED PLUS 1.5 T, SIGNA EXCITE EXPERT 1.5T, SIGNA EXCITE GS 1.5T 4 CHANNEL,SIGNA EXCITE II, SIGNA HISPD LX K4 1.5T FIXED, SIGNA HORIZON LX 1.5T FIXED, SIGNA LX 1.5 EXCITE II., SIGNA LX 1.5T EXCITE, SIGNA LX 1.5T EXCITE II, SIGNA LX1.5T HISPEED, SIGNA MR 1.5 ES EXCITE, SIGNA MR/I 1.5T EXCITE, SIGNA MR/I ECHO W/E,SIGNA MR/I ECHOESPEED W/EXCITE, SIGNA MR/I ECHOSPD, SIGNA MR/I ECHOSPD 1.5T FIXED, SIGNA MR/I ECHOSPD W/ EXCITE 1.5, SIGNA MR/I ECHOSPEED W/ EXCITE, SIGNA MR/I ECHOSPEED W/ EXCITE 1.5T, SIGNA MR/I ECHOSPEED W/EXCITE, SIGNA MR/I ECHOSPEED WITH EXCI, SIGNA MR/I ECHOSPEED WITH EXCITE***SIGNA, MR/I | GE Healthcare MRI Systems, Software Version 11.x | 2 | CDRH | 2018-10-15 |  | 2020-06-18 09:37 | GE Medical Systems, LLC | The latest on-site software version may not have been reinstalled at some sites after service activities were performed resulting in images possibly flipped left/right and/or there can be patient data mismatch. | Class 2 Device Recall GE Healthcare MRI Systems, Software Version 11.x Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GE Healthcare MRI Systems, Software Version 11.x FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GE Healthcare MRI Systems, Software Version 11.x Date Initiated by Firm June 07, 2018 Create Date October 15, 2018 Recall Status 1 Terminated 3 on June 18, 2020 Recall Number Z-0110-2019 Recall Event ID 80323 510(K)Number K012200 K013636 K030874 K040444 K041476 K052293 Product Classification System, nuclear magnetic resonance imaging - Product Code LNH Product GE Healthcare MRI Systems, Software Version 11.x, sold under the following product names: ************1.5T CXK4 LCC MAGNET,1.5T EXCITE, 1.5T EXCITE E/S, 1.5T EXCITE GS, ******1.5T EXCITE HYBRID, 1.5T EXCITE TWIN, 1.5T EXCITE UPGRAD, 1.5T HD ES/HS TR W/EXCITE, 1.5T MR MOBILE (SERIAL # R350), 1.5T SIGNA ECHOSPEED, 1.5T SIGNA EXCITE, 1.5T SIGNA EXCITE MRI, 1.5T SIGNA VIBRANT 16CH FIXED, 3.0T MRI TWINSPEED 8CH, 3.0T SIGNA EXCITE FIXED, 3.0T SIGNA EXCITE UPGRADE, 3T EXCITE, EX ES+8CH USED******EXCITE, EXCITE 1.5T, EXCITE 3.0T, EXCITE TWIN, GOLDSEAL 1.5T EXCITE 4CH, GOLDSEAL 1.5T EXCITE 8CH, GOLDSEAL 1.5T EXCITE REL******GOLDSEAL EXCITE 1.5T, GOLDSEAL EXCITE 1.5T 8CH, GS 1.5 EX ES+ 4CH NEW, GS 1.5 EX ES+4CH USED, GS 1.5 EX ES+8CH NEW, GS 1.5T EXCITE MRI, GS CLASS 1.5T ES+ EX 4CHN, GS SIGNA EXCITE II 1.5T, GS SIGNA MR/I 1.5T, LCC EXCITE 1.5T, MR 1.5 EX ES 8CH NEW, MR 1.5 EX ES+ 4CH, MR 1.5 EX ES+4CH MOB, MR 1.5 EX ES+8CH, MR 1.5EX8CH, MR 1.5EX8CH USE -HUNG, MR 1.5T EX8CH, MR 1.5T EXCITE 4CH, MR 1.5T EXCITE EXPERT MR, MR 1.5T EXCITE GS, MR 1.5T EXCITE GS MR, MR 1.5T GS ES+ EXCITE 4 CHANNE, MR 1.5T SIGNA EXCITE 8 CH GOLD SEAL, MR 1.5T SIGNA EXCITE II 4CH, MR 3.0T TWINSPEED, MR 3.0T SIGNA EXCI, MR 3.0T SIGNA EXCITE, MR 3.0T SIGNA EXCITE FIXED, MR EXCITE, MR EXCITE 1.5T, MR EXCITE 3, MR EXCITE ECHOSPEE, MR EXCITE ECHOSPEE UPG, MR EXCITE ECHOSPEED, MR EXCITE HISPEED******MR EXCITE TWINSPEE, MR EXCITE TWINSPEED, MR GOLD SEAL 1.5 EX ES+8CH NEW, MR GOLDSEAL 1.5 ES PL EX 4 CH, MR GOLDSEAL 1.5 EX,****** MR GOLDSEAL 1.5 EX, MR GOLDSEAL 1.5 EX 4CH, MR GOLDSEAL 1.5 EX 8CH, MR GOLDSEAL 1.5 EX ES+ 4CH, MR GOLDSEAL 1.5 EX ES+ 4CH NEW, MR GOLDSEAL 1.5 EX ES+4CH, MR GOLDSEAL 1.5 EX ES+4CH USD, MR GOLDSEAL 1.5 EX ES+4CH USE, MR GOLDSEAL 1.5 EX ES+4CH USED, R GOLDSEAL 1.5 EX ES+8CH, MR GOLDSEAL 1.5 EX ES+8CH NEW, MR GOLDSEAL 1.5 EX ES+8CH USED, MR GOLDSEAL 1.5 EXCITE S+ 4CH, MR GOLDSEAL 1.5 FIXED ACTIVE, MR GOLDSEAL 1.5EX4CH NEW -HUNG, MR GOLDSEAL 3.0T EXCITE HD FIX, MR GS 1.5 EX, MR GS 1.5 EX ES+8C, MR GS 1.5 EX ES+8CH, MR SIGNA 1.5 EXCIT, MR SIGNA 1.5 EXCITE EXPERT, MR SIGNA 1.5T EXCITE, MR SIGNA 1.5T EXCITE 4CH******MR SIGNA 1.5T EXCITE 8CH, MR SIGNA 1.5T EXCITE 8CH MOBIL, MR SIGNA 1.5T EXCITE EXPERT, MR SIGNA 1.5T EXCITE II, MR SIGNA 15T EXCITE 8CH, MR SIGNA 15T EXCITE 8CH MOBIL, MR SIGNA EXCITE 1.5 T GS (4 CHANNEL, MR SIGNA MR/I ECHOS W/EXCITE, MR SIGNA MR/I ECHOSPD W/ EXCITE, MR SIGNA MR/I ECHOSPD W/ EXCITE , MR SIGNA MR/I ECHOSPEED W/ EXCITE, MR SIGNA MR/I HISPEED W/ EXCITE******MR SIGNA MR/I TWISP TR W/EXCIT, MR TWINSPEED 1.5T W/ EXCITE, MR UPG 1.5 SIGNA LX TO EXCITE, MR UPG 3.0T94 ECLIPSE TO EXCIT, MR/I ECHOSPD W/ EXCITE, MR/I ECHOSPEED W/ EXCITE, MR/I SIGNA ECHOSPD W/ EXCITE, MRI EXCITE, MRI EXCITE 1.5T FIXED, MRI SIGNA EXCITE 1.5T, PREOWN 1.5T EXITE, SIGNA 1.5 EXCITE, SIGNA 1.5T , CHOSPEED W EXCITE, SIGNA 1.5T ECHOSPEED W/EXCITE, SIGNA 1.5T ES EXCITE, SIGNA 1.5T EXCITE, SIGNA 1.5T EXCITE 2, SIGNA 1.5T EXCITE EXPERT, SIGNA 1.5T EXCITE II, SIGNA 1.5T SMART SPEED, /EXCITE, SIGNA 1.5T W/ EXCITE, SIGNA 3.0T E, SIGNA 3.0T SIGNA EXCITE FIXED, SIGNA ECHOSPEED 1.5T FIXED, SIGNA ECHOSPEED EXCITE 1.5T, SIGNA ECHOSPEED LX 1.5T FIXED******SIGNA ECHOSPEED PLUS + EXCITE, SIGNA EXCITE, SIGNA EXCITE 1.5******SIGNA EXCITE 1.5 T HISPEED, SIGNA EXCITE 1.5T, SIGNA EXCITE 1.5T RELEASE 12.X/MR, SIGNA EXCITE 1.5T SMART SPEED, SIGNA EXCITE 11X, SIGNA EXCITE 3.0T, SIGNA EXCITE ECHOSPEED 1,5 T, SIGNA EXCITE ECHOSPEEED PLUS 1.5 T, SIGNA EXCITE EXPERT 1.5T, SIGNA EXCITE GS 1.5T 4 CHANNEL,SIGNA EXCITE II, SIGNA HISPD LX K4 1.5T FIXED, SIGNA HORIZON LX 1.5T FIXED, SIGNA LX 1.5 EXCITE II., SIGNA LX 1.5T EXCITE, SIGNA LX 1.5T EXCITE II, SIGNA LX1.5T HISPEED, SIGNA MR 1.5 ES EXCITE, SIGNA MR/I 1.5T EXCITE, SIGNA MR/I ECHO W/E,SIGNA MR/I ECHOESPEED W/EXCITE, SIGNA MR/I ECHOSPD, SIGNA MR/I ECHOSPD 1.5T FIXED, SIGNA MR/I ECHOSPD W/ EXCITE 1.5, SIGNA MR/I ECHOSPEED W/ EXCITE, SIGNA MR/I ECHOSPEED W/ EXCITE 1.5T, SIGNA MR/I ECHOSPEED W/EXCITE, SIGNA MR/I ECHOSPEED WITH EXCI, SIGNA MR/I ECHOSPEED WITH EXCITE******SIGNA, MR/I Code Information Model Number, Serial Number (System ID) : N/A, 002003-0530-08 (GOLD010252), 46-294231G1, 002004-0406-06 (HC5300MR02), 46-294231G1, 000004-0311-16 (IE1053MR01), 2377062-2, 00000963562YM4 (0850270477), 2377062, 00000983439YM5 (0850270470), 46-294231G1, 002002-1021-51 (0850270122), 46-294231G1, 0000099-0809-8 (482180SIGNA), 2377062, 00000941219YM8 (561863MR1), N/A, N/A (0910272059), 2226300, 000000000R1040 (PL1124MR02), N/A, 0002005-0922-4 (00491MRS01), N/A, N/A (083027200000913), 2377062-2, 963531YM9 (559713CMIMR), 7452, 000000HDMR6133 (IE1111MR01), 46-294231G1, 002004-1028-17 (2059393T), N/A, 002004-0823-18 (M4135138); 46-294231G1, 002004-1115-13 (0850270170); 46-294231G1, 04 0211-4 (912354MRI); 46-294231G1, 002004-1109-02 (44EXCITE2550834); 46-294231G1, 002004-0913-16 (YM4157); 46-294231G1, 002005-0914-20 (YM1414); 46-294231G1, 002003-0708-02 (YM2028); 2226300, 0000000R7722 (1337081); 2377062-2, 00000951060YM3 (MRR7674); 2226300, 000000000R3880 (083027602171311); 2357500-2, 00000268322MR5 (MRR0455); 2214149-2, 00000006147M23 (1347499); 2377062-2, 00000956341YM2 (2682089); 2357500-2, 00000281517MR3 (2641178); 2226300, 000000000R5825 (2682131); 2377062-2, 00000963279YM5 (083027846014410); 2226300, 000000000R5837 (1238483); 2226300, 000000000R5419 (48MRE2744667); N/A, N/A (299SIGEXC); N/A, N/A (H1787MR11); 2226300, 000000000R7325 (2850446); 2226300, 000000000R7402 (PPR94108; 2357500-2, 00000272881MR4 (UPW21614); 2377062-2, 00000963044YM3 (MPX39916); 2377062, 00000962662YM3 (A5912915); 2377062-2, 00000002522YR1 (YP567303); 2377062, 00000941242YM0 (FPG23603); 2226300, 000000000R7543 (FPG84701); 2377062, 00000951821YM8 (1292548); 2377062, 00000962780YM3 (2713328); 2226300, 000000000R6610 (2762011); N/A, 000000000R5956 (1238481); 2226300, 000000000R4173 (MRR4173); 2357500-3, 00000298377MR3 (PPR67317); 5127452, 000000HDMR5690 (CEN03406); 46-294231G1, 0002005-1109-4 (0853270072); 46-294231G1, 002004-0830-13 (082427040042); 46-294231G1, 002004-0830-10 (082427230015); N/A, N/A (083027600255513); 5139394-2, 00000006184YR6 (082427040067); 46-294231G1, 000006-0511-18 (7771MREXCITE3); 2357500-2, 00000254491MR4 (082427040056); N/A, 00000309690MR6 (082427160133); 2226300, 000000000R5262 (082427100072); 2226300, 000000000R4812 (082427160046); 2294302, 00000249253MR6 (082427130020); 2226300, 000000000R3913 (082427040044); 2226300, 000000000R7983 (MRR7983); 2226300, R1698 (678324MR); 2226300, 000000000R2220 (083027861003910); 2226300, 000000000R1928 (083027062176711); 2357500-3, 00000299538MR9 (ILSMR1); 2357500-3, 00000299960MR5 (MRR8022); 2357500-2, 00000282641MR0 (H17659MR1); N/A, N/A (205206MR1); 2357500-3, 00000298579MR4 (MRR8015); 2214149-2, 00000002656M27 (MRR8481); 2357500-3, 00000297885MR6 (771MR2891120); N/A, N/A (479936MRHOG); 2357500-3, 00000299516MR5 (MR259430MR7); N/A, N/A (928699MR); N/A, N/A (907VNSMR); 2226300, 000000000R5752 (RO1121MR01); N/A, N/A (916733MR2); 2214149-2, 00000002639M23 (MRR8461); 5122373-2, 00000086257DP3 (901516U3T); 2214149-2, 00000002793M28 (L2770MR01); N/A, N/A (083027002824690); 2214149-2, 00000002043M28 (PL3677MR01); 2377062-2, 00000963601YM0 (082427010016); 2226300, 000000000R6529 (PPR29305); 2357500-2, 00000258396MR1 (PPR10411); 2377062-2, 00000956334YM7 (UP389021); 2226300, 000000000R2739 (NPV89300); 2284300, 00000276193MR0 (UPW21611); 2377062-2, 00000943593YM4 (2709088); 2226300, 000000000R3694 (PER28125); 2189500-3, 00000204698MR5 (MP118827); 2226300, 000000000R4458 (FPG11803); N/A, 000000000R6087 (H17661MR01); 46-294231G1, 002003-0320-09 (FEG45915); N/A, 0002005-0725-8 (PPR47506); 2377062, 00000962577YM3 (0853270061); 2189500-3, 00000206006MR9 (RU3408MR01); 2284300, 00000253648MR0 (34148MRS01); 46-294231G1, 002005-0927-23 (0853270074); 46-294231G1, 0002003.111.06 (5606TWEX); N/A, N/A (082427140023); N/A, 00000254375MR9 (NPV26104); 46-294231G1, 002003-0915-23 (FPG13900); N/A, 00000963342YM1 (KE1000MR01); N/A, 000000000R4581 (B5387602); N/A, N/A (0910274031); 2377062, 00000951828YM3 (NPV24000); 46-294231G1, 002000-1003-14 (NP739709); N/A, N/A (083027598401216); 46-294231G1, 002003-0708-10 (0910273024); N/A, N/A (083027042101154); N/A, N/A (083027022100534); 46-294231G1, 000099-0223-10 (126292MR01); 2294302, 00000249254MR4 (2222EXITE); 2357500-2, 00000274050MR4 (2284SIG15T); 2377062, 00000962626YM8 (2774243SIGNA); 2377062-2, 00000963367YM8 (1228042-1); 46-294231G1, 000097-0929-12 (KW1007MR02); 46-294231G1, 000004-0311-15 (A5355209); N/A, 002004-0701-08 (409740MR01); 46-294231G1, 002005-0927-30 (0910272044); N/A, N/A (805648MR); N/A, 00001009BD655A (GOLD011048); N/A, N/A (209225VCMR); N/A, 002003-0320-18 (110807MR01); 2226300, 000000000R4130 (2564618SIGNA); 2226300, 000000000R3122 (MRR3122); 46-294231G1, 002005-0620-20 (H2135MR07); 2226300, 000000000R6040 (2672732SIGNA); 2226300, 000000000R6211 (2730910SIGNA); 2226300, 000000000R4428 (2585063SIGEX); 2226300, 000000000R3928 (A5769207); N/A, 002004-0406-16 (M4434864); 2226300, 000000000R5189 (H44432MR01); 2377062-2, 00000951044YM7 (H17660MR3); 2377062-2, 00000963349YM6 (4867SIGEXC1); 2226300, 000000000R5611 (H17658MR01); 46-294231G1, 002001-0904-14 (38880918MRCTR); N/A, N/A (504MCMR); N/A, N/A (205431MR1); 46-294231G1, 000000000R3717 (SY2012MR01); N/A, 00000010462GQ0 (GOLD10637M); 2226300, 000000000R1083 (GOLD0EU113); 2377062-61, 00000008358YR4 (GOLD0EU120); 46-294231G1, 0002002-0422-9 (A5162244); 46-294231G1, 002004-0830-12 (3882540976MRMET); N/A, N/A (441887MR01); 46-294231G1, 002004-0909-13 (2842546BEXCITE); N/A, N/A (083027202100984); N/A, N/A (626795MR); N/A, 000000000R5449 (GOLD011065); 46-294231G1, 0002005-0519-7 (EG1522MR01); 46-294231G1, 002004-1115-07 (06034MRS02); 46-294231G1, 002004-1130-12 (06166MRS01); 46-294231G1, 002005-0812-18 (RU2402MR01); 46-294231G1, 002005-0627-17 (HU1243MR01); 46-294231G1, 002002-1030-07 (INSIGHTMR1277); 2377062-2, 00000963389YM2 (619340MR1); 2294302, 00000241135MR3 (CC0820MR05); 46-294231G1, 0000099-0518-3 (HC0399MR01); 46-294231G1, 002004-0713-04 (SA1255MR01); N/A, 00000309602MR1 (RO4165MR01); 46-294231G1, 002003-1218-15 (RU2371MR01); 46-294231G1, 002004-1209-03 (34169MRS01); 2226300, 000000000R3948 (34421MRS01); 46-294231G1, 002003-1218-21 (RU2490MR01); 2226300, 000000000R3764 (RO1089MR01); 46-294231G1, 002003-0320-01 (225647MR); 2284300, 00000245980MR8 (242501602MR); 46-294231G1, 002004-0303-13 (ZA2119MR01); 2377062, 00000962741YM5 (770454EDMR); 2377062-2, 00000963075YM7 (INSIGHTMR1348); 46-294231G1, 002002-1204-07 (A5753803); 2226300, 000000000R2594 (A5761301); 46-294231G1, 002005-1003-16 (SA1022MR01); 2226300, 000000000R4200 (16041MRS03); N/A, N/A (38895718MRMET); N/A, N/A (225381MR); 2294302, 00000249260MR1 (53515EXCITE); 2214149-2, 00000009718M28 (BMALTA46); 46-294231G1, 002004-0416-11 (K3585); N/A, 0002002-0405-2 (MC6946MR05); 46-294231G1, 002001-0823-10 (0910274012); 2357500-2, 00000265611MR4 (11393838MR); 46-294231G1, 002003-0409-01 (ZTM0120); 2226300, 000000000R2823 (HU1530MR01); N/A, 00000310561MR6 (082427120138); 2377062, 00000962791YM0 (314644IMR1); 46-294231G1, 002003-1015-18 (941795MR); 46-294231G1, 002003-0530-15 (732493MR1); 46-294231G1, 0000098-0617-6 (501219AHMR); 2377062-2, 00000943653YM6 (PIKETWIN); 46-294231G1, 000098-0824-10 (719578MR). Recalling Firm/ Manufacturer GE Medical Systems, LLC 3200 N Grandview Blvd Waukesha WI 53188-1693 For Additional Information Contact 800-437-1171 Manufacturer Reason for Recall The latest on-site software version may not have been reinstalled at some sites after service activities were performed resulting in images possibly flipped left/right and/or there can be patient data mismatch. FDA Determined Cause 2 Software change control Action The firm, GE Healthcare, notified its customers on or before 06/07/2018 via a letter entitled "URGENT MEDICAL DEVICE CORRECTION". The letter described the product, problem and actions to be taken. The letter informed customers that GE Healthcare will inspect all affected systems and ensure the proper software version is installed. Instructions for customers to confirm the current software included the following steps: 1. Click on tools icon located on the top left of the screen; 2. Click on Service Browser Button; 3. The MR Service Desktop browser will launch (may take 1 minute to launch); 4. Look at the Release Information on the screen and compare both MrpApps Build Number and Service Pack Build Number on the screen to two tables provided in the safety issue section of the customer letter. ************If you have identified that your system has invalid software version, please contact GE Healthcare service representative immediately.****** ******GE Healthcare will advise you after your software version has been inspected and inform you if any correction is needed. If a correction is needed, a service representative will contact you to arrange for this correction.************Questions or concerns regarding this notification can be directed to GE Healthcare Service at 1-800-437-1171 or your local Service Representative. Quantity in Commerce 5,200 total units Distribution Worldwide Distribution: US (nationwide) including states of:: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV; and countries of: Algeria, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, CHINA, Colombia, Costa Rica, Croatia, CYPRUS, Czech Republic, DENMARK, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, Georgia, Greece, Guadeloupe, Guatemala, Guyana, Honduras, Hong Kong, Iceland, India, Indonesia, Iran, IRAQ, Ireland, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea (Republic of), Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Macedonia, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Pakistan, Paraguay, Philippines, Portugal, Puerto Rico, Russia, Saudi Arabia, SINGAPORE, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, United Arab Emirates, Uruguay, Uzbekistan, Venezuela, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LNH 510(K)s with Product Code = LNH 510(K)s with Product Code = LNH - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, GE Healthcare, notified its customers on or before 06/07/2018 via a letter entitled "URGENT MEDICAL DEVICE CORRECTION". The letter described the product, problem and actions to be taken. The letter informed customers that GE Healthcare will inspect all affected systems and ensure the proper software version is installed. Instructions for customers to confirm the current software included the following steps: 1. Click on tools icon located on the top left of the screen; 2. Click on Service Browser Button; 3. The MR Service Desktop browser will launch (may take 1 minute to launch); 4. Look at the Release Information on the screen and compare both MrpApps Build Number and Service Pack Build Number on the screen to two tables provided in the safety issue section of the customer letter. ************If you have identified that your system has invalid software version, please contact GE Healthcare service representative immediately.****** ******GE Healthcare will advise you after your software version has been inspected and inform you if any correction is needed. If a correction is needed, a service representative will contact you to arrange for this correction.************Questions or concerns regarding this notification can be directed to GE Healthcare Service at 1-800-437-1171 or your local Service Representative. | SoftwareChangeControl
69 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=165871 | Z-0112-2019 | GE Healthcare MRI Systems, Software Version 15.x or HD 16 sold under the following product names:*** ****ER*15T HDXT ECHOSPEED 16 CHANNEL, *OC*1.5T HDX ECHOSPEED 8 CHANNEL, *SC*1.5T HD TO HDX UPGRADE, *TR*15T HDXT 16CH, *TR*3.0T HDX TWINSPEED 16 CHANNEL, 1,5T HDXT, 1.5 HD TO HDX UPGRADE, 1.5 HDX, 1.5 HDX ECHOSPEED 16 CH, 1.5 HDX ECHOSPEED 8 CHANNEL, 1.5 HDXT, 1.5 HDXT-16 , 1.5 LX TO HDX 16C FORKLIFT UPG, 1.5 LX TO HDX 8C FORKLIFT UPG, 1.5 T UPGRADE TO HDX, 1.5EXCIT TO HDXT UPG, 1.5T HDX 16 CHANNEL, 1.5T 16CH UPG, 1.5T 6 CH HDX , UPGRADE, 1.5T 8CH HDXT TO 16CH HDXT UPG, 1.5T CRM LX TO HDX UPG, 1.5T ES HDX W/ONCOLOGY PACK, 1.5T EXC.TO HDX UPG, 1.5T EXC.TO HDXT UPG, 1.5T EXCITE HD TO HDX UPG, 1.5T EXCITE TO HDX UPG, 1.5T EXCITE TO HDX UPG, 1.5T EXCITE TO HDX UPG. W/ACGD, 1.5T EXCITE TO HDXT UPG, 1.5T EXCITE TO HDXT UPGRADE, 1.5T EX-HDX, 1.5T EX-HDXTWIN, 1.5T HD ES/HS MOBILE, 1.5T HD TO HDX UPG. 1.5T HD TO HDX UPGRADE, 1.5T HD TO HDXT UPG, 1.5T HD TO HDXT UPGRADE***1.5T HDI,1.5T HDI 16CH FIXED, 1.5T HDI 8CH FIXED, 1.5T HDI ECHOSPEED***1.5T HDI ECHOSPEED 8 CHANNEL, 1.5T HDI ECHOSPEED 8 CHANNEL***1.5T HDI ECHOSPEED 8 CHANNEL., 1.5T HDI ECHOSPEED 8 CHANNEL, ***1.5T HDI ECHOSPEED 8 CHAN¿SO2762285, 1.5T HDI ECHOSPEED 8CH, ***1.5T HDI ECHOSPEED MOBILE, 1.5T HDI ECSHOSPEED 8CH, 1.5T HDI TO HDXT UPGRADE, 1.5T HDX, 1.5T HDX - MEIRAV (BREAST), 1.5T HDX 8CM***1.5T HDX ECHOSPD 16, 1.5T HDX ECHOSPEED, 1.5T HDX ECHOSPEED ***1.5T HDX ECHOSPEED 16, 1.5T HDX ECHOSPEED 16 CHANNEL, 1.5T HDX ECHOSPEED 16 CHANNEL, 1.5T HDX ECHOSPEED 16 CHANNELS, 1.5T HDX ECHOSPEED 16CH¿SO2659508, 1.5T HDX ECHOSPEED 16CH-SO2718893, 1.5T HDX ECHOSPEED 32 CHANNEL, 1.5T HDX ECHOSPEED 8 CH, 1.5T HDX ECHOSPEED 8 CHANNEL, 1.5T HDX ECHOSPEED MOBILE***1.5T HDX ECHOSPEED MRI, 1.5T HDX ES, 1.5T HDX ES 16, 1.5T HDX TWINSPEED, 1.5T HDX TWINSPEED 16 CHANNEL, 1.5T HDXT***1.5T HDXT 16CH,1.5T HDXT 16CH CXK4, 1.5T HDXT 16CH ES FIXED***1.5T HDXT 8 CH, 1.5T HDXT 8CH ES FIXED, 1.5T HDXT 8CH UP, 1.5T HDXT ECHOSPEE, 1.5T HDXT ECHOSPEED, 1.5T HDXT ECHOSPEED 16 CH***1.5T HDXT ECHOSPEED 16 CHAN, 1.5T HDXT ECHOSPEED 16 CHANNEL***1.5T HDXT ECHOSPEED 8, 1.5T HDXT ECHOSPEED 8 CHANN, 1.5T HDXT ECHOSPEED 8 CHANNEL, 1.5T HDXT ECHOSPEED 8 CHANNEL , 1.5T HDXT ECHOSPEED 8_CHANNEL, 1.5T HDXT ECHOSPEED MOBILE, 1.5T HDXT ES 8 CH, 1.5T HDXT GS, 1.5T HDXT TWIN, 1.5T HDXT TWINSPEE, 1.5T HDXT TWINSPEED 16 CHAN,1.5T HDXT UPG, 1.5T HDXTID, 1.5T LX HDX, 1.5T LX TO 16 CH HDXT UPG, 1.5T LX TO 16 CHAN, 1.5T LX TO 16 CHANNEL HDX UPG, 1.5T LX TO 16 CHANNEL HDX UPG, 1.5T LX TO 16 CHANNEL HDX UPGRADE, 1.5T LX TO 16 CHANNEL HDXT, 1.5T LX TO 16 CHANNEL HDXT UPG, 1.5T LX TO 16CHANNEL HDX U/G2777896, 1.5T LX TO 8 CH HDX UPG, 1.5T LX TO 8 CH HDXT UPG, 1.5T LX TO 8 CHANNEL HDX UP, 1.5T LX TO 8 CHANNEL HDX UPG, 1.5T LX TO 8 CHANNEL HDX UPG, 1.5T LX TO 8 CHANNEL HDX UPG.827015, 1.5T LX TO 8 CHANNEL HDXT UPG, 1.5T LX TO 8 CHANNEL HDXT UPGRADE, 1.5T LX TO HDXT***1.5T NON LCC4X5X TO HDXT UPGD, 1.5T NON-LCC LX/HRZ 5.5 TO 1.5T HDX, 1.5T SIGNA HDX UPG, 1.5T SIGNA HDXT, 1.5TEXCITE TO HDX UPG W/O A, 1.5TEXCITE TO HDX UPG W/O ACGD, 1.5TNONLCC LXHZ TO HDXT UPG, 1.5TNONLCC LXHZ TO HDXUPGW/ACD, 1.5TNONLCC4X5X TO HDXUPGW/OACD, 1.5TNONLCCLXHZ TO HDXUPGW/O ACD, 1¿5T 8CH HDXT MRI, 1¿5T HDXT, 1¿5T HDXT MRI, 1-5T CRM LX TO HDX , G.W/OACGD, 1-5T CRM LX TO HDXT UPG, 1-5T EXCITE TO***15T EXCITE TO HDX UPG, 1-5T HDI ECHOSPEED 16 CHANNEL, 15T HDI ECHOSPEED 8 CHANNEL, 1-5T HDX ECHOSPEED 16 CHANN, 15T HDX ECHOSPEED 16 CHANNEL, 15T HDX ECHOSPEED 8 CHANNEL, 15T HDX ECHOSPEED MOBILE, 1-5T HDX ECHOSPEED MOBILE, 1-5T HDX ECHOSPEED TRANSPORT, 1-5T HDX TWINSPEED 8 CHANNEL, 1-5T HDXT ECHOSPEED 16, 1-5T HDXT ECHOSPEED 16 CHAN, 1-5T HDXT ECHOSPEED 16 CHAN-HD16.0, 1-5T HDXT ECHOSPEED 16 CHANNEL, ***1-5T HDXT ECHOSPEED 16 CHANNEL, 1-5T HDXT ECHOSPEED 16 UPGRADE, 1-5T HDXT ECHOSPEED 8 CHANN, 15T HDXT ECHOSPEED 8 CHANNEL, 1-5T HDXT ECHOSPEED MOBILE, 1-5T HDXT ECHOSPEED RELOCATABLE, 1-5T LX TO 16 CHANNEL HDX UPG, 1-5T LX TO 16 CHANNEL HDXT, 1-5T LX TO 8 CHANNEL HDXT U, 1-5T LX TO | GE Healthcare MRI Systems, Software Version 15.x or HD 16 | 2 | CDRH | 2018-10-15 |  | 2020-06-18 09:37 | GE Medical Systems, LLC | The latest on-site software version may not have been reinstalled at some sites after service activities were performed resulting in images possibly flipped left/right and/or there can be patient data mismatch. | Class 2 Device Recall GE Healthcare MRI Systems, Software Version 15.x or HD 16 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GE Healthcare MRI Systems, Software Version 15.x or HD 16 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GE Healthcare MRI Systems, Software Version 15.x or HD 16 Date Initiated by Firm June 07, 2018 Create Date October 15, 2018 Recall Status 1 Terminated 3 on June 18, 2020 Recall Number Z-0112-2019 Recall Event ID 80323 510(K)Number K012200 K013636 K030874 K040444 K041476 K052293 Product Classification System, nuclear magnetic resonance imaging - Product Code LNH Product GE Healthcare MRI Systems, Software Version 15.x or HD 16 sold under the following product names:****** *******ER*15T HDXT ECHOSPEED 16 CHANNEL, *OC*1.5T HDX ECHOSPEED 8 CHANNEL, *SC*1.5T HD TO HDX UPGRADE, *TR*15T HDXT 16CH, *TR*3.0T HDX TWINSPEED 16 CHANNEL, 1,5T HDXT, 1.5 HD TO HDX UPGRADE, 1.5 HDX, 1.5 HDX ECHOSPEED 16 CH, 1.5 HDX ECHOSPEED 8 CHANNEL, 1.5 HDXT, 1.5 HDXT-16 , 1.5 LX TO HDX 16C FORKLIFT UPG, 1.5 LX TO HDX 8C FORKLIFT UPG, 1.5 T UPGRADE TO HDX, 1.5EXCIT TO HDXT UPG, 1.5T HDX 16 CHANNEL, 1.5T 16CH UPG, 1.5T 6 CH HDX , UPGRADE, 1.5T 8CH HDXT TO 16CH HDXT UPG, 1.5T CRM LX TO HDX UPG, 1.5T ES HDX W/ONCOLOGY PACK, 1.5T EXC.TO HDX UPG, 1.5T EXC.TO HDXT UPG, 1.5T EXCITE HD TO HDX UPG, 1.5T EXCITE TO HDX UPG, 1.5T EXCITE TO HDX UPG, 1.5T EXCITE TO HDX UPG. W/ACGD, 1.5T EXCITE TO HDXT UPG, 1.5T EXCITE TO HDXT UPGRADE, 1.5T EX-HDX, 1.5T EX-HDXTWIN, 1.5T HD ES/HS MOBILE, 1.5T HD TO HDX UPG. 1.5T HD TO HDX UPGRADE, 1.5T HD TO HDXT UPG, 1.5T HD TO HDXT UPGRADE******1.5T HDI,1.5T HDI 16CH FIXED, 1.5T HDI 8CH FIXED, 1.5T HDI ECHOSPEED******1.5T HDI ECHOSPEED 8 CHANNEL, 1.5T HDI ECHOSPEED 8 CHANNEL******1.5T HDI ECHOSPEED 8 CHANNEL., 1.5T HDI ECHOSPEED 8 CHANNEL, ******1.5T HDI ECHOSPEED 8 CHANSO2762285, 1.5T HDI ECHOSPEED 8CH, ******1.5T HDI ECHOSPEED MOBILE, 1.5T HDI ECSHOSPEED 8CH, 1.5T HDI TO HDXT UPGRADE, 1.5T HDX, 1.5T HDX - MEIRAV (BREAST), 1.5T HDX 8CM******1.5T HDX ECHOSPD 16, 1.5T HDX ECHOSPEED, 1.5T HDX ECHOSPEED ******1.5T HDX ECHOSPEED 16, 1.5T HDX ECHOSPEED 16 CHANNEL, 1.5T HDX ECHOSPEED 16 CHANNEL, 1.5T HDX ECHOSPEED 16 CHANNELS, 1.5T HDX ECHOSPEED 16CHSO2659508, 1.5T HDX ECHOSPEED 16CH-SO2718893, 1.5T HDX ECHOSPEED 32 CHANNEL, 1.5T HDX ECHOSPEED 8 CH, 1.5T HDX ECHOSPEED 8 CHANNEL, 1.5T HDX ECHOSPEED MOBILE******1.5T HDX ECHOSPEED MRI, 1.5T HDX ES, 1.5T HDX ES 16, 1.5T HDX TWINSPEED, 1.5T HDX TWINSPEED 16 CHANNEL, 1.5T HDXT******1.5T HDXT 16CH,1.5T HDXT 16CH CXK4, 1.5T HDXT 16CH ES FIXED******1.5T HDXT 8 CH, 1.5T HDXT 8CH ES FIXED, 1.5T HDXT 8CH UP, 1.5T HDXT ECHOSPEE, 1.5T HDXT ECHOSPEED, 1.5T HDXT ECHOSPEED 16 CH******1.5T HDXT ECHOSPEED 16 CHAN, 1.5T HDXT ECHOSPEED 16 CHANNEL******1.5T HDXT ECHOSPEED 8, 1.5T HDXT ECHOSPEED 8 CHANN, 1.5T HDXT ECHOSPEED 8 CHANNEL, 1.5T HDXT ECHOSPEED 8 CHANNEL , 1.5T HDXT ECHOSPEED 8_CHANNEL, 1.5T HDXT ECHOSPEED MOBILE, 1.5T HDXT ES 8 CH, 1.5T HDXT GS, 1.5T HDXT TWIN, 1.5T HDXT TWINSPEE, 1.5T HDXT TWINSPEED 16 CHAN,1.5T HDXT UPG, 1.5T HDXTID, 1.5T LX HDX, 1.5T LX TO 16 CH HDXT UPG, 1.5T LX TO 16 CHAN, 1.5T LX TO 16 CHANNEL HDX UPG, 1.5T LX TO 16 CHANNEL HDX UPG, 1.5T LX TO 16 CHANNEL HDX UPGRADE, 1.5T LX TO 16 CHANNEL HDXT, 1.5T LX TO 16 CHANNEL HDXT UPG, 1.5T LX TO 16CHANNEL HDX U/G2777896, 1.5T LX TO 8 CH HDX UPG, 1.5T LX TO 8 CH HDXT UPG, 1.5T LX TO 8 CHANNEL HDX UP, 1.5T LX TO 8 CHANNEL HDX UPG, 1.5T LX TO 8 CHANNEL HDX UPG, 1.5T LX TO 8 CHANNEL HDX UPG.827015, 1.5T LX TO 8 CHANNEL HDXT UPG, 1.5T LX TO 8 CHANNEL HDXT UPGRADE, 1.5T LX TO HDXT******1.5T NON LCC4X5X TO HDXT UPGD, 1.5T NON-LCC LX/HRZ 5.5 TO 1.5T HDX, 1.5T SIGNA HDX UPG, 1.5T SIGNA HDXT, 1.5TEXCITE TO HDX UPG W/O A, 1.5TEXCITE TO HDX UPG W/O ACGD, 1.5TNONLCC LXHZ TO HDXT UPG, 1.5TNONLCC LXHZ TO HDXUPGW/ACD, 1.5TNONLCC4X5X TO HDXUPGW/OACD, 1.5TNONLCCLXHZ TO HDXUPGW/O ACD, 15T 8CH HDXT MRI, 15T HDXT, 15T HDXT MRI, 1-5T CRM LX TO HDX , G.W/OACGD, 1-5T CRM LX TO HDXT UPG, 1-5T EXCITE TO******15T EXCITE TO HDX UPG, 1-5T HDI ECHOSPEED 16 CHANNEL, 15T HDI ECHOSPEED 8 CHANNEL, 1-5T HDX ECHOSPEED 16 CHANN, 15T HDX ECHOSPEED 16 CHANNEL, 15T HDX ECHOSPEED 8 CHANNEL, 15T HDX ECHOSPEED MOBILE, 1-5T HDX ECHOSPEED MOBILE, 1-5T HDX ECHOSPEED TRANSPORT, 1-5T HDX TWINSPEED 8 CHANNEL, 1-5T HDXT ECHOSPEED 16, 1-5T HDXT ECHOSPEED 16 CHAN, 1-5T HDXT ECHOSPEED 16 CHAN-HD16.0, 1-5T HDXT ECHOSPEED 16 CHANNEL, ******1-5T HDXT ECHOSPEED 16 CHANNEL, 1-5T HDXT ECHOSPEED 16 UPGRADE, 1-5T HDXT ECHOSPEED 8 CHANN, 15T HDXT ECHOSPEED 8 CHANNEL, 1-5T HDXT ECHOSPEED MOBILE, 1-5T HDXT ECHOSPEED RELOCATABLE, 1-5T LX TO 16 CHANNEL HDX UPG, 1-5T LX TO 16 CHANNEL HDXT, 1-5T LX TO 8 CHANNEL HDXT U, 1-5T LX TO Code Information Model Number, Serial Number (System ID) : N/A, 000000000R7742 (MP697620); N/A, 00000834BD4B76 (M2865810); 46-294231G1, 0002005-0607-9 (80040MRS02); N/A, 00000293242MR4 (GOLD00EU79); 5127452, 00000274388MR8 (M4160077); 2226300, 000000000R7586 (PL2130MR01); 2226300, 000000000R2785 (A5106403); N/A, N/A (083027060000499); 2226300, 000000000R5870 (A5177224); N/A, 000000000R4933 (C4194430); 2226300, 000000000R7628 (PL1807MR01); 5911000-4 , 00001201FMM0FJ (083027829345912); 2226300, 000000000R5979 (818713MR1); 5308573, 00000292202MR9 (706378MR2); 46-294231G1, 002003-1211-30 (KW1005MR02); 2226300, 000000000R5829 (PL1677MR01); 46-294231G1, 000099-1119-10 (A5401507); N/A, N/A (083027040099711); 46-294231G1, 002001-0530-21 (0920279017); 2377062-61, 00000010633YR6 (808433TMR2); 46-294231G1, 002001-1001-91 (PC1961MR01); 5308573, 00000300778MR8 (0856270046); 2226300, 000000000R5217 (PL1649MR01); N/A, N/A (PL0195MR01); 2226300, 000000000R4456 (70772MRS90); 5127452, 000000HDMR6307 (DK1297MR01); 2226300, 000000000R1970 (M4018076); 46-294231G1, 0000094-1209-4 (178SIGNA15); 2377062-2, 00000963500YM4 (0910272057); 46-294231G1, 002004-0211-05 (PC2703MR02); 46-294231G1, 002001-0301-23 (YM0731); 46-294231G1, 002003-0128-22 (YM2020); 2226300, 000000000R5133 (KAISERVJOMR); 5127452-3, 00001011BD6645 (082427040110); 46-294231G1, 0000098-0617-2 (414SIGNAHOR); 5308573, 00000294505MR3 (34407MRS01); 5107849, 00000275094MR1 (DZ1004MR02); 2226300, 000000000R6384 (0850270266); N/A, 00000000R7803A (2900429); 2214149-2, 00000003388M26 (MRR8311); N/A, 000000000R6449 (110040MR01); 2214149-2, 00000001634M25 (MRR8049); /A, N/A (0910272067); 2226300, 000000000R6691 (A200581902); 2226300, 000000000R6343 (787831SIGNAHDX); 2226300, 000000000R6461 (SA2161MR01); 2226300, 000000000R7934 (78003MRS01); 5308573, 00000291238MR4 (MA1135MR01); 5308573, 00000299348MR3 (07025MRS01); 2226300, 000000000R7004 (PL1426MR02); 2226300, 000000000R7487 (1306777); 2226300, 000000000R6769 (IL1041MR08); 395001, 00000280692MR5 (CC0820MR09); 2226300, 000000000R6889 (PL1735MR01); N/A, 000000000R6609 (M4464817); 2226300, 000000000R6707 (083027207801078); N/A, 00000014680GQ3 (M4033099); 5127452, 000000HDMR3604 (01203MRS01); N/A, N/A (0910274028); 2226300, 000000000R6050 (10563MRS01); 2226300, 000000000R5325 (SA1056MR01); 2226300, 000000000R6042 (SA1047MR01); 46-294231G1, 0002001-0219-8 (281477MWMR); 2226300, 000000000R5823 (A5718006); 2226300, 000000000R6111 (DZ1250MR01); N/A, N/A (DK1228MR07); 46-294231G1, 0000098-1207-9 (IE1028MR02); 2226300, 000000000R5149 (00398MRS01); 2226300, 000000000R5639 (00441MRS01); 2226300, 000000000R7809 (CBSMR1); 2226300, 000000000R5836 (083027862100837); 46-294231G1, 000004-0528-17 (083027602200104); 2226300, 000000000759 (083027603014911); N/A, 002004-0323-16 (GOLD0EU131); 2226300, 000000000R7996 (MRR7977); 2226300, 000000000R7781 (PL1639MR02); N/A, 000000000R7452 (2869229); N/A, N/A (HC2821MR01); 5308573, 00000294659MR8 (083027848131849); 46-294231G1, 0002000-1003-3 (0835270003); 2226300, 000000000R7237 (M4212028); 2395001, 00000276676MR4 (083027842100686); N/A, N/A (0910272088); N/A, N/A (0859270005); 2226300, 000000000R7837 (2913994HDXT); 5127452-3, 00001011BD667C (DZ1005MR01); 2226300, 000000000R7563 (IR1906MR01); 5911000-4 , 00001239FMM01H (083027800039612); 5127452-3, 00000919BD5B8F (FI1024MR02); 2226300, 000000000R7940 (SA2214MR01); 5127452-3, 00000837BD4ESC (0850270435); 46-294231G1, 000006-0523-15 (0850270200); 5127452-3, 00001003BD6435 (083027601000510); 5308573, 00000302567MR3 (083027243783711); 2377062, 00000941265YM1 (600118MR02); N/A, N/A (0850270291); 2395001-3, 00000289020MR0 (083027202100488); 2226300, 000000000R2404 (PL1169MR01); N/A, N/A (083027601013710); N/A, N/A (0910274011); 46-294231G1, 0000099-0506-1 (0910272025); 5127452, 000000HDMR5710 (IE1038MR01); N/A, N/A (083027001018110); N/A, N/A (083027592052828); 2226300, 000000000R4321 (SA1220MR01); 46-294231G1, 002002-1015-30 (PC0114MR04); 46-294231G1, 002003-0820-07 (PL1394MR01); 46-294231G1, 002001-0619-24 (IE1006MR01); 2226300, 0000000000R891 (0920279013); N/A, N/A (0910272038); 46-294231G1, 002000-0418-17 (SA1001MR02); 46-294231G1, 0`2001-0219-13 (259170MR05); N/A, 0000099-1217-4 (10140MRM44); N/A, 00000303865MR0 (10331MRS01); 46-294231G1, 0000096-0614-8 (801224MR); N/A, 0000096-0520-5 (AE1025MR01); 46-294231G1, 002001-0709-20 (0856270038); 5308573, 00000301170MR7 (MRR8186); 5127452, 000000HDMR6250 (00213MRS01); 46-294231G1, 002004-0917-09 (600092MR03); N/A, N/A (0856270050); 46-294231G1, 0000097-0922-9 (213340MR); 46-294231G1, 0000094-0928-7 (4224SIGN15); 5127452, 000000HDMR5751 (70383MRS01); N/A, N/A (083027592002609); 5127452-3, 00000903BD55FD (083027002100119); 2226300, 000000000R7217 (083027848480709); 5127452, 00000HDMR10911 (801262MR2); 46-294231G1, 002001-0425-10 (765INNERMR); 2377062, 00000962578YM1 (519685MR); 2377062-61, 00000009061YR3 (DK1401MR01); 2226300, 000000000R6839 (501268MR); 5308573, 00000301743MR1 (45019MRS01); 5127452, 000000HDMR6310 (0834270012); 2226300, 000000000R6802 (H10649MR06); 5107849-2, 00000288195MR1 (H10652MR08); 2395001, 00000286008MR8 (21234MRM02); 2108800-2, 535651 (SHIMR02604); 2226300, 000000000R6903 (503763MR1); 46-294231G1, 002002-0627-18 (RADNETMR611); N/A, N/A (083027820006912); N/A, N/A (4954MR15T); N/A, 000000000R6911 (2805953); 46-294231G1, 0000091-1022-5 (205431MR2); 5911000-4 ,00001227FMM06G (083027200018712); N/A, N/A (083027590475005); N/A, N/A (1111395); 2214149-2, 00000005560M28 (IQ1394MR01); N/A, N/A (RADNETMRM01); 46-294231G1, 002005-0627-18 (ALLIANCESIG421); 46-294231G1, 0000099-0125-9 (907276ARMR); N/A, N/A (6444MRDICO); N/A, N/A (2809956); 46-294231G1, 0000001-0215-1 (INSIGHTMR1915); 2226300, 000000000R6838 (623445MR1); 2377062-2, 00000943572YM8 (201632MR); 46-294231G1, 0002002-0510-7 (ALLIANCESIG474); 46-294231G1, 002002-0125-22 (559585CCICMR2); 46-294231G1, 002002-0405-12 (205783TSMR); 5127452, 000000HDMR5415 (850484OMRI); 5308573, 00000300970MR1 (RADNETMR507); 5127452, 00000HDMR10886 (916423MR); 5127452, 000000HDMR3674 (501370CHMR); 5107849, 00000269580MR7 (PL3040MR01); 46-294231G1, 002004-1209-10 (YM1260); 2395001-2, 00000284023MR9 (083027002101567); 46-294231G1, 0000007-403-15 (083027802000517); 5127452, 000000HDMR5483 (1063THDX); 5127452, 000000HDMR3964 (SA1001MR03); 2395001-2, 00000280672MR7 (00440MRS02); N/A, 000000HDMR5233 (110037MR02); 5127452-3, 00001012BD66E0 (083027000143110); 2358431, 00000000WB0607 (AZ4013MR01); 5127452-3, 00001009BD6557 (BA4080MR01); 5308573-2, 00000299495MR2 (MRWB0540); N/A, N/A (5606MRS3T); 5127452-3, 00001009BD65F2 (083027009445910); N/A, N/A (773665MR3T); 5127452-3, 00001009BD6579 (2887688HDXT30T); 5127452, 00000275305MR1 (831423MR); N/A, N/A (2558948); 46-294231G1, 0002005-0607-4 (7067213TMR); 5127452-3, 000S0933BD5DFD (719542MR); N/A, 002003-0912-12 (DK1047MR04); 5127452, 000000HDMR5406 (0910273033); 5127452, 000000HDMR2642 (0850270218); 5421824-2, 00000310388MR4 (083027240009214); 2377062-2, 00000963651YM5 (IL1194MR01); 46-294231G1, 002005-1020-29 (HBLUCOMR01); N/A, 002004-0917-02 (00473MRS01); N/A, N/A (MR4013522); N/A, N/A (2679752); N/A, N/A (2541440); 2226300, 000000000R1733 (MRR1733); 2226300, 0000000000R273 (MRR273); 46-294231G1, 000099-0205-11 (0910272023); N/A, N/A (PL1681MR01); N/A, N/A (H4443MR04); 5308573, 00000301741MR5 (MRR8267); 5308573, 00000301223MR4 (2918179MR); 2214149-2, 00000002339M20 (MRR8325); 2214149-2, 00000003155M29 (MRR8701); N/A, 000096-0126-27 (A020MR02); N/A, 00000175300MR3 (10079MRM42); N/A, N/A (5272MRSIGNA4X); 2226300, 000000000R5436 (DEN00700); 2226300, 000000000R6601 (CE165329); N/A, N/A (0910273026); 46-294231G1, 002002-1126-23 (082427100148); 5308573, 00000300700MR2 (DEN00933); 46-294231G1, 000099-0122-13 (YE600210); 5127452, 000000HDMR5250 (DZ1009MR02); 2226300, 000000000R3823 (CS1275MR02); 2226300, 000000000R2680 (ZA2159MR02); 2214149-2, 00000014057GQ4 (082427030054); 2226300, 000000000R5680 (082427130073); 46-294231G1, 002005-1201-30 (HU1039MR03); 2226300, 000000000R6212 (DK1306MR02); 2214149-2 ,00000007668M27 (082427280013); 5308573, 00000301498MR2 (082427240016); 2226300, 000000000R7667 (082427050024); 5308573, 00000301500MR5 (MRR8257); 2226300, 000000000R6024 (CS1020MR01); 5308573, 00000291246MR7 (2761886); N/A, 000000000R6222 (2726027); 5308573, 00000301383MR6 (082427030063); 2226300, 000000000R7426 (082427100239); 46-294231G1, 002001-0718-35 (0910273020); 5308573, 00000301264MR8 (082427030061); 2226300, 000000000R6289 (082427160057); 2226300, 000000000R6200 (MPX19700); 2226300, 000000000R7028 (082427210051); 2226300, 000000000R6582 (DZ1259MR01); 2226300, 000000000R5895 (SA1129MR02); 2226300, 000000000R6955 (2807539); 2226300, 000000000R5724 (PPS18202); 2226300, 000000000R5733 (RU1550MR01); 2226300, 000000000R6139 (BG4538MR01); 2226300, 000000000R6439 (RU1279MR01); 2226300, 000000000R5352 (GE1004MR01); N/A, 000000000R6511 (2767032); 2214149-2, 00000004486M27 (082427080033); 2226300, 000000000R6806 (082427040088); 2226300, 000000000R5120 (0910273030); 2226300, 000000000R7718 (082427250022); 2226300, 000000000R5579 (RU1067MR02); N/A, N/A (082427040098); 5421824-2, 00000310944MR4 (082427080047); 2226300, 000000000R7870 (MRR7870); 46-294231G1, 000099-0817-11 (MPX47002); 5308573, 00000298557MR0 (0853270042IB1); 5421824, 00000311257MR0 (082427190081); 2214149-2, 00000002805M20 (082427040129); 2226300, 000000000R7050 (UPW00400); 2226300, 000000000R7884 (NPV85600); 2214149-2, 00000002750M28 (CS1338MR01); 2226300, 000000000R7342 (082427030049); 2214149-2, 00000003393M26 (RU1031MR01); 2226300, 000000000R7130 (CS1319MR01); 5127452-3, 00001012BD66A2 (0853270074IB1); 46-294231G1, 0002005-1201-2 (082427120126); N/A, N/A (082427160098); 46-294231G1, 0000099-0414-7 (CP223906); N/A, 000000000R3618 (BG4641MR01); 46-294231G1, 002002-1030-14 (H3943MR01); 46-294231G1, 002001-1018-36 (ZA2103MR01); 2214149-2, 00000005509M25 (082427230084); 5421824, 00000311572MR2 (082427140132); 5269997-9, 00000015025YR0 (0828270033); 2226300, 000000000R5260 (082427030103); 46-294231G1, 0000095-1120-5 (RU1089MR01); 5421824, 00000311574MR8 (082427230107); 2377062-61, 00000011084YR1 (27008MRS01); 2214149-2, 00000003870M23 (RU1457MR01); 5127452-3 ,00000838BD4ECB (35062MRS02); 7849-2, 00000283036MR2 (55027MRS01); 2395001, 00000285169MR9 (JO4109MR01); 5421824-2, 00000311280MR2 (082427230097); 507849-2, 00000288588MR7 (DK1393MR01); 2214149-2 , 00000007860M20 (082427210142MM3); 2214149-2, 00000004304M22 (H2294MR07); 2226300, 000000000R7091 (48021MRS01); 2214149-2, 00000004492M25 (KA2839MR01); 2214149-2, 00000003509M27 (RU1173MR01); 2226300, 000000000R3065 (082427160161); N/A, N/A (570844HCM3T); 5127452, 000000HDMR3447 (DPM32124); 5421834-2, 00000311145MR7 (082427130084); 5127452-3, 00000919BD5BB7 (082427220041); N/A, N/A (082427120070); 5127452, 000000HDMR1125 (082427040073); 2395001-2, 00000275466MR1 (DEN51500); 5127452, 000000HDMR5758 (082427040092); N/A, 0000000HDMR357 (CS1006MR03); 5127452, 000000HDMR1900 (0910274020); 5127452-3, 00001003BD640C (082427080026); 5127452-3, 00000938BD5F17 (082427190069); 5127452, 00000275290MR5 (DPC79014); N/A, N/A (082427100192); 5308573-2, 00000299275MR8 (082427200036); 5127452-3, 00001012BD66E1 (PER08626); 5127452-3, 00001003BD640D (PER28007); 5127452-3, 00000938BD5F14 (RO1189MR01); 5127452, 000000HDMR6035 (082427140120); 5107849-3, 00000302786MR9 (MRJ061); N/A, 00000293706MR8 (NPV58000); N/A, N/A (0910276003); N/A, N/A (626570MR); N/A, N/A (907561AIMR); N/A, N/A (309451MR); 46-294231G1, 000002-1003-15 (201967MR); 2226300, 000000000R5771 (0853270133GS); 46-294231G1, 0000099-0521-6 (ALLIANCESIG85); 2377062-61, 00000011265YR6 (MRR6469); N/A, N/A (083027067005818); N/A, N/A (520GVHMR); N/A, N/A (2939576HDXT); N/A, N/A (520872MR); 2226300, 000000000R1970 (H10650MR08); 2226300, 000000000R6733 (A197998701); 2226300, 000000000R6448 (0828270030); 2214149-2 , 00000015124GQ1 (0834270011); N/A, 002000-0907-10 (CPM80615); 5127452, 000000HDMR5960 (0834270010); 5127452-3 , 00001009BD6578 (0834270015); 46-294231G1, 000096-0126-15 (YP630007); 2226300, 000000000R6352 (YP600200); 2226300, 000000000R2306 (RU1086MR01); 46-294231G1, 0002002-0506-3 (FEC80061); 2377062, 4337424 (MEX34410); 5127452, 000000HDMR4309 (UPW28900); 2395001, 00000280327MR8 (UPW21605); 2226300, 000000000R5291 (YP601000); 46-294231G1, 002000-0426-18 (0856270029); 46-294231G1, 0000099-0722-1 (ME208331); 2226300, 000000000R4276 (FPF35809); 46-294231G1, 002004-1117-12 (0828270020IB01); N/A, 00000001400M21 (NPV84508); 2377062-61 , 00000011251YR6 (MEX71105); 2395001, 00000283516MR3 (MP118825); 46-294231G1, 002003-1027-20 (A5103135); 5396469-2, 00000310306MR6 (NO1005MR02); 46-294231G1, 0000095-0810-3 (A5278401); 46-294231G1, 002003-0903-12 (A5111732); 2226300, 000000000R4164 (A5684809); 2226300, 000000000R6773 (PER08305); 2226300, 000000000R6502 (2766544); 2226300, 000000000R5403 (H1063MR18); 2226300, 000000000R5912 (H1420MR14); 46-294231G1, 002002-0809-08 (CPZ09609); 46-294231G1, 002003-0708-01 (PER00417); 46-294231G1, 002000-1013-15 (A5632405); 2214149-2, 00000002229M23 (MRR8295); 2226300, 000000000R2375 (A5745509); N/A, N/A (0853270100); N/A, N/A (083027100973317); N/A, 002002-0814-05 (H1066MR16); 46-294231G1, 000093-0528-15 (DE528434); 2226300, 000000000R7136 (0910272086); 2395001, 00000280778MR2 (0843270015); N/A, N/A (0910274013); 46-294231G1, 0000098-0608-5 (0910272020); 2395001, 00002835171MR1 (083027040006015); 2226300, 000000000R6574 (2767673); 2226300, 000000000R5159 (3882647967MROCC); 46-294231G1, 0002005-0331-3 (H12041MR02); 46-294231G1, 0000000MRU0407 (IE1025MR01); 2226300, 000000000R3756 (PVIL01MR01); 46-294231G1, 000097-1107-10 (PNUE01MR03); 46-294231G1, 002005-0823-17 (HPAD01MR01); N/A, N/A (PL1774MR02); 2226300, 000000000R6092 (0859270002); 2226300, 000000000R6021 (0859270001); 5127452, 000000HDMR1019 (RU3023MR01); N/A, 000000HDMR5374 (PM1633MR03); 2226300, 000000000R4692 (PL1704MR01); N/A, N/A (083027242000437); 2226300, 000000000R1663 (PC0976MR01); 46-294231G1, 0002001-0323-1 (MA1152MR01); 46-294231G1, 0000095-0912-5 (PX2261MR02); 5127452, 1529759 (MC6946MR07); 2395001-2, 00000278549MR1 (RU1127MR02); 46-294231G1, 000006-0523-10 (RU1834MR01); 5127452, 3877 (00240MRS01); 46-294231G1, 002000-0816-10 (PC0396MR03); 46-294231G1, 0000098-0528-5 (A5566913); 5127452, 000000HDMR2483 (A171995610); 2226300, 000000000R2753 (DK1343MR01); 2226300, 000000000R6499 (A5192410); N/A, 000000000R7233 (1324655); 5308573, 00000302145MR8 (MRR8303); 2226300, 000000000R6460 (22MRHDX2762068); 2226300, 000000000R5214 (H17641MR12); N/A, 02001-12210-13 (A042MR01); 2226300, 000000000R7197 (2835341HDX); 2395001-2, 00000286650MR7 (H2193MR05); 2226300, 000000000R6520 (A51273121); 2226300, 000000000R2396 (MRR2396); 46-294231G1, 000007-0530-16 (UK28); 46-294231G1, 000097-0331-10 (DK1029MR01); 2214149-2 , 00000009304M27 (MRR8592); 2320288, 00000096546WH7 (10695MRM01); 2226300, 000000000R2753 (DK1394MR01); 2357500, 00000249262MR7 (16683THDX); 2226300, 000000000R7564 (A5146405); N/A, N/A (650742KSSFMR); 46-294231G1, 0000096-0419-4 (0920279006); N/A, N/A (916SRFMR); N/A, N/A (7366); 46-294231G1, 002002-0517-19 (IR1026MR01); N/A, N/A (0847270018); 46-294231G1, 0000099-0329-4 (A5127353); N/A, N/A (0847270025); 46-294231G1, 002001-1018-32 (PC7425MR01); N/A, 000000000R1116 (AH4511MR02); N/A, N/A (843549MR); 2226300, 000000000R1533 (LLS0120); 5308573, 00000293870MR2 (CC0820MR10); 46-294231G1, 002004-0713-22 (393115MR01); N/A, N/A (00200MRS03); N/A, N/A (ALLIANCESIG270); N/A, 0002000-1102-1 (DK1042MR01); 46-294231G1, 002004-0206-15 (397480MR01); N/A, N/A (707571KSRMR); N/A, N/A (8183MRLX); 2377062-2, 00000951076YM9 (805687MR2); 2226300, 0000000000R575 (HC1770MR01 ); N/A, N/A (A5183710); N/A, N/A (585786MRI); 2226300, 000000000R6712 (0847270023); 2226300, 000000000R5610 (0843270017); 46-294231G1, 002002-0315-13 0847270001); 46-294231G1, 0002000-0301-4 (SA1012MR01); N/A, N/A (570844TWINSPEED); N/A, N/A (422SIGNA15); N/A, 002000-1221-14 (DK1047MR03); 46-294231G1, 002003-0929-03 (N4192928); N/A, N/A (0910274008); 46-294231G1, 002003-0818-13 (34026MRS03); N/A, 00000275019MR8 (35045MRS03); 2226300, 000000000R5050 (082427130074); N/A, 00000310642MR4 (082427130082); 5421824, 00000311573MR0 (082427180036); 5453862-2, 308735MR0 (082427240048); N/A, N/A (PL3331MR01); N/A, N/A (409637LX1); N/A, N/A (618798SMR). Recalling Firm/ Manufacturer GE Medical Systems, LLC 3200 N Grandview Blvd Waukesha WI 53188-1693 For Additional Information Contact 800-437-1171 Manufacturer Reason for Recall The latest on-site software version may not have been reinstalled at some sites after service activities were performed resulting in images possibly flipped left/right and/or there can be patient data mismatch. FDA Determined Cause 2 Software change control Action The firm, GE Healthcare, notified its customers on or before 06/07/2018 via a letter entitled "URGENT MEDICAL DEVICE CORRECTION". The letter described the product, problem and actions to be taken. The letter informed customers that GE Healthcare will inspect all affected systems and ensure the proper software version is installed. Instructions for customers to confirm the current software included the following steps: 1. Click on tools icon located on the top left of the screen; 2. Click on Service Browser Button; 3. The MR Service Desktop browser will launch (may take 1 minute to launch); 4. Look at the Release Information on the screen and compare both MrpApps Build Number and Service Pack Build Number on the screen to two tables provided in the safety issue section of the customer letter. ************If you have identified that your system has invalid software version, please contact GE Healthcare service representative immediately.****** ******GE Healthcare will advise you after your software version has been inspected and inform you if any correction is needed. If a correction is needed, a service representative will contact you to arrange for this correction.************Questions or concerns regarding this notification can be directed to GE Healthcare Service at 1-800-437-1171 or your local Service Representative. Quantity in Commerce 5,200 total units Distribution Worldwide Distribution: US (nationwide) including states of:: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV; and countries of: Algeria, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, CHINA, Colombia, Costa Rica, Croatia, CYPRUS, Czech Republic, DENMARK, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, Georgia, Greece, Guadeloupe, Guatemala, Guyana, Honduras, Hong Kong, Iceland, India, Indonesia, Iran, IRAQ, Ireland, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea (Republic of), Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Macedonia, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Pakistan, Paraguay, Philippines, Portugal, Puerto Rico, Russia, Saudi Arabia, SINGAPORE, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, United Arab Emirates, Uruguay, Uzbekistan, Venezuela, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LNH 510(K)s with Product Code = LNH 510(K)s with Product Code = LNH - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, GE Healthcare, notified its customers on or before 06/07/2018 via a letter entitled "URGENT MEDICAL DEVICE CORRECTION". The letter described the product, problem and actions to be taken. The letter informed customers that GE Healthcare will inspect all affected systems and ensure the proper software version is installed. Instructions for customers to confirm the current software included the following steps: 1. Click on tools icon located on the top left of the screen; 2. Click on Service Browser Button; 3. The MR Service Desktop browser will launch (may take 1 minute to launch); 4. Look at the Release Information on the screen and compare both MrpApps Build Number and Service Pack Build Number on the screen to two tables provided in the safety issue section of the customer letter. ************If you have identified that your system has invalid software version, please contact GE Healthcare service representative immediately.****** ******GE Healthcare will advise you after your software version has been inspected and inform you if any correction is needed. If a correction is needed, a service representative will contact you to arrange for this correction.************Questions or concerns regarding this notification can be directed to GE Healthcare Service at 1-800-437-1171 or your local Service Representative. | SoftwareChangeControl
70 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164777 | Z-2266-2018 | CARESCAPE R860 ventilators with software version 10SP05******Product Usage:***The CARESCAPE R860 ventilator is designed to provide mechanical ventilation or support to neonatal, pediatric, and adult patients weighing 0.25 kg and above. The CARESCAPE R860 ventilator is a microprocessor based, electronically controlled, pneumatically driven ventilator that includes integrated monitoring of FiO2, airway pressure, flow, and volume. Additional respiratory gas monitoring capabilities are supported through the use of optional GE patient monitoring modules. Not all features are available for all patient types or product configurations. The CARESCAPE R860 ventilator is not a pulmonary function calculation device. The system is designed for facility use, including within facility transport, and should only be used under the orders of a clinician.*** | CARESCAPE R860 ventilators with software version 10SP05 | 2 | CDRH | 2018-06-20 | 2018-06-01 21:04 | 2021-04-29 14:13 | Datex-Ohmeda, Inc. | Potential for the display processor to experience an***unexpected failure. | Class 2 Device Recall CARESCAPE R860 ventilators with software version 10SP05 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall CARESCAPE R860 ventilators with software version 10SP05 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall CARESCAPE R860 ventilators with software version 10SP05 Date Initiated by Firm December 08, 2017 Date Posted June 01, 2018 Recall Status 1 Terminated 3 on April 29, 2021 Recall Number Z-2266-2018 Recall Event ID 80173 510(K)Number K142679 Product Classification Ventilator, continuous, facility use - Product Code CBK Product CARESCAPE R860 ventilators with software version 10SP05************Product Usage:******The CARESCAPE R860 ventilator is designed to provide mechanical ventilation or support to neonatal, pediatric, and adult patients weighing 0.25 kg and above. The CARESCAPE R860 ventilator is a microprocessor based, electronically controlled, pneumatically driven ventilator that includes integrated monitoring of FiO2, airway pressure, flow, and volume. Additional respiratory gas monitoring capabilities are supported through the use of optional GE patient monitoring modules. Not all features are available for all patient types or product configurations. The CARESCAPE R860 ventilator is not a pulmonary function calculation device. The system is designed for facility use, including within facility transport, and should only be used under the orders of a clinician. Code Information Lot / Serials #'s: 103841473, 103725811-1, 103725811-2, CBRW01546, CBRW01547, CBRU02110, CBRV00344, CBRV02569, CBRV00298, 103721045, 103808843, CBRV03050, CBRV03054, CBRV03056, CBRV03063, CBRV03064, CBRV03071, CBRV03072, CBRV03073, CBRW00417, CBRW00419, CBRW00424, CBRW00441, CBRW00443, CBRW00458, CBRW00460, CBRW00461, CBRW02306, CBRW02627, CBRW02629, 103721436, CBRU02450, CBRW01735, CBRW01044, CBRW01048, CBRW01049, CBRW01051, CBRW01052, CBRW01053, CBRW01056, CBRW01057, CBRW01059, CBRW01061, CBRW01062, CBRW01063, CBRW01075, CBRW01076, CBRW01077, CBRW01232, CBRW01814, CBRW01410, CBRW01412, 103798917, CBRW01918, CBRV02578, CBRW02442, CBRW02449, CBRW01830, CBRW01833, CBRW01838, CBRW01839, CBRW01849, CBRW01860, CBRW01889, CBRW01895, CBRW01896, CBRW01897, CBRW01898, CBRW01899, CBRW01900, CBRW01902, CBRW01903, CBRW01904, CBRW01905, CBRW01906, CBRW01909, CBRW01932, CBRV01459, 103730032, 103721104, CBRW01244, CBRW02420, CBRW01347, 103771244, CBRW02661, CBRU02452, 11495021, 70175736-1, 70175736-10, 70175736-100, 70175736-101, 70175736-102, 70175736-103, 70175736-104, 70175736-105, 70175736-106, 70175736-107, 70175736-108, 70175736-109, 70175736-11, 70175736-110, 70175736-111, 70175736-112, 70175736-113, 70175736-114, 70175736-115, 70175736-116, 70175736-117, 70175736-118, 70175736-119, 70175736-12, 70175736-120, 70175736-121, 70175736-122, 70175736-123, 70175736-124, 70175736-125, 70175736-126, 70175736-127, 70175736-128, 70175736-129, 70175736-13, 70175736-130, 70175736-131, 70175736-132, 70175736-133, 70175736-134, 70175736-135, 70175736-136, 70175736-137, 70175736-138, 70175736-139, 70175736-14, 70175736-140, 70175736-141, 70175736-142, 70175736-143, 70175736-144, 70175736-145, 70175736-146, 70175736-147, 70175736-148, 70175736-149, 70175736-15, 70175736-150, 70175736-151, 70175736-152, 70175736-153, 70175736-154, 70175736-155, 70175736-156, 70175736-157, 70175736-158, 70175736-159, 70175736-16, 70175736-160, 70175736-161, 70175736-162, 70175736-163, 70175736-164, 70175736-165, 70175736-166, 70175736-167, 70175736-168, 70175736-169, 70175736-17, 70175736-170, 70175736-171, 70175736-172, 70175736-173, 70175736-174, 70175736-175, 70175736-176, 70175736-177, 70175736-178, 70175736-179, 70175736-18, 70175736-180, 70175736-181, 70175736-182, 70175736-183, 70175736-184, 70175736-185, 70175736-186, 70175736-187, 70175736-188, 70175736-189, 70175736-19, 70175736-190, 70175736-191, 70175736-192, 70175736-193, 70175736-194, 70175736-195, 70175736-196, 70175736-197, 70175736-198, 70175736-199, 70175736-2, 70175736-20, 70175736-200, 70175736-21, 70175736-22, 70175736-23, 70175736-24, 70175736-25, 70175736-26, 70175736-27, 70175736-28, 70175736-29, 70175736-3, 70175736-30, 70175736-31, 70175736-32, 70175736-33, 70175736-34, 70175736-35, 70175736-36, 70175736-37, 70175736-38, 70175736-39, 70175736-4, 70175736-40, 70175736-41, 70175736-42, 70175736-43, 70175736-44, 70175736-45, 70175736-46, 70175736-47, 70175736-48, 70175736-49, 70175736-5, 70175736-50, 70175736-51, 70175736-52, 70175736-53, 70175736-54, 70175736-55, 70175736-56, 70175736-57, 70175736-58, 70175736-59, 70175736-6, 70175736-60, 70175736-61, 70175736-62, 70175736-63, 70175736-64, 70175736-65, 70175736-66, 70175736-67, 70175736-68, 70175736-69, 70175736-7, 70175736-70, 70175736-71, 70175736-72, 70175736-73, 70175736-74, 70175736-75, 70175736-76, 70175736-77, 70175736-78, 70175736-79, 70175736-8, 70175736-80, 70175736-81, 70175736-82, 70175736-83, 70175736-84, 70175736-85, 70175736-86, 70175736-87, 70175736-88, 70175736-89, 70175736-9, 70175736-90, 70175736-91, 70175736-92, 70175736-93, 70175736-94, 70175736-95, 70175736-96, 70175736-97, 70175736-98, 70175736-99, 103725706, CBRW02482, 103721074, CBRW01588, CBRW02287, 103721038, CBRW01151, CBRW01152, CBRW01155, CBRW01156, CBRW01157, CBRW01158, 103737388, CBRU02113, CBRU02114, CBRU02124, CBRU02126, CBRU02128, CBRU02135, CBRU02143, CBRU02454, CBRU02456, CBRV01435, CBRV01436, CBRV01438, CBRV01449, CBRV01451, CBRV01453, CBRV02039, CBRV02117, CBRV02156, CBRV02179, CBRV02188, CBRV02199, CBRV02540, CBRV02542, CBRV02563, CBRV02570, CBRV02571, CBRV02579, CBRV02583, CBRV02584, CBRV01458, CBRV02025, CBRW00442, CBRW00459, CBRW01055, CBRW01058, CBRW01074, 11494687, 103722840, 103771194, 103785995-1, 103785995-2, 103785995-3, 103785995-4, 103785995-5, 11466039-1, 11466039-2, 11466039-3, 11466039-4, 11466039-5, CBRW01437, CBRW01448, 11485242-1, 11485242-2, CBRW01112, CBRW01464, CBRW01797, CBRW01799, CBRW01800, CBRW01810, CBRW01831, CBRW01842, CBRW01846, CBRW01866, CBRW01867, CBRW01868, CBRW01869, CBRW01878, CBRW01879, CBRW01880, CBRW01881, CBRW01882, CBRW01883, CBRW01809, CBRW01934, CBRW01937, CBRW01946, CBRW01947, CBRW01948, 6527249, 6528751, 6526140-1 6526140-2, 6528030-1, 6528030-2, 6528030-3, CBRW01401, CBRW01439, CBRW01964, CBRW01965, CBRW01971, CBRW01972, CBRW01974, CBRW02707, CBRW02735, CBRW02753, CBRV02796, CBRW01328, CBRW01771, CBRW01783, CBRW01794, CBRW01795, CBRW01796, CBRW01798, CBRW01802, CBRW01807, CBRW01767, CBRW01768, CBRW01770, CBRW01772, CBRW02562, CBRW02573, CBRW02574, CBRW02575, CBRW02576, CBRW02577, CBRW02578, CBRW02579, CBRW02580, CBRW02581, CBRW02582, CBRW02583, CBRW02584, CBRW02585, CBRW02586, CBRW02587, CBRW02588, CBRW02598, CBRW02599, CBRW02600, CBRW02601, CBRW02602, CBRW02603, CBRW02614, CBRW02615, CBRW02617, CBRW02631, CBRW02632, CBRW02634, CBRW02635, CBRW02636, CBRW02637, CBRW02638, CBRW02639, CBRW02640, CBRW02641, CBRW02642, CBRW01147, CBRW01149, CBRW01573, CBRW01891, CBRW01892, CBRW01893, CBRW01894, CBRW02605, CBRW02606, CBRW02608, CBRW02609, CBRW02610, CBRW02612, CBRW02613, CBRW02616 CBRW02421, CBRW02422, CBRW02424, CBRW02425, CBRW01983, CBRW01992, CBRW01644, CBRW01645, CBRW01658, CBRW01413, CBRW01414, CBRW01582, CBRW01586, CBRW02301, CBRW02304, CBRW02305, CBRW02307, CBRW02308, CBRW02309, CBRW02310, CBRW02311, CBRW02312, CBRW02314, CBRW02506, CBRW02507, CBRW02563, CBRW02564, CBRW02659, CBRW02660, CBRW02696, CBRW02697, CBRW02698, CBRW02757, CBRW02758, CBRW02772, CBRW02773, CBRW02774, CBRW00934, CBRW00935, CBRW00936, CBRW00937, CBRW00938, CBRW00939, CBRW00940, CBRW01634, CBRW01669, CBRW01901, CBRW01907, CBRW01908, CBRW01910, CBRW01911, CBRW01912, CBRW01913, CBRW01914, CBRW01915, CBRW01916, CBRW01917, CBRW01919, CBRW01920, CBRW01921, CBRW01922, CBRW01923, CBRW01924, CBRW01925, CBRW01926, CBRW01927, CBRW00942, CBRW00943, CBRW00944, CBRW02713, CBRW02763, CBRW02764, CBRW02786, CBRW02810, CBRW02526, CBRU00173, 10190157, 10190159, 10191029, 10193770, 10194582, CBRU00498, CBRU00509, CBRU00512, CBRU00517, CBRU00522, CBRU00523, CBRU00540, CBRU00541, CBRU00550, CBRU00551, CBRU00552, CBRU00553, CBRU00554, CBRU00555, CBRU00559, CBRU00560, CBRU00561, CBRU00562, CBRU00563, CBRU00565, CBRU00566, CBRU00567, CBRU00568, CBRU00569, CBRU00570, CBRU00571, CBRU00572, CBRU00573, CBRU00574, CBRU00575, CBRU00577, CBRU00579, CBRU00580, CBRU00581, CBRU00593, CBRU00594, CBRU00595, CBRU00596, CBRU00597, CBRU00598, CBRU00599, CBRU00600, CBRU00601, CBRU00602, CBRU00603, CBRU00604, CBRU00605, CBRU00606, CBRU00607, CBRU00608, CBRU00613, CBRU00614, CBRU00619, CBRU00620, CBRU00621, CBRU00622, CBRU00624, CBRW00918, CBRW02249, CBRW00919, CBRW01284, CBRW01570, CBRW01630, 11484134, CBRW01604, CBRW01856, CBRW01857, CBRW01858, CBRW01375, CBRW01447, CBRW01619, CBRW01715, CBRW01742, CBRW01746, CBRW01753, CBRW01754, CBRW01755, CBRW01757, CBRW01758, CBRW01804, CBRW01805, CBRW01808, CBRW01820, CBRW01821, CBRW01823, CBRW01824, CBRW01825, CBRW01827, CBRW01834, CBRW01835, CBRW01836, CBRW02023, CBRW02046, CBRW02048, CBRW02049, CBRW02053, CBRW02055, CBRW02056, CBRW02060, CBRW02061, CBRW02063, CBRW02064, CBRW02065, CBRW02066, CBRW02067, CBRW02069, CBRW02070, CBRW02071, CBRW02072, CBRW02073, CBRW02074, CBRW02075, CBRW02076, CBRW02077, CBRW02078, CBRW02079, CBRW02080, CBRW02081, CBRW02082, CBRW02083, CBRW02084, CBRW02085, CBRW02086, CBRW02087, CBRW02088, CBRW02089, CBRW02090, CBRW02091, CBRW02092, CBRW02093, CBRW02094, CBRW02095, CBRW02096, CBRW02097, CBRW02098, CBRW02099, CBRW02100, CBRW02102, CBRW02105, CBRW02107, CBRW02108, CBRW02109, CBRW02110, CBRW02114, CBRW02115, CBRW02120, CBRW02122, CBRW02123, CBRW02124, CBRW02125, CBRW02126, CBRW02127, CBRW02128, CBRW02129, CBRW02130, CBRW02131, CBRW02144, CBRW02145, CBRW02146, CBRW02147, CBRW02216, CBRW02611, CBRW02633, CBRW02648, CBRW02652, CBRW02653, CBRW02654, CBRW02655, CBRW02656, CBRW02657, CBRW02658, CBRW02676, CBRW02677, CBRW02678, CBRW02679, CBRW02684, CBRW02686, CBRW02687, CBRW02688, CBRW02689, CBRW02690, CBRW02693, CBRW02694, CBRW02695, CBRW02699, CBRW02700, CBRW02702, CBRW02037, CBRW02051, CBRW02285, CBRW02298, CBRW02103,CBRW02104, CBRW02143, CBRW01578, CBRW01928, CBRW01929, CBRW01930, CBRW01933, CBRW01952, CBRW01954, CBRW02021, CBRW02044, CBRW02045, CBRW02047, CBRW02050, CBRW02052, CBRW02038, CBRW01991, CBRW02018, CBRW02028, CBRW02034, CBRW02035, CBRW02036, CBRW00982, CBRW00983, CBRW00984, CBRW00985, CBRW00986, CBRW01829, CBRW01850, CBRW01851, CBRW01855, CBRW01859, CBRW01568, CBRW01574, CBRW01108, CBRW02592, CBRW02593, CBRW02594, CBRW01078, CBRW01079, CBRW01080, CBRW01081, CBRW01082, CBRW01083, CBRW01084, CBRW01085, CBRW01086, CBRW01087, CBRW01614, CBRW01615, CBRW01617, CBRW01621, CBRW01628, CBRW01629, CBRW01637, CBRW01638, CBRW01639, CBRW01650, CBRW01663, CBRW01202, CBRW01203, CBRW01207, CBRW01211, CBRW01217, CBRW01218, CBRW01745, CBRW01585, CBRW01593, CBRW01269, CBRW01270, CBRW01271, CBRW01272, CBRW01273, CBRW01274, CBRW01275, CBRW01276, CBRW01277, CBRW01278, CBRW01279, CBRW01280, CBRW01281, CBRW01282, CBRW01688, CBRW01734, CBRW01748, CBRW01765, CBRW01766, CBRW01773, CBRW01781, CBRW01782, CBRW01784, CBRW01785, CBRW01786, CBRW01787, CBRW01789, CBRW01790, CBRW01791, CBRW01792, CBRW01793, CBRW01938, CBRW01941, CBRW01942, CBRW01943 CBRW01944 CBRW01945, CBRW01949, CBRW01956, CBRW01957, CBRW01958, CBRW02483, CBRW02484, CBRW02485, CBRW02486, CBRW02487, CBRW02490, CBRW02493, CBRW02494, CBRW02496, CBRW02497, CBRW02498, CBRW02499, CBRW02500, CBRW02501, CBRW02502, CBRW02503, CBRW02504, CBRW02511, CBRW02565, CBRW02566, CBRW02568, CBRW02590, CBRW02591, CBRW01551, CBRW01564, CBRW01565, CBRW01566, CBRW01567, CBRW01601, CBRV01842, 11489276, CBRW01571, CBRW01572, CBRW01959, CBRW01961, CBRW01710, CBRW02423, CBRW02433, CBRW02448, CBRW02548, CBRW02553, CBRW02559, CBRW02589, CBRW02595, CBRW02596, CBRW02618, CBRW02619, CBRW02622, CBRW02623, CBRW02625, CBRW02626, CBRW02646, CBRW02647, CBRW02668, CBRW02672, CBRW01285, CBRW01587, CBRW01589, CBRW01594, CBRW01596, CBRW01110, CBRW01120, CBRW01121, CBRW01138, CBRV01741, CBRW01045, CBRW01046, CBRW01047, CBRW01119, CBRW01136, CBRW01141, CBRW01713, CBRW01720, CBRW01732, CBRW01743, CBRW01744, CBRW01756, CBRW01761, CBRW01832, CBRW01862, CBRW01885, CBRW01886, CBRW02170, CBRW02329, CBRW02330, CBRW02331, CBRW02332, CBRW02333, CBRW02334, CBRW02335, CBRW02336, CBRW02337, CBRW02338, CBRW02339, CBRW02340, CBRW02341, CBRW02342, CBRW02343, CBRW02344, CBRW02345, CBRW02346, CBRW02347, CBRW02348, CBRW02351, CBRW02352, CBRW02353, CBRW02354, CBRW02356, CBRW02357, CBRW02358, CBRW02374, CBRW02378, CBRW02379, CBRW02380, CBRW02381, CBRW02382, CBRW02383, CBRW02384, CBRW02385, CBRW02387, CBRW02388, CBRW02390, CBRW02391, CBRW02392, CBRW02393, CBRW02394, CBRW02471, CBRW02472, CBRW02473, CBRW02474, CBRW02475, CBRW02489, CBRW02491, CBRW02492, CBRW02508, CBRW02510, CBRW02515, CBRW02516, CBRW02525, CBRW02528, CBRW02533, CBRW02543, CBRW02546, CBRW02572, CBRW02597, CBRW02620, CBRW02621, CBRW02711, CBRW02746, CBRW02747, CBRW02752, CBRW02754, CBRW02755, CBRW02803, CBRW02828, CBRW02844, CBRW02850, CBRW01840, CBRW01884, CBRW01580, CBRW01848, CBRW01865, CBRW02544, 5447395, 5449444, 5449882, 5450368, 5450806, 5453175, 5456247, 5458189, 5445615-1, 5445615-2, 5445657-1, 5445657-2, 5445657-3, 5445657-4, 5445705-1, 5445705-2, 5445723-1, 5445723-2, 5445815-1, 5445815-2, 5445815-3, 5445816-1, 5445816-2, 5447746-1, 5447746-2, 5447746-3, 5447746-4, 5447746-5, 5448449-1, 5448449-2, 5454469-1, 5454469-2, 5457167-1, 5457167-10, 5457167-11, 5457167-12, 5457167-13, 5457167-14, 5457167-15, 5457167-16, 5457167-17, 5457167-18, 5457167-19, 5457167-2, 5457167-3, 5457167-4, 5457167-5, 5457167-6, 5457167-7, 5457167-8, 5457167-9, 5457943-1, 5457943-2, 5457943-3, 5457943-4, CBRW01092, CBRW01129, CBRW01131, CBRW01132, CBRW01133, CBRT00331, CBRW01236, CBRW01981, CBRW02181, CBRW02182, CBRW02184, CBRW02187, CBRW02197, CBRW02198, CBRW02201, CBRW02203, CBRW02206, CBRW02208, 5447306, 5457519, 11017048, 11017065, 11017073, 11017091, 11017097, 11017109, 11017125, 11017186, 11017371, 11017406, 11017419, 11017424, 11017509, 11017510, 11017514, 11017670, 11017725, 11017795, 11017949, 11019263, 11017185-1, 11017185-2, 11017185-3, 11017185-4, 11017185-5, 11017488-1, 11017488-2, 11017488-3, 11017488-4, 11017488-5, BRW00964, CBRW00965, CBRW00966, CBRW00967, CBRW01216, CBRW01224, CBRW01283, CBRW01344, CBRW01345, CBRW01346, CBRW01352, CBRW01353, CBRW01680, CBRW01861, CBRW01863, CBRW01876, CBRW01877, CBRW01887, CBRW01888, CBRW01890, CBRW02010, CBRW02011, CBRW02013, CBRW02101, CBRW02303, CBRW02313, CBRW02316, CBRW02370, CBRW02371, CBRW02373, CBRW02375, CBRW02389, CBRW02395, CBRW02396, CBRW02397, CBRW02398, CBRW02399, CBRW02400, CBRW02401, CBRW02402, CBRW02407, CBRW02408, CBRW02409, CBRW02410, CBRW02411, CBRW02412, CBRW02413, CBRW02414, CBRW02415, CBRW02416, CBRW02417, CBRW02418, CBRW02419, CBRW02426, CBRW02427, CBRW02428, CBRW02430, CBRW02450, CBRW02538, CBRW02542, CBRW02771, CBRW02776, CBRW02794, CBRW02795, CBRW02796, CBRW02797, CBRW02798, CBRW02820, CBRW01548, CBRW01556, CBRW01558, CBRW01137, CBRW01139, CBRW01140, CBRW01145, CBRW01146, CBRW01122, CBRW01123, CBRW01124, CBRW01125, CBRW01126, CBRW01127, CBRW01128, CBRW01130, CBRU02111, CBRW01659, CBRW01391, CBRW01392, CBRW01517, CBRW01518, CBRW01519, CBRW01522, CBRW01523, CBRW01525, CBRW01527, CBRW01528, CBRW01529, CBRW01530, CBRU00192, CBRU03363, CBRW01837, CBRU02907, CBRW01198, CBRW01181, CBRW01182, CBRW01183, CBRU02631, CBRW01332, CBRW01374, CBRW01376, CBRW01377, CBRW01378, CBRW01386, CBRW01387, CBRW01388, CBRW01389, CBRW01390, CBRW01247, CBRW01427, CBRW01428, CBRV01710, CBRW01488, CBRW01489, CBRW01490, CBRW02728, CBRW02759, 6526298, 6526558, 6527101, 6529043, CBRW01205, CBRW01206, CBRW01215, CBRW01222, CBRW01223, CBRW01434, CBRW01435, CBRW01463, CBRW01467, CBRW01477, CBRW01503, CBRW01504, CBRW01616, CBRW01626, CBRW01670, CBRW01681, CBRW01683, CBRW01685, CBRW01691, CBRW01711, CBRW01714, CBRW01716, CBRW01717, CBRW01718, CBRW01719, CBRW01721, CBRW01806, CBRW01815, CBRW01816, CBRW01817, CBRW01818, CBRW01819, CBRW01822, CBRW01828, CBRW01843, CBRW01844, CBRW01845, CBRW01852, CBRW01853, CBRW01864, CBRW01870, CBRW01871, CBRW01872, CBRW01873, CBRW01874, CBRW01875, CBRW01935, CBRW01936, CBRW01939, CBRW01940, CBRW01950, CBRW01960, CBRW01966, CBRW01967, CBRW01968, CBRW01970, CBRW02137, CBRW02153, CBRW02154, CBRW02155, CBRW02156, CBRW02159, CBRW02163, CBRW02164, CBRW02167, CBRW02168, CBRW02169, CBRW02171, CBRW02172, CBRW02173, CBRW02177, CBRW02178, CBRW02188, CBRW02317, CBRW02320, CBRW02321, CBRW02322, CBRW02323, CBRW02324, CBRW02325, CBRW02326, CBRW02327, CBRW02328, CBRW00941, CBRW01444, CBRW01442, CBRW01443, CBRW01001, CBRW01002, CBRW01003, CBRW01004, CBRW01006, CBRW01007, CBRW01226, CBRW01228, CBRW01241, CBRW01230, CBRW01231, CBRW01234, CBRW01235, CBRW01237, CBRW01238, CBRW00981, CBRW00988, CBRW00989, CBRW00990, CBRW00998, CBRW00999, CBRW01117, CBRW01118, CBRW01415, CBRW01416, CBRW01417, CBRW01418, CBRW01060, CBRW01433, CBRW01436, CBRW01438, CBRW01440, CBRW01441, CBRW01395, CBRW01400, CBRW01398, CBRW01399, CBRW01173, CBRW01175, CBRW01176, CBRW01177, CBRW01178, CBRW01179, CBRW01419, CBRW01420, CBRW01421, CBRW01422, CBRW01423, CBRW01424, CBRW01425, CBRW01426, CBRW01199, CBRW01201, CBRW01379, CBRW01380, CBRW01381, CBRW01383, CBRW01384, CBRW01396, CBRW01200, CBRW01204, CBRW01618, CBRW01712, CBRW01736, CBRW01751, CBRW01759, CBRW01763, CBRW01764, CBRW01043, CBRW01050, CBRW01054, CBRW01106, CBRW01195, CBRW01196, CBRW01197, CBRW01208, CBRW01209, CBRW01213, CBRW01214, CBRW01219, CBRW01220, CBRW01225, CBRW01227, CBRW01229, CBRW01233, CBRW01239, CBRW01240, CBRW01847, CBRW01969, CBRW02003, CBRW02004, CBRW02006, CBRW02009, CBRW02012, CBRW02015, CBRW02016, CBRW02022, CBRW02027, CBRW02030, CBRW02031, CBRW02054, CBRW02057, CBRW02058, CBRW02059, CBRW02062, CBRW02111, CBRW02113, CBRW02118, CBRW00909, CBRW00910, CBRW00911, CBRW00912, CBRW00913, CBRW00915, CBRW00916,, CBRW00917, CBRW00920, CBRW00921, CBRW00922, CBRW00923, CBRW00924, CBRW00925,, CBRW00926, CBRW00927, CBRW00928, CBRW00929, CBRW00930, CBRW00931, CBRW00932, CBRW00933, CBRW00945, CBRW00946, CBRW00947, CBRW00948, CBRW00949, CBRW00950, CBRW00951, CBRW00952, CBRW00953, CBRW00954, CBRW00955, CBRW00956, CBRW00957, CBRW00958, CBRW00959, CBRW00960, CBRW01364, CBRW01365, CBRW01366, CBRW01367, CBRW01931, CBRW01445, CBRW01526, CBRW01549, CBRW01552, CBRW01598, CBRW01613, CBRW01620, CBRW01623, CBRW01627, CBRW01660, CBRW01667, CBRW01668, CBRW01245, CBRW01246, CBRW01248, CBRW01249, CBRW01250, CBRW01251, CBRW01252, CBRW01253, CBRW01254, CBRW01255, CBRW01256, CBRW01257, CBRW01258, CBRW01259, CBRW01260, CBRW01261, CBRW01262, CBRW02302, CBRW01987, CBRW01988, CBRW02005, CBRW02019, CBRW01243, CBRW01148, CBRW01150, CBRW01164, CBRW01185, CBRW01186, CBRW01187, CBRW01188, CBRW01189, CBRW01190, CBRW01191, CBRW01192, CBRW01193, CBRW01194, CBRW01550, CBRW01569, CBRW01576, CBRW01579, CBRW01597, CBRW01611, CBRW01612,, CBRW01641, CBRW01643, CBRW00997, CBRW02434, CBRW02435, CBRW02436, CBRW02437, CBRW02438, CBRW02439, CBRW02440, CBRW02441, CBRW02444, 6528877, CBRW01406, CBRW01407, CBRW01408, CBRW01409, CBRW01411, CBRW02232, CBRW02236, CBRW02238, CBRW00971, CBRW00972, CBRW00973, CBRW00974, CBRW00975, CBRW00976, CBRW00977, CBRW00978, CBRW00979, CBRW00980, CBRW01354, CBRW01473, CBRW01474, CBRW01476, CBRW01042, 6526204-1, 6526204-2, 6526204-3, CBRW0174, CBRW01184, CBRT00611 CBRW01554, CBRW01562, CBRW01575, CBRW01591, CBRW01605, CBRW01606, CBRW01607,, CBRW01635, CBRW01553 CBRW01975 CBRW01976 CBRW02527 CBRW01368 CBRW02199 CBRW02200 CBRV02046 CBRW02121 CBRW01098 CBRW01100 CBRW01101 CBRW01103 CBRW01286 CBRW01287 CBRW01487 CBRW01515 CBRW01524 CBRW02191 CBRW02207 CBRW02545 CBRW02558 CBRW01640 CBRW01653 CBRW01654 CBRW01656 CBRW01657 CBRW01664 CBRW01665 CBRW01666 CBRW02429 CBRW01242 CBRW02007 CBRW02008 CBRW01531 CBRW01532 CBRW01533 CBRW01534 CBRW01535 CBRW01536 CBRW01180 CBRW01446 CBRW01449 CBRW01450 CBRW01451 CBRW01452 CBRW01453 CBRW01454 CBRW01455 CBRW01456 CBRW01457 CBRW01458 CBRW01459 CBRW01460 CBRW01461 CBRW01462 CBRW01465 CBRW01466 CBRW01468 CBRW01469 CBRW01470 CBRW01471 CBRW01472 CBRW01475 CBRW01481 CBRW01483 CBRW01484 CBRW01973 CBRW01351 CBRW01491 CBRW01492 CBRW01493 CBRW01494 CBRW01495 CBRW01496 CBRW01497 CBRW01498 CBRW01499 CBRW01500 CBRW01501 CBRW01502 CBRW01513 CBRW01514 CBRW02202 CBRW01595 CBRW01672 CBRW01682 CBRW01686 CBRW01689 CBRW01692 CBRW01694 CBRW01695 CBRW01696 CBRW01697 CBRW01699 CBRW01708 CBRW00962 CBRW00963 CBRW00968 CBRW00969 CBRW00970 CBRW01625 CBRW01633 CBRW01636 CBRW01662 CBRW01671 CBRW01679 CBRW01698 CBRW01700 CBRW01701 CBRW01702 CBRW01703 CBRW01704 CBRW01705 CBRW01706 CBRW01707 CBRW01263 CBRW01264 CBRW01265 CBRW01266 CBRW01267 CBRW01268 CBRU02180 CBRV02754 CBRW00641 5454414 5454413-1 5454413-2 6527173-1 6527173-2 6527173-3 6527173-4 6527173-5 6527173-6 6528692-1 6528692-2 CBRV01531 CBRV01534 CBRV01535 CBRW01951 CBRW01977 CBRW01978 CBRW01985 CBRW01989 CBRW01995 CBRW01996 CBRW01997 CBRW01998 CBRW01999 CBRW02276 CBRW02277 CBRW02278 CBRW02279 CBRW02294 CBRW02295 CBRW02297 CBRW02522 CBRW02532 CBRW02664 CBRW02670 CBRW02673 CBRW02682 CBRW02683 CBRW02685 CBRW02709 CBRW02710 CBRW02715 CBRW02721 CBRW02722 CBRW02723 CBRW02729 CBRW02741 CBRW02742 CBRW02745 CBRW02751 CBRW02756 CBRW02760 CBRW02761 CBRW02767 CBRW02770 CBRW02789 CBRW02811 CBRW02827 CBRW02842 CBRW01540 CBRW01545 CBRW01642 CBRW01622 CBRW01624 CBRW01693 CBRW01728 CBRW01729 CBRW01730 CBRW01775 CBRW01776 CBRW01777 CBRW01778 CBRW01779 CBRW01780 CBRW01984 CBRW01994 CBRW02000 CBRW02001 CBRW02002 CBRW02296 CBRW01762 CBRW02189 CBRW00961 6526481-1 6526481-2 CBRW01813 CBRW02106 CBRW02119 CBRW02607 CBRW02680 CBRW02681 CBRW02724 CBRW02727 CBRW02782 CBRW02847 CBRW02848 CBRW02869 CBRW01429 CBRW01430 CBRW01431 CBRW01432 CBRW02020 CBRW02024 CBRW02029 CBRW02032 CBRW02033 CBRW01337 CBRW02116 CBRW02117 CBRW01980 CBRW02014 CBRW02017 CBRW02025 CBRW02026 CBRW02039 CBRW02040 CBRW01955 CBRW01962 CBRW01963 CBRW01979 CBRW01537 CBRW01538 CBRW01539 CBRW01541 CBRW01542 CBRW01543 CBRW02315 CBRW02319 6529674 CBRW00987 CBRW00992 CBRW00993 CBRW00994 CBRW00995 CBRW00996 CBRW01000 CBRW01005 CBRW01008 CBRW01011 CBRW01012 CBRW01013 CBRW01014 CBRW01015 CBRW01017 CBRW01025 CBRW01397 CBRW01402 CBRW01403 CBRW01404 CBRW01405 CBRW01982 CBRW01986 CBRW01990 CBRW01993 CBRW02160 CBRW02174 CBRW02175 CBRW02176 CBRW02243 CBRW01212 CBRW01221 CBRW01649 CBRW01655 CBRW01675 CBRW01676 CBRW01677 CBRW01678 CBRW01651 CBRW02258 CBRW02259 CBRW02260 CBRW02261 CBRW02262 CBRW02263 CBRW02265 CBRW02266 CBRW01557 CBRW01563 CBRW01577 CBRW01583 CBRW01590 CBRW01600 CBRW01602 CBRW01603 CBRW01610 CBRW01632 CBRW01646 CBRW01647 CBRW01648 CBRW01652 CBRW01673 CBRW01674 CBRW01363 CBRW01369 CBRW01370 CBRW01371 CBRW01372 CBRW01373 CBRW01544 CBRW01555 CBRW01599 CBRW01608 CBRW01609 CBRW02431 CBRW02432 CBRW02443 CBRW02445 CBRW02447 CBRW02547 CBRW02549 CBRW02550 CBRW02551 CBRW02552 CBRW02556 CBRW02557 CBRW02567 CBRW02570 CBRW02571 CBRW01095 CBRW01099 CBRW01288 CBRW01289 CBRW01290 CBRW01291 CBRW01296 CBRW01788 CBRW01803 CBRW01826 CBRW02554 CBRW02555 CBRW02560 CBRW02561 CBRW02644 CBRW02662 CBRW02691 CBRW02692 CBRW02708 CBRW02730 CBRW02731 CBRW02732 CBRW02733 CBRW02734 CBRW02736 CBRW02737 CBRW02738 CBRW02739 CBRW01159 CBRW01760 CBRW01811 CBRW01812 CBRW02604 CBRW01210 CBRW02299 CBRW02300 CBRW01631 CBRW01661 CBRW01687 CBRW01690 CBRW01709 CBRW01722 CBRW01723 CBRW01724 CBRW01725 CBRW01726 CBRW01731 CBRW01747 CBRW01749 CBRW01750 CBRW01752 CBRW01801 CBRW01841 CBRW01854 CBRW02041 CBRW02042 CBRW02043 CBRW00991 CBRW01009 CBRW01010 CBRW01016 CBRW01018 CBRW01026 CBRW01027 CBRW01028 CBRW01029 CBRW01030 CBRW01037 CBRW01038 CBRW01039 CBRW01040 CBRW01041 CBRW01088 CBRW01089 CBRW01090 CBRW01091 CBRW01093 CBRW01094 CBRW01292 CBRW01293 CBRW01294 CBRW01295 CBRW01297 CBRW01298 CBRW01299 CBRW01300 CBRW01301 CBRW01302 CBRW01303 CBRW01304 CBRW01305 CBRW01306 CBRW01307 CBRW01308 CBRW01309 CBRW01310 CBRW01311 CBRW01312 CBRW01313 CBRW01314 CBRW01315 CBRW01316 CBRW01317 CBRW01318 CBRW01319 CBRW01320 CBRW01321 CBRW01322 CBRW01323 CBRW01324 CBRW01325 CBRW01326 CBRW01327 CBRW01329 CBRW01330 CBRW01331 CBRW01333 CBRW01338 CBRW01727 CBRW01733 CBRW01738 CBRW01740 CBRW01769 CBRW01774 CBRW02068 CBRW02112 CBRW02349 CBRW02350 CBRW02355 CBRW02359 CBRW02360 CBRW02361 CBRW02362 CBRW02363 CBRW02364 CBRW02365 CBRW02366 CBRW02367 CBRW02368 CBRW02369 CBRW02372 CBRW02376 CBRW02377 CBRW02386 CBRW02403 CBRW02404 CBRW02446 CBRW02451 CBRW02452 CBRW02453 CBRW02454 CBRW02455 CBRW02456 CBRW02457 CBRW02458 CBRW02459 CBRW02460 CBRW02461 CBRW02462 CBRW02463 CBRW02464 CBRW02465 CBRW02466 CBRW02467 CBRW02468 CBRW02480 CBRW02481 CBRW02512 CBRW02513 CBRW02624 CBRW02628 CBRW02630 CBRW02643 CBRW02645 CBRW02649 CBRW02650 CBRW02651 CBRW02669 CBRW02671 CBRW02674 CBRW02675 CBRW02701 CBRW02703 CBRW02704 CBRW02705 CBRW02706 CBRW02712 CBRW02717 CBRW02718 CBRW02719 CBRW02720 CBRW02725 CBRW02726 CBRW02740 CBRW02743 CBRW02744 CBRW02765 CBRW02766 CBRW02768 CBRW02769 CBRW02775 CBRW02777 CBRW02778 CBRW02779 CBRW02780 CBRW02781 CBRW02783 CBRW02784 CBRW02785 CBRW02790 CBRW02791 CBRW02792 CBRW02793 CBRW02800 CBRW02801 CBRW02802 CBRW02804 CBRW02806 CBRW02809 CBRW02812 CBRW02813 CBRW02814 CBRW02815 CBRW02816 CBRW02817 CBRW02818 CBRW02819 CBRW02821 CBRW02822 CBRW02823 CBRW02824, CBRW02830, CBRW02831, CBRW02832, CBRW02833, CBRW02835, CBRW02836, CBRW02837, CBRW02838, CBRW02839, CBRW02843, CBRW02849, CBRW02851, CBRW02852, CBRW02853, CBRW02854, CBRW02855, CBRW02856, CBRW02857, CBRW02858, CBRW02859, CBRW02860, CBRW02861, CBRW02862, CBRW02868, CBRW02870, CBRW02871, CBRW02872, CBRW02874, CBRW02878, CBRW02880, CBRW02881, CBRW02882, CBRW02883, CBRW02884, CBRW02885, CBRW02886, CBRW02887, CBRW02888, CBRW02889, CBRW02890, CBRW02891, CBRW02892, CBRW02893, CBRW02895, CBRW02896, CBRW02902, CBRW02903, CBRW02907, CBRW02909, CBRW03005, CBRW03006, CBRW03007, CBRW03013, CBRW03014, CBRW01478, CBRW01479, CBRW01684, CBRW01737, CBRW01739, CBRW01142, CBRW01143, CBRW01144, CBRW01334, CBRW01335, CBRW01336, CBRW01339, CBRW01340, CBRW01341, CBRW01342, CBRW01343, CBRW01348, CBRW01560, CBRW01561, CBRW01581, CBRW01584, CBRW01592, CBRW01953, CBRW02318, CBRW01096, CBRW01097, CBRW01102, CBRW01104, CBRW01105, CBRW01107, CBRW01109, CBRW01111, CBRW01113, CBRW01114, CBRW01115, CBRW01116, CBRW01134, CBRW01135, CBRW01153, CBRW01154, CBRW01160, CBRW01161, CBRW01162, CBRW01163, CBRW01165, CBRW01166, CBRW01167, CBRW01168, CBRW01169, CBRW01170, CBRW01171, CBRW01172, CBRW01174, CBRW01349, CBRW01350, CBRW01355, CBRW01356, CBRW01357, CBRW01358, CBRW01359, CBRW01360, CBRW01361, CBRW01362, CBRW01382, CBRW01385, CBRW01393, CBRW01394, CBRW01480, CBRW01482, CBRW01485, CBRW01486, CBRW01505, CBRW01506, CBRW01507, CBRW01508, CBRW01509, CBRW01510, CBRW01511, CBRW01512, CBRW01516, CBRW01520, CBRW01521, CBRW02132, CBRW02133, CBRW02134, CBRW02135, CBRW02136, CBRW02138, CBRW02139, CBRW02140, CBRW02141, CBRW02142, CBRW02148, CBRW02149, CBRW02150, CBRW02151, CBRW02152, CBRW02157, CBRW02158, CBRW02161, CBRW02162, CBRW02165, CBRW02166, CBRW02179, CBRW02180, CBRW02183, CBRW02185, CBRW02186, CBRW02190, CBRW02192, CBRW02193, CBRW02194, CBRW02195, CBRW02196, CBRW02204, CBRW02205, CBRW02209, CBRW02210, CBRW02211, CBRW02212, CBRW02213, CBRW02214, CBRW02215, CBRW02217, CBRW02218, CBRW02219, CBRW02220, CBRW02221, CBRW02222, CBRW02223, CBRW02224, CBRW02225, CBRW02226, CBRW02227, CBRW02228, CBRW02229, CBRW02230 CBRW02231, CBRW02233, CBRW02234, CBRW02235, CBRW02237, CBRW02239, CBRW02240, CBRW02241, CBRW02242, CBRW02244, CBRW02245, CBRW02246, CBRW02247, CBRW02248, CBRW02250, CBRW02251, CBRW02252, CBRW02253, CBRW02254, CBRW02255, CBRW02256, CBRW02257, CBRW02264, CBRW02267, CBRW02268, CBRW02269, CBRW02270, CBRW02271, CBRW02272, CBRW02273, CBRW02274, CBRW02275, CBRW02280, CBRW02281, CBRW02282, CBRW02283, CBRW02284, CBRW02286, CBRW02288, CBRW02289, CBRW02290, CBRW02291, CBRW02292, CBRW02293, CBRW02405, CBRW02406, CBRW02569, CBRW02663, CBRW02665, CBRW02666, CBRW02667, Recalling Firm/ Manufacturer Datex-Ohmeda, Inc. 3030 Ohmeda Dr Madison WI 53718-6704 For Additional Information Contact Service 800-437-1171 Manufacturer Reason for Recall Potential for the display processor to experience an******unexpected failure. FDA Determined Cause 2 Software change control Action On December 12, 2017 GE Healthcare sent letters to their consignees stating that There is a potential for the display processor to experience an unexpected failure. If this failure occurs, the device screen will temporarily go blank and will then display an Error Code ***EC01 XX***. The blank display results in loss of monitored data being displayed to the user. There have been no injuries reported as a result of this issue. You may continue to use the system. Device performance is not affected on devices where no such display failure occurs.******In the event this issue occurs, an audible alarm will sound. Mechanical ventilation will continue at previous settings. A device restart is required to clear the alarm condition. Consider using appropriate alternative means of ventilation such as a self-inflating bag or another ventilator to prevent patient injury while restarting the device.************GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction.************If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. Quantity in Commerce 349 (US) + 2,051 (OUS) Distribution Worldwide Distribution - USA (nationwide) Distribution to the states of : AL, AZ, CA, GA, ID, IL, IN, MS, NC, NJ, NK, NV, OH, OK, RI, TX, and WI., and to the countries of : Albania, Argentina, Australia, Bahrain, Bangladesh, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Ecuador, El Salvador, Estonia, Ethiopia, Egypt, Finland, France, Germany, Ghana, Hungary, India, Indonesia, Italy, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Malaysia, Mexico, Morocco, Myanmar, Nepal, New Zealand, Nigeria, Norway, Pakistan, Palestine, Peru, Philippines, Poland, Russia, Saudi Arabia, Senegal, Singapore, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Vietnam Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = CBK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On December 12, 2017 GE Healthcare sent letters to their consignees stating that There is a potential for the display processor to experience an unexpected failure. If this failure occurs, the device screen will temporarily go blank and will then display an Error Code ***EC01 XX***. The blank display results in loss of monitored data being displayed to the user. There have been no injuries reported as a result of this issue. You may continue to use the system. Device performance is not affected on devices where no such display failure occurs.******In the event this issue occurs, an audible alarm will sound. Mechanical ventilation will continue at previous settings. A device restart is required to clear the alarm condition. Consider using appropriate alternative means of ventilation such as a self-inflating bag or another ventilator to prevent patient injury while restarting the device.************GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction.************If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. | SoftwareChangeControl
71 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164900 | Z-2109-2018 | 12C (AdaPTinsight)***Affected component: AdaPTinsight software******Product Usage:***12C is used with a charged particle or photon radiation therapy system for localization of the patient position with respect to the therapy equipment and to provide correction feedback to the radiation therapy system.*** | AdaPTinsight | 2 | CDRH | 2018-06-07 |  | 2019-08-09 10:37 | Ion Beam Applications S.A. | IBA is initiating this recall to address an issue identified with AdaPTinsight software and to reduce risk related to this problem. | Class 2 Device Recall AdaPTinsight Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall AdaPTinsight FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall AdaPTinsight Date Initiated by Firm October 28, 2016 Create Date June 07, 2018 Recall Status 1 Terminated 3 on August 09, 2019 Recall Number Z-2109-2018 Recall Event ID 80219 510(K)Number K132847 Product Classification System, radiation therapy, charged-particle, medical - Product Code LHN Product 12C (AdaPTinsight)******Affected component: AdaPTinsight software************Product Usage:******12C is used with a charged particle or photon radiation therapy system for localization of the patient position with respect to the therapy equipment and to provide correction feedback to the radiation therapy system. Code Information Affected units: SAT.123, SAT.126, PAT.115, SAT.122 and SBF.101 Recalling Firm/ Manufacturer Ion Beam Applications S.A. Chemin du Cyclotron, 3 Louvain La Neuve Belgium Manufacturer Reason for Recall IBA is initiating this recall to address an issue identified with AdaPTinsight software and to reduce risk related to this problem. FDA Determined Cause 2 Software change control Action IBA mailed an Urgent Field Safety Notice to affected customers on 10/28/2016 to inform them of the issue. The issue will be solved in AdaPTinsights versions starting from Al 2.0. The solution was to be deployed on impacted IBA Proton Therapy sites by September 2017. Quantity in Commerce 5 units (12C) Distribution Worldwide Distribution - US Nationwide in the states of LA, TX, Italy, Sweden and France. Each country's National Competent Authorities were notified of the firm's Field Safety Notice. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LHN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action IBA mailed an Urgent Field Safety Notice to affected customers on 10/28/2016 to inform them of the issue. The issue will be solved in AdaPTinsights versions starting from Al 2.0. The solution was to be deployed on impacted IBA Proton Therapy sites by September 2017. | SoftwareChangeControl
72 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=162647 | Z-1170-2018 | Proclaim DRG Implantable Pulse Generator, Model Number 3664 | Proclaim DRG Implantable Pulse Generator | 2 | CDRH | 2018-03-23 | 2018-03-23 21:03 | 2020-11-03 09:50 | St. Jude Medical, Inc. | The firm received complaints of error messages that occurred during routine impedance checks on Proclaim DRG IPGs, model 3664. Some complaints were also associated with transient over stimulation which created discomfort for the patients. | Class 2 Device Recall Proclaim DRG Implantable Pulse Generator Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Proclaim DRG Implantable Pulse Generator FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Proclaim DRG Implantable Pulse Generator Date Initiated by Firm March 09, 2018 Date Posted March 23, 2018 Recall Status 1 Terminated 3 on November 03, 2020 Recall Number Z-1170-2018 Recall Event ID 79560 PMA Number P150004S002 Product Classification Implantable pulse generator, programmer and battery charger - Product Code PMP Product Proclaim DRG Implantable Pulse Generator, Model Number 3664 Code Information Lot Number 6210849 Serial Number AVM174.1 Lot Number 6210852 Serial Numbers: AVA560.1, AVA214.1, AVL027.1 Lot Number 6499208 Serial Numbers: AVD426.1, AVA922.1 Lot Number 6263128 Serial Numbers: AVB119.1, AVD400.1, AVA622.1, AVD431.1, AVE633.1, AVK678.1, AVN837.1 Lot Number 6268723 Serial Numbers: AUV770.1, AVW805.1 Lot Number 6278155 Serial Number AVK896.1 Recalling Firm/ Manufacturer St. Jude Medical, Inc. 6901 Preston Rd Plano TX 75024-2508 For Additional Information Contact Justin Paquette 651-756-6293 Manufacturer Reason for Recall The firm received complaints of error messages that occurred during routine impedance checks on Proclaim DRG IPGs, model 3664. Some complaints were also associated with transient over stimulation which created discomfort for the patients. FDA Determined Cause 2 Software change control Action Notification material in the form of a Customer Letter was hand-delivered to two******consignees in the US beginning 09 Mar 2018. The Physician Letter was hand-delivered to impacted implanting physicians, beginning 14 Mar 2018. ******A Patient Letter was delivered via certified mailing to all active implanted patients in the US, beginning 14 Mar 2018. Quantity in Commerce 17 units Distribution US Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = PMP - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Notification material in the form of a Customer Letter was hand-delivered to two******consignees in the US beginning 09 Mar 2018. The Physician Letter was hand-delivered to impacted implanting physicians, beginning 14 Mar 2018. ******A Patient Letter was delivered via certified mailing to all active implanted patients in the US, beginning 14 Mar 2018. | SoftwareChangeControl
73 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153395 | Z-1816-2017 | TotalCare Bed Systems, Model Numbers P1900, PR1900, and U1900 with 3-Level Patient Position Monitoring and Scale Upgrade Kits, Model Numbers P1947A and P1948A). | TotalCare | 2 | CDRH | 2017-04-12 |  | 2017-09-19 | Hill-Rom, Inc. | The bed exit alarm may not function properly if the alarm has been armed for a period longer than seventy-two (72) hours without the alarm being triggered by patient movement or removal of patient weight within that time period. | Class 2 Device Recall TotalCare Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TotalCare FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TotalCare Date Initiated by Firm February 23, 2017 Create Date April 12, 2017 Recall Status 1 Terminated 3 on September 19, 2017 Recall Number Z-1816-2017 Recall Event ID 76506 510(K)Number K962942 Product Classification Bed, ac-powered adjustable hospital - Product Code FNL Product TotalCare Bed Systems, Model Numbers P1900, PR1900, and U1900 with 3-Level Patient Position Monitoring and Scale Upgrade Kits, Model Numbers P1947A and P1948A). Code Information P1900 SN: K328AM6065 to Q245AM0691 PR1900 SN: L082AL7316 to R343AL4516 U1900 SN: L257AM0541 to O311AM7169 Scale Upgrade Kit P1957A produced after March 13, 2013 (not serialized). Recalling Firm/ Manufacturer Hill-Rom, Inc. 1069 State Road 46 E Batesville IN 47006-7520 For Additional Information Contact Hill-Rom Technical Support 800-445-3720 Manufacturer Reason for Recall The bed exit alarm may not function properly if the alarm has been armed for a period longer than seventy-two (72) hours without the alarm being triggered by patient movement or removal of patient weight within that time period. FDA Determined Cause 2 Software change control Action On 2/23/2017, ***URGENT: MEDICAL DEVICE CORRECTION*** notifications were sent to the affected customers via courier. ******The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. ************Customers without a service contract received the following instructions:******Action to be taken by the user: ******A software update for the TotalCare bed***s GCI has been released to resolve this issue. Enclosed with this letter are a USB stick with the updated software, instructions on how to update the GCI software, and a response form to be returned within 30 days. ******1) Identify any affected TotalCare beds with 3-Level Patient Position Monitoring in your facility by using the criteria above. ******2) Please use the enclosed instructions and update the software on each potentially affected TotalCare bed immediately. ******3) Record the serial numbers of the beds being updated on the response form and return to Hill-Rom within 30 days of receipt of this letter. ************Additional USB sticks with the updated software may be obtained by contacting Technical Support at 1-800-445-3720 (select Option 2, then Option 2 again). ******Until each of the affected TotalCare beds are updated with the new software, we recommend you do one or more of the below to avoid any potential occurrences of bed exit not functioning. ************1. Power cycle (unplug and plug in the bed mains power cord) once every 48 hours. ******2. Remove the Patient and re-zero the scale once every 48 hours. ******3. Turn off and do not use the Bed Exit Alarm functionality. ************Continued use of the bed exit alarm without the mitigations listed above may cause risk to patient. ************If you have an upgrade kit (P1957A) that has not been installed, please contact Hill-Rom for return and replacement. ************Customers with a service contract received the following instructions:******Action to be taken by the user: ******A Hill-Rom representative per Quantity in Commerce 3037 Distribution US Nationwide. ******AUSTRALIA******BANGLADESH******BRAZIL******CANADA******CHINA******GERMANY******INDIA******ITALY******KUWAIT******PAKISTAN******SAUDI ARABIA******THAILAND******UNITED ARAB EMIRATES******CYPRUS******GREECE******NETHERLANDS Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = FNL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 2/23/2017, ***URGENT: MEDICAL DEVICE CORRECTION*** notifications were sent to the affected customers via courier. ******The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. ************Customers without a service contract received the following instructions:******Action to be taken by the user: ******A software update for the TotalCare bed***s GCI has been released to resolve this issue. Enclosed with this letter are a USB stick with the updated software, instructions on how to update the GCI software, and a response form to be returned within 30 days. ******1) Identify any affected TotalCare beds with 3-Level Patient Position Monitoring in your facility by using the criteria above. ******2) Please use the enclosed instructions and update the software on each potentially affected TotalCare bed immediately. ******3) Record the serial numbers of the beds being updated on the response form and return to Hill-Rom within 30 days of receipt of this letter. ************Additional USB sticks with the updated software may be obtained by contacting Technical Support at 1-800-445-3720 (select Option 2, then Option 2 again). ******Until each of the affected TotalCare beds are updated with the new software, we recommend you do one or more of the below to avoid any potential occurrences of bed exit not functioning. ************1. Power cycle (unplug and plug in the bed mains power cord) once every 48 hours. ******2. Remove the Patient and re-zero the scale once every 48 hours. ******3. Turn off and do not use the Bed Exit Alarm functionality. ************Continued use of the bed exit alarm without the mitigations listed above may cause risk to patient. ************If you have an upgrade kit (P1957A) that has not been installed, please contact Hill-Rom for return and replacement. ************Customers with a service contract received the following instructions:******Action to be taken by the user: ******A Hill-Rom representative per | SoftwareChangeControl
74 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153003 | Z-1494-2017 | Final Fit Software Version 1.11 and 1.12;***PC Based software installed outside Nidek EC-5000 Excimer Laser System.*********Ophthalmic laser system use for correction of corneal refraction and ablation of the corneal surface. The system is composed of a laser generator which produces an excimer laser radiation of wavelength 193nm, a beam delivery unit, an optical system for observation, a gas system and a computer for system control. | Nidek EC-5000 Excimer Laser System, Final Fit Software Version 1.11 and 1.12 | 2 | CDRH | 2017-03-20 |  | 2017-12-05 | Nidek Inc | During treatment planning, the procedure was programmed with an unintended (wrong) correction. | Class 2 Device Recall Nidek EC5000 Excimer Laser System, Final Fit Software Version 1.11 and 1.12 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Nidek EC5000 Excimer Laser System, Final Fit Software Version 1.11 and 1.12 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Nidek EC5000 Excimer Laser System, Final Fit Software Version 1.11 and 1.12 Date Initiated by Firm February 01, 2017 Create Date March 20, 2017 Recall Status 1 Terminated 3 on December 05, 2017 Recall Number Z-1494-2017 Recall Event ID 76376 PMA Number P970053 Product Classification Excimer laser system - Product Code LZS Product Final Fit Software Version 1.11 and 1.12;******PC Based software installed outside Nidek EC-5000 Excimer Laser System.******************Ophthalmic laser system use for correction of corneal refraction and ablation of the corneal surface. The system is composed of a laser generator which produces an excimer laser radiation of wavelength 193nm, a beam delivery unit, an optical system for observation, a gas system and a computer for system control. Code Information Software version 1.11 and 1.12. Serial numbers of Affected devices: MNE0009; NVE0002; OH00295; NC00185; CAE0016; COE0008; MI00131; GA00241; CAE0007; VAE0004; MI00131; CAE0031; AZ00167; COE0011; CA00561; PAE0003; TN00168; MNE0009; WA00246; CAE0047; OHE0005; WAE0008; MI00244; NY00473; CAE0007; WAE0008; NVE0002; WA00233; CAE0016. Recalling Firm/ Manufacturer Nidek Inc 47651 Westinghouse Dr Fremont CA 94539-7474 For Additional Information Contact Preeti Gandhi 510-353-7718 Manufacturer Reason for Recall During treatment planning, the procedure was programmed with an unintended (wrong) correction. FDA Determined Cause 2 Software change control Action Nidek sent a Recall notification letter dated February 1, 2017 to all affected consignees by certified mail method beginning February 2, 2017. ******Letters advised of the issue received from a complaint and the steps the user took that created the issue. To prevent the issue, the letters provide steps that should be taken. Nidek's Customer Support team will contact customers by phone or email. Customers with questions were instructed to contact Customer Service at 1-800-223-9044 or email ff_support@nidek.com. Quantity in Commerce 29 Distribution Nationwide Distribution to MN, NV, OH, NC, CA, CO, CA, GA, VA, MI, AZ, PA, TN, WA, TX, NY. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZS - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Nidek sent a Recall notification letter dated February 1, 2017 to all affected consignees by certified mail method beginning February 2, 2017. ******Letters advised of the issue received from a complaint and the steps the user took that created the issue. To prevent the issue, the letters provide steps that should be taken. Nidek's Customer Support team will contact customers by phone or email. Customers with questions were instructed to contact Customer Service at 1-800-223-9044 or email ff_support@nidek.com. | SoftwareChangeControl
75 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152488 | Z-1045-2017 | Ambra PACS******UDI: +AMBRHEALTHSOLUTIONS0/$$+7_V3.16.13.0R***Software Version Number: V3.16.13.0******Intended for use as a primary diagnostic and analysis tool for diagnostic images. | Ambra PACS | 2 | CDRH | 2017-01-18 | 2017-01-18 21:00 | 2017-04-07 | DICOM GRID INC | A software error caused the window/level to become the same in one series regardless if the image had different levels; image results have a washed-out grey appearance. | Class 2 Device Recall Ambra PACS Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Ambra PACS FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Ambra PACS Date Initiated by Firm December 16, 2016 Date Posted January 18, 2017 Recall Status 1 Terminated 3 on April 07, 2017 Recall Number Z-1045-2017 Recall Event ID 76187 510(K)Number K152977 Product Classification System, image processing, radiological - Product Code LLZ Product Ambra PACS************UDI: +AMBRHEALTHSOLUTIONS0/$$+7_V3.16.13.0R******Software Version Number: V3.16.13.0************Intended for use as a primary diagnostic and analysis tool for diagnostic images. Code Information N/A Recalling Firm/ Manufacturer DICOM GRID INC 4530 E Muirwood Dr Ste. 113 Phoenix AZ 85048-7639 For Additional Information Contact 602-538-0061 Manufacturer Reason for Recall A software error caused the window/level to become the same in one series regardless if the image had different levels; image results have a washed-out grey appearance. FDA Determined Cause 2 Software change control Action On 12/16/16 a fix is the software issue was conducted via remote cloud system. On 01/11/17 customers were sent Release Notes that addressed the software correction. The issue was resolved and required no further action. Quantity in Commerce 209 Distribution U.S. distribution to the following; TX, FL, NE, CA, WI, AZ, NH, MO, MA, CO, PA, NV, PR, IL, OH, NY, MD, NJ, MN, MI, CT, VA, GA, UT, KS, WA, MI, HI.************Foreign distribution to the following; Canada, Spain, Israel, Singapore, Kosovo, Austria, Brazil, Malaysia, Australia. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 12/16/16 a fix is the software issue was conducted via remote cloud system. On 01/11/17 customers were sent Release Notes that addressed the software correction. The issue was resolved and required no further action. | SoftwareChangeControl
76 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150430 | Z-0700-2017 | Model 37751 Recharger******Product Usage:***The Medtronic 37751 Recharger is a hand-held device used to charge the battery in a patient s neurostimulator. It includes a display to provide information on the charging system. The charging system consists of the Model 37751 Recharger, an AC power supply and power cord (Model 37761) and an antenna.*** | Recharger | 2 | CDRH | 2016-12-03 |  | 2022-05-04 09:36 | Medtronic Neuromodulation | Medtronic has identified an increased number of complaints from customers involving reports of Rechargers (Medtronic Model 37751 Recharger, which is included in Models 37754 and 97754 Charging Systems,Spinal Cord Stimulation)that are in an unresponsive error state, where the Recharger is non-functional with a blank display screen and is beeping every 5 seconds. Medtronic has determined all Rechargers manufactured starting in November 2014 (indicated by serial numbers beginning with NKA4 or NKU4 ) are more susceptible to this error state. | Class 2 Device Recall Recharger Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Recharger FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Recharger Date Initiated by Firm October 03, 2016 Create Date December 03, 2016 Recall Status 1 Terminated 3 on May 04, 2022 Recall Number Z-0700-2017 Recall Event ID 75458 PMA Number P960009S051 Product Classification Stimulator, electrical, implanted, for parkinsonian tremor - Product Code MHY Product Model 37751 Recharger************Product Usage:******The Medtronic 37751 Recharger is a hand-held device used to charge the battery in a patient s neurostimulator. It includes a display to provide information on the charging system. The charging system consists of the Model 37751 Recharger, an AC power supply and power cord (Model 37761) and an antenna. Code Information serial numbers beginning with ***NKA4*** or ***NKU4 Recalling Firm/ Manufacturer Medtronic Neuromodulation 7000 Central Ave Minneapolis MN 55421-1241 For Additional Information Contact Medtronic Technical Services 800-707-0933 Manufacturer Reason for Recall Medtronic has identified an increased number of complaints from customers involving reports of Rechargers (Medtronic Model 37751 Recharger, which is included in Models 37754 and 97754 Charging Systems,Spinal Cord Stimulation)that are in an unresponsive error state, where the Recharger is non-functional with a blank display screen and is beeping every 5 seconds. Medtronic has determined all Rechargers manufactured starting in November 2014 (indicated by serial numbers beginning with ***NKA4*** or ***NKU4***) are more susceptible to this error state. FDA Determined Cause 2 Software change control Action Consignees were sent a series of documents which included an introductory Medtronic and "Urgent Medical Device Correction" letters dated September 2016. The urgent letters included one for Deep Brain Stimulation (DBS) and one for Spinal Cord Stimulation(SCS) . Each letter was addressed to Healthcare Professional. Only physicians associated with both SCS and DBS therapy received the cover letter. ******The letters provided information on "Background, Issues Mitigation, Recommendation, Additional Information and Patient Notification (DBS).******Requested consignees to complete and return the Physician /HCP Reply Form.******For questions contact Medtronic Technical Services at 1-800-707-0933 weekdays 7 am to 6pm, CT. Quantity in Commerce 42,887 Distribution Worldwide Distribution - US Nationwide including DC************OUS: Canada and others. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = MHY - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Consignees were sent a series of documents which included an introductory Medtronic and "Urgent Medical Device Correction" letters dated September 2016. The urgent letters included one for Deep Brain Stimulation (DBS) and one for Spinal Cord Stimulation(SCS) . Each letter was addressed to Healthcare Professional. Only physicians associated with both SCS and DBS therapy received the cover letter. ******The letters provided information on "Background, Issues Mitigation, Recommendation, Additional Information and Patient Notification (DBS).******Requested consignees to complete and return the Physician /HCP Reply Form.******For questions contact Medtronic Technical Services at 1-800-707-0933 weekdays 7 am to 6pm, CT. | SoftwareChangeControl
77 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148728 | Z-2711-2016 | SPECULAR MICROSCOPE CEM-530; Software version 1.08 and 1.09.******Opthalmic: The NIDEK Specular Microscope CEM-530 provides non-contact, high magnification image capture of endothelium enabling observation of the size and shape of cells. Information such as the number of endothelial cells, cell density, and cell area is analyzed through the captured images. The captured***images and analysis results of endothelium are used for intraocular or corneal surgery, postoperative follow-up and corneal observation such as for endothelial disorders or the corneal state of patients who wear extended-wear contact lenses.*** | SPECULAR MICROSCOPE CEM-530 | 2 | CDRH | 2016-09-01 | 2016-09-01 21:00 | 2017-01-09 | Nidek Inc | Software version 1.08 and 1.09 for the Specular Microscope CEM 530 included a change of analysis results feature that was not reviewed and approved by the FDA. | Class 2 Device Recall SPECULAR MICROSCOPE CEM530 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SPECULAR MICROSCOPE CEM530 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SPECULAR MICROSCOPE CEM530 Date Initiated by Firm August 08, 2016 Date Posted September 01, 2016 Recall Status 1 Terminated 3 on January 09, 2017 Recall Number Z-2711-2016 Recall Event ID 74900 510(K)Number K151706 Product Classification Microscope, specular - Product Code NQE Product SPECULAR MICROSCOPE CEM-530; Software version 1.08 and 1.09.************Opthalmic: The NIDEK Specular Microscope CEM-530 provides non-contact, high magnification image capture of endothelium enabling observation of the size and shape of cells. Information such as the number of endothelial cells, cell density, and cell area is analyzed through the captured images. The captured******images and analysis results of endothelium are used for intraocular or corneal surgery, postoperative follow-up and corneal observation such as for endothelial disorders or the corneal state of patients who wear extended-wear contact lenses. Code Information Microscopes with serial numbers: 120004,120005, 120006, 120007, 120009, 120010, 120011,120012, 120013, 120014, 0120016, 120019, 120021, 120023, 120024, 120026, 120030, 120036, 0120038, 120039, 120040, 120042, 120044,120045, 120046, 120047, 120048, 120049, 120050. Recalling Firm/ Manufacturer Nidek Inc 47651 Westinghouse Dr Fremont CA 94539-7474 For Additional Information Contact Preeti Gandhi Bhatia 510-353-7718 Manufacturer Reason for Recall Software version 1.08 and 1.09 for the Specular Microscope CEM 530 included a change of analysis results feature that was not reviewed and approved by the FDA. FDA Determined Cause 2 Software change control Action The firm, Nidek, sent a Dear Doctor letters entitled "RECALL NOTIFICATION" dated August 15, 2016 to affected sites by Federal Express. The letter informs users of the issue and stated that Nidek field service representatives will visit to install updated software.************If you have any questions, please call Customer Service at 1-800-722-0219; Monday-Friday 8am to 5pm PST or email: cem530update@noritsuservice.com. Quantity in Commerce 29 affected devices Distribution US Distribution to states of: TX, KS, FL, NV, CA, PA, KY, OR, TN, AL, VA, MO, OK, CT, NY. LA, NM, and IN. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = NQE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, Nidek, sent a Dear Doctor letters entitled "RECALL NOTIFICATION" dated August 15, 2016 to affected sites by Federal Express. The letter informs users of the issue and stated that Nidek field service representatives will visit to install updated software.************If you have any questions, please call Customer Service at 1-800-722-0219; Monday-Friday 8am to 5pm PST or email: cem530update@noritsuservice.com. | SoftwareChangeControl
78 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144871 | Z-2370-2016 | Brilliance CT 16 Power Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Brilliance CT 16 Power Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host. | Class 2 Device Recall Brilliance CT 16 Power Computed Tomography Xray systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Brilliance CT 16 Power Computed Tomography Xray systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Brilliance CT 16 Power Computed Tomography Xray systems Date Initiated by Firm March 24, 2016 Create Date August 09, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2370-2016 Recall Event ID 73750 510(K)Number K012009 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Brilliance CT 16 Power Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Model #: 728240; Serial #s: 3100, 3236, 5043, 5047, 5280, 5345, 5390, 5395, 5413, 5430, 5476, 5495, 5499, 5500, 5512, 6007, 6009, 6013, 6032, 6041, 6045, 6049, 6051, 6058, 6062, 6064, 6066, 6067, 6070, 6071, 6077, 6080, 6081, 6084, 6089, 6095, 6098, 6100, 6106, 6109, 6113, 6115, 6116, 6119, 6120, 6123, 6124, 6131, 6132, 6134, 6141, 6144, 6155, 6161, 6162. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall Software defect causing intermittently slow response of Host. FDA Determined Cause 2 Software change control Action On 3/24/2016 the firm mailed Customer Information letters to their customers. Quantity in Commerce 55 units Distribution Product was sent to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV & WY.************Product was also sent to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Cuba, Denmark, Djibouti, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, R¿union, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 3/24/2016 the firm mailed Customer Information letters to their customers. | SoftwareChangeControl
79 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144868 | Z-2367-2016 | Brilliance BigBore Oncology Computed Tomography X-ray systems, intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Brilliance BigBore Oncology Computed Tomography X-ray systems, | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host. | Class 2 Device Recall Brilliance BigBore Oncology Computed Tomography Xray systems, Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Brilliance BigBore Oncology Computed Tomography Xray systems, FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Brilliance BigBore Oncology Computed Tomography Xray systems, Date Initiated by Firm March 24, 2016 Create Date August 09, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2367-2016 Recall Event ID 73750 510(K)Number K033357 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Brilliance BigBore Oncology Computed Tomography X-ray systems, intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Model #: 728243; Serial #s: 7063, 7116, 7169, 7205, 7309, 7314, 7354, 7611. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall Software defect causing intermittently slow response of Host. FDA Determined Cause 2 Software change control Action On 3/24/2016 the firm mailed Customer Information letters to their customers. Quantity in Commerce 8 units Distribution Product was sent to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV & WY.************Product was also sent to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Cuba, Denmark, Djibouti, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, R¿union, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 3/24/2016 the firm mailed Customer Information letters to their customers. | SoftwareChangeControl
80 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144870 | Z-2369-2016 | Brilliance 16 Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Brilliance 16 Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host. | Class 2 Device Recall Brilliance 16 Computed Tomography Xray systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Brilliance 16 Computed Tomography Xray systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Brilliance 16 Computed Tomography Xray systems Date Initiated by Firm March 24, 2016 Create Date August 09, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2369-2016 Recall Event ID 73750 510(K)Number K012009 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Brilliance 16 Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Model #: 728246; Serial #s: 3010, 3017, 3031, 3040, 3043, 3070, 3072, 3085, 3090, 3117, 3123, 3149, 3179, 3216, 3247, 3293, 3314, 3324, 3411, 3465, 3490, 5006, 5009, 5012, 5018, 5020, 5024, 5026, 5028, 5030, 5031, 5035, 5044, 5045, 5048, 5052, 5053, 5054, 5055, 5058, 5059, 5060, 5064, 5069, 5070, 5075, 5077, 5081, 5086, 5088, 5091, 5093, 5097, 5099, 5101, 5103, 5107, 5108, 5110, 5111, 5113, 5115, 5116, 5118, 5120, 5123, 5126, 5127, 5128, 5131, 5135, 5138, 5140, 5147, 5155, 5160, 5161, 5163, 5166, 5181, 5184, 5186, 5188, 5190, 5191, 5192, 5193, 5195, 5197, 5208, 5209, 5216, 5219, 5224, 5231, 5235, 5236, 5237, 5270, 5272, 5279, 5282, 5285, 5287, 5288, 5294, 5295, 5315, 5316, 5347, 5348, 5350, 5354, 5357, 5360, 5368, 5394, 5405, 5412, 5433, 5434, 5443, 5453, 5457, 5474, 5480, 5487, 5501, 5502, 5505, 5516, 5519, 5523, 5526, 5554, 5572, 5573, 5576, 5609, 5615, 5633, 5647, 5653, 5674, 5683, 5703, 5723, 5758, 5762, 5764, 5771, 5811, 5897, 5929, 5937, 5956, 5974, 5995, 6021, 6022, 6031, 6033, 6048, 6053, 6072, 6154, 6220, 6229, 6243, 6280, 6283, 6298, 6299, 6309, 6321, 6360, 6375, 6405, 6426, 6441, 6465, 6474, 6484, 6492, 6503, 6511, 6517, 6519, 6523, 6549, 6555, 6559, 30011, 30049, 30110, 32287, 32366, 50011, 50053, 50087, 50180, 50181, 5238A. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall Software defect causing intermittently slow response of Host. FDA Determined Cause 2 Software change control Action On 3/24/2016 the firm mailed Customer Information letters to their customers. Quantity in Commerce 203 units Distribution Product was sent to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV & WY.************Product was also sent to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Cuba, Denmark, Djibouti, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, R¿union, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 3/24/2016 the firm mailed Customer Information letters to their customers. | SoftwareChangeControl
81 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144867 | Z-2366-2016 | Brilliance 64 Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Brilliance 64 Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host. | Class 2 Device Recall Brilliance 64 Computed Tomography Xray systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Brilliance 64 Computed Tomography Xray systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Brilliance 64 Computed Tomography Xray systems Date Initiated by Firm March 24, 2016 Create Date August 09, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2366-2016 Recall Event ID 73750 510(K)Number K033326 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Brilliance 64 Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Model #: 728231; Serial #s: 9009, 9029, 9033, 9037, 9038, 9043, 9045, 9051, 9052, 9056, 9058, 9064, 9065, 9067, 9068, 9070, 9072, 9073, 9075, 9086, 9089, 9092, 9104, 9105, 9106, 9107, 9108, 9112, 9120, 9139, 9151, 9152, 9156, 9190, 9243, 9259, 9502, 9503, 9504, 9507, 9508, 9511, 9512, 9515, 9516, 9522, 9525, 9526, 9529, 9530, 9531, 9533, 9534, 9545, 9547, 9550, 9551, 9556, 9561, 9564, 9572, 9577, 9580, 9584, 9591, 9594, 9596, 9598, 9603, 9606, 9607, 9608, 9612, 9613, 9614, 9618, 9620, 9624, 9625, 9626, 9628, 9630, 9633, 9634, 9635, 9639, 9640, 9641, 9642, 9646, 9648, 9650, 9655, 9663, 9666, 9667, 9672, 9676, 9678, 9685, 9686, 9690, 9691, 9692, 9694, 9695, 9697, 9700, 9705, 9706, 9709, 9713, 9716, 9726, 9729, 9732, 9735, 9736, 9738, 9740, 9750, 9751, 9752, 9753, 9754, 9757, 9759, 9760, 9761, 9775, 9777, 9780, 9781, 9788, 9790, 9793, 9794, 9802, 9805, 9809, 9811, 9812, 9821, 9823, 9826, 9830, 9833, 9835, 9843, 9846, 9855, 9857, 9860, 9861, 9866, 9869, 9870, 9873, 9875, 9876, 9878, 9879, 9881, 9882, 9885, 9888, 9889, 9890, 9894, 9895, 9902, 9906, 9910, 9916, 9917, 9918, 9924, 9936, 9937, 9938, 9942, 9948, 9951, 9954, 9960, 9961, 9965, 9966, 9968, 9969, 9974, 9975, 9978, 9980, 9982, 9985, 9992, 9993, 9994, 9999, 10001, 10005, 10014, 10018, 10026, 10030, 10031, 10035, 10036, 10037, 10041, 10045, 10046, 10051, 10054, 10065, 10067, 10070, 10072, 10079, 10080, 10085, 10086, 10094, 10095, 10096, 10098, 10103, 10111, 10112, 10113, 10120, 10122, 10125, 10141, 10148, 10190, 10194, 10200, 10202, 10203, 10211, 10224, 10256, 10263, 10291, 10292, 10310, 10332, 10353, 10359, 10363, 10365, 10370, 10371, 10372, 10373, 10380, 10381, 10386, 10388, 10413, 10417, 10424, 10426, 10473, 10486, 10505, 10513, 10527, 10535, 10561, 10570, 10645, 10651, 10655, 10673, 10686, 10733, 10743, 10773, 10776, 29005, 29007, 29016, 29018, 29025, 29027, 29035, 29037, 29053, 29054, 29068, 29097, 64015, 90031, 90047, 90062, 90063, 90128, 95003, 95008, 95022, 95057, 95058, 95068, 95149, 95197, 95210, 95348, 95357, 95377, 95392, 95403, 95811, 95341A, 9851T, 9172. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall Software defect causing intermittently slow response of Host. FDA Determined Cause 2 Software change control Action On 3/24/2016 the firm mailed Customer Information letters to their customers. Quantity in Commerce 318 units Distribution Product was sent to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV & WY.************Product was also sent to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Cuba, Denmark, Djibouti, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, R¿union, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 3/24/2016 the firm mailed Customer Information letters to their customers. | SoftwareChangeControl
82 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144869 | Z-2368-2016 | Brilliance BigBore Radiology Computed Tomography X-ray systems, intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Brilliance BigBore Radiology Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host. | Class 2 Device Recall Brilliance BigBore Radiology Computed Tomography Xray systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Brilliance BigBore Radiology Computed Tomography Xray systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Brilliance BigBore Radiology Computed Tomography Xray systems Date Initiated by Firm March 24, 2016 Create Date August 09, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2368-2016 Recall Event ID 73750 510(K)Number K033357 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Brilliance BigBore Radiology Computed Tomography X-ray systems, intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Model #: 728244; Serial #s: 7052, 7172, 7200, 7244, 7252, 7375, 7380, 7384, 7390, 7514, 7533, 7546, 7614, 7633, 7708, 7781, 7789. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall Software defect causing intermittently slow response of Host. FDA Determined Cause 2 Software change control Action On 3/24/2016 the firm mailed Customer Information letters to their customers. Quantity in Commerce 17 units Distribution Product was sent to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV & WY.************Product was also sent to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Cuba, Denmark, Djibouti, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, R¿union, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 3/24/2016 the firm mailed Customer Information letters to their customers. | SoftwareChangeControl
83 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144872 | Z-2371-2016 | Ingenuity Core Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Ingenuity Core Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host. | Class 2 Device Recall Ingenuity Core Computed Tomography Xray systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Ingenuity Core Computed Tomography Xray systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Ingenuity Core Computed Tomography Xray systems Date Initiated by Firm March 24, 2016 Create Date August 09, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2371-2016 Recall Event ID 73750 510(K)Number K033326 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Ingenuity Core Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Model #: 728321; Serial #s: 310004, 310005, 310012, 310016, 310020, 310021, 310027, 310028, 310033, 310034, 310041, 310044, 310045, 310057, 310061, 310065, 310072, 310074, 310076, 310077, 310078, 310094, 310099, 310100, 310103, 310105, 310117, 310120, 310135. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall Software defect causing intermittently slow response of Host. FDA Determined Cause 2 Software change control Action On 3/24/2016 the firm mailed Customer Information letters to their customers. Quantity in Commerce 29 units Distribution Product was sent to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV & WY.************Product was also sent to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Cuba, Denmark, Djibouti, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, R¿union, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 3/24/2016 the firm mailed Customer Information letters to their customers. | SoftwareChangeControl
84 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144873 | Z-2372-2016 | Ingenuity Core 128 Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Ingenuity Core 128 Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host. | Class 2 Device Recall Ingenuity Core 128 Computed Tomography Xray systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Ingenuity Core 128 Computed Tomography Xray systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Ingenuity Core 128 Computed Tomography Xray systems Date Initiated by Firm March 24, 2016 Create Date August 09, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2372-2016 Recall Event ID 73750 510(K)Number K033326 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Ingenuity Core 128 Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Model #: 728323; Serial #s: 310067, 320002, 320004, 320053, 320056. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall Software defect causing intermittently slow response of Host. FDA Determined Cause 2 Software change control Action On 3/24/2016 the firm mailed Customer Information letters to their customers. Quantity in Commerce 5 units Distribution Product was sent to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV & WY.************Product was also sent to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Cuba, Denmark, Djibouti, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, R¿union, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 3/24/2016 the firm mailed Customer Information letters to their customers. | SoftwareChangeControl
85 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144201 | Z-1394-2016 | VERION Reference Unit (Vision Planner)******Product Usage:***The VERION Image Guided system is an advanced surgical planning, imaging and guidance technology designed to provide greater accuracy and efficiency during cataract surgery. It allows eye surgeons to capture a high-resolution, diagnostic reference image of the patients eye pre-operatively, quickly determine an optimized surgical plan that enables surgeons to see all inclusions and alignment in real-time.*** | VERION Image Guided System | 2 | CDRH | 2016-04-14 | 2016-04-14 21:04 | 2017-04-13 | Alcon Research, Ltd. | Alcon is conducting a voluntary medical device correction of all VERION Reference Units (Vision Planner) that are shared with the Alcon LenSx Laser System after receiving reports concerning the inclusion of unplanned arcuates on printed, saved (.pdf) or exported surgical plans. | Class 2 Device Recall VERION Image Guided System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VERION Image Guided System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VERION Image Guided System Date Initiated by Firm March 01, 2016 Date Posted April 14, 2016 Recall Status 1 Terminated 3 on April 13, 2017 Recall Number Z-1394-2016 Recall Event ID 73430 Product Classification Device, communications, images, ophthalmic - Product Code NFG Product VERION Reference Unit (Vision Planner)************Product Usage:******The VERION Image Guided system is an advanced surgical planning, imaging and guidance technology designed to provide greater accuracy and efficiency during cataract surgery. It allows eye surgeons to capture a high-resolution, diagnostic reference image of the patient***s eye pre-operatively, quickly determine an optimized surgical plan that enables surgeons to see all inclusions and alignment in real-time. Code Information Serial Numbers: 1017, 1119, 1123, 1125, 1126, 1128, 1133, 1135, 1136, 1137, 1138, 1139, 1142, 1147, 1150, 1153, 1155, 1161, 1165, 1168, 1172, 1178, 1181, 1182, 1183, 1188, 1190, 1191, 1193, 1196, 1197, 1198, 1199, 1201, 1203, 1206, 1207, 1220, 1226, 1228, 1230, 1232, 1250, 1256, 1258, 1259, 1421, 1433, 1435, 1439, 1451, 1452, 1506, 1508, 1509, 1523, 1536, 1548, 1559, 1562, 1565, 1568, 1581, 1597, 1606, 1628, 1632, 1649, 1661, 1663, 1668, 1671, 1694, 1696, 1705, 1716, 1726, 1729, 1798, 1816, 1818, 1820, 1824, 1829, 1921, 1922, 1934, 1937, 1938, 1940, 1942, 1954, 1957, 2057, 2069, 2086, 2092 Recalling Firm/ Manufacturer Alcon Research, Ltd. 6201 South Fwy Fort Worth TX 76134-2099 For Additional Information Contact Alcon 800-862-5266 Manufacturer Reason for Recall Alcon is conducting a voluntary medical device correction of all VERION Reference Units (Vision Planner) that are shared with the Alcon LenSx Laser System after receiving reports concerning the inclusion of unplanned arcuates on printed, saved (.pdf) or exported surgical plans. FDA Determined Cause 2 Software change control Action Alcon sent an Urgent Medical Device Correction letter dated February 29, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that an Alcon Technical Services Engineers will contact customers, schedule Service Calls, and perform onsite corrections. For questions call 1-800-862-5266. Quantity in Commerce 97 units Distribution Worldwide Distribution - US Nationwide throughout the US and the countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Czech Republic, Chile, China, Colombia, Costa Rica, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, France, Georgia, Greece, Guatlemala, Hungary, India, Indoesia, Isreal, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Nicaragua, Norway, Panama, Philippines, Portugal, Romania, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, UK, Ukrain, United Arab Emirates Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Alcon sent an Urgent Medical Device Correction letter dated February 29, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that an Alcon Technical Services Engineers will contact customers, schedule Service Calls, and perform onsite corrections. For questions call 1-800-862-5266. | SoftwareChangeControl
86 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=141582 | Z-0319-2016 | Syngo Imaging VB36D_HF02. Radiological image processing system. | Syngo Imaging | 2 | CDRH | 2015-11-20 | 2015-11-20 21:01 | 2016-10-06 | Siemens Medical Solutions USA, Inc | To provide supplementary information regarding the release of the syngo Imaging Software Version VB36D_HF02. The software provides improvements for all syngo Imaging installations running on the software Version VB36D. | Class 2 Device Recall Syngo Imaging Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Syngo Imaging FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Syngo Imaging Date Initiated by Firm October 28, 2015 Date Posted November 20, 2015 Recall Status 1 Terminated 3 on October 06, 2016 Recall Number Z-0319-2016 Recall Event ID 72576 510(K)Number K071114 K081734 Product Classification System, image processing, radiological - Product Code LLZ Product Syngo Imaging VB36D_HF02. Radiological image processing system. Code Information model# 10014063 and 10014064 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact 610-219-6300 Manufacturer Reason for Recall To provide supplementary information regarding the release of the syngo Imaging Software Version VB36D_HF02. The software provides improvements for all syngo Imaging installations running on the software Version VB36D. FDA Determined Cause 2 Software change control Action Siemens sent an Customer Information Letter dated October 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.************If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions, please call: (610) 219-6300. Quantity in Commerce 3 Distribution US Distribution to the states of : NC, NE and OH. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens sent an Customer Information Letter dated October 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.************If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions, please call: (610) 219-6300. | SoftwareChangeControl
87 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139574 | Z-0117-2016 | NicoletOne Software v5.94, Catalog/Part Number: 482-649600.***The NicoletOne data acquisition and review system, either with or without synchronous digital video. The system is intended for medical purposes to record, measure, store and display full band (FbEEG) cerebral EEG and extracerebral activity for Clinical EEG, Electrocorticography (EcOG), Long Term Monitoring (LTM), Intensive Care Unit (ICU) monitoring and Polysomnography (PSG) Sleep studies. While the Nicolet Neurodiagnostic systems are capable of displaying signals, such as Sp02 and EKG, the system is NOT intended for monitoring such signals for the preservation of life, The Nicolet Neurodiagnostic systems are intended to acquire, analyze, and display data | NicoletOne Software v5.94 | 2 | CDRH | 2015-10-13 | 2015-10-13 21:00 | 2016-03-08 | Natus Neurology Inc | When using the NicoletOne v5.94 software, after exiting the impedance check function and returning to the EEG screen, the impedance check signal remains active in waveform, obscuring the EEG signals. | Class 2 Device Recall NicoletOne Software v5.94 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall NicoletOne Software v5.94 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall NicoletOne Software v5.94 Date Initiated by Firm September 11, 2015 Date Posted October 13, 2015 Recall Status 1 Terminated 3 on March 08, 2016 Recall Number Z-0117-2016 Recall Event ID 71991 510(K)Number K964280 Product Classification Full-montage Standard Electroencephalograph - Product Code GWQ Product NicoletOne Software v5.94, Catalog/Part Number: 482-649600.******The NicoletOne data acquisition and review system, either with or without synchronous digital video. The system is intended for medical purposes to record, measure, store and display full band (FbEEG) cerebral EEG and extracerebral activity for Clinical EEG, Electrocorticography (EcOG), Long Term Monitoring (LTM), Intensive Care Unit (ICU) monitoring and Polysomnography (PSG) Sleep studies. While the Nicolet Neurodiagnostic systems are capable of displaying signals, such as Sp02 and EKG, the system is NOT intended for monitoring such signals for the preservation of life, The Nicolet Neurodiagnostic systems are intended to acquire, analyze, and display data Code Information Serial numbers include TS150456M, TR150171M, TR150172M, 1TR150173M, TR150174M, TS150451M, TS150452M, TR150170M, UH150119M, TR150169M, UR150226M, UR150227M, UR150228M, UR150229M, TS150455M, TS150457M, TS150458M, TS150454M, TS150453M, UC150184M, PN151700M, SJ150272M, TT150134M, PN151699M, UR150230M, UR150231M, and SJ150271M. Recalling Firm/ Manufacturer Natus Neurology Inc 3150 Pleasant View Rd Middleton WI 53562-4800 For Additional Information Contact Toni Gorman 608-829-8600 Manufacturer Reason for Recall When using the NicoletOne v5.94 software, after exiting the impedance check function and returning to the EEG screen, the impedance check signal remains active in waveform, obscuring the EEG signals. FDA Determined Cause 2 Software change control Action Natus Neurology sent an Ugent Medical Device Corrrection Notification letter dated September 9, 2015 to all affected consignees. The letter described the problem and the product involved in the recall. Consignees were informed consignees that a Field Service Technician will contact them to install the new software. Once the software is installed they will ask the Recipient or Field service Technician to complete and return the Verification Form.******For questions regarding this recall call 608-829-8600. Quantity in Commerce 101 (88 US, 13 OUS) Distribution Worldwide distribution. US nationwide, INDIA, ROMANIA, DENMARK, JAPAN, AUSTRALIA, RUSSIA, BANGLADESH, SRI LANKA, BRUNEI, and UNITED ARAB EMIRATES. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = GWQ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Natus Neurology sent an Ugent Medical Device Corrrection Notification letter dated September 9, 2015 to all affected consignees. The letter described the problem and the product involved in the recall. Consignees were informed consignees that a Field Service Technician will contact them to install the new software. Once the software is installed they will ask the Recipient or Field service Technician to complete and return the Verification Form.******For questions regarding this recall call 608-829-8600. | SoftwareChangeControl
88 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139556 | Z-2784-2015 | Merlin@home RF Remote Monitoring Transmitter Model EX1150.******Intended to aid in the remote support of SJM implantable pacemakers and ICD's. | Merlin@home RF Remote Monitoring Transmitter Model EX1150 | 2 | CDRH | 2015-09-17 | 2015-09-28 21:00 | 2016-05-19 | St Jude Medical Cardiac Rhythm Management Division | Some devices exhibit backup VVI operation due to a backup reset. | Class 2 Device Recall Merlinhome RF Remote Monitoring Transmitter Model EX1150 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Merlinhome RF Remote Monitoring Transmitter Model EX1150 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Merlinhome RF Remote Monitoring Transmitter Model EX1150 Date Initiated by Firm August 18, 2015 Date Posted September 28, 2015 Recall Status 1 Terminated 3 on May 19, 2016 Recall Number Z-2784-2015 Recall Event ID 71995 PMA Number P910023S346 P910023S323 P910023S307 P910023S017 Product Classification Pulse generator, permanent, implantable - Product Code NVZ Product Merlin@home RF Remote Monitoring Transmitter Model EX1150.************Intended to aid in the remote support of SJM implantable pacemakers and ICD's. Code Information Model: EX1150 Serial #: 02902036 05103228 12200027 12200047 12200063 12200065 12200071 12200093 12200100 12200111 12200120 12200122 12200157 12200165 12200168 12200174 12200187 12200235 12200245 12200250 12200331 12200333 12200351 12200367 12200369 12200371 12200399 12200432 12200449 12200459 12200466 12200471 12200478 12200522 12200539 12200574 12200582 12200621 12200641 12200658 12200661 12200679 12200680 12200700 12200738 12200749 12200754 12200758 12200771 12200772 12200788 12200791 12200800 12200832 12200845 12200850 12200851 12200864 12200868 12200877 12200891 12200892 12200906 12200914 12200945 12200963 12200965 12200966 12200967 12200972 12200975 12200976 12200977 12200978 12200980 12200987 12200991 12200992 12200997 12200999 12201000 12201001 12201002 12201005 12201006 12201007 12201008 12201010 12201011 12201016 12201019 12201020 12201025 12201026 12201036 12201037 12201038 12201039 12201040 12201042 12201043 12201048 12201054 12201055 12201056 12201060 12201064 12201066 12201067 12201068 12201071 12201073 12201074 12201096 12201116 12201117 12201122 12201128 12201132 12201178 12201190 12201194 12201204 12201213 12201230 12201231 12201244 12201247 12201249 12201252 12201265 12201270 12201281 12201290 12201308 12201330 12201372 12201381 12201392 12201394 12201408 12201414 12201440 12201445 12201464 12201465 12201466 12201469 12201471 12201486 12201490 12201491 12201492 12201493 12201498 12201499 12201500 12201501 12201510 12201517 12201522 12201523 12201525 12201529 12201533 12201538 12201539 12201540 12201543 12201546 12201549 12201550 12201551 12201558 12201565 12201567 12201568 12201570 12201575 12201581 12201588 12201589 12201591 12201597 12201598 12201599 12201604 12201606 12201608 12201610 12201613 12201617 12201618 12201621 12201622 12201625 12201638 12201640 12201641 12201643 12201646 12201650 12201651 12201654 12201660 12201662 12201663 12201665 12201666 12201669 12201670 12201671 12201673 12201676 12201678 12201679 12201681 12201684 12201685 12201689 12201691 12201692 12201694 12201695 12201696 12201698 12201700 12201701 12201703 12201706 12201709 12201710 12201711 12201713 12201714 12201715 12201716 12201717 12201719 12201720 12201724 12201726 12201728 12201730 12201734 12201736 12201737 12201738 12201773 12201780 12201790 12201799 12201805 12201808 12201812 12201826 12201827 12201838 12201846 12201847 12201856 12201869 12201870 12201897 12201920 12201931 12201934 12201944 12201960 12201976 12202002 12202006 12202012 12202016 12202022 12202023 12202044 12202048 12202057 12202058 12202077 12202091 12202093 12202103 12202112 12202123 12202144 12202146 12202149 12202157 12202160 12202174 12202193 12202196 12202197 12202202 12202222 12202226 12202228 12202233 12202262 12202283 12202296 12202297 12202304 12202312 12202326 12202338 12202350 12202356 12202362 12202372 12202382 12202385 12202400 12202413 12202418 12202432 12202458 12202459 12202466 12202469 12202472 12202477 12202484 12202488 12202496 12202502 12202526 12202549 12202553 12202577 12202580 12202588 12202600 12202609 12202616 12202623 12202631 12202643 12202688 12202730 12202732 12202735 12202792 12202795 12202796 12202798 12202856 12202862 12202866 12202870 12202885 12202888 12202906 12202925 12202932 12202933 12202934 12202949 12202961 12202980 12202982 12202986 12202987 12203035 12203051 12203060 12203062 12203066 12203070 12203080 12203081 12203157 12203167 12203193 12203195 12203201 12203213 12203216 12203223 12203224 12203245 12203258 12203262 12203263 12203264 12203273 12203280 12203281 12203285 12203320 12203328 12203329 12203335 12203357 12203359 12203363 12203375 12203378 12203381 12203383 12203388 12203407 12203409 12203417 12203431 12203479 12203481 12203485 12203486 12203501 12203519 12203554 12203566 12203608 12203621 12203642 12203651 12203667 12203668 12203679 12203693 12203695 12203708 12203739 12203783 12203795 12203797 12203825 12203838 12203846 12203853 12203859 12203881 12203884 12203885 12203888 12203890 12203896 12203915 12203923 12203926 12203928 12203955 12203956 12203961 12203973 12203977 12203989 12204011 12204015 12204020 12204023 12204031 12204057 12204061 12204071 12204119 12204123 12204142 12204150 12204161 12204164 12204167 12204171 12204237 12204241 12204249 12204265 12204274 12204279 12204280 12204288 12204291 12204297 12204308 12204346 12204360 12204361 12204366 12204381 12204391 12204395 12204398 12204402 12204416 12204428 12204430 12204433 12204470 12204491 12204497 12204498 12204499 12204507 12204511 12204527 12204534 12204546 12204552 12204573 12204610 12204630 12204631 12204636 12204637 12204640 12204651 12204676 12204689 12204702 12204712 12204717 12204743 12204752 12204759 12204764 12204768 12204781 12204782 12204789 12204790 12204791 12204811 12204844 12204845 12204863 12204877 12204900 12204905 12204908 12204911 12204917 12204920 12204927 12204952 12204976 12204988 12204991 12205000 12205016 12205019 12205028 12205038 12205040 12205060 12205068 12205079 12205107 12205116 12205127 12205128 12205142 12205162 12205213 12205230 12205234 12205236 12205245 12205273 12205335 12205364 12205374 12205391 12205401 12205407 12205432 12205436 12205442 12205446 12205448 12205460 12205492 12205500 12205501 12205511 12205516 12205532 12205544 12205545 12205581 12205583 12205585 12205592 12205617 12205618 12205746 12205779 12205800 12205801 12205814 12205821 12205826 12205833 12205835 12205845 12205851 12205867 12205879 12205881 12205886 12205911 12205913 12205914 12205928 12205948 12205949 12205966 12205970 12205971 12205974 12205977 12206001 12206024 12206027 12206030 12206063 12206064 12206078 12206105 12206119 12206120 12206158 12206164 12206184 12206191 12206197 12206203 12206219 12206221 12206234 12206235 12206236 12206254 12206259 12206265 12206302 12206303 12206312 12206319 12206325 12206346 12206368 12206371 12206402 12206409 12206414 12206467 12206470 12206471 12206473 12206485 12206512 12206518 12206525 12206609 12206612 12206613 12206635 12206667 12206722 12206725 12206752 12206779 12206804 12206822 12206824 12206858 12206872 12206900 12206971 12206996 12206997 12207082 12207119 12207146 12207160 12207162 12207163 12207193 12207218 12207224 12207230 12207233 12207236 12207257 12207262 12207285 12207297 12207360 12207372 12207373 12207384 12207386 12207426 12207429 12207442 12207475 12207476 12207482 12207521 12207539 12207543 12207544 12207564 12207575 12207579 12207619 12207623 12207652 12207666 12207683 12207701 12207710 12207719 12207721 12207725 12207733 12207834 12207838 12207853 12207867 12207879 12207887 12207888 12207906 12207929 12207978 12207986 12207991 12208001 12208026 12208057 12208068 12208075 12208088 12208156 12208163 12208172 12208173 12208177 12208187 12208195 12208209 12208235 12208242 12208274 12208325 12208333 12208335 12208338 12208348 12208360 12208364 12208395 12208404 12208414 12208417 12208443 12208449 12208469 12208489 12208511 12208519 12208520 12208522 12208544 12208553 12208560 12208599 12208605 12208643 12208654 12208655 12208666 12208676 12208706 12208708 12208710 12208713 12208732 12208736 12208746 12208758 12208794 12208801 12208836 12208843 12208855 12208856 12208864 12208869 12208877 12208898 12208924 12208935 12208944 12208947 12208960 12208976 12208991 12209014 12209022 12209040 12209056 12209132 12209133 12209142 12209149 12209154 12209189 12209190 12209202 12209248 12209261 12209265 12209267 12209268 12209277 12209284 12209291 12209311 12209316 12209328 12209329 12209332 12209350 12209353 12209356 12209364 12209379 12209386 12209414 12209431 12209463 12209513 12209517 12209521 12209527 12209528 12209533 12209545 12209553 12209566 12209579 12209597 12209619 12209628 12209651 12209665 12209668 12209673 12209678 12209704 12209710 12209717 12209738 12209740 12209783 12209796 12209807 12209820 12209830 12209832 12209838 12209841 12209846 12209909 12209932 12209933 12209936 12209946 12209949 12209955 12209972 12209973 12209991 12209992 12210006 12210030 12210034 12210041 12210054 12210057 12210059 12210065 12210078 12210081 12210093 12210110 12210113 12210122 12210159 12210196 12210201 12210225 12210243 12210263 12210283 12210284 12210285 12210287 12210294 12210300 12210306 12210322 12210325 12210329 12210339 12210346 12210362 12210364 12210372 12210392 12210402 12210425 12210426 12210441 12210453 12210466 12210469 12210476 12210477 12210500 12210518 12210530 12210532 12210569 12210583 12210589 12210594 12210600 12210776 12210795 12210822 12210834 12210845 12210856 12210859 12210862 12210880 12210898 12210912 12210920 12210928 12210943 12210948 12210953 12210957 12210960 12211000 12211012 12211057 12211064 12211080 12211081 12211096 12211099 12211106 12211122 12211125 12211126 12211141 12211143 12211144 12211146 12211161 12211182 12211206 12211216 12211222 12211225 12211239 12211244 12211261 12211280 12211364 12211422 12211427 12211438 12211441 12211443 12211477 12211487 12211495 12211499 12211509 12211535 12211536 12211548 12211574 12211584 12211602 12211610 12211622 12211644 12211647 12211662 12211663 12211665 12211666 12211669 12211678 12211679 12211685 12211698 12211704 12211727 12211758 12211766 12211784 12211787 12211798 12211808 12211841 12211849 12211850 12211869 12211877 12211905 12211940 12211952 12211954 12211957 12211966 12211981 12211999 12212005 12212007 12212008 12212024 12212032 12212037 12212041 12212042 12212054 12212078 12212094 12212095 12212179 12212213 12212218 12212221 12212225 12212230 12212235 12212250 12212257 12212265 12212301 12212325 12212334 12212337 12212346 12212361 12212362 12212397 12212398 12212406 12212412 12212420 12212457 12212473 12212490 12212491 12212492 12212509 12212513 12212519 12212545 12212565 12212570 12212579 12212581 12212592 12212613 12212619 12212656 12212657 12212658 12212666 12212675 12212691 12212702 12212745 12212756 12212764 12212766 12212768 12212769 12212776 12212789 12212796 12212804 12212811 12212828 12212830 12212839 12212863 12212867 12212877 12212889 12212908 12212912 12212930 12212944 12212952 12212958 12212959 12212967 12212971 12212972 12212979 12212985 12213020 12213024 12213029 12213038 12213067 12213068 12213073 12213077 12213078 12213079 12213087 12213093 12213096 12213125 12213127 12213137 12213154 12213161 12213167 12213193 12213199 12213205 12213209 12213210 12213214 12213221 12213234 12213235 12213239 12213245 12213251 12213272 12213308 12213314 12213331 12213353 12213356 12213383 12213386 12213390 12213396 12213398 12213410 12213436 12213439 12213449 12213451 12213460 12213462 12213463 12213464 12213468 12213474 12213488 12213494 12213498 12213504 12213527 12213545 12213546 12213549 12213553 12213571 12213577 12213579 12213583 12213588 12213592 12213595 12213598 12213600 12213601 12213610 12213611 12213639 12213650 12213685 12213695 12213697 12213710 12213725 12213764 12213767 12213772 12213780 12213795 12213799 12213802 12213809 12213812 12213828 12213839 12213862 12213865 12213872 12213903 12213905 12213914 12213916 12213932 12213934 12213949 12213963 12213970 12213974 12213986 12214001 12214005 12214022 12214031 12214033 12214038 12214066 12214101 12214107 12214119 12214124 12214128 12214132 12214172 12214180 12214182 12214188 12214189 12214204 12214216 12214236 12214248 12214262 12214268 12214271 12214279 12214280 12214283 12214290 12214298 12214301 12214315 12214326 12214333 12214342 12214383 12214411 12214430 12214444 12214463 12214470 12214503 12214537 12214540 12214558 12214583 12214586 12214591 12214597 12214616 12214642 12214648 12214650 12214661 12214702 12214706 12214708 12214719 12214720 12214724 12214728 12214737 12214746 12214748 12214754 12214758 12214760 12214769 12214772 12214793 12214797 12214800 12214830 12214880 12214891 12214892 12214895 12214909 12214912 12214920 12214954 12214956 12214963 12214981 12214983 12215002 12215003 12215006 12215008 12215010 12215014 12215026 12215031 12215032 12215035 12215041 12215083 12215090 12215133 12215150 12215157 12215168 12215194 12215197 12215221 12215222 12215238 12215244 12215297 12215301 12215312 12215313 12215323 12215329 12215341 12215353 12215358 12215361 12215364 12215370 12215372 12215380 12215421 12215427 12215485 12215502 12215540 12215545 12215563 12215570 12215589 12215599 12215612 12215655 12215656 12215657 12215658 12215675 12215677 12215681 12215686 12215700 12215705 12215709 12215714 12215723 12215737 12215754 12215780 12215784 12215788 12215830 12215836 12215845 12215857 12215868 12215869 12215879 12215885 12215893 12215895 12215897 12215901 12215912 12215938 12215941 12215949 12215951 12215954 12215972 12215988 12215990 12216000 12216034 12216042 12216056 12216058 12216063 12216080 12216086 12216109 12216118 12216121 12216126 12216129 12216154 12216163 12216190 12216194 12216195 12216203 12216207 12216208 12216240 12216242 12216262 12216270 12216283 12216284 12216288 12216300 12216304 12216315 12216324 12216353 12216360 12216363 12216375 12216392 12216399 12216416 12216433 12216439 12216451 12216473 12216487 12216510 12216514 12216518 12216532 12216569 12216588 12216596 12216601 12216603 12216618 12216659 12216660 12216696 12216711 12216738 12216761 12216769 12216776 12216780 12216792 12216796 12216803 12216806 12216808 12216811 12216819 12216827 12216851 12216856 12216864 12216876 12216885 12216886 12216937 12216943 12216960 12216963 12216971 12216976 12216980 12216986 12216994 12216995 12216998 12217002 12217009 12217010 12217019 12217067 12217074 12217081 12217089 12217091 12217101 12217108 12217109 12217115 12217116 12217120 12217125 12217134 12217136 12217151 12217153 12217161 12217164 12217188 12217189 12217191 12217223 12217228 12217259 12217315 12217330 12217429 12217430 12217436 12217454 12217467 12217479 12217488 12217497 12217500 12217510 12217548 12217566 12217573 12217578 12217611 12217614 12217663 12217691 12217693 12217713 12217714 12217720 12217721 12217726 12217727 12217728 12217749 12217772 12217774 12217777 12217783 12217786 12217789 12217791 12217793 12217794 12217796 12217813 12217818 12217833 12217834 12217836 12217844 12217847 12217864 12217879 12217889 12217898 12217928 12217962 12218002 12218008 12218014 12218030 12218033 12218037 12218041 12218044 12218052 12218070 12218081 12218111 12218128 12218138 12218146 12218156 12218161 12218167 12218177 12218181 12218194 12218195 12218199 12218209 12218220 12218250 12218252 12218273 12218280 12218309 12218318 12218331 12218346 12218354 12218356 12218361 12218371 12218395 12218398 12218415 12218452 12218453 12218454 12218470 12218473 12218482 12218485 12218488 12218501 12218515 12218530 12218543 12218545 12218570 12218575 12218583 12218584 12218600 12218614 12218626 12218637 12218643 12218645 12218665 12218671 12218698 12218718 12218719 12218725 12218728 12218730 12218740 12218741 12218744 12218759 12218760 12218762 12218767 12218779 12218800 12218814 12218816 12218831 12218833 12218834 12218836 12218839 12218867 12218871 12218892 12218895 12218896 12218918 12218931 12218960 12218966 12218978 12218982 12219014 12219018 12219021 12219055 12219071 12219074 12219075 12219110 12219117 12219123 12219129 12219154 12219176 12219187 12219190 12219191 12219195 12219200 12219207 12219217 12219219 12219224 12219230 12219264 12219286 12219290 12219308 12219314 12219317 12219319 12219326 12219369 12219370 12219375 12219377 12219387 12219407 12219420 12219424 12219433 12219437 12219448 12219449 12219466 12219478 12219493 12219498 12219514 12219515 12219522 12219527 12219533 12219560 12219581 12219595 12219604 12219630 12219670 12219705 12219714 12219726 12219730 12219732 12219737 12219772 12219775 12219791 12219798 12219801 12219814 12219824 12219869 12219870 12219874 12219880 12219889 12219894 12219902 12219939 12219969 12219990 12220053 12220066 12220075 12220150 12220155 12220180 12220183 12220203 12220214 12220223 12220246 12220247 12220253 12220258 12220263 12220271 12220273 12220286 12220298 12220312 12220325 12220373 12220384 12220386 12220398 12220405 12220417 12220418 12220425 12220429 12220432 12220442 12220453 12220454 12220463 12220467 12220474 12220486 12220491 12220500 12220515 12220540 12220554 12220560 12220584 12220592 12220610 12220620 12220623 12220627 12220637 12220638 12220648 12220655 12220667 12220668 12220669 12220677 12220685 12220687 12220738 12220742 12220759 12220770 12220788 12220798 12220807 12220819 12220821 12220828 12220829 12220846 12220851 12220862 12220871 12220876 12220883 12220886 12220889 12220899 12220926 12220934 12220939 12220970 12221027 12221040 12221043 12221044 12221060 12221064 12221087 12221100 12221107 12221110 12221120 12221128 12221170 12221183 12221184 12221186 12221192 12221231 12221237 12221245 12221269 12221273 12221306 12221308 12221313 12221333 12221358 12221423 12221434 12221436 12221439 12221441 12221454 12221459 12221479 12221482 12221512 12221534 12221553 12221570 12221587 12221602 12221614 12221619 12221622 12221644 12221651 12221660 12221698 12221709 12221734 12221743 12221746 12221788 12221792 12221798 12221800 12221814 12221839 12221852 12221859 12221926 12221931 12221948 12221968 12221978 12221996 12221997 12222000 12222006 12222012 12222014 12222034 12222037 12222044 12222060 12222071 12222091 12222097 12222116 12222123 12222146 12222159 12222162 12222164 12222167 12222174 12222221 12222223 12222257 12222262 12222269 12222297 12222317 12222322 12222341 12222368 12222378 12222382 12222384 12222387 12222391 12222441 12222443 12222468 12222477 12222498 12222511 12222540 12222543 12222544 12222591 12222605 12222621 12222640 12222696 12222714 12222719 12222720 12222730 12222731 12222738 12222739 12222768 12222769 12222824 12222841 12222896 12222908 12222929 12222947 12222958 12222964 12222968 12223008 12223011 12223048 12223079 12223091 12223109 12223110 12223120 12223123 12223145 12223150 12223175 12223203 12223242 12223251 12223258 12223284 12223360 12223429 12223468 12223471 12223475 12223505 12223514 12223534 12223539 12223544 12223549 12223560 12223580 12223589 12223599 12223603 12223620 12223625 12223645 12223647 12223653 12223654 12223657 12223679 12223683 12223727 12223734 12223735 12223774 12223779 12223782 12223803 12223811 12223822 12223888 12223901 12223906 12223908 12223909 12223912 12223935 12223937 12223951 12223959 12223976 12223987 12223988 12224032 12224039 12224043 12224061 12224071 12224076 12224086 12224107 12224109 12224110 12224114 12224116 12224156 12224186 12224187 12224193 12224204 12224207 12224214 12224216 12224229 12224243 12224296 12224298 12224312 12224330 12224350 12224353 12224361 12224375 12224384 12224388 12224390 12224403 12224404 12224412 12224429 12224433 12224450 12224462 12224490 12224507 12224523 12224537 12224564 12224567 12224582 12224598 12224612 12224627 12224652 12224653 12224663 12224670 12224681 12224707 12224708 12224750 12224764 12224772 12224775 12224780 12224804 12224812 12224849 12224853 12224884 12224887 12224909 12224923 12224938 12224942 12224948 12224962 12224964 12224979 12225000 12225036 12225038 12225074 12225080 12225081 12225084 12225132 12225172 12225182 12225204 12225207 12225224 12225232 12225275 12225282 12225287 12225291 12225334 12225343 12225384 12225392 12225400 12225403 12225407 12225410 12225411 12225422 12225425 12225461 12225484 12225514 12225527 12225541 12225556 12225574 12225580 12225586 12225592 12225606 12225610 12225616 12225624 12225626 12225630 12225639 12225660 12225668 12225680 12225693 12225706 12225734 12225745 12225771 12225772 12225782 12225787 12225804 12225829 12225832 12225870 12225882 12225885 12225891 12225893 12225912 12225921 12225953 12225995 12226052 12226133 12226136 12226146 12226154 12226181 12226183 12226198 12226199 12226204 12226213 12226214 12226226 12226227 12226247 12226254 12226256 12226264 12226282 12226295 12226328 12226359 12226384 12226402 12226410 12226413 12226423 12226429 12226456 12226486 12226488 12226489 12226499 12226517 12226518 12226588 12226620 12226642 12226658 12226659 12226662 12226694 12226699 12226734 12226750 12226754 12226766 12226773 12226775 12226784 12226791 12226799 12226801 12226811 12226834 12226842 12226869 12226898 12226909 12226919 12226940 12226968 12227047 12227075 12227094 12227125 12227141 12227155 12227175 12227192 12227202 12227221 12227276 12227333 12227340 12227346 12227358 12227394 12227395 12227398 12227401 12227405 12227457 12227474 12227485 12227497 12227509 12227531 12227557 12227563 12227581 12227598 12227609 12227639 12227642 12227646 12227668 12227683 12227710 12227732 12227735 12227736 12227740 12227742 12227783 12227793 12227813 12227816 12227826 12227839 12227852 12227854 12227858 12227870 12227873 12227887 12227889 12227892 12227899 12227919 12227932 12227933 12227948 12227956 12227974 12227995 12228000 12228020 12228029 12228051 12228066 12228072 12228085 12228090 12228093 12228098 12228100 12228111 12228113 12228119 12228120 12228135 12228137 12228143 12228148 12228200 12228231 12228239 12228246 12228247 12228257 12228265 12228273 12228287 12228301 12228321 12228332 12228379 12228381 12228396 12228401 12228422 12228424 12228429 12228441 12228460 12228465 12228484 12228495 12228496 12228517 12228549 12228564 12228566 12228592 12228612 12228616 12228617 12228621 12228648 12228659 12228663 12228674 12228677 12228678 12228692 12228700 12228717 12228718 12228721 12228722 12228725 12228727 12228730 12228731 12228735 12228741 12228745 12228755 12228772 12228775 12228778 12228787 12228789 12228790 12228794 12228799 12228817 12228847 12228862 12228864 12228879 12228881 12228887 12228909 12228925 12228930 12228959 12228984 12228989 12228990 12229005 12229006 12229010 12229019 12229064 12229090 12229092 12229099 12229102 12229108 12229119 12229140 12229177 12229229 12229241 12229261 12229276 12229307 12229326 12229342 12229361 12229367 12229368 12229369 12229380 12229388 12229402 12229425 12229430 12229437 12229473 12229506 12229507 12229514 12229518 12229523 12229530 12229532 12229536 12229564 12229579 12229591 12229611 12229632 12229642 12229649 12229709 12229721 12229727 12229745 12229806 12229813 12229816 12229821 12229830 12229834 12229842 12229854 12229873 12229883 12229898 12229979 12230008 12230013 12230019 12230026 12230029 12230050 12230065 12230070 12230094 12230107 12230120 12230162 12230198 12230232 12230252 12230261 12230269 12230279 12230281 12230311 12230359 12230365 12230395 12230396 12230408 12230435 12230448 12230462 12230469 12230484 12230490 12230516 12230871 12230929 12230930 12230941 12230945 12230951 12230960 12230970 12230994 12231027 12231029 12231113 12231114 12231123 12231133 12231158 12231161 12231173 12231195 12231236 12231267 12231293 12231313 12231323 12231333 12231336 12231359 12231364 12231373 12231378 12231382 12231413 12231421 12231440 12231460 12231465 12231473 12231476 12231499 12231505 12231507 12231509 12231543 12231561 12231598 12231612 12231623 12231654 12231674 12231683 12231717 12231737 12231751 12231754 12231760 12231767 12231783 12231792 12231822 12231858 12231865 12231881 12231888 12231891 12231909 12231911 12231951 12231959 12231965 12231970 12231983 12231993 12231997 12232003 12232013 12232018 12232019 12232027 12232030 12232038 12232041 12232052 12232070 12232075 12232093 12232096 12232097 12232102 12232104 12232107 12232125 12232138 12232154 12232175 12232202 12232218 12232231 12232234 12232235 12232240 12232263 12232273 12232279 12232282 12232290 12232297 12232305 12232318 12232328 12232340 12232348 12232362 12232372 12232377 12232403 12232416 12232426 12232429 12232486 12232502 12232508 12232541 12232548 12232556 12232560 12232561 12232604 12232613 12232652 12232662 12232674 12232678 12232689 12232698 12232734 12232736 12232740 12232741 12232769 12232771 12232800 12232828 12232830 12232832 12232879 12232889 12232895 12232900 12232916 12232921 12232941 12232944 12232993 12233007 12233050 12233079 12233117 12233119 12233126 12233161 12233174 12233181 12233182 12233190 12233197 12233200 12233218 12233230 12233234 12233238 12233264 12233291 12233297 12233319 12233327 12233333 12233345 12233348 12233361 12233374 12233381 12233453 12233454 12233469 12233483 12233495 12233504 12233515 12233537 12233580 12233591 12233604 12233611 12233641 12233654 12233660 12233675 12233696 12233767 12233771 12233776 12233807 12233833 12233852 12233861 12233869 12233870 12233877 12233880 12233942 12233948 12233952 12233959 12233963 12233969 12234001 12234020 12234028 12234039 12234050 12234060 12234069 12234088 12234089 12234108 12234148 12234178 12234226 12234233 12234246 12234262 12234281 12234308 12234315 12234317 12234321 12234324 12234338 12234358 12234369 12234388 12234391 12234403 12234408 12234419 12234431 12234455 12234470 12234536 12234538 12234546 12234551 12234555 12234580 12234613 12234624 12234667 12234672 12234714 12234733 12234744 12234748 12234772 12234790 12234800 12234803 12234810 12234824 12234832 12234860 12234912 12234921 12234952 12234953 12234960 12234976 12234988 12235002 12235005 12235034 12235035 12235065 12235097 12235101 12235108 12235124 12235200 12235215 12235226 12235244 12235263 12235276 12235318 12235320 12235325 12235327 12235339 12235359 12235451 12235499 12235516 12235532 12235564 12235612 12235624 12235642 12235656 12235659 12235668 12235683 12235732 12235770 12235803 12235813 12235818 12235838 12235855 12235868 12235887 12235892 12235922 12235937 12236017 12236022 12236040 12236064 12236069 12236084 12236093 12236118 12236124 12236158 12236161 12236167 12236207 12236212 12236218 12236230 12236238 12236259 12236267 12236268 12236304 12236353 12236368 12236376 12236384 12236401 12236423 12236434 12236469 12236488 12236513 12236573 12236589 12236590 12236607 12236619 12236620 12236661 12236680 12236706 12236736 12236831 12236845 12236864 12236906 12236908 12236946 12237036 12237048 12237060 12237120 12237169 12237204 12237257 12237273 12237304 12237327 12237329 12237365 12237377 12237409 12237412 12237447 12237461 12237468 12237486 12237489 12237491 12237510 12237526 12237542 12237553 12237569 12237574 12237593 12237598 12237603 12237613 12237619 12237629 12237674 12237681 12237696 12237702 12237735 12237738 12237763 12237767 12237794 12237796 12237811 12237854 12237865 12237868 12237874 12237908 12237915 12237919 12237921 12237924 12237935 12237941 12237959 12238077 12238161 12238177 12238204 12238215 12238254 12238295 12238304 12238326 12238355 12238360 12238366 12238378 12238414 12238419 12238445 12238481 12238489 12238509 12238514 12238533 12238563 12238590 12238611 12238618 12238631 12238632 12238634 12238635 12238639 12238661 12238666 12238676 12238681 12238685 12238699 12238740 12238775 12238794 12238822 12238826 12238899 12238905 12238954 12238965 12238982 12239038 12239062 12239068 12239077 12239128 12239134 12239143 12239161 12239183 12239235 12239240 12239247 12239255 12239300 12239329 12239339 12239341 12239353 12239372 12239405 12239450 12239451 12239452 12239487 12239489 12239563 12239569 12239587 12239597 12239601 12239669 12239675 12239747 12239757 12239770 12239776 12239781 12239798 12239805 12239816 12239830 12239840 12239846 12239903 12239905 12239945 12239949 12239950 12240016 12240033 12240041 12240052 12240060 12240063 12240071 12240079 12240084 12240107 12240137 12240167 12240192 12240221 12240224 12240252 12240255 12240261 12240263 12240287 12240294 12240299 12240326 12240331 12240335 12240345 12240357 12240360 12240411 12240425 12240449 12240546 12240558 12240560 12240584 12240603 12240605 12240608 12240631 12240636 12240644 12240651 12240670 12240672 12240674 12240676 12240679 12240682 12240690 12240696 12240705 12240710 12240713 12240729 12240733 12240744 12240745 12240749 12240796 12240807 12240822 12240828 12240830 12240848 12240897 12240911 12240913 12240928 12240958 12240966 12241001 12241010 12241029 12241035 12241042 12241051 12241058 12241089 12241094 12241132 12241133 12241161 12241162 12241177 12241190 12241208 12241300 12241310 12241317 12241333 12241338 12241343 12241386 12241393 12241406 12241407 12241443 12241453 12241493 12241494 12241534 12241535 12241537 12241541 12241632 12241634 12241653 12241670 12241687 12241688 12241711 12241735 12241736 12241765 12241769 12241784 12241788 12241807 12241811 12241843 12241851 12241874 12241883 12241951 12241973 12241996 12241999 12242024 12242025 12242038 12242053 12242094 12242096 12242127 12242133 12242144 12242163 12242178 12242219 12242262 12242265 12242352 12242377 12242425 12242428 12242433 12242440 12242453 12242476 12242526 12242545 12242548 12242549 12242555 12242568 12242570 12242585 12242606 12242631 12242649 12242650 12242680 12242692 12242699 12242702 12242714 12242734 12242798 12242804 12242808 12242815 12242829 12242836 12242882 12242885 12242907 12242916 12242918 12242925 12242929 12242955 12242966 12242979 12242985 12242999 12243002 12243052 12243086 12243093 12243098 12243099 12243104 12243122 12243134 12243154 12243177 12243183 12243185 12243188 12243192 12243193 12243194 12243201 12243214 12243220 12243222 12243229 12243291 12243298 12243299 12243312 12243318 12243321 12243333 12243372 12243376 12243413 12243419 12243427 12243442 12243447 12243457 12243463 12243472 12243478 12243481 12243491 12243526 12243542 12243551 12243566 12243596 12243615 12243631 12243636 12243650 12243661 12243667 12243671 12243678 12243685 12243690 12243699 12243701 12243726 12243729 12243735 12243742 12243746 12243749 12243754 12243798 12243822 12243848 12243851 12243873 12243912 12243939 12243968 12243972 12243973 12243975 12243988 12243995 12244022 12244035 12244042 12244056 12244080 12244090 12244091 12244145 12244226 12244244 12244249 12244258 12244270 12244291 12244306 12244363 12244374 12244388 12244389 12244400 12244431 12244439 12244555 12244608 12244694 12244699 12244706 12244707 12244710 12244721 12244729 12244732 12244735 12244747 12244764 12244778 12244790 12244929 12244939 12244962 12245013 12245292 12245353 12245360 12245364 12245458 12245544 12245576 12246062 12246195 12246400 12246961 12246963 12246997 12247021 12247037 12247040 12247057 12247085 12247178 12247199 12247299 12247309 12247312 12247315 12247331 12247371 12247399 12247400 12247414 12247421 12247429 12247443 12247466 12247467 12247470 12247479 12247494 12247508 12247516 12247545 12247554 12247564 12247575 12247578 12247593 12247614 12247619 12247622 12247627 12247633 12247659 12247684 12247727 12247728 12247740 12247758 12247798 12247799 12247809 12247823 12247868 12247885 12247894 12247907 12247999 12248029 12248043 12248075 12248088 12248098 12248099 12248109 12248112 12248146 12248151 12248172 12248176 12248195 12248207 12248209 12248279 12248404 12248406 12248453 12248459 12248473 12248488 12248489 12248531 12248547 12248559 12248580 12248610 12248645 12248652 12248655 12248694 12248712 12248721 12248727 12248758 12248759 12248779 12248798 12248801 12248819 12248831 12248844 12248855 12248868 12248869 12248908 12248914 12248948 12248966 12248968 12248974 12249018 12249019 12249054 12249081 12249087 12249121 12249124 12249127 12249133 12249153 12249188 12249193 12249198 12249203 12249212 12249227 12249265 12249283 12249298 12249302 12249320 12249321 12249384 12249399 12249403 12249406 12249431 12249434 12249460 12249471 12249484 12249512 12249516 12249518 12249553 12249591 12249592 12249665 12249675 12249699 12249718 12249766 12249776 12249779 12249805 12249806 12249825 12249836 12249847 12249881 12249882 12249928 12249933 12249953 12249968 12249982 12250003 12250021 12250027 12250029 12250038 12250057 12250079 12250112 12250164 12250176 12250177 12250180 12251449 12251474 12251490 12251505 122 |  | SoftwareChangeControl
89 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138620 | Z-2258-2015 | Computed Tomography X-ray Systems Ingenuity Core 128.******Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Computed Tomography X-ray Systems Ingenuity Core 128 | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict | Class 2 Device Recall Computed Tomography Xray Systems Ingenuity Core 128 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Computed Tomography Xray Systems Ingenuity Core 128 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Computed Tomography Xray Systems Ingenuity Core 128 Date Initiated by Firm February 05, 2015 Date Posted July 24, 2015 Recall Status 1 Terminated 3 on June 26, 2018 Recall Number Z-2258-2015 Recall Event ID 71124 510(K)Number K033326 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Computed Tomography X-ray Systems Ingenuity Core 128.************Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Model #728323; Serial #s: 32010, 32013, 32015, 32016, 32018, 32021, 32024, 32026, 32032, 32033, 32034, 32036, 32037, 32041, 32042, 32043, 32044, 32046, 32047, 32048, 32049, 32051, 32052, 32053, 32054, 32055, 32057, 32059, 32060, 32062, 32064, 32065, 32067, 32068, 32069, 32070, 32072, 32073, 32074, 32075, 32077, 32078, 32079, 32080, 32082, 32083, 32084, 32085, 32087, 32089, 32090, 32091, 32092, 32094, 32095, 32100, 32101, 32103, 32104, 32106, 300137, 320003, 320005, 320006, 320018, 320024, 320029, 320030, 320033, 320037, 320054, 320069, 320073, 320074, 320078, 320079, 320081, 320082, 320083, 320084, 320088, 320090, 320092, 320096, 320099, 320101, 320103, 320107, 320111, 320114, 320115, 320119, 320122, 320123, 320125, 320126, 320129, 320131, 320132, 320133, 320134, 320138, 320144, 320146, 320147, 336011, 336012, 336013, 336016, 336018, 336028, 336029, 336034, 336044, 336046, 336059, 336061, 336064, 336066, 336067, 336068, 336069, 336070, 336071, 336072, 336073, 336075, 336076, 336077, 336078, 336080, 336082, 336083, 336086, 336087, 336088, 336089, 336090 & 336092. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mr. Yuchol Kim 440-483-2997 Manufacturer Reason for Recall 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict FDA Determined Cause 2 Software change control Action On 2/5/2015 the firm sent Customer Information Letters to their customers. Quantity in Commerce 139 units Distribution Worldwide Distribution-US (nationwide) including the states of AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV & WY, and to the following countries: Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 2/5/2015 the firm sent Customer Information Letters to their customers. | SoftwareChangeControl
90 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138615 | Z-2255-2015 | Computed Tomography X-ray systems Brilliance iCT, Philips Healthcare System.******Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Computed Tomography X-ray systems Brilliance iCT | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict | Class 2 Device Recall Computed Tomography Xray systems Brilliance iCT Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Computed Tomography Xray systems Brilliance iCT FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Computed Tomography Xray systems Brilliance iCT Date Initiated by Firm February 05, 2015 Date Posted July 24, 2015 Recall Status 1 Terminated 3 on June 26, 2018 Recall Number Z-2255-2015 Recall Event ID 71124 510(K)Number K060937 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Computed Tomography X-ray systems Brilliance iCT, Philips Healthcare System.************Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Brilliance iCT: Model #728306; Serial #***s: 85012, 85013, 85014, 100029, 100038, 100040, 100053, 100058, 100094, 100099, 100133, 100140, 100162, 100178, 100236, 100244, 100373, 100399, 100408, 100411, 100416, 100441, 100450, 100454, 100461, 100463, 100465, 100466, 100498, 100503, 100506, 100507, 100508, 100509, 100511, 100512, 100517 & 100518. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mr. Yuchol Kim 440-483-2997 Manufacturer Reason for Recall 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict FDA Determined Cause 2 Software change control Action On 2/5/2015 the firm sent Customer Information Letters to their customers. Quantity in Commerce 38 units Distribution Worldwide Distribution-US (nationwide) including the states of AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV & WY, and to the following countries: Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 2/5/2015 the firm sent Customer Information Letters to their customers. | SoftwareChangeControl
91 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138618 | Z-2256-2015 | Computed Tomography X-ray Systems Brilliance iCT SP.******Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Computed Tomography X-ray Systems Brilliance iCT SP | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict | Class 2 Device Recall Computed Tomography Xray Systems Brilliance iCT SP Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Computed Tomography Xray Systems Brilliance iCT SP FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Computed Tomography Xray Systems Brilliance iCT SP Date Initiated by Firm February 05, 2015 Date Posted July 24, 2015 Recall Status 1 Terminated 3 on June 26, 2018 Recall Number Z-2256-2015 Recall Event ID 71124 510(K)Number K060937 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Computed Tomography X-ray Systems Brilliance iCT SP.************Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Model #728311; Serial #s: 200047 & 200121 Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mr. Yuchol Kim 440-483-2997 Manufacturer Reason for Recall 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict FDA Determined Cause 2 Software change control Action On 2/5/2015 the firm sent Customer Information Letters to their customers. Quantity in Commerce 2 units Distribution Worldwide Distribution-US (nationwide) including the states of AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV & WY, and to the following countries: Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 2/5/2015 the firm sent Customer Information Letters to their customers. | SoftwareChangeControl
92 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=136169 | Z-2254-2015 | Computed Tomography X-ray Systems Brilliance CT 64-channel with Essence technology.******Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Computed Tomography X-ray Systems Brilliance CT | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict | Class 2 Device Recall Computed Tomography Xray Systems Brilliance CT Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Computed Tomography Xray Systems Brilliance CT FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Computed Tomography Xray Systems Brilliance CT Date Initiated by Firm February 05, 2015 Date Posted July 24, 2015 Recall Status 1 Terminated 3 on June 26, 2018 Recall Number Z-2254-2015 Recall Event ID 71124 510(K)Number K033326 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Computed Tomography X-ray Systems Brilliance CT 64-channel with Essence technology.************Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Brilliance CT 64-channel with Essence technology: Model #728231; Serial #s: 9654, 95130, 95157 & 95504. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mr. Yuchol Kim 440-483-2997 Manufacturer Reason for Recall 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict FDA Determined Cause 2 Software change control Action On 2/5/2015 the firm sent Customer Information Letters to their customers. Quantity in Commerce 4 Units Distribution Worldwide Distribution-US (nationwide) including the states of AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV & WY, and to the following countries: Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 2/5/2015 the firm sent Customer Information Letters to their customers. | SoftwareChangeControl
93 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138621 | Z-2259-2015 | Computed Tomography X-ray Systems Ingenuity CT scanners.******Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Computed Tomography X-ray Systems Ingenuity CT | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict | Class 2 Device Recall Computed Tomography Xray Systems Ingenuity CT Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Computed Tomography Xray Systems Ingenuity CT FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Computed Tomography Xray Systems Ingenuity CT Date Initiated by Firm February 05, 2015 Date Posted July 24, 2015 Recall Status 1 Terminated 3 on June 26, 2018 Recall Number Z-2259-2015 Recall Event ID 71124 510(K)Number K033326 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Computed Tomography X-ray Systems Ingenuity CT scanners.************Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Model #728326; Serial #s: 30003, 30004, 30005, 30006, 30007, 30008, 30009, 30011, 30012, 30013, 30015, 30018, 30019, 30020, 30025, 32086, 32093, 32107, 52023, 52044, 52051, 300003, 300004, 300005, 300010, 300011, 300012, 300013, 300014, 300016, 300017, 300019, 300020, 300021, 300024, 300025, 300027, 300028, 300029, 300031, 300032, 300033, 300034, 300035, 300036, 300037, 300039, 300040, 300041, 300042, 300043, 300044, 300045, 300047, 300049, 300050, 300051, 300052, 300053, 300054, 300055, 300056, 300057, 300058, 300059, 300060, 300061, 300062, 300063, 300065, 300066, 300068, 300069, 300070, 300071, 300072, 300073, 300074, 300075, 300077, 300078, 300079, 300080, 300081, 300082, 300084, 300086, 300087, 300088, 300089, 300090, 300091, 300092, 300093, 300094, 300095, 300096, 300097, 300098, 300099, 300100, 300101, 300102, 300103, 300104, 300105, 300106, 300107, 300108, 300110, 300111, 300112, 300113, 300114, 300115, 300116, 300117, 300118, 300119, 300120, 300121, 300122, 300123, 300124, 300125, 300126, 300127, 300128, 300129, 300130, 300131, 300132, 300135, 300136, 300139, 300140, 300142, 300143, 300144, 300145, 300146, 300147, 300148, 300149, 300150, 300151, 300156, 300157, 300158, 300159, 300160, 300161, 300162, 300163, 300164, 300165, 300166, 310058 & 30004. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mr. Yuchol Kim 440-483-2997 Manufacturer Reason for Recall 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict FDA Determined Cause 2 Software change control Action On 2/5/2015 the firm sent Customer Information Letters to their customers. Quantity in Commerce 159 units Distribution Worldwide Distribution-US (nationwide) including the states of AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV & WY, and to the following countries: Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 2/5/2015 the firm sent Customer Information Letters to their customers. | SoftwareChangeControl
94 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138619 | Z-2257-2015 | Computed Tomography X-ray Systems Ingenuity Core.******Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes.****** | Computed Tomography X-ray Systems Ingenuity Core | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict | Class 2 Device Recall Computed Tomography Xray Systems Ingenuity Core Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Computed Tomography Xray Systems Ingenuity Core FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Computed Tomography Xray Systems Ingenuity Core Date Initiated by Firm February 05, 2015 Date Posted July 24, 2015 Recall Status 1 Terminated 3 on June 26, 2018 Recall Number Z-2257-2015 Recall Event ID 71124 510(K)Number K033326 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Computed Tomography X-ray Systems Ingenuity Core.************Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information Model #728321; Serial #s: 31001, 52007, 52008, 52010, 52012, 52013, 52018, 52020, 52021, 52024, 52026, 52030, 52031, 52032, 52033, 52034, 52035, 52036, 52037, 52038, 52039, 52042, 52043, 52046, 52052, 52053, 52054, 52055, 52057, 52058, 310001, 310002, 310003, 310009, 310017, 310052, 310055, 310059, 310064, 310069, 310070, 310071, 310081, 310087, 310088, 310096, 310111, 310113, 310119, 310121, 310129, 310133, 310134, 310141, 310143, 310144, 310149, 310151, 310157, 310160, 310170, 333002, 333009, 333014, 333015, 333030, 333031, 333037, 333042, 333043, 333044, 333045, 333046, 333051, 333052, 333053, 333055, 333056, 333057, 333059, 333060 & 333061. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mr. Yuchol Kim 440-483-2997 Manufacturer Reason for Recall 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict FDA Determined Cause 2 Software change control Action On 2/5/2015 the firm sent Customer Information Letters to their customers. Quantity in Commerce 82 units Distribution Worldwide Distribution-US (nationwide) including the states of AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV & WY, and to the following countries: Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan & Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 2/5/2015 the firm sent Customer Information Letters to their customers. | SoftwareChangeControl
95 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=133315 | Z-1125-2015 | Alive ECG App 2.1.2 ( a medical device application for the Apple iOS operating system, intended to be used with the AliveCor Heart Monitor. The AliveCor Heart Monitor is intended to record, store and transfer single-channel electrocardiogram (ECG) rhythms. The AliveCor Heart Monitor also displays ECG rhythms and detects the presence of atrial fibrillation and normal sinus rhythm (when prescribed or used under the care of a physician). | AliveECG App Version 2.1.2 used with AliveCor Heart Monitor | 3 | CDRH | 2015-02-18 | 2015-02-18 21:00 | 2015-02-18 | Alivecor SFO | Alive ECG App version 2.1.2 (intended to be used with the AliveCor Heart Monitor) crashed upon use of the application. | Class 3 Device Recall AliveECG App Version 2.1.2 used with AliveCor Heart Monitor Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall AliveECG App Version 2.1.2 used with AliveCor Heart Monitor FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall AliveECG App Version 2.1.2 used with AliveCor Heart Monitor Date Initiated by Firm January 09, 2015 Date Posted February 18, 2015 Recall Status 1 Terminated 3 on February 18, 2015 Recall Number Z-1125-2015 Recall Event ID 70363 510(K)Number K140933 Product Classification Transmitters and receivers, electrocardiograph, telephone - Product Code DXH Product Alive ECG App 2.1.2 ( a medical device application for the Apple iOS operating system, intended to be used with the AliveCor Heart Monitor. The AliveCor Heart Monitor is intended to record, store and transfer single-channel electrocardiogram (ECG) rhythms. The AliveCor Heart Monitor also displays ECG rhythms and detects the presence of atrial fibrillation and normal sinus rhythm (when prescribed or used under the care of a physician). Code Information v2.1.2 Recalling Firm/ Manufacturer Alivecor SFO 30 Maiden Ln Suite 600 San Francisco CA 94108-5429 For Additional Information Contact Albert Boniske 855-338-8800 Manufacturer Reason for Recall Alive ECG App version 2.1.2 (intended to be used with the AliveCor Heart Monitor) crashed upon use of the application. FDA Determined Cause 2 Software change control Action AliveCor posted information on their website, Facebook Page and Twitter to alert users of the issue. Quantity in Commerce 5600 active users with Alive ECG app for iOS . Distribution Downloaded by Apple users - locations not shared by Apple. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = DXH - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action AliveCor posted information on their website, Facebook Page and Twitter to alert users of the issue. | SoftwareChangeControl
96 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131698 | Z-0856-2015 | Lite DM (Lite Data Manager)***Software Version(s): Versions 2.0.0.19 and lower***Product Model #: D00A00224D******The Lite DM (Lite Data Manager) is a software application developed by HORIBA***Medical intended to receive, store, and manage Startup, Patient, and Quality Control***(QC) results from an ABX Micros 60 Hematology Analyzer or ABX Micros CRP 200***Hematology Analyzer. Only one instrument connection is allowed per each Lite DM***workstation. The Lite DM application runs on a PC within a Windows Operating***System platform. Data transmitted from the instrument to the Lite DM workstation is a Serial RS-232 unidirectional ABX Format communication. The Lite DM can act as a standalone data manager for data received from the instrument, or allow transmission of this data to a Host (LIS  Laboratory Information System or EMR  Electronic Medical Record system).******Following the completion of each blood sample analysis performed on the connected HORIBA Medical hematology analyzer, results are then communicated (uni-directional) to the Lite DM workstation. Results received will display to the user in a specific color coding format indicating any normal or panic reference range violations. Results are then rejected or accepted by the user for final reporting - formatted to guidelines stipulated by CLIA. At no time are any results interpreted; only reported for review and appropriate action. | Lite DM (Lite Data Manager) | 2 | CDRH | 2014-12-23 | 2014-12-23 21:00 | 2015-06-22 | Horiba Instruments Inc | HORIBA Medical is informing all HORIBA Medical Lite DM customers using software version 2.0.0.19 and lower of a software bug related to the potential for results received being applied to the wrong order in the Lite DM Worklist. This bug has not been seen in software versions above V2.0.0.19******IMPACT: Potential for results received being applied to the wrong order in the Lite DM Worklist. The is | Class 2 Device Recall Lite DM (Lite Data Manager) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Lite DM (Lite Data Manager) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Lite DM (Lite Data Manager) Date Initiated by Firm November 19, 2014 Date Posted December 23, 2014 Recall Status 1 Terminated 3 on June 22, 2015 Recall Number Z-0856-2015 Recall Event ID 69847 Product Classification Medical device data system - Product Code OUG Product Lite DM (Lite Data Manager)******Software Version(s): Versions 2.0.0.19 and lower******Product Model #: D00A00224D************The Lite DM (Lite Data Manager) is a software application developed by HORIBA******Medical intended to receive, store, and manage Startup, Patient, and Quality Control******(QC) results from an ABX Micros 60 Hematology Analyzer or ABX Micros CRP 200******Hematology Analyzer. Only one instrument connection is allowed per each Lite DM******workstation. The Lite DM application runs on a PC within a Windows Operating******System platform. Data transmitted from the instrument to the Lite DM workstation is a Serial RS-232 unidirectional ABX Format communication. The Lite DM can act as a standalone data manager for data received from the instrument, or allow transmission of this data to a Host (LIS *** Laboratory Information System or EMR *** Electronic Medical Record system).************Following the completion of each blood sample analysis performed on the connected HORIBA Medical hematology analyzer, results are then communicated (uni-directional) to the Lite DM workstation. Results received will display to the user in a specific color coding format indicating any normal or panic reference range violations. Results are then rejected or accepted by the user for final reporting - formatted to guidelines stipulated by CLIA. At no time are any results interpreted; only reported for review and appropriate action. Code Information Lot/Serial #(s): N/A Software Version(s): Versions 2.0.0.19 and lower. Expiration Date (s): N/A Recalling Firm/ Manufacturer Horiba Instruments Inc 9755 Research Dr Irvine CA 92618-4626 For Additional Information Contact 949-453-0500 Manufacturer Reason for Recall HORIBA Medical is informing all HORIBA Medical Lite DM customers using software version 2.0.0.19 and lower of a software bug related to the potential for results received being applied to the wrong order in the Lite DM Worklist. This bug has not been seen in software versions above V2.0.0.19************IMPACT: Potential for results received being applied to the wrong order in the Lite DM Worklist. The is FDA Determined Cause 2 Software change control Action ACTION/RESOLUTION:******Before accepting any results please verify the Sample ID in the Worklist matches the Sample ID of the Patient information displayed for the results (Figure 1 below). If the two numbers do not match please contact HORIBA Medical Technical Support.************Your Lite DM should be upgraded to V3.0.0.21. Please contact HORIBA Medical Technical Support to schedule an upgrade. Please call the following phone # 1(888) 903-5001 prompt # 3.************Please share this information with your laboratory staff and retain this notification as part of your laboratory documentation. It is mandatory for you to complete and return the enclosed response form within 10 days, in order for us to verify your******understanding of this Field Information Notice.************If you have any questions or concerns, please contact your local HORIBA Medical representative. We sincerely apologize for any inconvenience this may cause your laboratory. Quantity in Commerce 322 (5 in U.S. Military) Distribution Nationwide within the U.S. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action ACTION/RESOLUTION:******Before accepting any results please verify the Sample ID in the Worklist matches the Sample ID of the Patient information displayed for the results (Figure 1 below). If the two numbers do not match please contact HORIBA Medical Technical Support.************Your Lite DM should be upgraded to V3.0.0.21. Please contact HORIBA Medical Technical Support to schedule an upgrade. Please call the following phone # 1(888) 903-5001 prompt # 3.************Please share this information with your laboratory staff and retain this notification as part of your laboratory documentation. It is mandatory for you to complete and return the enclosed response form within 10 days, in order for us to verify your******understanding of this Field Information Notice.************If you have any questions or concerns, please contact your local HORIBA Medical representative. We sincerely apologize for any inconvenience this may cause your laboratory. | SoftwareChangeControl
97 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=129827 | Z-2708-2014 | Dako Autostainer Link 48 with software version Dako Link 4.0.3, Model No. AS480******Product Usage: For in vitro diagnostic use. The Dako Autostainer Link 48 is an automated slide processing system compatible with reagents for the staining of paraffin-embedded and frozen tissue sections, cytospins, cell smears and fine needle aspirates. The system is designed to automate manual staining methods routinely used in immunohistochemistry and cytochemistry, enabling the transfer of established protocols from the bench to the Autostainer Link 48 automated workstations operating either singly or in networked configurations. | Dako Autostainer Link 48 software version 4.0.3 | 2 | CDRH | 2014-09-25 | 2014-09-26 21:00 | 2015-03-03 | Dako North America Inc. | Dako is recalling the Dako Link 4.0.3 software on the Dako Autostainer Link 48 because Dako has determined that an upgrade to this version of the software may cause premature rejection of reagent bottles or inadequate aspiration of reagent if the instrument has also had a preventative maintenance or waste pump replacement, and the baseplate is not completely aligned. | Class 2 Device Recall Dako Autostainer Link 48 software version 4.0.3 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Dako Autostainer Link 48 software version 4.0.3 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Dako Autostainer Link 48 software version 4.0.3 Date Initiated by Firm August 29, 2014 Date Posted September 26, 2014 Recall Status 1 Terminated 3 on March 03, 2015 Recall Number Z-2708-2014 Recall Event ID 69191 Product Classification Slide stainer, automated - Product Code KPA Product Dako Autostainer Link 48 with software version Dako Link 4.0.3, Model No. AS480************Product Usage: For in vitro diagnostic use. The Dako Autostainer Link 48 is an automated slide processing system compatible with reagents for the staining of paraffin-embedded and frozen tissue sections, cytospins, cell smears and fine needle aspirates. The system is designed to automate manual staining methods routinely used in immunohistochemistry and cytochemistry, enabling the transfer of established protocols from the bench to the Autostainer Link 48 automated workstations operating either singly or in networked configurations. Code Information Serial Numbers: AS1101D0902 AS1722D1009 AS2404D1203 AS2710D1211 AS1200D0904 AS1298D0906 AS1420D0908 AS1606D1004 AS1608D1004 AS1610D1004 AS1406D0908 AS1498D0910 AS1511D0910 AS1542D1001 AS1566D1002 AS1684D1008 AS1685D1008 AS1708D1008 AS1767D1010 AS1779D1010 AS1802D1010 AS1865D1011 AS2002D1105 AS2039D1106 AS2040D1106 AS2041D1106 AS2042D1106 AS2044D1106 AS2052D1106 AS2057D1107 AS2058D1107 AS2071D1107 AS2701D1211 AS2707D1211 AS48-0363-01 AS1251D0905 AS1149D0903 AS1216D0904 AS1270D0905 AS1271D0905 AS1311D0906 AS1314D0906 AS1356D0907 AS1363D0907 AS1691D1008 AS1883D1101 AS1934D1103 AS2109D1108 AS2111D1108 AS2141D1108 AS2158D1109 AS2159D1109 AS2170D1109 AS2182D1109 AS2192D1110 AS2300D1111 AS2358D1202 AS2359D1202 AS2361D1202 AS2416D1204 AS2504D1206 AS2516D1206 AS2678D1210 AS48-0196-01 AS48-0233-01 AS48-0234-01 AS48-0236-01 AS48-0237-01 AS48-0306-01 AS48-0332-01 AS48-0344-01 AS1025D0809 AS1027D0809 AS1029D0809 AS1030D0809 AS1032D0809 AS1072D0811 AS1185D0904 AS1186D0904 AS1187D0904 AS1418D0908 AS1220D0904 AS1292D0906 AS1354D0907 AS1472D0909 AS1473D0909 AS1556D1001 AS1817D1011 AS1904D1102 AS2036D1106 AS2353D1202 AS2357D1202 AS48-0273-01 AS48-0337-01 AS1266D0905 AS1624D1005 AS1695D1008 AS1560D1001 AS1603D1004 AS1636D1005 AS1941D1103 AS2230D1110 AS2234D1110 AS2235D1110 AS2236D1110 AS2308D1111 AS2542D1207 AS1342D0907 AS1446D0909 AS48-0114-01 AS1293D0906 AS2110D1108 AS1161D0903 AS1162D0903 AS1427D0909 AS2147D1109 AS2148D1109 AS2418D1204 AS1431D0909 AS1014D0809 AS1093D0901 AS1190D0904 AS1210D0904 AS48-0309-01 AS48-0310-01 AS48-0317-01 AS48-0318-01 AS48-0326-01 AS1052D0811 AS1094D0901 AS1109D0902 AS1123D0903 AS1127D0903 AS1138D0903 AS1140D0903 AS1238D0904 AS1239D0904 AS1240D0905 AS1301D0906 AS1320D0906 AS1322D0906 AS1331D0907 AS1332D0907 AS1376D0907 AS1508D0910 AS1520D0910 AS1528D0911 AS1551D1001 AS1637D1006 AS1663D1007 AS1681D1007 AS1682D1007 AS1712D1008 AS1732D1009 AS1749D1009 AS1765D1010 AS1775D1010 AS1778D1010 AS1800D1010 AS1837D1011 AS1839D1011 AS1848D1011 AS1850D1011 AS1855D1011 AS1866D1012 AS1867D1012 AS1880D1101 AS1975D1104 AS1993D1105 AS2030D1106 AS2055D1106 AS48-0150-01 AS48-0152-01 AS48-0156-01 AS48-0166-01 AS48-0167-01 AS48-0177-01 AS48-0192-01 AS48-0205-01 AS48-0207-01 AS48-0213-01 AS48-0248-01 AS48-0259-01 AS48-0296-01 AS48-0315-01 AS2761D1301 AS48-0129-01 AS2169D1109 AS1820D1011 AS1584D1003 AS2014D1105 Recalling Firm/ Manufacturer Dako North America Inc. 6392 Via Real Carpinteria CA 93013-2921 For Additional Information Contact 805-566-6655 Manufacturer Reason for Recall Dako is recalling the Dako Link 4.0.3 software on the Dako Autostainer Link 48 because Dako has determined that an upgrade to this version of the software may cause premature rejection of reagent bottles or inadequate aspiration of reagent if the instrument has also had a preventative maintenance or waste pump replacement, and the baseplate is not completely aligned. FDA Determined Cause 2 Software change control Action Dako North America sent an Urgent Medical Device Correction letter to all customers who were upgraded to the Dako Link 4.0.3 software that is used on the Dako Autostainer Link 48 and who also had preventative maintenance or a waste pump replacement. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed that Dako will contact customers to set up an appointment for a service visit in order to ensure that the instruments are running safely or reliably. Customers with questions are instructed to contact Dako. Quantity in Commerce 191 units total (57 units in US) Distribution Worldwide Distribution - US Nationwide in the states of AT, AU, BE, CA, CN, DE, DK, ES, Fl, FR, GB, IE, IT, NL, NO, SE, US, DZ, GR, IS, OM, SG, UY. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Dako North America sent an Urgent Medical Device Correction letter to all customers who were upgraded to the Dako Link 4.0.3 software that is used on the Dako Autostainer Link 48 and who also had preventative maintenance or a waste pump replacement. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed that Dako will contact customers to set up an appointment for a service visit in order to ensure that the instruments are running safely or reliably. Customers with questions are instructed to contact Dako. | SoftwareChangeControl
98 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=128604 | Z-2109-2014 | QLAB 10 Core Module PN 453561704771. QLAB Quantification Software is a software application package designed to view and quantify image data acquired on Philips Medical Systems ultrasound products. | Philips Ultrasound QLAB | 2 | CDRH | 2014-07-23 | 2014-07-23 21:00 | 2015-11-06 | Philips Ultrasound, Inc. | The QLAB a2DQ and aCMQ features have a defect that may use incorrect measurement values under specific conditions. | Class 2 Device Recall Philips Ultrasound QLAB Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Ultrasound QLAB FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Ultrasound QLAB Date Initiated by Firm July 09, 2014 Date Posted July 23, 2014 Recall Status 1 Terminated 3 on November 06, 2015 Recall Number Z-2109-2014 Recall Event ID 68756 510(K)Number K152165 Product Classification System, image processing, radiological - Product Code LLZ Product QLAB 10 Core Module PN 453561704771. QLAB Quantification Software is a software application package designed to view and quantify image data acquired on Philips Medical Systems ultrasound products. Code Information *****SYSTEM BASED ID OUTSIDE THE US********** 120F-P213FA35-V41E, 279E-Q21384E8-P41D, 3FBF-X2649F25-J4BF, 47FD-Q1FD112B-J3F0, 4A8F-U24EA0B5-R493, 5570-Q20D1AB2-W410, 5D8C-G1FC9BDE-N3EF, 684B-V2122501-R41A, 6DA2-G232AD5C-S45B, 731B-Q1FC2691-V3EE, 808E-P20C3018-I40E, 88AA-T1FBB144-W3ED, A0CB-I24CCB81-M48F, A0ED-H1F5BC5B-P3E1, AB9B-S236CD11-W463, BE98-V1E4C83A-P3BF, C6A3-H1FFD0F9-X3F5, CBE9-M24BE0E7-Y48D, EEFC-Y22FED8E-S455, EF0D-I20465FB-I3FE, F475-U1F966C3-X3E8, and V49C-R203F0AE-I3FE. ***********Two units shipped in US and one unit shipped outside the US to distributors do not have serial numbers. There are licensing agreements with these distributors that allows the distributors to create their own licensing and serial numbers and/or tracking information. Recalling Firm/ Manufacturer Philips Ultrasound, Inc. 22100 Bothell Everett Hwy Bothell WA 98021-8431 For Additional Information Contact Technical Support 425-482-8466 Manufacturer Reason for Recall The QLAB a2DQ and aCMQ features have a defect that may use incorrect measurement values under specific conditions. FDA Determined Cause 2 Software change control Action Philips Medical Systems notified customers of the voluntary recall by sending a certified URGENT-Medical Device Correction letter, dated 2014 JUL 09, to all US and Canadian consignees. Philips representatives will distribute the letter internationally. ************Follow-up letters to wholesale customers will be provided and wholesalers are requested to forward copies of recall letters and response forms to their customers. Philips Medical Systems will also have Customer Reply Cards returned for tracking purposes. Customers are directed to call local Philips representative or Philips Customer Service at 1-866-767-7822. Quantity in Commerce 25 units total (2 units in the US and 23 units outside the US) Distribution Distributed in the states of CA, IA, KY, MD, NY, PA, UT, VA, WA and WI.******Also distributed in Australia, Belgium, Canada, India, Italy, Japan, South Korea, Netherlands, New Zealand, Spain, Sweden, and Switzerland. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Medical Systems notified customers of the voluntary recall by sending a certified URGENT-Medical Device Correction letter, dated 2014 JUL 09, to all US and Canadian consignees. Philips representatives will distribute the letter internationally. ************Follow-up letters to wholesale customers will be provided and wholesalers are requested to forward copies of recall letters and response forms to their customers. Philips Medical Systems will also have Customer Reply Cards returned for tracking purposes. Customers are directed to call local Philips representative or Philips Customer Service at 1-866-767-7822. | SoftwareChangeControl
99 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=128605 | Z-2110-2014 | QLAB 10.1 Core Module PN 453561728001. QLAB Quantification Software is a software application package designed to view and quantify image data acquired on Philips Medical Systems ultrasound products. | Philips Ultrasound QLAB | 2 | CDRH | 2014-07-23 | 2014-07-23 21:00 | 2015-11-06 | Philips Ultrasound, Inc. | The QLAB a2DQ and aCMQ features have a defect that may use incorrect measurement values under specific conditions. | Class 2 Device Recall Philips Ultrasound QLAB Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Ultrasound QLAB FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Ultrasound QLAB Date Initiated by Firm July 09, 2014 Date Posted July 23, 2014 Recall Status 1 Terminated 3 on November 06, 2015 Recall Number Z-2110-2014 Recall Event ID 68756 510(K)Number K152165 Product Classification System, image processing, radiological - Product Code LLZ Product QLAB 10.1 Core Module PN 453561728001. QLAB Quantification Software is a software application package designed to view and quantify image data acquired on Philips Medical Systems ultrasound products. Code Information ****SYSTEM BASED ID IN THE US*********************** 12DD-14B7BC8B-295F, 25EB-143A3021-2864, 3E1D-145FC2CB-28AF, 5914-14544E46-2898, 5936-13FD3F20-27EA, 91B6-1437E5A0-285F, 9C64-1478F656-28E1, A1DD-14426F8B-2874, AC8B-14838041-28F6, C501-13FAF49F-27E5, CCFB-144184F1-2872, and J57B-147C2B71-28E8. *****SYSTEM BASED ID OUTSIDE THE US************************ 189A-13D31774-2796, 1B4E-13CD97D8-278B, 2B31-1486400F-28FC, 2B42-145AB87C-28A5, 2E29-13D2A227-2795, 3099-147B40D7-28E6, 3B7A-1439BAD4-2863, 4BD4-13C1AE06-2773, 4E88-13BC2E6A-2768, 50E7-149054AD-2910, 512B-13E23661-27B4, 539B-148AD511-2905, 564F-14855575-28FA, 5682-1402BEBC-27F5, 5E5A-14A05E34-2930, 5E6B-1474D6A1-28D9, 611F-146F5705-28CE, 66CB-13B63981-275C, 6BCD-14B067BB-2950, 6ED6-13D14240-2792, 73E9-149FE8E7-292F, 76F2-13C0C36C-2771, 7951-1494E9AF-2919, 7C05-148F6A13-290E, 7C16-1463E280-28B7, 7C49-13E14BC7-27B2, 7C5A-13B5C434-275B, 7EA8-14B5720A-295A, 8465-13D0CCF3-2791, 8C5F-14175D45-281E, 8F35-13BACE83-2765, 91E9-13B54EE7-275A, A46F-1493FF15-2917, B74A-1499H964-2921, C4CE-147D8B58-28EA, D798-14AE1D3A-294B, D7A9-148295A7-28F4, DA5D-147D160B-28E9, DFB4-149D9E66-292A, DFC5-147216D3-28D3, E268-14981ECA-291F, E7BF-14B8A725-2960, EA95-145C1863-28A7, ED7C-13D4H20E-2797, F598-13C3833A-2776, FDA3-13DE8BF9-27AC, and XAF4-1445A4A6-287B. Recalling Firm/ Manufacturer Philips Ultrasound, Inc. 22100 Bothell Everett Hwy Bothell WA 98021-8431 For Additional Information Contact Technical Support 425-482-8466 Manufacturer Reason for Recall The QLAB a2DQ and aCMQ features have a defect that may use incorrect measurement values under specific conditions. FDA Determined Cause 2 Software change control Action Philips Medical Systems notified customers of the voluntary recall by sending a certified URGENT-Medical Device Correction letter, dated 2014 JUL 09, to all US and Canadian consignees. Philips representatives will distribute the letter internationally. ************Follow-up letters to wholesale customers will be provided and wholesalers are requested to forward copies of recall letters and response forms to their customers. Philips Medical Systems will also have Customer Reply Cards returned for tracking purposes. Customers are directed to call local Philips representative or Philips Customer Service at 1-866-767-7822. Quantity in Commerce 59 units total (12 units in the US and 47 units outside the US) Distribution Distributed in the states of CA, IA, KY, MD, NY, PA, UT, VA, WA and WI.******Also distributed in Australia, Belgium, Canada, India, Italy, Japan, South Korea, Netherlands, New Zealand, Spain, Sweden, and Switzerland. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Medical Systems notified customers of the voluntary recall by sending a certified URGENT-Medical Device Correction letter, dated 2014 JUL 09, to all US and Canadian consignees. Philips representatives will distribute the letter internationally. ************Follow-up letters to wholesale customers will be provided and wholesalers are requested to forward copies of recall letters and response forms to their customers. Philips Medical Systems will also have Customer Reply Cards returned for tracking purposes. Customers are directed to call local Philips representative or Philips Customer Service at 1-866-767-7822. | SoftwareChangeControl
100 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122026 | Z-0032-2014 | FSE Image Upgrade Kit, Catalog # 92019C******This kit is used to install and/or upgrade software used with the EVOLIS Microplate System. | Image Upgrade Kit, Catalog # 92019C | 3 | CDRH | 2013-10-21 | 2013-10-21 21:01 | 2013-11-13 | Bio-Rad Laboratories | The FSE Image Upgrade kit (part 92019C) is recalled due to issue related to the export of data from the EVOLIS analyzer to Laboratory Information Systems (LIS) when using the MONOLISA Anti-HBs EIA***Quantitative Determination in EVOLIS APF version 6.0. | Class 3 Device Recall Image Upgrade Kit, Catalog 92019C Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall Image Upgrade Kit, Catalog 92019C FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall Image Upgrade Kit, Catalog 92019C Date Initiated by Firm August 20, 2013 Date Posted October 21, 2013 Recall Status 1 Terminated 3 on November 13, 2013 Recall Number Z-0032-2014 Recall Event ID 66293 PMA Number P050048S004 Product Classification Test, hepatitis b (b core, be antigen, be antibody, b core igm) - Product Code LOM Product FSE Image Upgrade Kit, Catalog # 92019C************This kit is used to install and/or upgrade software used with the EVOLIS Microplate System. Code Information Lot 122822 Recalling Firm/ Manufacturer Bio-Rad Laboratories 6565 185th Ave NE Redmond WA 98052-5039 For Additional Information Contact Cindi Jinneman 800-224-6723 Manufacturer Reason for Recall The FSE Image Upgrade kit (part 92019C) is recalled due to issue related to the export of data from the EVOLIS analyzer to Laboratory Information Systems (LIS) when using the MONOLISA Anti-HBs EIA******Quantitative Determination in EVOLIS APF version 6.0. FDA Determined Cause 2 Software change control Action BioRad has sent an Urgent Product Correction letter dated August 19, 2013, to customers who had the EVOLIS FSE Upgrade Image Kit installed on the EVOLIS at their site by a Bio-Rad Laboratories Field Service Engineer (FSE).************This issue impacts laboratory only if you are running the MONOLISA Anti-HBs EIA Quantitative Determination method with EVOLIS APF version 6.0 and TimeLiner Data Management Software. If you use this Quantitative Determination method with EVOLIS APF version 6.0, please contact Bio-Rad Laboratories for additional information or assistance.************Bio-Rad Laboratories is currently developing and qualifying a software update CD consisting of the EVOLIS APF V6.0b and Timeliner Data Management Software to correct the issue. The software update CD will be installed by Bio-Rad Field Service Engineers at each of the customer sites when******installing the EVOLIS APF V6.0b.************Customers should contact Bio-Rad Laboratories at 1-800-224-6723, Option******2 and then Option 3 for any questions about this product correction. Quantity in Commerce 23 kits Distribution US Distribution including the states of FL, IA, ME, MO, MS, NE, NY, TN, TX, and WA Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LOM - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BioRad has sent an Urgent Product Correction letter dated August 19, 2013, to customers who had the EVOLIS FSE Upgrade Image Kit installed on the EVOLIS at their site by a Bio-Rad Laboratories Field Service Engineer (FSE).************This issue impacts laboratory only if you are running the MONOLISA Anti-HBs EIA Quantitative Determination method with EVOLIS APF version 6.0 and TimeLiner Data Management Software. If you use this Quantitative Determination method with EVOLIS APF version 6.0, please contact Bio-Rad Laboratories for additional information or assistance.************Bio-Rad Laboratories is currently developing and qualifying a software update CD consisting of the EVOLIS APF V6.0b and Timeliner Data Management Software to correct the issue. The software update CD will be installed by Bio-Rad Field Service Engineers at each of the customer sites when******installing the EVOLIS APF V6.0b.************Customers should contact Bio-Rad Laboratories at 1-800-224-6723, Option******2 and then Option 3 for any questions about this product correction. | SoftwareChangeControl
101 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121962 | Z-2318-2013 | McKesson Horizon Medical Imaging (HMI) Radiological Image Processing System, product versions 4.6.1 to 11.6., is a medical image and information management application intended to receive, transmit, store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from medical imaging systems. | McKesson Horizon Medical Imaging | 2 | CDRH | 2013-09-30 | 2013-09-30 21:00 | 2017-02-06 | Mckesson Information Solutions LLC | Software malfunction may occur resulting in a certain combination of server configurations which has the potential to cause image loss. | Class 2 Device Recall McKesson Horizon Medical Imaging Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall McKesson Horizon Medical Imaging FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall McKesson Horizon Medical Imaging Date Initiated by Firm September 04, 2013 Date Posted September 30, 2013 Recall Status 1 Terminated 3 on February 06, 2017 Recall Number Z-2318-2013 Recall Event ID 66301 510(K)Number K043146 Product Classification System, image processing, radiological - Product Code LLZ Product McKesson Horizon Medical Imaging (HMI) Radiological Image Processing System, product versions 4.6.1 to 11.6., is a medical image and information management application intended to receive, transmit, store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from medical imaging systems. Code Information Products: HMI 4.6.1; HMI 5.x (5.0.1, 5.0.5, 5.0.6, 5.0.7, 5.0.8); HMI 11.0x (11.0.5, 11.0.6, 11.0.6 FR EP2, 11.0.7, 11.0.8); HMI 11.5x (11.5, 11.5.1, 11.5.2) and HMI 11.6. Recalling Firm/ Manufacturer Mckesson Information Solutions LLC 5995 Winward Plaza Information Technology Business Alpharetta GA 30005 For Additional Information Contact Paul Sumner 404-338-2376 Manufacturer Reason for Recall Software malfunction may occur resulting in a certain combination of server configurations which has the potential to cause image loss. FDA Determined Cause 2 Software change control Action Consignees were notified via Urgent Field Safety Notice dated September 4. MMIC Customer support will be contacting customers known to have a multi-server configuration and will implement the configuration change. Quantity in Commerce 329 units Distribution Worldwide Distribution - United States including the territories of Guam and Puerto Rico and the countries of Canada and France. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Consignees were notified via Urgent Field Safety Notice dated September 4. MMIC Customer support will be contacting customers known to have a multi-server configuration and will implement the configuration change. | SoftwareChangeControl
102 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120033 | Z-1916-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 NO. Part number: 414565.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1916-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 NO. Part number: 414565.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
103 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120031 | Z-1914-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 IT. Part number: 414563.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1914-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 IT. Part number: 414563.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
104 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120032 | Z-1915-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 JA. Part number: 414564.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1915-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 JA. Part number: 414564.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
105 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120037 | Z-1920-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 ZH. Part number: 414569.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1920-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 ZH. Part number: 414569.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
106 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120038 | Z-1921-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit V2S RP5800XPE TO 6.01 WES7. Part number: 6201773.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. *** | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1921-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit V2S RP5800XPE TO 6.01 WES7. Part number: 6201773.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
107 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120028 | Z-1911-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 EN. Part number: 414560.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. *** | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1911-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 EN. Part number: 414560.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
108 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120027 | Z-1910-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 DE. Part number: 414539.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1910-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 DE. Part number: 414539.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
109 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120025 | Z-1908-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01ML1. Part number: 414537.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. *** | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1908-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01ML1. Part number: 414537.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
110 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120036 | Z-1919-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 EL, TR. Part number: 414568.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1919-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 EL, TR. Part number: 414568.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
111 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120026 | Z-1909-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01ML2. Part number: 414538.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1909-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01ML2. Part number: 414538.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
112 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120029 | Z-1912-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 ES. Part number: 414561.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1912-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 ES. Part number: 414561.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
113 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120030 | Z-1913-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 FR. Part number: 414562.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1913-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 FR. Part number: 414562.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
114 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120035 | Z-1918-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 DA, SV. Part number: 414567.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1918-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 DA, SV. Part number: 414567.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
115 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=119016 | Z-1907-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, V2S 6.01 RP5800-WES7 Media. Part number: 413861.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1907-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, V2S 6.01 RP5800-WES7 Media. Part number: 413861.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
116 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120034 | Z-1917-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 PT. Part number: 414566.******The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall bioMerieux Vitek 2 Systems Software version 6.01 software Date Initiated by Firm June 14, 2013 Date Posted August 06, 2013 Recall Status 1 Terminated 3 on June 01, 2015 Recall Number Z-1917-2013 Recall Event ID 65470 Product Classification Susceptibility test cards, antimicrobial - Product Code LTW Product bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 PT. Part number: 414566.************The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. Code Information version 6.01 Recalling Firm/ Manufacturer Biomerieux Inc 595 Anglum Rd Hazelwood MO 63042-2320 For Additional Information Contact Ryan LeMelle 314-731-8582 Manufacturer Reason for Recall A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table FDA Determined Cause 2 Software change control Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. Quantity in Commerce 12,384 total kits Distribution Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.************If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. | SoftwareChangeControl
117 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=119832 | Z-1814-2013 | Picis ED Pulsecheck - EMR Software Application - 2125, Software Versions: 5.2 and 5.3. The application stores patient information in a database, and it may analyze and/or display the data in different formats for evaluation by healthcare professionals for informational purposes. | ED PulseCheck | 2 | CDRH | 2013-07-29 | 2013-07-29 21:00 | 2014-12-19 | Picis Inc. | Notes associated with prescription are not printed to the prescription or to the patient chart. | Class 2 Device Recall ED PulseCheck Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ED PulseCheck FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ED PulseCheck Date Initiated by Firm June 21, 2013 Date Posted July 29, 2013 Recall Status 1 Terminated 3 on December 19, 2014 Recall Number Z-1814-2013 Recall Event ID 65666 Product Classification Software, transmission & storage, patient data - Product Code NSX Product Picis ED Pulsecheck - EMR Software Application - 2125, Software Versions: 5.2 and 5.3. The application stores patient information in a database, and it may analyze and/or display the data in different formats for evaluation by healthcare professionals for informational purposes. Code Information Software Versions 5.2 and 5.3 Recalling Firm/ Manufacturer Picis Inc. 100 Quannapowitt Parkway Suite 405 Wakefield MA 01880 For Additional Information Contact Support Representative 781-557-3000 Manufacturer Reason for Recall Notes associated with prescription are not printed to the prescription or to the patient chart. FDA Determined Cause 2 Software change control Action Initial customer notifications were sent via email on June 21, 2013 informing consignees of the recall and providing further instruction regarding the software solution. Consignees were provided with recommended actions until they receive the necessary update. Quantity in Commerce 35 Distribution Nationwide Distribution, including the states of: AK, AR, AZ, CA, CO, DC, DE, FL, GA, ID, IN, MA, MD, MO, NH, NJ, OH, OR, SC, TN, WA, and WV. 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Initial customer notifications were sent via email on June 21, 2013 informing consignees of the recall and providing further instruction regarding the software solution. Consignees were provided with recommended actions until they receive the necessary update. | SoftwareChangeControl
118 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=118569 | Z-1560-2013 | Tecan Freedom EVO with Touch Tools Suite Version 3.0 ******The intended use of the Touch Tools Suite is to run previously defined scripts and processes on a Freedom EVO platform (Touch Tools Suite is an optional software add-on for Freedom EVOware (the base operational software for the Freedom EVO 2 instrument. | Freedom EVO 2 with TouchTools Suite Version 3.0 | 2 | CDRH | 2013-06-19 | 2013-06-19 21:00 | 2013-06-20 | Tecan US, Inc. | WIth TouchTools 3.0 after the execution of a Rich User Prompt in Display and Wait mode, door lock errors are ignored. If an operator presses "Continue" following a Rich User Prompt and the door has been left open, the script will progress (continue without warning). This could lead to the instrument moving unexpectedly and potentially cause an injury to the operator if they are physically inside t | Class 2 Device Recall Freedom EVO 2 with TouchTools Suite Version 3.0 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Freedom EVO 2 with TouchTools Suite Version 3.0 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Freedom EVO 2 with TouchTools Suite Version 3.0 Date Initiated by Firm February 19, 2013 Date Posted June 19, 2013 Recall Status 1 Terminated 3 on June 20, 2013 Recall Number Z-1560-2013 Recall Event ID 65331 Product Classification Station, pipetting and diluting, for clinical use - Product Code JQW Product Tecan Freedom EVO with Touch Tools Suite Version 3.0 ************The intended use of the Touch Tools Suite is to run previously defined scripts and processes on a Freedom EVO platform (Touch Tools Suite is an optional software add-on for Freedom EVOware (the base operational software for the Freedom EVO 2 instrument. Code Information Instrument Freedom EVO 100 Base Unit: Material number: 10641100, Serial numbers: 1211001569, 1212000203 and 1212000205; Instrument Freedom EVO 100 MCA: Material number: 30020010, Serial number: 1301007643; Instrument Freedom EVO 150 Base Unit: Material number10641150, Serial number: 1301006708; Instrument Freedom EVO 150 MCA: Material number: 30020015, Serial number: 1211007442; Instrument Freedom EVO 150 MCA384: Material number: 30032015, Serial number: 1211011594 and 1109003270; Instrument Freedom EVO 200 Base Unit: Material number: 1641200, Serial number: 1106003923 and 1211011578 and Instrument Freedom EVO 200 MCA: Material number: 30020020, Serial number: 1208007288 and 1208007284. Recalling Firm/ Manufacturer Tecan US, Inc. 9401 Globe Center Drive Suite 140 Morrisville NC 27560 For Additional Information Contact Ms Melody Humbles 919-361-5200 Ext. 19519 Manufacturer Reason for Recall WIth TouchTools 3.0 after the execution of a Rich User Prompt in Display and Wait mode, door lock errors are ignored. If an operator presses "Continue" following a Rich User Prompt and the door has been left open, the script will progress (continue without warning). This could lead to the instrument moving unexpectedly and potentially cause an injury to the operator if they are physically inside t FDA Determined Cause 2 Software change control Action The firm, Tecan, sent an "Urgent Field Corrective Action" letter dated February 15, 2013 to its consignees/customers and also contact them by phone February 19, 2013. The letter describes the product, problem and actions to be taken. The customers were instructed to not use Rich User Prompts in Display and Wait mode until TouchTools 3.0 Patch 1 or a higher version has been installed. Tecan is currently testing and releasing a patch that will correct this issue. The patch will be available by the end of February. Tecan will contact you upon its release as installation of the patch is required. ************If you have any further questions, please contact your local Tecan Helpdesk for assistance or call 919-361-5200 ext 19519. Quantity in Commerce 12 units Distribution Worldwide distribution: US (nationwide) including states of: CA, KY, NC, OH, TX, and WI; and countries of: France, Germany and Italy. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, Tecan, sent an "Urgent Field Corrective Action" letter dated February 15, 2013 to its consignees/customers and also contact them by phone February 19, 2013. The letter describes the product, problem and actions to be taken. The customers were instructed to not use Rich User Prompts in Display and Wait mode until TouchTools 3.0 Patch 1 or a higher version has been installed. Tecan is currently testing and releasing a patch that will correct this issue. The patch will be available by the end of February. Tecan will contact you upon its release as installation of the patch is required. ************If you have any further questions, please contact your local Tecan Helpdesk for assistance or call 919-361-5200 ext 19519. | SoftwareChangeControl
119 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=117984 | Z-1419-2013 | RayStation Therapy Treatment Planning System, Model Numbers 2.0, 2.5, and 3.0; CLASSIFICATION NAME: System, Planning, Radiation Therapy Treatment.******RayStation is a software system designed for treatment planning and analysis of radiation therapy. | RayStation Therapy Treatment Planning System | 2 | CDRH | 2013-05-28 | 2013-05-28 21:00 | 2017-02-17 | RAYSEARCH LABORATORIES AB | Under some circumstances, imported CT, MR and PET images are offset from their true positions by one pixel in one or two directions. Users must be aware of this information to avoid incorrect dose calculations during treatment planning. | Class 2 Device Recall RayStation Therapy Treatment Planning System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall RayStation Therapy Treatment Planning System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall RayStation Therapy Treatment Planning System Date Initiated by Firm February 15, 2013 Date Posted May 28, 2013 Recall Status 1 Terminated 3 on February 17, 2017 Recall Number Z-1419-2013 Recall Event ID 65136 510(K)Number K120387 Product Classification System,planning,radiation therapy treatment - Product Code MUJ Product RayStation Therapy Treatment Planning System, Model Numbers 2.0, 2.5, and 3.0; CLASSIFICATION NAME: System, Planning, Radiation Therapy Treatment.************RayStation is a software system designed for treatment planning and analysis of radiation therapy. Code Information Software version and build numbers 2.0.0.15, 2.5.1.89, and 3.0.0.251 (to determine if the version used is affected, bring up the About box in the RayStation application and check the build number reported0. Recalling Firm/ Manufacturer RAYSEARCH LABORATORIES AB Sveavaegen 9 Stockholm Sweden Manufacturer Reason for Recall Under some circumstances, imported CT, MR and PET images are offset from their true positions by one pixel in one or two directions. Users must be aware of this information to avoid incorrect dose calculations during treatment planning. FDA Determined Cause 2 Software change control Action RaySearch Laboratories sent an Urgent Field Safety Notice, Medical Device Correction letter dated January 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to educate their planning staff and all users about the recommended workarounds contained in the letter. Customers were also asked to inspect their product and identify all installed units that were infected and then complete the attached Reply Form and e-mail or scan a copy of the completed form to the firm. Customers were instructed to pass along this information to all those who need to be aware of this Medical Device Correction Customers with questions were instructed to contact the RaySearch support email at support@raysearchlabs.******For questions regarding this recall call +46 8 545 06130. Quantity in Commerce Domestic: 70 units; Foreign: 27 units Distribution Worldwide Distribution - USA including Connecticut, Florida, Massachusetts, Missouri, New York, Ohio, and foreign distribution to(Canada, Italy, Australia, Germany, Netherlands, and Switzerland. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MUJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action RaySearch Laboratories sent an Urgent Field Safety Notice, Medical Device Correction letter dated January 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to educate their planning staff and all users about the recommended workarounds contained in the letter. Customers were also asked to inspect their product and identify all installed units that were infected and then complete the attached Reply Form and e-mail or scan a copy of the completed form to the firm. Customers were instructed to pass along this information to all those who need to be aware of this Medical Device Correction Customers with questions were instructed to contact the RaySearch support email at support@raysearchlabs.******For questions regarding this recall call +46 8 545 06130. | SoftwareChangeControl
120 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=117878 | Z-1339-2013 | LATITUDE¿ Patient Management System, Model 6488, Version 7.3.******The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. | LATITUDE¿ Patient Management System | 2 | CDRH | 2013-05-20 | 2013-05-20 21:00 | 2013-11-26 | Boston Scientific CRM Corp | Boston Scientific CRM is conducting a recall on the LATITUDE Patient Management System Model 6488 Version 7.3 because between March 4 and 27, 2013, a Siebel release caused ~5000 distributed PMR clinic faxes to display with illegible characters and not the required, Patient non-compliance information that was intended. | Class 2 Device Recall LATITUDE Patient Management System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall LATITUDE Patient Management System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall LATITUDE Patient Management System Date Initiated by Firm April 10, 2013 Date Posted May 20, 2013 Recall Status 1 Terminated 3 on November 26, 2013 Recall Number Z-1339-2013 Recall Event ID 65096 PMA Number P910077S123 Product Classification Implantable cardioverter defibrillator (non-CRT) - Product Code LWS Product LATITUDE Patient Management System, Model 6488, Version 7.3.************The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. Code Information n/a Recalling Firm/ Manufacturer Boston Scientific CRM Corp 4100 Hamline Ave N Saint Paul MN 55112-5700 For Additional Information Contact 651-582-4000 Manufacturer Reason for Recall Boston Scientific CRM is conducting a recall on the LATITUDE Patient Management System Model 6488 Version 7.3 because between March 4 and 27, 2013, a Siebel release caused ~5000 distributed PMR clinic faxes to display with illegible characters and not the required, Patient non-compliance information that was intended. FDA Determined Cause 2 Software change control Action Boston Scientific CRM Corp began re-faxing the 4,912 previously illegible PMR notifications on April 10, 2013 and completed on April 15, 2013, successfully fulfilling labeling expectations that these notifications be communicated to the clinic.******For questions regarding this recall call 651-582-4000. Quantity in Commerce 1 Distribution Nationwide Distribution including DC and PR Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LWS - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Boston Scientific CRM Corp began re-faxing the 4,912 previously illegible PMR notifications on April 10, 2013 and completed on April 15, 2013, successfully fulfilling labeling expectations that these notifications be communicated to the clinic.******For questions regarding this recall call 651-582-4000. | SoftwareChangeControl
121 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116845 | Z-1149-2013 | MAQUET CARDIOHELP Support System: CARDIOHELP-I Software (up to version 3.3.0.0)******Product Usage: The Cardiohelp System is a blood oxygenation and carbon dioxide removal system used to pump blood through the extracorporeal bypass circuit for circulatory and/or pulmonary support during procedures requiring cardiopulmonary bypass (for periods up to six hours). It is also intended to provide circulatory and/or pulmonary support during procedures not requiring cardiopulmonary bypass (for periods up to six hours). | CARDIOHELP-i | 2 | CDRH | 2013-04-18 | 2013-04-18 21:00 | 2013-08-22 | Maquet Cardiovascular, LLC | It has come to the attention of MAQUET that there have been episodes of brief unexpected shutdown, followed by an automatic device restart of the human machine interface (touch screen) of the CARDIOHELP device. The human machine interface (HMI) is the central display of the CARDIOHELP device, where measured values, settings and alarms are displayed and adjusted/confirmed, with the exception of bl | Class 2 Device Recall CARDIOHELPi Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall CARDIOHELPi FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall CARDIOHELPi Date Initiated by Firm March 07, 2013 Date Posted April 18, 2013 Recall Status 1 Terminated 3 on August 22, 2013 Recall Number Z-1149-2013 Recall Event ID 64641 510(K)Number K102726 Product Classification Console, heart-lung machine, cardiopulmonary bypass - Product Code DTQ Product MAQUET CARDIOHELP Support System: CARDIOHELP-I Software (up to version 3.3.0.0)************Product Usage: The Cardiohelp System is a blood oxygenation and carbon dioxide removal system used to pump blood through the extracorporeal bypass circuit for circulatory and/or pulmonary support during procedures requiring cardiopulmonary bypass (for periods up to six hours). It is also intended to provide circulatory and/or pulmonary support during procedures not requiring cardiopulmonary bypass (for periods up to six hours). Code Information Catalog Number: 70104.8012 Cardiohelp-i. Multiple serial numbers. Recalling Firm/ Manufacturer Maquet Cardiovascular, LLC 45 Barbour Pond Dr Wayne NJ 07470-2094 For Additional Information Contact Mr. Phillip Freed 973-709-7498 Manufacturer Reason for Recall It has come to the attention of MAQUET that there have been episodes of brief unexpected shutdown, followed by an automatic device restart of the human machine interface (touch screen) of the CARDIOHELP device. The human machine interface (HMI) is the central display of the CARDIOHELP device, where measured values, settings and alarms are displayed and adjusted/confirmed, with the exception of bl FDA Determined Cause 2 Software change control Action MAQUET sent an Urgent Medical Device Field Correction letters dated March 7, 2013 via Federal Express to all affected customers. The letter identifies the affected product, problem and actions to be taken. Customers are asked to complete and submit the included Customer Fax-back Form. Maquet Service will exchange all affected products with the updated version. For questions MAQUET Customer Service at 888-627-8383 (press option 2, followed by option 2). Quantity in Commerce 154 units affected Distribution Worldwide Distribution - USA Nationwide and the following countries: Australia, Austria, Belarus, Belgium, Canada, Columbia, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Iceland, Iran, Ireland, Italy, Jordan, Kazakhstan, Korea, Kuwait, Luxembourg, The Netherlands, Norway, Oman, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates and the United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = DTQ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action MAQUET sent an Urgent Medical Device Field Correction letters dated March 7, 2013 via Federal Express to all affected customers. The letter identifies the affected product, problem and actions to be taken. Customers are asked to complete and submit the included Customer Fax-back Form. Maquet Service will exchange all affected products with the updated version. For questions MAQUET Customer Service at 888-627-8383 (press option 2, followed by option 2). | SoftwareChangeControl
122 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=112676 | Z-0123-2013 | UniCel¿ DxC 600/800 with***Software v5.0.11, Part Number B12218.******The UniCel¿ DxC Synchron¿ Clinical Systems are fully automated, computercontrolled clinical chemistry analyzers designed for the in vitro determination of a variety of general chemistries, therapeutic drugs, and other chemistries. Analysis can be performed on serum, plasma, urine, or cerebrospinal fluid (CSF) and whole blood (sample type is chemistry dependent). | UniCel¿ DxC 600/800 with | 2 | CDRH | 2012-10-24 | 2012-10-24 21:00 | 2014-02-14 | Beckman Coulter Inc. | Beckman Coulter has confirmed that the Shuttle to Barcode Reader and Cap Piercer alignment***procedure cannot be performed on certain DxC configurations with v5.0.11 software installed. If the Shuttle to Barcode Reader becomes misaligned, the Operator will not be able to perform this***alignment procedure on the DxC systems listed above. Software v5.0 .11 does not allow access to the***Shuttle to Bar | Class 2 Device Recall UniCel DxC 600/800 with Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall UniCel DxC 600/800 with FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall UniCel DxC 600/800 with Date Initiated by Firm September 03, 2012 Date Posted October 24, 2012 Recall Status 1 Terminated 3 on February 14, 2014 Recall Number Z-0123-2013 Recall Event ID 63143 510(K)Number K042291 K060256 K103842 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product UniCel DxC 600/800 with******Software v5.0.11, Part Number B12218.************The UniCel DxC Synchron Clinical Systems are fully automated, computercontrolled clinical chemistry analyzers designed for the in vitro determination of a variety of general chemistries, therapeutic drugs, and other chemistries. Analysis can be performed on serum, plasma, urine, or cerebrospinal fluid (CSF) and whole blood (sample type is chemistry dependent). Code Information Part Number B12218. Serial numbers 2329, 5477, 5479. Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 Manufacturer Reason for Recall Beckman Coulter has confirmed that the Shuttle to Barcode Reader and Cap Piercer alignment******procedure cannot be performed on certain DxC configurations with v5.0.11 software installed. If the Shuttle to Barcode Reader becomes misaligned, the Operator will not be able to perform this******alignment procedure on the DxC systems listed above. Software v5.0 .11 does not allow access to the******Shuttle to Bar FDA Determined Cause 2 Software change control Action Beckman Coulter sent an Urgent Field Safety Notice dated August 30, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that their Beckman Coulter Service Representative would contac them to reinstall the software. Customers were asked to complete and return the enclosed Response Form. For questions customers were instructed to call their Customers Support Center on the Beckman Coulter website at www.BeckmanCoulter.com/CustomerSupport/SupportRequest.******For questions regarding this recall call 714-993-5321. Quantity in Commerce 7 units total (both US and International) Distribution Worldwide Distribution - USA including TX and OH. Interntionally to China and Turkey Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Beckman Coulter sent an Urgent Field Safety Notice dated August 30, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that their Beckman Coulter Service Representative would contac them to reinstall the software. Customers were asked to complete and return the enclosed Response Form. For questions customers were instructed to call their Customers Support Center on the Beckman Coulter website at www.BeckmanCoulter.com/CustomerSupport/SupportRequest.******For questions regarding this recall call 714-993-5321. | SoftwareChangeControl
123 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=112243 | Z-2404-2012 | The Brilliance CT 16 System******Product Usage:***The Brilliance 16 is a Whole Body Computed Tomography X-ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from the same axial plane taken at different angles.*** | Brilliance 16 Slice (Air) | 3 | CDRH | 2012-09-18 | 2012-09-19 21:01 | 2013-12-24 | Philips Medical Systems (Cleveland) Inc | Philips was notified that the system logout in software version 2.3.6 is now longer and may sometimes fail to logout, requiring a forced system restart or hard shutdown with the power button. | Class 3 Device Recall Brilliance 16 Slice (Air) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall Brilliance 16 Slice (Air) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall Brilliance 16 Slice (Air) Date Initiated by Firm August 28, 2012 Date Posted September 19, 2012 Recall Status 1 Terminated 3 on December 24, 2013 Recall Number Z-2404-2012 Recall Event ID 63006 510(K)Number K012009 Product Classification System, x-ray, tomography, computed - Product Code JAK Product The Brilliance CT 16 System************Product Usage:******The Brilliance 16 is a Whole Body Computed Tomography X-ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from the same axial plane taken at different angles. Code Information System Code #, 728246, Serial #, 6462, 6447, 6495, 6514, 6467, 6482, 6473, 6544, 6439, 6453, 6436, 6536, 6448, 6449, 6471, 6498, 6502, 6504, 6452, 6510, 6508, 6438, 6513, 6456, 6505, 6457, 6442, 6454, 6461, 6469, 6485, 6458, 6475, 6487, 6531, 6520, 6490, 6470, 6437, 6497, 6481, 6489, 6488, 6455, 6543, 6499, 6493, 6443, 6463, 6476, 6477, 6478, 6526, 6545, 6546, 6518, 6491, 6450, 6451, 6468, 6464, 6466, 6459, 6472, 6444, 6445, 6529, 6530, 6440, 6532, 6515, 6486, 6507, 6512, 6483, 6524, 6480, 6479, 6509, 6460, 6441, 6465, 6474, 6484, 6492, 6503, 6511, 6517, 6519, 6523, 6525, 6549, 6555 Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Road Cleveland OH 44143-2131 For Additional Information Contact Kumudini J. Carter 440-483-7000 Manufacturer Reason for Recall Philips was notified that the system logout in software version 2.3.6 is now longer and may sometimes fail to logout, requiring a forced system restart or ***hard*** shutdown with the power button. FDA Determined Cause 2 Software change control Action Philips sent an Urgent - Medical Device Correction letter dated August 24, 2012 via certified mail to all US consignees. Philips Key Market representatives distributed the letter internationally. The letter identified the affected product, problem and actions to be taken by customer/user. Philips is implementing a software update to correct the issue free of charge. A Philips representative will contact affect sites to schedule the software installation. For further information or support contact your Philips representative or local Philips Healthcare office. Quantity in Commerce 93 units Distribution Worldwide Distribution - US Nationwide including the states of: AL, CA, GA, MA, MD, NC, OH, PA, TX, and WA; and to the following countries of: Algeria, Oman, Australia, Brazil, China, Saudi Arabia, United Arab Emirates, and India. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips sent an Urgent - Medical Device Correction letter dated August 24, 2012 via certified mail to all US consignees. Philips Key Market representatives distributed the letter internationally. The letter identified the affected product, problem and actions to be taken by customer/user. Philips is implementing a software update to correct the issue free of charge. A Philips representative will contact affect sites to schedule the software installation. For further information or support contact your Philips representative or local Philips Healthcare office. | SoftwareChangeControl
124 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=111260 | Z-2184-2012 | IMPAX Cardiovascular (CV) Reporting******Reporting tool used for structured reporting and structured report outputs (including XML style sheet customization and PDF) | IMPAX Cardiovascular (CV) Reporting | 2 | CDRH | 2012-08-10 | 2012-08-10 21:00 | 2015-01-13 | AGFA Corp. | Content entered into the "Conclusions" free text box on the Report Writer screen was not represented on the final printed report. | Class 2 Device Recall IMPAX Cardiovascular (CV) Reporting Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall IMPAX Cardiovascular (CV) Reporting FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall IMPAX Cardiovascular (CV) Reporting Date Initiated by Firm July 27, 2012 Date Posted August 10, 2012 Recall Status 1 Terminated 3 on January 13, 2015 Recall Number Z-2184-2012 Recall Event ID 62646 510(K)Number K050228 Product Classification System, image processing, radiological - Product Code LLZ Product IMPAX Cardiovascular (CV) Reporting************Reporting tool used for structured reporting and structured report outputs (including XML style sheet customization and PDF) Code Information Model # L9M2100 - All Software versions of IMPAX CV Reporting Recalling Firm/ Manufacturer AGFA Corp. 10 S Academy St Greenville SC 29601-2632 For Additional Information Contact Debbie Huff 864-421-1754 Manufacturer Reason for Recall Content entered into the "Conclusions" free text box on the Report Writer screen was not represented on the final printed report. FDA Determined Cause 2 Software change control Action AGFA Healthcare sent an URGENT FIELD SAFETY NOTICE dated July 27, 2012, to all affected customers. The letter described the safety alert and mitigation. Customer were asked to complete the Acknowledgment form and fax back to 864-421-1664 that the information was received and understood.******For questions customers were instructed tol call 401-604-2180.******For questions regarding this recall call 864-421-1754. Quantity in Commerce 280 Distribution Worldwide Distribution - USA (nationwide) and Canada Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action AGFA Healthcare sent an URGENT FIELD SAFETY NOTICE dated July 27, 2012, to all affected customers. The letter described the safety alert and mitigation. Customer were asked to complete the Acknowledgment form and fax back to 864-421-1664 that the information was received and understood.******For questions customers were instructed tol call 401-604-2180.******For questions regarding this recall call 864-421-1754. | SoftwareChangeControl
125 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=109499 | Z-1818-2012 | Philips Digital Diagnost Software, software release 2.0.2 and 2.0.2SP1 including mirror icon******Stationary Fluoroscopic system | Philips Healthcare Imaging | 2 | CDRH | 2012-06-15 | 2012-06-15 21:00 | 2017-02-24 | Philips Healthcare Inc. | With Digital Diagnost (DIDI) R 2.0.2 a mirrored "R in a circle" is burnt into the image (if mirrored) in the lower right corner of the image. When such an image is mirrored back inside the PACS system, an unmirrored "R in a circle" appears in the lower left corner. This can be mistaken for a "right patient side" marker, although this can appear on the left patient side | Class 2 Device Recall Philips Healthcare Imaging Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare Imaging FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare Imaging Date Initiated by Firm May 22, 2012 Date Posted June 15, 2012 Recall Status 1 Terminated 3 on February 24, 2017 Recall Number Z-1818-2012 Recall Event ID 61926 510(K)Number K982795 Product Classification Peripheral electromagnetic field (pemf) to aid wound healing - Product Code MQB Product Philips Digital Diagnost Software, software release 2.0.2 and 2.0.2SP1 including mirror icon************Stationary Fluoroscopic system Code Information Software release 2.0.2 and 2.0.2SP1 including mirror icon Recalling Firm/ Manufacturer Philips Healthcare Inc. 3000 Minuteman Road Andover MA 01810 For Additional Information Contact 978-687-1501 Manufacturer Reason for Recall With Digital Diagnost (DIDI) R 2.0.2 a mirrored "R in a circle" is burnt into the image (if mirrored) in the lower right corner of the image. When such an image is mirrored back inside the PACS system, an unmirrored "R in a circle" appears in the lower left corner. This can be mistaken for a "right patient side" marker, although this can appear on the left patient side FDA Determined Cause 2 Software change control Action Philips Healthcare issued a Field Safety Notice on 5/22/12 advising users of the error. A software update (Rel 2.1.2) will issue to correct the problem. Contact your local Philips representative. 1-800-722-9377.******For questions regarding this recall call 978-687-1501. Quantity in Commerce 752 units Distribution Worldwide Distribution - USA (nationwide) and the countries of:******Nationwide******Foreign: Canada******AFGHANISTAN******ALBANIA******ALGERIA******ANGOLA******ARGENTINA******ARMENIA******ARUBA******AUSTRALIA******AUSTRIA******AZERBAIJAN******BAHAMAS******BAHRAIN******BANGLADESH******BARBADOS******BELARUS******BELGIUM******BENIN******BHUTAN******BOLIVIA******BOSNIA&HERZEGOVINA******BRAZIL******BRUNEI DARUSSALAM******BULGARIA******BURUNDI******CAMBODIA******CAMEROON******CANARY ISLANDS******CHAD******CHILE******CHINA******COLOMBIA******CONGO******CONGO,DEMOCRATIC******COSTA RICA******COTE D'IVOIRE******CROATIA******CUBA******CYPRUS******CZECH REPUBLIC******DENMARK******DJIBOUTI******DOMINICAN REPUBLIC******ECUADOR******EGYPT******EL SALVADOR******EQUATORIAL GUINEA******ERITREA******ESTONIA******ETHIOPIA******fake test ssd******FINLAND******FRANCE******FRENCH GUIANA******GABON******GAMBIA******GEORGIE******GERMANY******GHANA******GREECE******GUADELOUPE******GUAM******GUATEMALA******GUINEA******GUYANA******HONDURAS******HONG KONG******HUNGARY******ICELAND******INDIA******INDONESIA******IRAN******IRAQ******IRELAND******ISRAEL******ITALY******JAMAICA******JAPAN******JORDAN******KAZAKSTAN******KENYA******KUWAIT******KYRGYZSTAN******LAO PEOPLE'S DEM.REP******LATVIA******LEBANON******LIBERIA******LIBYAN ARAB JAMAHIRI******LIECHTENSTEIN******LITHUANIA******LUXEMBOURG******MACAU******MACEDONIA******MALAWI******MALAYSIA******MALI******MALTA******MARTINIQUE******MAURITANIA******MAURITIUS******METRACOM******MEXICO******MOLDOVA******MONGOLIA******MOROCCO******MOZAMBIQUE******NAMIBIA******NEPAL******NETHERLANDS******NETHERLANDS ANTILLES******NEW CALEDONIA******NEW ZEALAND******NICARAGUA******NIGERIA******NORTH KOREA******NORWAY******OMAN******PAKISTAN******PALESTINE******PANAMA******PARAGUAY******PERU******PHC BEST******PHILIPPINES******PMCC BU Canada******PMCC BU USA******POLAND******PORTUGAL******PUERTO RICO******QATAR******REUNION******ROMANIA******RUSSIA******RWANDA******SAUDI ARABIA******SENEGAL******SERBIA&MONTE NEGRO******SIERRA LEONE******SINGAPORE******SLOVAKIA******SLOVENIA******SOMALIA******SOUTH AFRICA******SOUTH KOREA******SPAIN******SRI LANKA******SUDAN******SURINAME******SWEDEN******SWITZERLAND******SYRIAN ARAB REPUBLIC******TAIWAN******TAJIKISTAN******TANZANIA******THAILAND******TRINIDAD AND TOBAGO******TUNISIA******TURKEY******TURKMENISTAN******UGANDA******UKRAINE******Unidentified Country******UNITED ARAB EMIRATES******UNITED KINGDOM****** URUGUAY******UZBEKISTAN******VENEZUELA******VIET NAM******YEMEN******ZAMBIA******ZIMBABWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MQB - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare issued a Field Safety Notice on 5/22/12 advising users of the error. A software update (Rel 2.1.2) will issue to correct the problem. Contact your local Philips representative. 1-800-722-9377.******For questions regarding this recall call 978-687-1501. | SoftwareChangeControl
126 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=109143 | Z-1685-2012 | Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial***compounding feature Software Versions 3.0, 3.0A and 3.1. (VB20, VB20A, and VB21A);***Model #: 10440017 - ACUSON S2000 Refurb. Mainframe; Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Mountain View, CA******The ACUSON S2000 Ultrasound System is a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or Harmonic Imaging and 3D/4D Imaging on a Flat Panel Display. The ACUSON S2000 ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides the ability to measure anatomical structures {fetal, abdominal, intraoperative, intraoperative neurological, pediatric, small organ, neonatal cephalic, adult cephalic, cardiac, trans-esophageal, transrectal, transvaginal, peripheral vessel, musculoskeletal (conventional), musculo-skeletal (superficial) and neonatal cardiac} and calculation packages that provide information that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. *** | Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial | 2 | CDRH | 2012-05-30 | 2012-05-30 21:00 | 2013-07-16 | Siemens Medical Solutions USA, Inc. | When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale. | Class 2 Device Recall Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial Date Initiated by Firm April 23, 2012 Date Posted May 30, 2012 Recall Status 1 Terminated 3 on July 16, 2013 Recall Number Z-1685-2012 Recall Event ID 61808 510(K)Number K111674 Product Classification System, imaging, pulsed echo, ultrasonic - Product Code IYO Product Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial******compounding feature Software Versions 3.0, 3.0A and 3.1. (VB20, VB20A, and VB21A);******Model #: 10440017 - ACUSON S2000 Refurb. Mainframe; Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Mountain View, CA************The ACUSON S2000 Ultrasound System is a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or Harmonic Imaging and 3D/4D Imaging on a Flat Panel Display. The ACUSON S2000 ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides the ability to measure anatomical structures {fetal, abdominal, intraoperative, intraoperative neurological, pediatric, small organ, neonatal cephalic, adult cephalic, cardiac, trans-esophageal, transrectal, transvaginal, peripheral vessel, musculoskeletal (conventional), musculo-skeletal (superficial) and neonatal cardiac} and calculation packages that provide information that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. Code Information All units of these models, including: 200109, 200113, 200162, 200164, 200146, 200175, 200177, 210001, 200145, 200198, 200214, 200219, 200225, 200284, 200315, 200354, 200404, 200407, 200435, 200467, 200364, 200539, 200373, 200574, 200719, 200853, 200858, 200875, 200909, 201006, 201040, 201043, 201113, 201096, 201070, 201071, 201112, 201121, 201168, 201212, 201215, 201298, 201226, 201397, 201306, 201452, 201442, 201510, 201527, 201631, 201678, 201673, 201674, 201675, 201699, 201731, 201727, 201728, 201755, 201816, 201821, 201891, 201928, 201943, 202019, 202030, 202102, 202098, 202111, 202137, 202119, 202131, 202139, 202148, 202167, 202158, 202240, 202267, 202250, 202266, 202312, 202314, 202401, 202412, 202500, 202364, 202420, 202385, 202383, 202374, 202391, 202465, 202384, 202445, 202514, 202448, 202523, 202524, 202460, 202456, 202534, 202379, 202566, 202554, 202570, 202574, 202608, 202610, 202581, 202731, 202744, 202649, 202787, 202679, 202797, 202799, 202803, 202806, 202812, 202834, 202841, 202844, 202857, 202869, 202888, 202902, 202929, 202963, 202583, 202976, 202986, 203012, 203017, 203037, 202955, 202935, 203069, 203086, 203092, 203095, 203113, 203143, 203150, 203199, 203211, 203222, 203224, 203245, 203248, 203249, 203259, 203275, 203289, 203329, 203330, 203351, 203355, 203380, 203381, 203391, 203396, 203401, 203403, 203422, 203424, 203426, 203430, 203437, 203446, 203458, 203462, 203474, 205007, 205015, 205020, 203527, 203539, 203593, 203595, 203582, 203605, 203608, 203626, 203623, 203625, 205034, 205031, 205027, 201013, 205032, 202100, 205126, 203643, 205049, 205039, 205172, 205082, 205083, 205142, 205067, 205043, 205072, 205200, 205075, 205134, 205085, 205106, 205133, 205057, 205102, 205078, 205056, 205147, 205110, 205283, 205279, 205137, 205112, 205270, 205281, 205269, 205100, 205178, 203632, 205104, 205152, 205153, 205293, 205294, 205296, 205164, 205084, 205161, 205300, 205166, 205046, 205155, 205165, 205188, 205222, 205223, 205113, 205169, 205170, 205229, 205402, 205316, 205237, 205232, 205256, 205258, 205291, 205347, 205295, 205179, 205249, 205305, 205326, 205303, 205245, 205250, 205337, 205343, 205180, 205204, 205181, 205255, 205192, 205208, 205261, 205148, 205058, 205132, 205215, 205328, 205265, 205335, 205139, 205230, 205271, 205186, 205240, 205187, 205189, 205263, 205238, 205210, 205216, 205217, 205234, 201677, 205266, 205393, 205177, 205251, 205322, 205330, 205174, 205312, 205429, 205182, 205205, 205206, 205213, 205371, 205430, 201170, 205292, 205306, 205317, 205362, 205396, 205241, 205309, 205356, 205390, 205236, 205235, 205389, 205353, 205298, 205397, 205308, 205324, 205359, 205310, 205253, 205191, 205360, 205368, 205198, 205214, 205321, 205323, 205345, 201233, 202370, 200432, 205370, 205351, 202776, 205302, 205364, 205354, 205338, 205329, 205372, 205355, 205401, 205443, 205352, 205332, 205442, 205375, 201363, 205331, 205434, 202458, 202489, 205394, 205315, 205311, 205340, 205218, 205319, 205307, 205334, 205325, 205185, 202072, 205348, 205383, 205399, 205380, 202450, 201773, 205482, 205350, 205475, 205484, 205486, 205485, 205508, 205488, 205131, 205365, 205385, 205384, 205367, 205379, 205400, 205398, 205373, 205366, 205357, 205286, 205378, 201540, 205436, 205435, 205427, 205423, 205422, 205420, 205418, 205403, 205327, 205369, 205457, 205454, 205503, 205382, 205341, 205510, 205444, 205438, 205437, 205455, 205424, 205496, 205358, 205493, 205392, 202490, 200319, 205417, 205377, 205349, 205492, 205419, 200669, 205473, 205491, 205411, 205453, 202954, 202382, 202738, 205495, 205490, 205504, 201364, 205044, 205481, 205458, 205465, 205415, 205487, 205426, 205439, 205474, 205440, 205408, 205471, 205498, 205446, 205467, 205499, 205460, 202277, 201191, 205410, 205428, 205409, 205404, 205407, 205413, 205500, 205483, 205466, 205469, 205470, 205472, 205478, 205480, 202142, 205520, 205524, 205527, 205431, 205529, 202629, 202953, 205462, 205406, 205412, 205414, 205421, 205468, 205511, 205433, 205432, 205512, 205513, 205516, 205448, 201225, 200252, 205452, 205532, 205526, 205459, 205533, 205531, 202010, 200183, 205464, 205416, 205447, 205534, 202754, 205450, 205314, 205502, 205515, 205517, 205518, 205528, 205519, 205522, 205525, 205463, 205538, 202546, 205543, 205547, 205544, 205548, 205553, 205559, 205562, 205566, 205567, 205497, 205536, 205539, 205541, 205549, 205550, 205554, 205425, 202943, 205445, 205545, 205667, 205676, 200979, 201086, 203238, 201159, 205476, 202940, 205514, 205561, 205617, 205620, 205660, 201230, 202360, 205555, 205697, 205663, 205542, 205693, 205695, 205670, 205672, 205683, 205684, 205692, 205573, 205574, 205576, 205664, 205584, 205665, 205668, 205590, 205587, 205591, 201293, 205608, 205592, 205609, 205611, 205615, 205621, 205622, 205618, 205619, 205589, 205626, 205627, 205629, 205632, 205633, 205711, 205714, 205718, 205593, 205585, 205610, 205586, 202757, 205723, 205655, 205579, 205606, 205623, 205624, 205653, 202832, 205595, 205599, 205658, 203193, 201294, 201427, 205582, 202427, 205696, 205564, 205575, 205578, 200227, 205647, 205570, 205594, 205588, 205565, 205648, 205613, 205640, 205642, 205643, 205634, 205635, 205639, 205644, 205654, 205645, 205749, 200725, 205646, 205659, 205577, 202299, 205656, 205571, 205706, 205748, 205731, 205702, 205728, 205602, 205607, 205612, 205637, 205744, 205638, 203084, 205712, 205600, 205736, 205735, 202308, 205160, 205203, 205207, 205219, 200468, 205267, 202756, 205246, 203267, 200286, 205339, 205733, 205603, 205739, 202918, 205740, 205745, 205719, 205690, 205757, 205681, 205752, 205729, 205703, 205738, 205671, 205687, 205688, 205734, 205669, 205710, 205725, 205767, 205769, 205699, 205768, 205727, 205686, 205746, 205754, 205765, 205641, 205713, 205724, 205743, 205760, 205772, 205774, 205762, 205764, 205705, 205202, 205755, 205780, 205781, 205784, 205792, 205787, 205833, 205798, 205822, 205823, 205825, 205827, 205830, 202625, 205838, 205842, 205813, 205846, 205850, 205790, 205852, 205853, 205857, 202965, 205860, 205680, 205770, 205144, 205777, 203229, 205815, 205753, 205794, 205766, 205763, 205716, 205596, 205799, 205802, 205834, 205809, 205837, 205872, 205847, 205839, 205806, 205868, 205776, 205076, 205817, 205832, 205836, 205855, 205694, 205810, 205874, 205717, 205720, 205816, 205854, 205831, 205829, 205844, 205922, 205906, 205904, 205895, 205883, 206009, 205913, 205917, 206017, 205845, 205879, 205819, 205912, 206005, 205867, 205835, 205849, 205841, 205863, 205843, 205824, 205891, 205878, 205915, 205882, 201330, 205888, 205862, 205884, 205861, 205869, 206018, 205818, 205925, 206008, 206000, 205924, 205870, 205865, 205803, 205661, 205919, 205900, 206027, 206003, 206023, 206016, 205892, 205875, 206036, 206034, 206026, 205698, 206035, 206010, 206045, 205957, 206042, 205907, 205885, 206052, 205905, 205911, 205947, 206074, 205928, 206069, 206059, 205800, 206043, 205826, 205805, 205828, 206014, 205877, 205318, 205903, 205916, 205936, 205859, 205989, 205986, 206071, 206102, 205650, 205996, 205930, 205995, 206011, 206076, 205894, 202138, 206110, 206093, 203214, 200128, 201211, 202928 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc. 685 E Middlefield Rd Mountain View CA 94043-4045 For Additional Information Contact Sheila Pickering 650-969-9112 Manufacturer Reason for Recall When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale. FDA Determined Cause 2 Software change control Action Siemens initiated a Customer Safety Advisory Notification letter Field correction on April 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers' local Customer Service Engineer will install an update to their system as soon as it becomes available. ******For questions customers were instructed to contact their local service support person. Customers were also asked to pass the notification leter to all those within their organization who need to be aware of the issue until the corrective action is completed.******For questions regarding this recall call 650-694-5747. Distribution Worldwide Distribution - USA (nationwide) and the countries of Argentina, Austria, Australia, Belgium, Canada, Switzerland, Chile, China, Germany, Denmark, Ecuador, Spain, France, Great Britain, Greece, Hong Kong, Hungary, Indonesia, Israel, India, Italy, Jordan, Japan, South Korea, Kuwait, Lebanon, Myanmar, Mexico, Malaysia, Netherlands, Norway, New Zealand, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Sweden, Slovenia, El Salvador, and Thailand. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYO - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: |  | SoftwareChangeControl
127 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=109142 | Z-1684-2012 | Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial***compounding feature Software Versions 3.0, 3.0A and 3.1. (VB20, VB20A, and VB21A); Model #: 10041461 - ACUSON S2000 Mainframe; Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Mountain View, CA******The ACUSON S2000 Ultrasound System is a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or Harmonic Imaging and 3D/4D Imaging on a Flat Panel Display. The ACUSON S2000 ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides the ability to measure anatomical structures {fetal, abdominal, intraoperative, intraoperative neurological, pediatric, small organ, neonatal cephalic, adult cephalic, cardiac, trans-esophageal, transrectal, transvaginal, peripheral vessel, musculoskeletal (conventional), musculo-skeletal (superficial) and neonatal cardiac} and calculation packages that provide information that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. *** | Siemens brand ACUSON S2000 Ultrasound Systemsand the Advanced SieClearTM spatial, | 2 | CDRH | 2012-05-30 | 2012-05-30 21:00 | 2013-07-16 | Siemens Medical Solutions USA, Inc. | When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale. | Class 2 Device Recall Siemens brand ACUSON S2000 Ultrasound Systemsand the Advanced SieClearTM spatial, Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Siemens brand ACUSON S2000 Ultrasound Systemsand the Advanced SieClearTM spatial, FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Siemens brand ACUSON S2000 Ultrasound Systemsand the Advanced SieClearTM spatial, Date Initiated by Firm April 23, 2012 Date Posted May 30, 2012 Recall Status 1 Terminated 3 on July 16, 2013 Recall Number Z-1684-2012 Recall Event ID 61808 510(K)Number K111674 Product Classification System, imaging, pulsed echo, ultrasonic - Product Code IYO Product Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial******compounding feature Software Versions 3.0, 3.0A and 3.1. (VB20, VB20A, and VB21A); Model #: 10041461 - ACUSON S2000 Mainframe; Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Mountain View, CA************The ACUSON S2000 Ultrasound System is a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or Harmonic Imaging and 3D/4D Imaging on a Flat Panel Display. The ACUSON S2000 ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides the ability to measure anatomical structures {fetal, abdominal, intraoperative, intraoperative neurological, pediatric, small organ, neonatal cephalic, adult cephalic, cardiac, trans-esophageal, transrectal, transvaginal, peripheral vessel, musculoskeletal (conventional), musculo-skeletal (superficial) and neonatal cardiac} and calculation packages that provide information that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. Code Information All units of these models, including: 200109, 200113, 200162, 200164, 200146, 200175, 200177, 210001, 200145, 200198, 200214, 200219, 200225, 200284, 200315, 200354, 200404, 200407, 200435, 200467, 200364, 200539, 200373, 200574, 200719, 200853, 200858, 200875, 200909, 201006, 201040, 201043, 201113, 201096, 201070, 201071, 201112, 201121, 201168, 201212, 201215, 201298, 201226, 201397, 201306, 201452, 201442, 201510, 201527, 201631, 201678, 201673, 201674, 201675, 201699, 201731, 201727, 201728, 201755, 201816, 201821, 201891, 201928, 201943, 202019, 202030, 202102, 202098, 202111, 202137, 202119, 202131, 202139, 202148, 202167, 202158, 202240, 202267, 202250, 202266, 202312, 202314, 202401, 202412, 202500, 202364, 202420, 202385, 202383, 202374, 202391, 202465, 202384, 202445, 202514, 202448, 202523, 202524, 202460, 202456, 202534, 202379, 202566, 202554, 202570, 202574, 202608, 202610, 202581, 202731, 202744, 202649, 202787, 202679, 202797, 202799, 202803, 202806, 202812, 202834, 202841, 202844, 202857, 202869, 202888, 202902, 202929, 202963, 202583, 202976, 202986, 203012, 203017, 203037, 202955, 202935, 203069, 203086, 203092, 203095, 203113, 203143, 203150, 203199, 203211, 203222, 203224, 203245, 203248, 203249, 203259, 203275, 203289, 203329, 203330, 203351, 203355, 203380, 203381, 203391, 203396, 203401, 203403, 203422, 203424, 203426, 203430, 203437, 203446, 203458, 203462, 203474, 205007, 205015, 205020, 203527, 203539, 203593, 203595, 203582, 203605, 203608, 203626, 203623, 203625, 205034, 205031, 205027, 201013, 205032, 202100, 205126, 203643, 205049, 205039, 205172, 205082, 205083, 205142, 205067, 205043, 205072, 205200, 205075, 205134, 205085, 205106, 205133, 205057, 205102, 205078, 205056, 205147, 205110, 205283, 205279, 205137, 205112, 205270, 205281, 205269, 205100, 205178, 203632, 205104, 205152, 205153, 205293, 205294, 205296, 205164, 205084, 205161, 205300, 205166, 205046, 205155, 205165, 205188, 205222, 205223, 205113, 205169, 205170, 205229, 205402, 205316, 205237, 205232, 205256, 205258, 205291, 205347, 205295, 205179, 205249, 205305, 205326, 205303, 205245, 205250, 205337, 205343, 205180, 205204, 205181, 205255, 205192, 205208, 205261, 205148, 205058, 205132, 205215, 205328, 205265, 205335, 205139, 205230, 205271, 205186, 205240, 205187, 205189, 205263, 205238, 205210, 205216, 205217, 205234, 201677, 205266, 205393, 205177, 205251, 205322, 205330, 205174, 205312, 205429, 205182, 205205, 205206, 205213, 205371, 205430, 201170, 205292, 205306, 205317, 205362, 205396, 205241, 205309, 205356, 205390, 205236, 205235, 205389, 205353, 205298, 205397, 205308, 205324, 205359, 205310, 205253, 205191, 205360, 205368, 205198, 205214, 205321, 205323, 205345, 201233, 202370, 200432, 205370, 205351, 202776, 205302, 205364, 205354, 205338, 205329, 205372, 205355, 205401, 205443, 205352, 205332, 205442, 205375, 201363, 205331, 205434, 202458, 202489, 205394, 205315, 205311, 205340, 205218, 205319, 205307, 205334, 205325, 205185, 202072, 205348, 205383, 205399, 205380, 202450, 201773, 205482, 205350, 205475, 205484, 205486, 205485, 205508, 205488, 205131, 205365, 205385, 205384, 205367, 205379, 205400, 205398, 205373, 205366, 205357, 205286, 205378, 201540, 205436, 205435, 205427, 205423, 205422, 205420, 205418, 205403, 205327, 205369, 205457, 205454, 205503, 205382, 205341, 205510, 205444, 205438, 205437, 205455, 205424, 205496, 205358, 205493, 205392, 202490, 200319, 205417, 205377, 205349, 205492, 205419, 200669, 205473, 205491, 205411, 205453, 202954, 202382, 202738, 205495, 205490, 205504, 201364, 205044, 205481, 205458, 205465, 205415, 205487, 205426, 205439, 205474, 205440, 205408, 205471, 205498, 205446, 205467, 205499, 205460, 202277, 201191, 205410, 205428, 205409, 205404, 205407, 205413, 205500, 205483, 205466, 205469, 205470, 205472, 205478, 205480, 202142, 205520, 205524, 205527, 205431, 205529, 202629, 202953, 205462, 205406, 205412, 205414, 205421, 205468, 205511, 205433, 205432, 205512, 205513, 205516, 205448, 201225, 200252, 205452, 205532, 205526, 205459, 205533, 205531, 202010, 200183, 205464, 205416, 205447, 205534, 202754, 205450, 205314, 205502, 205515, 205517, 205518, 205528, 205519, 205522, 205525, 205463, 205538, 202546, 205543, 205547, 205544, 205548, 205553, 205559, 205562, 205566, 205567, 205497, 205536, 205539, 205541, 205549, 205550, 205554, 205425, 202943, 205445, 205545, 205667, 205676, 200979, 201086, 203238, 201159, 205476, 202940, 205514, 205561, 205617, 205620, 205660, 201230, 202360, 205555, 205697, 205663, 205542, 205693, 205695, 205670, 205672, 205683, 205684, 205692, 205573, 205574, 205576, 205664, 205584, 205665, 205668, 205590, 205587, 205591, 201293, 205608, 205592, 205609, 205611, 205615, 205621, 205622, 205618, 205619, 205589, 205626, 205627, 205629, 205632, 205633, 205711, 205714, 205718, 205593, 205585, 205610, 205586, 202757, 205723, 205655, 205579, 205606, 205623, 205624, 205653, 202832, 205595, 205599, 205658, 203193, 201294, 201427, 205582, 202427, 205696, 205564, 205575, 205578, 200227, 205647, 205570, 205594, 205588, 205565, 205648, 205613, 205640, 205642, 205643, 205634, 205635, 205639, 205644, 205654, 205645, 205749, 200725, 205646, 205659, 205577, 202299, 205656, 205571, 205706, 205748, 205731, 205702, 205728, 205602, 205607, 205612, 205637, 205744, 205638, 203084, 205712, 205600, 205736, 205735, 202308, 205160, 205203, 205207, 205219, 200468, 205267, 202756, 205246, 203267, 200286, 205339, 205733, 205603, 205739, 202918, 205740, 205745, 205719, 205690, 205757, 205681, 205752, 205729, 205703, 205738, 205671, 205687, 205688, 205734, 205669, 205710, 205725, 205767, 205769, 205699, 205768, 205727, 205686, 205746, 205754, 205765, 205641, 205713, 205724, 205743, 205760, 205772, 205774, 205762, 205764, 205705, 205202, 205755, 205780, 205781, 205784, 205792, 205787, 205833, 205798, 205822, 205823, 205825, 205827, 205830, 202625, 205838, 205842, 205813, 205846, 205850, 205790, 205852, 205853, 205857, 202965, 205860, 205680, 205770, 205144, 205777, 203229, 205815, 205753, 205794, 205766, 205763, 205716, 205596, 205799, 205802, 205834, 205809, 205837, 205872, 205847, 205839, 205806, 205868, 205776, 205076, 205817, 205832, 205836, 205855, 205694, 205810, 205874, 205717, 205720, 205816, 205854, 205831, 205829, 205844, 205922, 205906, 205904, 205895, 205883, 206009, 205913, 205917, 206017, 205845, 205879, 205819, 205912, 206005, 205867, 205835, 205849, 205841, 205863, 205843, 205824, 205891, 205878, 205915, 205882, 201330, 205888, 205862, 205884, 205861, 205869, 206018, 205818, 205925, 206008, 206000, 205924, 205870, 205865, 205803, 205661, 205919, 205900, 206027, 206003, 206023, 206016, 205892, 205875, 206036, 206034, 206026, 205698, 206035, 206010, 206045, 205957, 206042, 205907, 205885, 206052, 205905, 205911, 205947, 206074, 205928, 206069, 206059, 205800, 206043, 205826, 205805, 205828, 206014, 205877, 205318, 205903, 205916, 205936, 205859, 205989, 205986, 206071, 206102, 205650, 205996, 205930, 205995, 206011, 206076, 205894, 202138, 206110, 206093, 203214, 200128, 201211, 202928 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc. 685 E Middlefield Rd Mountain View CA 94043-4045 For Additional Information Contact Sheila Pickering 650-969-9112 Manufacturer Reason for Recall When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale. FDA Determined Cause 2 Software change control Action Siemens initiated a Customer Safety Advisory Notification letter Field correction on April 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers' local Customer Service Engineer will install an update to their system as soon as it becomes available. ******For questions customers were instructed to contact their local service support person. Customers were also asked to pass the notification leter to all those within their organization who need to be aware of the issue until the corrective action is completed.******For questions regarding this recall call 650-694-5747. Distribution Worldwide Distribution - USA (nationwide) and the countries of Argentina, Austria, Australia, Belgium, Canada, Switzerland, Chile, China, Germany, Denmark, Ecuador, Spain, France, Great Britain, Greece, Hong Kong, Hungary, Indonesia, Israel, India, Italy, Jordan, Japan, South Korea, Kuwait, Lebanon, Myanmar, Mexico, Malaysia, Netherlands, Norway, New Zealand, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Sweden, Slovenia, El Salvador, and Thailand. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYO - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: |  | SoftwareChangeControl
128 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=86301 | Z-1487-2012 | Insight rev H software. ******The InSIGHT Gastrointestinal Motility System is intended for use by gastroenterologists, surgeons, and medically trained personnel for esophageal, biliary, and anorectal studies as an aid in documenting and diagnosing digestive motility disorders. | Insight | 2 | CDRH | 2012-05-07 | 2012-05-07 21:00 | 2012-05-07 | Sandhill Scientific, Inc | Software error may produce incorrect esophageal contraction readings. | Class 2 Device Recall Insight Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Insight FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Insight Date Initiated by Firm February 23, 2009 Date Posted May 07, 2012 Recall Status 1 Terminated 3 on May 07, 2012 Recall Number Z-1487-2012 Recall Event ID 53340 510(K)Number K012232 Product Classification System, gastrointestinal motility (electrical) - Product Code FFX Product Insight rev H software. ************The InSIGHT Gastrointestinal Motility System is intended for use by gastroenterologists, surgeons, and medically trained personnel for esophageal, biliary, and anorectal studies as an aid in documenting and diagnosing digestive motility disorders. Code Information S98-1925-2H; S98-1925-4H All lots and shipments of the software. Recalling Firm/ Manufacturer Sandhill Scientific, Inc 9150 Commerce Center Cir Ste 500 Highlands Ranch CO 80129-1563 For Additional Information Contact Guy Harris 800-468-4556 Ext. 1021 Manufacturer Reason for Recall Software error may produce incorrect esophageal contraction readings. FDA Determined Cause 2 Software change control Action Sandhill Scientific sent a Urgent Product Recall letter dated February 23, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. ************The customers were informed All Suite revision H, BioVIEW Analysis Software rev 5.5.1 are affected. Replacement of the Suite revision H, BioVIEW Analysis Software rev 5.5.1, will be provided at no cost to you. Sandhill Scientific will pay for shipping/handling charges associated with replacement of this software.************We anticipate that by the first week of March, we will be able to provide you with the Software that resolves this problem.************For further questions please call toll free 800-468-4556. Quantity in Commerce 60 programs Distribution Worldwide Distribution--USA (nationwide) including the states of CA, CO, CT, FL, IA, IL, MA, MD, MI, MN, NC, ND, NE, NM, NY, OH, OR, PA, SC, SD, TX, UT, VA and WA and the countries of Belgium, Canada, Malaysia, Mexico, New Zealand, South-Korea, Switzerland and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = FFX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Sandhill Scientific sent a Urgent Product Recall letter dated February 23, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. ************The customers were informed All Suite revision H, BioVIEW Analysis Software rev 5.5.1 are affected. Replacement of the Suite revision H, BioVIEW Analysis Software rev 5.5.1, will be provided at no cost to you. Sandhill Scientific will pay for shipping/handling charges associated with replacement of this software.************We anticipate that by the first week of March, we will be able to provide you with the Software that resolves this problem.************For further questions please call toll free 800-468-4556. | SoftwareChangeControl
129 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=107808 | Z-1474-2012 | Philips eCareManager System, Model 865325, software versions 3.7, 3.7.1 and 3.8. ******Product Usage: The Philips eCareManager System software is intended for use in data collection, storage, and clinical information management with independent bedside devices and ancillary systems that are connected either directly or through networks. The eCareManager system provides patient information and surveillance of monitored patients at the point of care location and at a remote, supplementary care location through wide area networking technology and dedicated phone lines and is solely intended for use in a hospital environment. It is not intended to be used in a home environment. | Philips eCareManager System | 2 | CDRH | 2012-04-27 | 2012-04-27 21:00 | 2012-07-06 | Visicu, Inc. | Three versions of software for the eCareManager failed to display anticoagulants on the medications screen. The problem arose due to an unexpected change in the National Drug Data File. | Class 2 Device Recall Philips eCareManager System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips eCareManager System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips eCareManager System Date Initiated by Firm February 22, 2012 Date Posted April 27, 2012 Recall Status 1 Terminated 3 on July 06, 2012 Recall Number Z-1474-2012 Recall Event ID 61328 510(K)Number K001972 K012171 Product Classification System,network and communication,physiological monitors - Product Code MSX Product Philips eCareManager System, Model 865325, software versions 3.7, 3.7.1 and 3.8. ************Product Usage: The Philips eCareManager System software is intended for use in data collection, storage, and clinical information management with independent bedside devices and ancillary systems that are connected either directly or through networks. The eCareManager system provides patient information and surveillance of monitored patients at the point of care location and at a remote, supplementary care location through wide area networking technology and dedicated phone lines and is solely intended for use in a hospital environment. It is not intended to be used in a home environment. Code Information v3.7, v3.7.1, v3.8 Recalling Firm/ Manufacturer Visicu, Inc. 217 E Redwood St Ste 1900 Baltimore MD 21202-3315 For Additional Information Contact Ken Sterner 410-276-1960 Manufacturer Reason for Recall Three versions of software for the eCareManager failed to display anticoagulants on the medications screen. The problem arose due to an unexpected change in the National Drug Data File. FDA Determined Cause 2 Software change control Action Philips notified all affected customers via a preliminary email message advising them of the issue. An URGENT - Field Safety Notice follow-up letter was mailed to all customers on February 22, 2012. This letter identified the affected product, described the problem and circumstances in which it can occur; recommend actions the customer should take to prevent risks for patients, and list the actions planned by Philips to correct the problem. For questions contact Philips eICU Customer Support 877-374-2872. Quantity in Commerce 12 Distribution Nationwide Distribution (USA) including states of MA, MO, WI, AR, AZ, PA, SD, VA, and IN. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MSX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips notified all affected customers via a preliminary email message advising them of the issue. An URGENT - Field Safety Notice follow-up letter was mailed to all customers on February 22, 2012. This letter identified the affected product, described the problem and circumstances in which it can occur; recommend actions the customer should take to prevent risks for patients, and list the actions planned by Philips to correct the problem. For questions contact Philips eICU Customer Support 877-374-2872. | SoftwareChangeControl
130 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=108198 | Z-1454-2012 | Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4******The Dimension Vista(R) System is an in vitro diagnostic device intended to duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids. Dimension Vista(R) chemical and immunochemical applications use photometric, turbidimetric, chemiluminescence, nephelometric and integrated ion-selective multisensory technology for clinical use. | Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4 | 2 | CDRH | 2012-04-25 | 2012-04-25 21:01 | 2014-03-05 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed customer complaints on Vista(R) software version 3.4 for Below Manufacturer Assay Range error flags associated with QC and patient test results that are within the assay range as defined in the Instructions For Use for the respective method. This issue is isolated to Vista(R) software version 3.4. Customers who have not made any changes to the default method configurations | Class 2 Device Recall Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4 Date Initiated by Firm January 16, 2012 Date Posted April 25, 2012 Recall Status 1 Terminated 3 on March 05, 2014 Recall Number Z-1454-2012 Recall Event ID 61482 510(K)Number K051087 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4************The Dimension Vista(R) System is an in vitro diagnostic device intended to duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids. Dimension Vista(R) chemical and immunochemical applications use photometric, turbidimetric, chemiluminescence, nephelometric and integrated ion-selective multisensory technology for clinical use. Code Information Vista software version 3.4 Recalling Firm/ Manufacturer Siemens Healthcare Diagnostics, Inc. 500 Gbc Dr Ms 514 PO BOX 6101 Newark DE 19702-2466 For Additional Information Contact Robert J. King 800-441-9250 Manufacturer Reason for Recall Siemens has confirmed customer complaints on Vista(R) software version 3.4 for Below Manufacturer Assay Range error flags associated with QC and patient test results that are within the assay range as defined in the Instructions For Use for the respective method. This issue is isolated to Vista(R) software version 3.4. Customers who have not made any changes to the default method configurations FDA Determined Cause 2 Software change control Action Siemens sent an Urgent Field Safety Notice dated December 19, 2011 to all affected customers alerting them to Invalid Below Manufacturer Assay Range Flags on Dimension Vista(R) Systems with Software Version 3.4. The UFSN included instructions on how to correct the situation and also included an effectiveness check form to be faxed to the firm at 302-631-8467. On January 16, 2012, the firm issued an additional UFSN alerting consignees to Invalid Above and Below Manufacturer Assay Range Flags. This notice also included directions and a Field Correction Effectiveness Check form to be completed and faxed to 302-631-8467.******For technical questions customers were instruced to call 800-441-9250.******For questions regarding this recall call 302-631-6299. Quantity in Commerce 1,683 Distribution Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland & Baltics, France, Germany, Italy, Japan, South Korea, Malaysia, Netherlands, New Zealand, Norway, Portugal, Puerto Rico, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens sent an Urgent Field Safety Notice dated December 19, 2011 to all affected customers alerting them to Invalid Below Manufacturer Assay Range Flags on Dimension Vista(R) Systems with Software Version 3.4. The UFSN included instructions on how to correct the situation and also included an effectiveness check form to be faxed to the firm at 302-631-8467. On January 16, 2012, the firm issued an additional UFSN alerting consignees to Invalid Above and Below Manufacturer Assay Range Flags. This notice also included directions and a Field Correction Effectiveness Check form to be completed and faxed to 302-631-8467.******For technical questions customers were instruced to call 800-441-9250.******For questions regarding this recall call 302-631-6299. | SoftwareChangeControl
131 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=107919 | Z-1285-2012 | Whole body Computed Tomography X-Ray System (Brilliance 64 and Ingenuity CT), Brilliance 64 M/N 728231, Ingenuity CT M/N 728326, Dist. By Philips Medical Systems, Cleveland, OH | Brilliance 64 and Ingenuity CT | 2 | CDRH | 2012-03-23 | 2012-03-23 21:00 | 2013-02-22 | Philips Medical Systems (Cleveland) Inc | A number of non-conformances were identified during internal and external validation testing of the Brilliance 64 and Ingenuity CT systems with software version Beta 3 (version 4.0.0.xx368). | Class 2 Device Recall Brilliance 64 and Ingenuity CT Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Brilliance 64 and Ingenuity CT FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Brilliance 64 and Ingenuity CT Date Initiated by Firm February 17, 2012 Date Posted March 23, 2012 Recall Status 1 Terminated 3 on February 22, 2013 Recall Number Z-1285-2012 Recall Event ID 61363 510(K)Number K033326 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Whole body Computed Tomography X-Ray System (Brilliance 64 and Ingenuity CT), Brilliance 64 M/N 728231, Ingenuity CT M/N 728326, Dist. By Philips Medical Systems, Cleveland, OH Code Information System Code # 728231, Serial # 91003, 95130, 95157, 95504 & 300030; System Code # 728326, Serial # 300010 & 300018. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Road Cleveland OH 44143-2131 For Additional Information Contact Joseph Vinhais 440-483-7000 Manufacturer Reason for Recall A number of non-conformances were identified during internal and external validation testing of the Brilliance 64 and Ingenuity CT systems with software version Beta 3 (version 4.0.0.xx368). FDA Determined Cause 2 Software change control Action Philips Healthcare sent an Urgent Medical Device Correction letter dated February 17, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all memebers of their staff who need to be aware of the contents in the communication. Customers were asked to retain a copy of the communication with the equipment instruction for Use. ******For further information or support concerning this issue, customers should contact their local Philips representative or local Philips Healthcare office.******For North America and Canada contact the Customer Care Solutions Center at 1-800-722-9377, option 5.******For questions regarding this recall call 440-483-7000. Quantity in Commerce 7 units Distribution Worldwide Distribution - USA including OH and the following countries: Australia, China, Belgium, Israel & Sweden. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare sent an Urgent Medical Device Correction letter dated February 17, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all memebers of their staff who need to be aware of the contents in the communication. Customers were asked to retain a copy of the communication with the equipment instruction for Use. ******For further information or support concerning this issue, customers should contact their local Philips representative or local Philips Healthcare office.******For North America and Canada contact the Customer Care Solutions Center at 1-800-722-9377, option 5.******For questions regarding this recall call 440-483-7000. | SoftwareChangeControl
132 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=107347 | Z-1134-2012 | Philips IntelliVue Patient Monitors: ***The monitor is intended to be used for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults, pediatrics, and neonates in health care facilities. The monitor is to be used by trained health care professionals. Physiological Monitoring, Patient Monitor , Arrhythmia detector and alarm (including ST -segment measurement and alarm. | Philips IntelliVue Patient Monitors | 2 | CDRH | 2012-03-01 | 2012-03-01 21:00 | 2017-02-15 | Philips Healthcare Inc. | Central station monitors may not alarm (either visual or audible) when an alarm is announced at the affected patient monitor | Class 2 Device Recall Philips IntelliVue Patient Monitors Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips IntelliVue Patient Monitors FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips IntelliVue Patient Monitors Date Initiated by Firm February 09, 2012 Date Posted March 01, 2012 Recall Status 1 Terminated 3 on February 15, 2017 Recall Number Z-1134-2012 Recall Event ID 61127 510(K)Number K110622 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product Philips IntelliVue Patient Monitors: ******The monitor is intended to be used for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults, pediatrics, and neonates in health care facilities. The monitor is to be used by trained health care professionals. Physiological Monitoring, Patient Monitor , Arrhythmia detector and alarm (including ST -segment measurement and alarm. Code Information Models with SW Revision H (up to and including H.15.36) are affected, upgraded between October 6, 2010 and January 23, 2012 MP20 (M8001A), MP30 (M8002A), MP40 (M8003A), MP50 (M8004A), MP60 (M8005A), MP70 (M8007A), MP80 (M8008A), MP90 (M8010A), D80 (M8016A), MX600 (865242), MX700 (865241) MX800 (865240 Recalling Firm/ Manufacturer Philips Healthcare Inc. 3000 Minuteman Road Andover MA 01810 For Additional Information Contact SAME 978-687-1501 Manufacturer Reason for Recall Central station monitors may not alarm (either visual or audible) when an alarm is announced at the affected patient monitor FDA Determined Cause 2 Software change control Action Philips Healthcare sent a Urgent Medical Device Correction letter dated February 7, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. ************Philips has recently received reports that, under certain circumstances, alarms announced at the patient monitor are not announced (either visual or audible) at the central station. This issue affects the following IntelliVue patient monitor models with SW rev. H.xx.xx******MP20 (M8001A), MP30 (M8002A), MP40 (M8003A), MP50 (M8004A), MP60 (M8005A), MP70 (M8007A), MP80 (M8008A), MP90 (M8010A), Intelligent Display D80 (M8016A), MX600 (865242), MX700 (865241) and MX800 (865240)******If this issue occurs, the primary alarm function at the bedside monitor is not affected.************Philips is conducting this voluntary correction to upgrade software on affected devices. Please refer to the following pages, which provide instructions for actions to be taken. Follow the Action to be taken by Customer/User section of the instructions. This issue has been reported to the appropriate regulatory agencies.************Ensuring that you have the highest quality medical devices, accessories and supporting documentation is our top priority. Your satisfaction with Philips products is very important to us.************Review this information with all staff members who interface with the central station and need to be aware of the contents of this communication.******The Product Number and Serial Number is contained on the devices product label, located on the front of the device. The SW revision can be accessed via the Revision Screen at the bedside monitor. charge.A Philips Healthcare representative will contact customers with affected devices contact Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this correction to arrange an upgrade of the Intellivue software. Quantity in Commerce 83,534 units Distribution Worldwide Distribution -- USA (nationwide) and the countries of ALBANIA, ANGOLA, ANTILLES, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL, BYELORUSSIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, ESTONIA, FINLAND, FRANCE, GABON, GERMANY, GHANA, GREECE, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KUWAIT, LEBANON, LIBYAN, ARAB, JAMAHIRI, LITHUANIA, LUXEMBOURG,MACAU, MALAYSIA, MOROCCO, MAURITIUS, MEXICO, MOROCCO, NAMIBIA, NEPAL, NETHERLANDS, ANTILLES, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PALESTINE, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB, EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIET NAM, YEMEN, ARAB REP and ZAMBIA Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare sent a Urgent Medical Device Correction letter dated February 7, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. ************Philips has recently received reports that, under certain circumstances, alarms announced at the patient monitor are not announced (either visual or audible) at the central station. This issue affects the following IntelliVue patient monitor models with SW rev. H.xx.xx******MP20 (M8001A), MP30 (M8002A), MP40 (M8003A), MP50 (M8004A), MP60 (M8005A), MP70 (M8007A), MP80 (M8008A), MP90 (M8010A), Intelligent Display D80 (M8016A), MX600 (865242), MX700 (865241) and MX800 (865240)******If this issue occurs, the primary alarm function at the bedside monitor is not affected.************Philips is conducting this voluntary correction to upgrade software on affected devices. Please refer to the following pages, which provide instructions for actions to be taken. Follow the Action to be taken by Customer/User section of the instructions. This issue has been reported to the appropriate regulatory agencies.************Ensuring that you have the highest quality medical devices, accessories and supporting documentation is our top priority. Your satisfaction with Philips products is very important to us.************Review this information with all staff members who interface with the central station and need to be aware of the contents of this communication.******The Product Number and Serial Number is contained on the devices product label, located on the front of the device. The SW revision can be accessed via the Revision Screen at the bedside monitor. charge.A Philips Healthcare representative will contact customers with affected devices contact Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this correction to arrange an upgrade of the Intellivue software. | SoftwareChangeControl
133 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=106994 | Z-0927-2012 | AXIOM Artis and AXIOM Artis zee modular angiographic systems******Product Usage: Angiographic x-ray systems*** | AXIOM Artis and AXIOM Artis zee modular angiographic systems | 2 | CDRH | 2012-01-27 | 2012-01-27 21:00 | 2014-01-31 | Siemens Medical Solutions USA, Inc | There is a potential misalignment in the connection of the remote foot switch and the stationary unit. This potential malfunction may occur if either the wireless foot switch or the stationary unit is equipped with improper firmware installed in the bluetooth interface performing the communication between foot switch and system. This is not a general issue but sporadic as a vendor unsystematical | Class 2 Device Recall AXIOM Artis and AXIOM Artis zee modular angiographic systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall AXIOM Artis and AXIOM Artis zee modular angiographic systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall AXIOM Artis and AXIOM Artis zee modular angiographic systems Date Initiated by Firm January 16, 2012 Date Posted January 27, 2012 Recall Status 1 Terminated 3 on January 31, 2014 Recall Number Z-0927-2012 Recall Event ID 60989 510(K)Number K052202 Product Classification System, x-ray, fluoroscopic, image-intensified - Product Code JAA Product AXIOM Artis and AXIOM Artis zee modular angiographic systems************Product Usage: Angiographic x-ray systems Code Information Model numbers 5904433, 5904441, 5904649, 5904656, 5904466, 7007755, 7728350, 7412807, 7727717, 7728392, 7413078, 7555373, 7555365, 7555357, 7008605, 5917054, 10094135, 10094137, 10094139, 10094141, 10094142, 10094143, and 10280959. Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 51 Valley Stream Pkwy Malvern PA 19355 For Additional Information Contact Anastasia Mason 610-219-4834 Manufacturer Reason for Recall There is a potential misalignment in the connection of the remote foot switch and the stationary unit. This potential malfunction may occur if either the wireless foot switch or the stationary unit is equipped with improper firmware installed in the bluetooth interface performing the communication between foot switch and system. This is not a general issue but sporadic as a vendor unsystematical FDA Determined Cause 2 Software change control Action Siemens sent an Urgent Field Safety Notice letter on January 16, 2016, to all affected customers The letter identified the product, problem, potential risk, and actions to be taken. Customers were instructed to immediately pass this information to all those who need to be aware and instruct personnel accordingly. Siemens ask that if the device/equipment is no longer in the customers possession to forward this safety notice to the new owner of the device/equipment. Quantity in Commerce 45 Distribution Nationwide Distribution (USA) including the states of: CA, CT, FL, GA, HI, IL, IN, KY, ME, MD, MI, MS, MO, NH, NJ, NY, NC, OH, OR, TN, TX, UT, VA, WI and Puerto Rico. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAA - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens sent an Urgent Field Safety Notice letter on January 16, 2016, to all affected customers The letter identified the product, problem, potential risk, and actions to be taken. Customers were instructed to immediately pass this information to all those who need to be aware and instruct personnel accordingly. Siemens ask that if the device/equipment is no longer in the customers possession to forward this safety notice to the new owner of the device/equipment. | SoftwareChangeControl
134 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=105388 | Z-0278-2012 | SCC Soft Computer, Softlab Laboratory information System with DORA versions: 2.0.0.4 with Hot Fix 1.2808, 1.2809, 1.2956, or 1.3476; 2.0.0.5; 2.0.0.6; 2.0.0.7 with Hot Fix 1.6595; and 2.1.0.3. | Softlab | 2 | CDRH | 2011-11-23 | 2011-11-23 21:00 | 2012-10-29 | SCC Soft Computer | Potential for patient data to be sent to the wrong patient record. A software change affected how the Softlab handles recovery from certain types of table access conflicts. When such conflicts occur, this function has the potential to cause faulty writing to the database. | Class 2 Device Recall Softlab Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Softlab FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Softlab Date Initiated by Firm March 22, 2011 Date Posted November 23, 2011 Recall Status 1 Terminated 3 on October 29, 2012 Recall Number Z-0278-2012 Recall Event ID 60430 Product Classification Laboratory information System - Product Code LNX Product SCC Soft Computer, Softlab Laboratory information System with DORA versions: 2.0.0.4 with Hot Fix 1.2808, 1.2809, 1.2956, or 1.3476; 2.0.0.5; 2.0.0.6; 2.0.0.7 with Hot Fix 1.6595; and 2.1.0.3. Code Information Softlab with DORA versions: 2.0.0.4 with Hot Fix 1.2808, 1.2809, 1.2956, or 1.3476; 2.0.0.5; 2.0.0.6; 2.0.0.7 with Hot Fix 1.6595; and 2.1.0.3. Recalling Firm/ Manufacturer SCC Soft Computer 5400 Tech Data Drive Clearwater FL 33760 For Additional Information Contact Kathryn Branca 727-789-0100 Manufacturer Reason for Recall Potential for patient data to be sent to the wrong patient record. A software change affected how the Softlab handles recovery from certain types of table access conflicts. When such conflicts occur, this function has the potential to cause faulty writing to the database. FDA Determined Cause 2 Software change control Action Clients were notified through SCC Soft Computer's proprietary Task Management System On 03/22/201. All clients have received a hot fix that corrected the issue or made the hosparam change that prevents the issue. Quantity in Commerce 12 Distribution Sold throughout US and Canada. 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Clients were notified through SCC Soft Computer's proprietary Task Management System On 03/22/201. All clients have received a hot fix that corrected the issue or made the hosparam change that prevents the issue. | SoftwareChangeControl
135 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98566 | Z-0146-2012 | Labeling for RIO Robot Unit is comprised of three main components: 20399 RIO SURGICAL ARM, 201251 RIO GUIDANCE MODULE, 200294 RIO CAMERA STAND ASSEMBLY***PN 203999 SN ROB 125 2010-12 V 100/120/230 A 9.6/8.0/4.2 Hz 50/60 Class I Equipment. Conforms to IEC 60601-1/A2: 1995, EN 60601-1/A2: 1995 UL 60601-1: 2003, CAN/CSA-C22.2 No. 601.1-M90.***Manufactured in USA***MAKO SURGICAL CORP. 2555 DAVIE ROAD, FT. LAUDERDALE, FL 33317.******RIO System - The Tactile Guidance System v2.0 is intended to assist the surgeon in providing software defined spatial boundaries for orientation and reference information to anatomical structures during orthopedic procedures. The Tactile Guidance System v2.0 is indicated for use in surgical knee procedures, in which the use of stereotactic surgery may be approriate, and where reference to rigid anatomical bony structures can be identified relative to a CT base model of the anatomy. These procedures include unicondylar knee replacement and/or patellofemoral knee replacement. | RIO Robot Arm Interactive Orthopedic System | 2 | CDRH | 2011-11-01 | 2011-11-01 21:00 | 2011-12-13 | Mako Surgical Corporation | MAKO Surgical Corp. is recalling their RIO Robotic Arm Interactive Orthopedic System (RIO) due to software issue that exist that could potentially result in a bone resection. No adverse events reported. | Class 2 Device Recall RIO Robot Arm Interactive Orthopedic System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall RIO Robot Arm Interactive Orthopedic System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall RIO Robot Arm Interactive Orthopedic System Date Initiated by Firm January 20, 2011 Date Posted November 01, 2011 Recall Status 1 Terminated 3 on December 13, 2011 Recall Number Z-0146-2012 Recall Event ID 58162 510(K)Number K093425 Product Classification Orthopedic stereotaxic instrument - Product Code OLO Product Labeling for RIO Robot Unit is comprised of three main components: 20399 RIO SURGICAL ARM, 201251 RIO GUIDANCE MODULE, 200294 RIO CAMERA STAND ASSEMBLY***PN 203999 SN ROB 125 2010-12 V 100/120/230 A 9.6/8.0/4.2 Hz 50/60 Class I Equipment. Conforms to IEC 60601-1/A2: 1995, EN 60601-1/A2: 1995 UL 60601-1: 2003, CAN/CSA-C22.2 No. 601.1-M90.***Manufactured in USA***MAKO SURGICAL CORP. 2555 DAVIE ROAD, FT. LAUDERDALE, FL 33317.************RIO System - The Tactile Guidance System v2.0 is intended to assist the surgeon in providing software defined spatial boundaries for orientation and reference information to anatomical structures during orthopedic procedures. The Tactile Guidance System v2.0 is indicated for use in surgical knee procedures, in which the use of stereotactic surgery may be approriate, and where reference to rigid anatomical bony structures can be identified relative to a CT base model of the anatomy. These procedures include unicondylar knee replacement and/or patellofemoral knee replacement. Code Information Lot numbers: ROB034, ROB035, ROB036, ROB037, ROB038, ROB042, ROB044, ROB045, ROB046, ROB047, ROB048, ROB049, ROB050, ROB051, ROB053, ROB054, ROB055, ROB056, ROB057, ROB059, ROB060, ROB061, ROB062, ROB062, ROB063, ROB066, ROB067, ROB069, ROB070, ROB071, ROB072, ROB073, ROB074, ROB075, ROB076, ROB077, ROB078, ROB079, ROB080, ROB081, ROB082, ROB083, ROB084, ROB085, ROB086, ROB088, ROB093, ROB094, ROB096, ROB097, ROB098, ROB099, ROB100, ROB101, ROB102, ROB103, ROB104, ROB107, ROB108, ROB109, ROB110, ROB112, ROB114, ROB116, ROB117, ROB118, ROB120, ROB121, ROB122, ROB123, ROB124. Recalling Firm/ Manufacturer Mako Surgical Corporation 2555 Davie Rd Ste 110 Davie FL 33317-7424 For Additional Information Contact Rachele Riviere 954-927-2044 Ext. 605 Manufacturer Reason for Recall MAKO Surgical Corp. is recalling their RIO Robotic Arm Interactive Orthopedic System (RIO) due to software issue that exist that could potentially result in a bone resection. No adverse events reported. FDA Determined Cause 2 Software change control Action The firm, MAKO Surgical Corp., sent an "RECALL NOTICE 2.4. SITES" letter dated February 1, 2011 to its customers. The letter describes the product, problem, actions to be taken which includes a workaround and an updated user guide identifying the appropriate workflow. The customers were instructed to follow the user guide and workflow and to complete and return the enclosed confirmation memo to their MAKOplasty Specialist or return it in the pre-addressed, pre-paid envelope provided. MAKOplasty included with the letter new steps in the surgical workflow in the use of the RIO system to ensure error value accuracy; worked on a software update that will not allow the user to proceed with error values that do not support the required accuracy-this software was to be supplied to customers in February 2011, and trained MAKOplasty Specialist will make available written updates to the RIO use instructions as well.************If you have any questions or concerns, please do not hesitate to speak with your MAKOplasty Specialist or contact the Manager of Customer Relations at 954.927.4022 x447. Quantity in Commerce 70 Distribution Worldwide distribution: USA including states of: AZ, CA, CO, FL, GA, IA, IL, LA, MI, MO, MS, NC, NJ, NV, NY, OK, OH, OR, PA, RI, TX, UT, VA, WA, WI, and WV; and countries of: Scotland, Italy and Korea. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = OLO - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, MAKO Surgical Corp., sent an "RECALL NOTICE 2.4. SITES" letter dated February 1, 2011 to its customers. The letter describes the product, problem, actions to be taken which includes a workaround and an updated user guide identifying the appropriate workflow. The customers were instructed to follow the user guide and workflow and to complete and return the enclosed confirmation memo to their MAKOplasty Specialist or return it in the pre-addressed, pre-paid envelope provided. MAKOplasty included with the letter new steps in the surgical workflow in the use of the RIO system to ensure error value accuracy; worked on a software update that will not allow the user to proceed with error values that do not support the required accuracy-this software was to be supplied to customers in February 2011, and trained MAKOplasty Specialist will make available written updates to the RIO use instructions as well.************If you have any questions or concerns, please do not hesitate to speak with your MAKOplasty Specialist or contact the Manager of Customer Relations at 954.927.4022 x447. | SoftwareChangeControl
136 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=97478 | Z-3091-2011 | Software version 3.3 for Rapid Programmer devices, Model 3832.******Software version 3.3 for Rapid Programmer devices, a personal digital assistant (PDA) based programmer supporting implantable pulse generators of the Eon, Genesis, and Renew family. | Rapid Programmer | 2 | CDRH | 2011-08-30 | 2011-08-30 21:00 | 2012-09-27 | Advanced Neuromodulation Systems | Product programmed with unapproved software was distributed. | Class 2 Device Recall Rapid Programmer Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Rapid Programmer FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Rapid Programmer Date Initiated by Firm June 02, 2009 Date Posted August 30, 2011 Recall Status 1 Terminated 3 on September 27, 2012 Recall Number Z-3091-2011 Recall Event ID 57806 PMA Number P010032 Product Classification Stimulator, spinal-cord, totally implanted for pain relief - Product Code LGW Product Software version 3.3 for Rapid Programmer devices, Model 3832.************Software version 3.3 for Rapid Programmer devices, a personal digital assistant (PDA) based programmer supporting implantable pulse generators of the Eon, Genesis, and Renew family. Code Information Serial numbers: 101565, 101544, 101555, 101548, 101561, 101551, 101563, 101546, 101556, 101553, 101554, 101558, 101545, and 101566 Recalling Firm/ Manufacturer Advanced Neuromodulation Systems 6901 Preston Rd Plano TX 75024-2508 For Additional Information Contact 972-309-2154 Manufacturer Reason for Recall Product programmed with unapproved software was distributed. FDA Determined Cause 2 Software change control Action St. Jude Medical Neuromodulation contacted all affected customers by telephone between June 2, 2009 and June 5, 2009. Customers were given information about the product, problem and action to be taken.************All affected devices were returned to the firm as instructed. Quantity in Commerce 14 units Distribution Distributed in New Hampshire and Puerto Rico. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LGW - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action St. Jude Medical Neuromodulation contacted all affected customers by telephone between June 2, 2009 and June 5, 2009. Customers were given information about the product, problem and action to be taken.************All affected devices were returned to the firm as instructed. | SoftwareChangeControl
137 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=102758 | Z-2964-2011 | Lenstar LS 900 EyeSuite Biometry version i2.000, Four (4) Protective Cardboard Containers Per Each System.***Haag-Streit USA, Inc.******A non-invasive, non-contact OLCR (Optical Low Coherence Reflectometer) Biometer used for obtaining ocular measurements and performing calculations to assist in determination of the appropriate power and type of IOL (Intraocular Lens) for implantation after removal of the natural crystalline lens following cataract removal. | Lenstar LS 900 EyeSuite Biometry version i2.000, | 2 | CDRH | 2011-08-09 | 2011-08-09 21:00 | 2013-03-08 | Haag-Streit USA Inc | A patient record may be linked to an incorrect/inconsistent set of calculation results under specific circumstances when using EyeSuite Software version i2.000. | Class 2 Device Recall Lenstar LS 900 EyeSuite Biometry version i2.000, Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Lenstar LS 900 EyeSuite Biometry version i2.000, FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Lenstar LS 900 EyeSuite Biometry version i2.000, Date Initiated by Firm June 20, 2011 Date Posted August 09, 2011 Recall Status 1 Terminated 3 on March 08, 2013 Recall Number Z-2964-2011 Recall Event ID 59426 510(K)Number K082891 Product Classification Biometer; Instrument Used to Measure Corneal Radius. - Product Code HJB Product Lenstar LS 900 EyeSuite Biometry version i2.000, Four (4) Protective Cardboard Containers Per Each System.******Haag-Streit USA, Inc.************A non-invasive, non-contact OLCR (Optical Low Coherence Reflectometer) Biometer used for obtaining ocular measurements and performing calculations to assist in determination of the appropriate power and type of IOL (Intraocular Lens) for implantation after removal of the natural crystalline lens following cataract removal. Code Information Part Number 7220162. Serial Numbers: 21, 127, 223, 224, 284, 285, 316, 318, 319, 320, 322, 323, 324, 325, 326, 327, 328, 330, 331, 332, 333, 364, 365, 366, 368, 369, 371, 372, 376, 378, 379, 381, 384, 385, 386, 405, 407, 408, 409, 411, 412, 413, 414, 415, 416, 417, 418, 419, 420, 421, 423, 424, 443, 446, 447, 448, 449, 450, 452, 453, 454, 455, 456, 457, 458, 459, 460, 463, 469, 470, 481, 482, 483, 484, 485, 486, 487, 489, 490, 491, 492, 493, 494, 495, 496, 497, 499, 500, 527, 528, 529, 531, 532, 533, 535, 536, 537, 538, 539, 540, 541, 542, 543, 544, 545, 546, 547, 548, 549, 550, 551, 552, 558, 559, 560, 561, 563, 564, 565, 566, 567, 568, 569, 570, 571, 572, 573, 574, 576, 578, 579 581, 582, 583, 584, 585, 587, 588, 589, 590, 618, 619, 620, 621, 622, 624, 625, 626, 628, 630, 631, 632, 633, 634, 635, 636, 638, 639, 640, 641, 643, 644, 645, 647, 648, 649, 650, 651, 652, 653, 654, 656, 657, 658, 659, 660, 661, 662, 663, 665, 667, 668, 669, 670, 671, 672, 674, 675, 676, 677, 721, 723, 724, 725, 726, 727, 728, 729, 735, 824, 825, 922, 946, 966, 967, 968, 970, 1018, 1019, 1021, 1022, 1024, 1025, 1114, 1143 & 1195. Recalling Firm/ Manufacturer Haag-Streit USA Inc 3535 Kings Mills Rd Mason OH 45040 For Additional Information Contact Russell F. Wolfe 513-398-3937 Ext. 1167 Manufacturer Reason for Recall A patient record may be linked to an incorrect/inconsistent set of calculation results under specific circumstances when using EyeSuite Software version i2.000. FDA Determined Cause 2 Software change control Action Haag-Streit USA, Inc. sent an "Urgent Field Safety Notices" dated June 20, 2011 to all affected customers.************The letter included a description of the affected product and problem. The firm instructed their customers to make sure all new and current templates are saved and upgrade to EyeSuite version i4.000. Customers are asked to complete and return the Software Update Verification Form.************For additional information contact the Haag-Streit at (513) 398-3937. Quantity in Commerce 216 Distribution Nationwide (USA) Distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = HJB - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Haag-Streit USA, Inc. sent an "Urgent Field Safety Notices" dated June 20, 2011 to all affected customers.************The letter included a description of the affected product and problem. The firm instructed their customers to make sure all new and current templates are saved and upgrade to EyeSuite version i4.000. Customers are asked to complete and return the Software Update Verification Form.************For additional information contact the Haag-Streit at (513) 398-3937. | SoftwareChangeControl
138 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=99926 | Z-2733-2011 | SJM Confirm, Implantable Cardiac Monitor (ICM) Models: DM2100, DM2102 (IDE).******This product is a lead less non-therapeutic implantable cardiac monitor and is implanted to aid clinicians in diagnosing intermittent symptoms which may be due to arrhythmia | SJM Confirm, Implantable Cardiac Monitor (ICM) | 2 | CDRH | 2011-07-05 | 2011-07-05 21:00 | 2011-07-19 | St Jude Medical CRMD | The recall was initiated because St. Jude Medical made new software available for the SJM Confirm DM2100 implantable cardiac monitor (ICM) that incorporates enhancements to the devices sensing algorithm. These enhancements are intended to provide improved sensitivity and specificity for diagnosing arrhythmias. The user is presented with the option to upgrade new software into the SJM Confirm ICM | Class 2 Device Recall SJM Confirm, Implantable Cardiac Monitor (ICM) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SJM Confirm, Implantable Cardiac Monitor (ICM) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SJM Confirm, Implantable Cardiac Monitor (ICM) Date Initiated by Firm March 31, 2011 Date Posted July 05, 2011 Recall Status 1 Terminated 3 on July 19, 2011 Recall Number Z-2733-2011 Recall Event ID 58686 510(K)Number K081365 Product Classification Implantable Cardiac Monitor (without arrythmia detection) - Product Code MXC Product SJM Confirm, Implantable Cardiac Monitor (ICM) Models: DM2100, DM2102 (IDE).************This product is a lead less non-therapeutic implantable cardiac monitor and is implanted to aid clinicians in diagnosing intermittent symptoms which may be due to arrhythmia Code Information Confirm DM2100, DM2102 (IDE). Software Version 10.1.1.2 Recalling Firm/ Manufacturer St Jude Medical CRMD 15900 Valley View Ct Sylmar CA 91342-3577 For Additional Information Contact 818-362-6822 Manufacturer Reason for Recall The recall was initiated because St. Jude Medical made new software available for the SJM Confirm DM2100 implantable cardiac monitor (ICM) that incorporates enhancements to the devices sensing algorithm. These enhancements are intended to provide improved sensitivity and specificity for diagnosing arrhythmias. The user is presented with the option to upgrade new software into the SJM Confirm ICM FDA Determined Cause 2 Software change control Action St. Jude Cardiac Rhythm Management Division sent an "Important Product Correction" letter on March 25, 2011 to all affected customers.************The letter explains the reason for the recall and provided recommendations including a software upgrade by SJM Field Clinical Engineer or Technical Services. Customers were recommended to fill out and return a "Physician Acknowledgement Form".************For additional information please contact your local St.Jude Medical Sales Representative, Field Clinical Engineer or the Technical Services Department at 800-722-3774 Quantity in Commerce 83 units Distribution Worldwide distribution: (USA) Nationwide including the States of CA, WA, FL, WV, NY, MD and LA, and the country of Germany Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MXC - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action St. Jude Cardiac Rhythm Management Division sent an "Important Product Correction" letter on March 25, 2011 to all affected customers.************The letter explains the reason for the recall and provided recommendations including a software upgrade by SJM Field Clinical Engineer or Technical Services. Customers were recommended to fill out and return a "Physician Acknowledgement Form".************For additional information please contact your local St.Jude Medical Sales Representative, Field Clinical Engineer or the Technical Services Department at 800-722-3774 | SoftwareChangeControl
139 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98797 | Z-2063-2011 | Optical Guidance Platform, version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03899; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA ******Optical Guidance Platform is for use with a charged particle accelerator to perform precise positioning of treatment target for stereotactic radiosurgery or radiotherapy treatments on cranial or extracranial lesions. | Optical Guidance Platform | 2 | CDRH | 2011-04-25 | 2011-04-25 21:00 | 2012-06-19 | Varian Medical Systems, Inc. Oncology Systems | The optical guidance platform may not be properly enforcing a 24 hour time limit between optical camera recalibrations. | Class 2 Device Recall Optical Guidance Platform Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Optical Guidance Platform FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Optical Guidance Platform Date Initiated by Firm March 03, 2011 Date Posted April 25, 2011 Recall Status 1 Terminated 3 on June 19, 2012 Recall Number Z-2063-2011 Recall Event ID 58246 510(K)Number K071360 Product Classification Accelerator, linear, medical - Product Code IYE Product Optical Guidance Platform, version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03899; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA ************Optical Guidance Platform is for use with a charged particle accelerator to perform precise positioning of treatment target for stereotactic radiosurgery or radiotherapy treatments on cranial or extracranial lesions. Code Information HZ10003 ***HZ10151 ***HZ10305 ***HZ10507 ***HZ10670 ***HZ11015 ***HZ12040 ***HZ12090 ***HZ13050 ***HZ17041 ***HZ18062 HZ10008 ***HZ10153 ***HZ10311 ***HZ10508 ***HZ10674 ***HZ11016 ***HZ12041 ***HZ12091 ***HZ13052 ***HZ17042 ***HZ18066 HZ10013 ***HZ10154 ***HZ10313 ***HZ10509 ***HZ10675 ***HZ11017 ***HZ12042 ***HZ12092 ***HZ13053 ***HZ17043 ***HZ18068 HZ10014 ***HZ10155 ***HZ10316 ***HZ10513 ***HZ10676 ***HZ11018 ***HZ12043 ***HZ12093 ***HZ13054 ***HZ17044 ***HZ18069 HZ10015 ***HZ10157 ***HZ10326 ***HZ10528 ***HZ10679 ***HZ11019 ***HZ12044 ***HZ12094 ***HZ13055 ***HZ17046 ***HZ18070 HZ10017 ***HZ10159 ***HZ10330 ***HZ10534 ***HZ10681 ***HZ11020 ***HZ12045 ***HZ12095 ***HZ13056 ***HZ17048 ***HZ18076 HZ10019 ***HZ10165 ***HZ10333 ***HZ10537 ***HZ10684 ***HZ11021 ***HZ12046 ***HZ12096 ***HZ13057 ***HZ17049 ***HZ18077 HZ10023 ***HZ10170 ***HZ10335 ***HZ10540 ***HZ10685 ***HZ11022 ***HZ12047 ***HZ12097 ***HZ13058 ***HZ17051 ***HZ18079 HZ10024 ***HZ10172 ***HZ10337 ***HZ10541 ***HZ10689 ***HZ11023 ***HZ12048 ***HZ12098 ***HZ13059 ***HZ17057 ***HZ18085 HZ10025 ***HZ10179 ***HZ10343 ***HZ10542 ***HZ10693 ***HZ11024 ***HZ12049 ***HZ12099 ***HZ13060 ***HZ17058 ***HZ18088 HZ10026 ***HZ10181 ***HZ10347 ***HZ10544 ***HZ10698 ***HZ11025 ***HZ12050 ***HZ12100 ***HZ13062 ***HZ17059 ***HZ18118 HZ10027 ***HZ10182 ***HZ10352 ***HZ10545 ***HZ10700 ***HZ12001 ***HZ12051 ***HZ12101 ***HZ13063 ***HZ17061 ***HZ18120 HZ10029 ***HZ10183 ***HZ10357 ***HZ10547 ***HZ10703 ***HZ12002 ***HZ12052 ***HZ12102 ***HZ13064 ***HZ17062 ***HZ18126 HZ10044 ***HZ10185 ***HZ10368 ***HZ10552 ***HZ10706 ***HZ12003 ***HZ12054 ***HZ12103 ***HZ15006 ***HZ17065 ***HZ18131 HZ10046 ***HZ10187 ***HZ10372 ***HZ10553 ***HZ10712 ***HZ12004 ***HZ12056 ***HZ12104 ***HZ15070 ***HZ17067 ***HZ18132 HZ10048 ***HZ10188 ***HZ10376 ***HZ10557 ***HZ10715 ***HZ12005 ***HZ12057 ***HZ12107 ***HZ15118 ***HZ17068 ***HZ18134 HZ10051 ***HZ10190 ***HZ10378 ***HZ10560 ***HZ10717 ***HZ12006 ***HZ12058 ***HZ12108 ***HZ15480 ***HZ17069 ***HZ18142 HZ10052 ***HZ10194 ***HZ10379 ***HZ10562 ***HZ10722 ***HZ12007 ***HZ12059 ***HZ12125 ***HZ16006 ***HZ17070 ***HZ18143 HZ10054 ***HZ10196 ***HZ10387 ***HZ10568 ***HZ10724 ***HZ12008 ***HZ12060 ***HZ12126 ***HZ16069 ***HZ17071 ***HZ18155 HZ10055 ***HZ10201 ***HZ10392 ***HZ10580 ***HZ10728 ***HZ12009 ***HZ12061 ***HZ12154 ***HZ16070 ***HZ17072 ***HZ18175 HZ10058 ***HZ10203 ***HZ10394 ***HZ10584 ***HZ10730 ***HZ12010 ***HZ12062 ***HZ12188 ***HZ17001 ***HZ17074 ***HZ18188 HZ10061 ***HZ10205 ***HZ10398 ***HZ10591 ***HZ10731 ***HZ12011 ***HZ12063 ***HZ12190 ***HZ17002 ***HZ17078 ***HZ18193 HZ10063 ***HZ10209 ***HZ10401 ***HZ10592 ***HZ10733 ***HZ12012 ***HZ12064 ***HZ13006 ***HZ17003 ***HZ17079 ***HZ19001 HZ10065 ***HZ10210 ***HZ10403 ***HZ10594 ***HZ10737 ***HZ12013 ***HZ12065 ***HZ13009 ***HZ17006 ***HZ17081 HZ19003 HZ10068 ***HZ10220 ***HZ10404 ***HZ10598 ***HZ10740 ***HZ12015 ***HZ12066 ***HZ13013 ***HZ17007 ***HZ17106 ***HZ19004 HZ10070 ***HZ10228 ***HZ10408 ***HZ10601 ***HZ10743 ***HZ12016 ***HZ12067 ***HZ13019 ***HZ17010 ***HZ17118 ***HZ19005 HZ10073 ***HZ10237 ***HZ10409 ***HZ10607 ***HZ10744 ***HZ12018 ***HZ12068 ***HZ13020 ***HZ17011 ***HZ17126 ***HZ19006 HZ10076 ***HZ10239 ***HZ10411 ***HZ10609 ***HZ10759 ***HZ12019 ***HZ12069 ***HZ13021 ***HZ17012 ***HZ17137 ***HZ19007 HZ10083 ***HZ10240 ***HZ10412 ***HZ10611 ***HZ10762 ***HZ12020 ***HZ12070 ***HZ13022 ***HZ17013 ***HZ17142 ***HZ19010 HZ10084 ***HZ10242 ***HZ10425 ***HZ10612 ***HZ10764 ***HZ12021 ***HZ12071 ***HZ13023 ***HZ17016 ***HZ17175 ***HZ19012 HZ10086 ***HZ10245 ***HZ10426 ***HZ10613 ***HZ10766 ***HZ12022 ***HZ12072 ***HZ13024 ***HZ17017 ***HZ18001 ***HZ19013 HZ10088 ***HZ10247 ***HZ10430 ***HZ10617 ***HZ10771 ***HZ12023 ***HZ12073 ***HZ13026 ***HZ17018 ***HZ18002 ***HZ19014 HZ10090 ***HZ10249 ***HZ10434 ***HZ10620 ***HZ10774 ***HZ12024 ***HZ12074 ***HZ13029 ***HZ17019 ***HZ18004 ***HZ19015 HZ10092 ***HZ10254 ***HZ10438 ***HZ10627 ***HZ10776 ***HZ12026 ***HZ12075 ***HZ13030 ***HZ17020 ***HZ18006 ***HZ19017 HZ10095 ***HZ10257 ***HZ10442 ***HZ10629 ***HZ10778 ***HZ12027 ***HZ12076 ***HZ13031 ***HZ17021 ***HZ18013 ***HZ19018 HZ10098 ***HZ10258 ***HZ10450 ***HZ10632 ***HZ11001 ***HZ12028 ***HZ12077 ***HZ13032 ***HZ17023 ***HZ18018 ***HZ19019 HZ10102 ***HZ10260 ***HZ10453 ***HZ10640 ***HZ11002 ***HZ12029 ***HZ12078 ***HZ13038 ***HZ17024 ***HZ18020 ***HZ19020 HZ10108 ***HZ10271 ***HZ10460 ***HZ10644 ***HZ11003 ***HZ12030 ***HZ12080 ***HZ13039 ***HZ17027 ***HZ18022 ***HZ19022 HZ10110 ***HZ10273 ***HZ10467 ***HZ10646 ***HZ11004 ***HZ12031 ***HZ12081 ***HZ13041 ***HZ17028 ***HZ18024 ***HZ19023 HZ10116 ***HZ10274 ***HZ10479 ***HZ10647 ***HZ11005 ***HZ12032 ***HZ12082 ***HZ13042 ***HZ17030 ***HZ18026 ***HZ19025 HZ10117 ***HZ10283 ***HZ10489 ***HZ10649 ***HZ11006 ***HZ12033 ***HZ12083 ***HZ13043 ***HZ17032 ***HZ18035 ***HZ19026 HZ10121 ***HZ10293 ***HZ10493 ***HZ10650 ***HZ11007 ***HZ12034 ***HZ12084 ***HZ13044 ***HZ17034 ***HZ18040 ***HZ19028 HZ10126 ***HZ10294 ***HZ10495 ***HZ10652 ***HZ11008 ***HZ12035 ***HZ12085 ***HZ13045 ***HZ17035 ***HZ18041 ***HZ19029 HZ10138 ***HZ10295 ***HZ10496 ***HZ10656 ***HZ11010 ***HZ12036 ***HZ12086 ***HZ13046 ***HZ17036 ***HZ18042 ***HZ19033 HZ10144 ***HZ10297 ***HZ10497 ***HZ10659 ***HZ11011 ***HZ12037 ***HZ12087 ***HZ13047 ***HZ17037 ***HZ18043 ***HZ19034 HZ10146 ***HZ10299 ***HZ10500 ***HZ10663 ***HZ11013 ***HZ12038 ***HZ12088 ***HZ13048 ***HZ17039 ***HZ18046 ***HZ19035 HZ10148 ***HZ10303 ***HZ10501 ***HZ10665 ***HZ11014 ***HZ12039 ***HZ12089 ***HZ13049 ***HZ17040 ***HZ18048 ***HZ19036 HZ19040 ***HZ19041 ***HZ19043 ***HZ19044 ***HZ19045 ***HZ19046 ***HZ19051 ***HZ19052 ***HZ19053 ***HZ19055 ***HZ19056 HZ19059 ***HZ19060 ***HZ19062 ***HZ19065 ***HZ19066 ***HZ19068 ***HZ19069 ***HZ19071 ***HZ19072 ***HZ19076 ***HZ19078 HZ19079 ***HZ19086 ***HZ19087 ***HZ19088 ***HZ19089 ***HZ19091 ***HZ19095 ***HZ19096 ***HZ19098 ***HZ19099 ***HZ19104 HZ19106 ***HZ19109 ***HZ19113 ***HZ19114 ***HZ19115 ***HZ19117 ***HZ19118 ***HZ19119 ***HZ19121 ***HZ19125 ***HZ19126 HZ19128 ***HZ19129 ***HZ19130 ***HZ19131 ***HZ19133 ***HZ19134 ***HZ19135 ***HZ19140 ***HZ19141 ***HZ19142 ***HZ19143 HZ19155 ***HZ19158 ***HZ19159 ***HZ19162 ***HZ19167 ***HZ19170 ***HZ19171 *** HZ19172 ***HZ19179 ***HZ19186 ***HZ19189 HZ19190 ***HZ19193 ***HZ19194 ***HZ19195 ***HZ19201 ***HZ19226 ***HZ19227 ***HZ19228 ***HZ19229 ***HZ19235 ***HZ19236 HZ19611 ***HZ19926 ***HZ19996 ***HZ19997 ***HZ19998 Recalling Firm/ Manufacturer Varian Medical Systems, Inc. Oncology Systems 911 Hansen Way Palo Alto CA 94304-1028 For Additional Information Contact 650-424-5731 Manufacturer Reason for Recall The optical guidance platform may not be properly enforcing a 24 hour time limit between optical camera recalibrations. FDA Determined Cause 2 Software change control Action The firm, Varian, sent an "URGENT MEDICAL DEVICE CORRECTION-URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check the Calibration timeout on their optical guidance platform. If this exceeds 24 hour, please set the factory default of 16 hours to ensure that the camera will be calibrated daily. The instructions are located in the document. The customers were also instructed to download the most recent version of the Optical Guidance Platform Instructions for Use at www.MyVarian.com or request a copy from their local Varian Customer Support District or Regional Manager. In addition, the customers were instructed to advise the appropriate personnel working in their radiotherapy department of the content of this letter. ************Varian is preparing a software patch which will be applied to affected units to correct this issue.************If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager; USA and Canada -1888-VARIAN5 (888-827-4265) or Europe- +41 41 749 8844. Quantity in Commerce 599 units Distribution Worldwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, Varian, sent an "URGENT MEDICAL DEVICE CORRECTION-URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check the Calibration timeout on their optical guidance platform. If this exceeds 24 hour, please set the factory default of 16 hours to ensure that the camera will be calibrated daily. The instructions are located in the document. The customers were also instructed to download the most recent version of the Optical Guidance Platform Instructions for Use at www.MyVarian.com or request a copy from their local Varian Customer Support District or Regional Manager. In addition, the customers were instructed to advise the appropriate personnel working in their radiotherapy department of the content of this letter. ************Varian is preparing a software patch which will be applied to affected units to correct this issue.************If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager; USA and Canada -1888-VARIAN5 (888-827-4265) or Europe- +41 41 749 8844. | SoftwareChangeControl
140 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=80569 | Z-2000-2011 | Siemens LANTIS 6.1 Oncology Information System with MEDOne Module, Linear Medical Accelerators; Product Numbers: 4503178, 5835991, 5851055; Product is manufactured by IMPAC Medical, Inc., Sunnyvale, CA.******Medical charged-particle radiation therapy system. | LANTIS 6.1 Oncology Information System with MEDOne Module | 2 | CDRH | 2011-04-15 | 2011-04-15 21:00 | 2011-04-15 | Siemens Medical Solutions USA, Inc | The product has a software problem; it does not calculate the Creatinine Clearance correctly for patients less than two years old. | Class 2 Device Recall LANTIS 6.1 Oncology Information System with MEDOne Module Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall LANTIS 6.1 Oncology Information System with MEDOne Module FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall LANTIS 6.1 Oncology Information System with MEDOne Module Date Initiated by Firm March 06, 2009 Date Posted April 15, 2011 Recall Status 1 Terminated 3 on April 15, 2011 Recall Number Z-2000-2011 Recall Event ID 51504 510(K)Number K060226 Product Classification Linear Medical Accelerator - Product Code IYE Product Siemens LANTIS 6.1 Oncology Information System with MEDOne Module, Linear Medical Accelerators; Product Numbers: 4503178, 5835991, 5851055; Product is manufactured by IMPAC Medical, Inc., Sunnyvale, CA.************Medical charged-particle radiation therapy system. Code Information All units of these product numbers. Software version 6.1 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 4040 Nelson Ave Concord CA 94520-1200 For Additional Information Contact Mark Burgess 925-246-8407 Manufacturer Reason for Recall The product has a software problem; it does not calculate the Creatinine Clearance correctly for patients less than two years old. FDA Determined Cause 2 Software change control Action On 3/6/2009, a Safety Advisory Letter was distributed to all affected users, with a description of the problem and identifying the affected product. Customers were instructed not to use the LANTIS to calculate Creatine Clearance for any patient under two years of age. Customers should direct their questions with regards to the letter to their local Siemens Service organization. A Siemens Representative will update the systems. Quantity in Commerce 17 Distribution Worldwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 3/6/2009, a Safety Advisory Letter was distributed to all affected users, with a description of the problem and identifying the affected product. Customers were instructed not to use the LANTIS to calculate Creatine Clearance for any patient under two years of age. Customers should direct their questions with regards to the letter to their local Siemens Service organization. A Siemens Representative will update the systems. | SoftwareChangeControl
141 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=88565 | Z-1932-2011 | Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program, 1600 and 2000 tests Cartridge Sets. Catalog Numbers: 270-2417 and 270-2415; Distributed and Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CA***The Intended Use is as follows: The Bior-Rad II TURBO Hemoglobin A1c Program is intended for the percent determination of hemoglobin Alc in human whole blood using ion-exchange high-performance liquid chromatography (HPLC). The Bio-Rad VARIANT II. TURBO Hemoglobin Alc Program is intended for Professional Use Only--For In Vitro Diagnostic Use. Reagent Kit Information: The VARIANT II TURBO Hemoglobin A1c program utilizes principles of ion¿exchange high performance liquid chromatography (HPLC). The VARIANT II Hemoglobin A1c Program is based on chromatographic separation of HbAlc on a cation exchange cartridge. System Overview: The VARIANT II TURBO Hemoglobin testing System provides an integrated method for sample preparation, separation and the percent determination of percent hemoglobin Alc in EDTA human whole blood. Technology Overview: The V ARIANT II TURBO Hemoglobin Testing is a fully automated, high-¿throughput hemoglobin analyzer. It consists of two modules- the VARIANT II Chromatographic Station (VCS) and the VARIANT II Sampling Station (VSS). In addition, a personal computer is used to control the VARIANT II TURBO System using Clinical Data Management (CDM) software. CDM is a dedicated software package of instrument control, data control, results analysis, calibration, quality control and service software. | Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program | 2 | CDRH | 2011-04-11 | 2011-04-11 21:00 | 2011-08-30 | Bio-Rad Laboratories Inc | Some customers have reported the occurrence of some ramping baselines on the chromatograms. The ramping baseline can affect hemoglobin test results. | Class 2 Device Recall Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program Date Initiated by Firm December 01, 2009 Date Posted April 11, 2011 Recall Status 1 Terminated 3 on August 30, 2011 Recall Number Z-1932-2011 Recall Event ID 54483 Product Classification Assay, glycoslylated hemoglobin - Product Code LCP Product Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program, 1600 and 2000 tests Cartridge Sets. Catalog Numbers: 270-2417 and 270-2415; Distributed and Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CA******The Intended Use is as follows: The Bior-Rad II TURBO Hemoglobin A1c Program is intended for the percent determination of hemoglobin Alc in human whole blood using ion-exchange high-performance liquid chromatography (HPLC). The Bio-Rad VARIANT II. TURBO Hemoglobin Alc Program is intended for Professional Use Only--For In Vitro Diagnostic Use. Reagent Kit Information: The VARIANT II TURBO Hemoglobin A1c program utilizes principles of ionexchange high performance liquid chromatography (HPLC). The VARIANT II Hemoglobin A1c Program is based on chromatographic separation of HbAlc on a cation exchange cartridge. System Overview: The VARIANT II TURBO Hemoglobin testing System provides an integrated method for sample preparation, separation and the percent determination of percent hemoglobin Alc in EDTA human whole blood. Technology Overview: The V ARIANT II TURBO Hemoglobin Testing is a fully automated, high-throughput hemoglobin analyzer. It consists of two modules- the VARIANT II Chromatographic Station (VCS) and the VARIANT II Sampling Station (VSS). In addition, a personal computer is used to control the VARIANT II TURBO System using Clinical Data Management (CDM) software. CDM is a dedicated software package of instrument control, data control, results analysis, calibration, quality control and service software. Code Information Catalog No. 270-2415: LOTS: Lot 70291500 Exp. June 30, 2010; Lot 70292239 Exp. June 30, 2010; and Lot 70292240 Exp. June 30, 2010. Catalog No. 270-2417: LOTS: Lot 70291520 Exp. Dec. 2, 2009; Lot 70292241 Exp.Dec.2, 2009; Lot 70291520R Exp. Jan. 31, 2010; Lot 70292242 Exp.Jan.31, 2010; and Lot 70292318 Exp. Jan. 31, 2010 Recalling Firm/ Manufacturer Bio-Rad Laboratories Inc 4000 Alfred Nobel Dr Hercules CA 94547-1803 For Additional Information Contact Jackie Buckley 510-724-7000 Manufacturer Reason for Recall Some customers have reported the occurrence of some ramping baselines on the chromatograms. The ramping baseline can affect hemoglobin test results. FDA Determined Cause 2 Software change control Action Bio Rad Laboratories, Inc. sent an Urgent Medical Device Correction letter dated December 1, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediatey examine their records to identify all customers who received one of the affected lots.******Collect a Customer Medical Device Correction Response Form from every customer.******Once forms were received from all applicable customers, complete the following Medical Device Correction Response Form and fax it to Bio-Rad CSD Regulatory Affairs at 510-741-3954 or pdf to CSD Recall Coordinator (excluding USSD).******For questions regarding this recall call 510-724-7000. Quantity in Commerce 1099 Distribution Worldwide Distribution - USA (nationwide) and the countries of Australia, New Zealand, Taiwan, Singapore, India, China, S. Korea, Hong Kong, France, Sweden, United Kingdom, Italy, Hungary, Germany, Brazil, Mexico and Colombia. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Bio Rad Laboratories, Inc. sent an Urgent Medical Device Correction letter dated December 1, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediatey examine their records to identify all customers who received one of the affected lots.******Collect a Customer Medical Device Correction Response Form from every customer.******Once forms were received from all applicable customers, complete the following Medical Device Correction Response Form and fax it to Bio-Rad CSD Regulatory Affairs at 510-741-3954 or pdf to CSD Recall Coordinator (excluding USSD).******For questions regarding this recall call 510-724-7000. | SoftwareChangeControl
142 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98675 | Z-1901-2011 | Prismaflex System, Catalog Number 107493, ******Indicated for continuous solute and/or fluid removal in patients with acute renal failure or fluid overload. | Prismaflex | 2 | CDRH | 2011-04-06 | 2011-04-06 21:00 | 2011-09-01 | Gambro Renal Products, Incorporated | Prismaflex touch screen may become unresponsive during Continuous Renal Replacement Therapy procedure. | Class 2 Device Recall Prismaflex Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Prismaflex FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Prismaflex Date Initiated by Firm March 14, 2011 Date Posted April 06, 2011 Recall Status 1 Terminated 3 on September 01, 2011 Recall Number Z-1901-2011 Recall Event ID 58208 510(K)Number K072093 Product Classification Dialyzer, high permeability with or without sealed dialysate system - Product Code KDI Product Prismaflex System, Catalog Number 107493, ************Indicated for continuous solute and/or fluid removal in patients with acute renal failure or fluid overload. Code Information All units with Version 3.20 software. Recalling Firm/ Manufacturer Gambro Renal Products, Incorporated 14143 Denver West Pkwy Lakewood CO 80401-3266 For Additional Information Contact 303-222-6566 Manufacturer Reason for Recall Prismaflex touch screen may become unresponsive during Continuous Renal Replacement Therapy procedure. FDA Determined Cause 2 Software change control Action The firm, GAMBRO Renal Products, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated March 14, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were provided with response procedures to use if the frozen screen problem occurs while using the product. The customers were instructed to closely monitor the return line for potential air and/or clots in the circuit; review the instructions provided in the attached copy of Chapter 5: Troubleshooting, "Softkeys won't work." of the Prismaflex Operators Manual in conjunction with the attached Troubleshooting Checklist; be aware if performing Manual Termination with Blood Return, air detection is not provided, and complete and return the Customer Reply Form via fax to: Regulatory Affairs at 1-866-447-4471 or 303-222-6916, or by email to: GRPRegulatory.Affairs@us.gambro.com. ************If you have any questions or concerns, please contact for Clinical questions: ICON 800-554-1312 (available 24/7) or your local Gambro Representative; for Technical questions: Gambro Customer Support 800-525-2623, option 2 (Available 6:00am to 5:30pm MT, Monday thru Friday). Quantity in Commerce 1300 units Distribution Nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = KDI - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, GAMBRO Renal Products, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated March 14, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were provided with response procedures to use if the frozen screen problem occurs while using the product. The customers were instructed to closely monitor the return line for potential air and/or clots in the circuit; review the instructions provided in the attached copy of Chapter 5: Troubleshooting, "Softkeys won't work." of the Prismaflex Operators Manual in conjunction with the attached Troubleshooting Checklist; be aware if performing Manual Termination with Blood Return, air detection is not provided, and complete and return the Customer Reply Form via fax to: Regulatory Affairs at 1-866-447-4471 or 303-222-6916, or by email to: GRPRegulatory.Affairs@us.gambro.com. ************If you have any questions or concerns, please contact for Clinical questions: ICON 800-554-1312 (available 24/7) or your local Gambro Representative; for Technical questions: Gambro Customer Support 800-525-2623, option 2 (Available 6:00am to 5:30pm MT, Monday thru Friday). | SoftwareChangeControl
143 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=80551 | Z-1900-2011 | Siemens brand ACUSON S2000 Ultrasound System with software versions:***VA10, VA10A, VA10B, VA10C, VA15; Model Number: 10041461;***Product is manufactured and distributed by Siemens Medical Solutions USA, Inc. Ultrasound Business Unit, Mountain View, CA******The ACUSON SC2000 ultrasound imaging system is intended for the following applications: Cardiac, Neo-natal and Fetal cardiac, Pediatric, Transesophageal, Adult Cephalic, Peripheral Vessel, Intraoperative Neurological, Musculo-skeletal Conventional, and Musculo-skeletal Superficial applications. | ACUSON S2000 Ultrasound System | 2 | CDRH | 2011-04-06 | 2011-04-06 21:00 | 2011-04-06 | Siemens Medical Solutions USA, Inc. | The product has a software problem in which previous patient measurement data gets associated with another patient's image. | Class 2 Device Recall ACUSON S2000 Ultrasound System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ACUSON S2000 Ultrasound System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ACUSON S2000 Ultrasound System Date Initiated by Firm January 14, 2009 Date Posted April 06, 2011 Recall Status 1 Terminated 3 on April 06, 2011 Recall Number Z-1900-2011 Recall Event ID 51497 510(K)Number K082142 Product Classification Counter, differential cell - Product Code GKZ Product Siemens brand ACUSON S2000 Ultrasound System with software versions:******VA10, VA10A, VA10B, VA10C, VA15; Model Number: 10041461;******Product is manufactured and distributed by Siemens Medical Solutions USA, Inc. Ultrasound Business Unit, Mountain View, CA************The ACUSON SC2000 ultrasound imaging system is intended for the following applications: Cardiac, Neo-natal and Fetal cardiac, Pediatric, Transesophageal, Adult Cephalic, Peripheral Vessel, Intraoperative Neurological, Musculo-skeletal Conventional, and Musculo-skeletal Superficial applications. Code Information All units of this model. Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc. 1230 Shorebird Way P.O. Box 7393 Mountain View CA 94043 For Additional Information Contact 650-694-5993 Manufacturer Reason for Recall The product has a software problem in which previous patient measurement data gets associated with another patient's image. FDA Determined Cause 2 Software change control Action The firm, SIEMENS, issued a "CUSTOMER SAFETY ADVISORY NOTIFICATION" letter to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that the first image in Exam Review is from the current patient; ensure that the Heart Cycle measurement tool is not selected on the Cale menu when entering Exam Review; be aware of this issue as described in the letter and perform four (4) or less measurements of the same measurement label. ************Should you have any questions, please call (650) 694-5398. Distribution Worldwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = GKZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: |  | SoftwareChangeControl
144 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94615 | Z-1837-2011 | ACUSON S2000 ultrasound systems (all versions) with cardiovascular options installed;***ACUSON S2000: 10041461;***S2000 Cardio Option: 10041853;***Product is manufactured and distributed by ***Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA***The ACUSON S2000 Ultrasound System is a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or Harmonic Imaging and 3D/4D Imaging on a Flat Panel Display. The ACUSON S2000 ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides the ability to measure anatomical structures {fetal, abdominal, intraoperative, intraoperative neurological, pediatric, small organ, neonatal cephalic, adult cephalic, cardiac, trans-esophageal, transrectal, transvaginal, peripheral vessel, musculoskeletal (conventional), musculo-skeletal (superficial) and neonatal cardiac} and calculation packages that provide information that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. *** | ACUSON S2000 ultrasound systems | 2 | CDRH | 2011-03-28 | 2011-03-28 21:00 | 2012-06-18 | Siemens Medical Solutions USA, Inc., Mountain View, CA | There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow. | Class 2 Device Recall ACUSON S2000 ultrasound systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ACUSON S2000 ultrasound systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ACUSON S2000 ultrasound systems Date Initiated by Firm August 17, 2010 Date Posted March 28, 2011 Recall Status 1 Terminated 3 on June 18, 2012 Recall Number Z-1837-2011 Recall Event ID 56840 510(K)Number K082142 Product Classification System, imaging, pulsed dopper, ultrasonic - Product Code IYO Product ACUSON S2000 ultrasound systems (all versions) with cardiovascular options installed;******ACUSON S2000: 10041461;******S2000 Cardio Option: 10041853;******Product is manufactured and distributed by ******Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA******The ACUSON S2000 Ultrasound System is a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or Harmonic Imaging and 3D/4D Imaging on a Flat Panel Display. The ACUSON S2000 ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides the ability to measure anatomical structures {fetal, abdominal, intraoperative, intraoperative neurological, pediatric, small organ, neonatal cephalic, adult cephalic, cardiac, trans-esophageal, transrectal, transvaginal, peripheral vessel, musculoskeletal (conventional), musculo-skeletal (superficial) and neonatal cardiac} and calculation packages that provide information that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. Code Information All units of these models. Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc., Mountain View, CA 1230 Shorebird Way P.O. Box 7393 Mountain View CA 94043 For Additional Information Contact Ms. Sheila Pickering 650-694-5993 Manufacturer Reason for Recall There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow. FDA Determined Cause 2 Software change control Action Siemens Medical Solutions USA, Inc., sent a Customer Safety Advisory Notification letter on August 17, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. ******Customers were instructed to pass the notice on to all those within their organization who need to aware of the issue until the corrective action was completed. ******For questions customers were to contact their local service support person. ******For questions regarding this recall call 650-694-5993. Distribution Worldwide Distribution - USA (nationwide)and the countries of Antares distribution includes: United Arab Emirates-AE, Austria-AT, Australia-AU, Azerbaijan-AZ, Brazil-BR, Belarus-BY, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Egypt-EG, Spain-ES, France-FR, Great Britain-GB, Guatemala-GT, Croatia-HR, Hungary-HU, Indonesia-ID, Ireland-IE, India-IN, Italy-IT, Jordan-JO, Japan-JP, South Korea-KR, Moldova-MD, Mexico-MX, Malaysia-MY, Netherlands ***NL, Norway-NO, New Zealand-NZ, Poland-PL, Portugal-PT, Romania-RO, Serbia-RS, Russian Federation-RU, Saudi Arabia-SA, Sweden-SE, Singapore-SG, Slovak Republic-SK, Suriname-SR, Turkey-TR, Taiwan-TW, Ukraine-UA, United States-US, Uruguay-UY, Uzbekistan-UZ, Venezuela-VE, Vietnam-VN, South Africa-ZA S2000 distribution includes: Austria-AT, Australia-AU, Brazil-BR, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Spain-ES, Finland-FI, France-FR, Great Britain-GB, Hong Kong-HK, Indonesia-ID, India-IN, Italy-IT, Japan-JP, South Korea-KR, Kazakhstan-KZ, Lebanon-LB, Mexico-MX, Norway-NO, Philippines-PH, Poland-PL, Portugal-PT, Romania-RO, Russian Federation-RU, Sweden-SE, Slovak Republic-SK, Turkmenistan-TM, Turkey-TR, Taiwan-TW, United States-US X300 distribution includes: AE (United Arab Emirates) , AL (Albania) , AN (Netherlands Antilles), AO (Angola), AR (Argentina), AT (Austria), AU (Australia), AZ (Azerbaijan), BA (Bosnia-Herzegovina), BD (Bangladesh), BE (Belgium), BG (Bulgaria), BN (Brunei), BO (Bolivia), BR (Brazil), BY (Belarus), CA (Canada), CH (Switzerland), CL (Chile), CN (China), CO (Colombia), CV (Cape Verde), CZ (Czech Republic), DE (Germany), DK (Denmark), DZ (Algeria), EC (Ecuador), EG (Egypt), ES (Spain), ET (Ethiopia), FJ (Fiji), FR (France), GB (United Kingdom), GE (Georgia), GR (Greece), HK (Hong Kong), HN (Honduras), HR (Croatia), HU (Hungary), ID (Indonesia), IE (Ireland), IN (India), IR (Iran), IT (Italy), JO (Jordan), JP (Japan), KE (Kenya), KR (Korea, South), KW (Kuwait), KZ (Kazakhstan), LT (Lithuania), LV (Latvia), LY (Libya), MD (Moldova), ME (Montenegro), MK (Macedonia), MM (Myanmar), MU (Mauritius), MX (Mexico), MY (Malaysia), NG (Nigeria), NI (Nicaragua), NL (Netherlands), NO (Norway), NP (Nepal), NZ (New Zealand), PE (Per¿), PH (Philippines), PK (Pakistan), PL (Poland), PS (Palestine), PT (Portugal), QA (Qatar), RS (Serbia), RO (Romania), RU (Russia), RW (Rwanda), SA (Saudi Arabia), SC (Seychelles), SE (Sweden), SG (Singapore), SI (Slovenia), SK (Slovakia), SV (El Salvador), SY (Syria), TH (Thailand), TM (Turkmenistan), TN (Tunisia), TR (Turkey), TT (Trinidad/Tobago), TW (Taiwan, ROC), UA (Ukraine), US (United States of America), UY (Uruguay), UZ (Uzbekistan), VE (Venezuela), VN (Viet Nam), YE (Yemen), ZA and (South Africa) Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYO - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: |  | SoftwareChangeControl
145 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94614 | Z-1836-2011 | ACUSON and SONOLINE Antares and Antares PE ultrasound systems at product versions 4.0 and 5.0 with cardiovascular options installed;***ACUSON Antares 4.0 & 5.0: 10032747, 10037592, 10040728;***ACUSON Antares PE 4.0 & 5.0: 10032746, 10037591, 10038202, 10040729;***SONOLINE Antares 4.0 & 5.0: 05936518, 08653771, 10037593;***Antares Cardio Options: 10036613, 10036588, 10036585, 10436250;***Product is manufactured and distributed by ***Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA***The ACUSON Antare is a general purpose Diagnostic Ultrasound system. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in: B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, Harmonic Imaging, or 3D/4D imaging on a FPD display. The Antares ultrasound imaging system is intended for the following applications: Abdominal, Intraoperative, Small Parts, Transcranial, OB/GYN, Cardiac, Transesophageal, Pelvic, Neonatal/Adult Cephalic, Vascular, Intravascular, Musculoskeletal, Superficial Musculoskeletal, Great Vessel, and Peripheral Vascular applications. The system also provides for the measurement of anatomical structures that provide information that is used for clinical diagnosis purposes. *** | ACUSON and SONOLINE Antares and Antares PE ultrasound systems | 2 | CDRH | 2011-03-28 | 2011-03-28 21:00 | 2012-06-18 | Siemens Medical Solutions USA, Inc., Mountain View, CA | There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow. | Class 2 Device Recall ACUSON and SONOLINE Antares and Antares PE ultrasound systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ACUSON and SONOLINE Antares and Antares PE ultrasound systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ACUSON and SONOLINE Antares and Antares PE ultrasound systems Date Initiated by Firm August 17, 2010 Date Posted March 28, 2011 Recall Status 1 Terminated 3 on June 18, 2012 Recall Number Z-1836-2011 Recall Event ID 56840 510(K)Number K063803 Product Classification System, imaging, pulsed doppler ultrasonic - Product Code IYN Product ACUSON and SONOLINE Antares and Antares PE ultrasound systems at product versions 4.0 and 5.0 with cardiovascular options installed;******ACUSON Antares 4.0 & 5.0: 10032747, 10037592, 10040728;******ACUSON Antares PE 4.0 & 5.0: 10032746, 10037591, 10038202, 10040729;******SONOLINE Antares 4.0 & 5.0: 05936518, 08653771, 10037593;******Antares Cardio Options: 10036613, 10036588, 10036585, 10436250;******Product is manufactured and distributed by ******Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA******The ACUSON Antare is a general purpose Diagnostic Ultrasound system. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in: B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, Harmonic Imaging, or 3D/4D imaging on a FPD display. The Antares ultrasound imaging system is intended for the following applications: Abdominal, Intraoperative, Small Parts, Transcranial, OB/GYN, Cardiac, Transesophageal, Pelvic, Neonatal/Adult Cephalic, Vascular, Intravascular, Musculoskeletal, Superficial Musculoskeletal, Great Vessel, and Peripheral Vascular applications. The system also provides for the measurement of anatomical structures that provide information that is used for clinical diagnosis purposes. Code Information All units of these models. Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc., Mountain View, CA 1230 Shorebird Way P.O. Box 7393 Mountain View CA 94043 For Additional Information Contact Ms. Sheila Pickering 650-694-5993 Manufacturer Reason for Recall There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow. FDA Determined Cause 2 Software change control Action Siemens Medical Solutions USA, Inc., sent a Customer Safety Advisory Notification letter on August 17, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. ******Customers were instructed to pass the notice on to all those within their organization who need to aware of the issue until the corrective action was completed. ******For questions customers were to contact their local service support person. ******For questions regarding this recall call 650-694-5993. Distribution Worldwide Distribution - USA (nationwide)and the countries of Antares distribution includes: United Arab Emirates-AE, Austria-AT, Australia-AU, Azerbaijan-AZ, Brazil-BR, Belarus-BY, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Egypt-EG, Spain-ES, France-FR, Great Britain-GB, Guatemala-GT, Croatia-HR, Hungary-HU, Indonesia-ID, Ireland-IE, India-IN, Italy-IT, Jordan-JO, Japan-JP, South Korea-KR, Moldova-MD, Mexico-MX, Malaysia-MY, Netherlands ***NL, Norway-NO, New Zealand-NZ, Poland-PL, Portugal-PT, Romania-RO, Serbia-RS, Russian Federation-RU, Saudi Arabia-SA, Sweden-SE, Singapore-SG, Slovak Republic-SK, Suriname-SR, Turkey-TR, Taiwan-TW, Ukraine-UA, United States-US, Uruguay-UY, Uzbekistan-UZ, Venezuela-VE, Vietnam-VN, South Africa-ZA S2000 distribution includes: Austria-AT, Australia-AU, Brazil-BR, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Spain-ES, Finland-FI, France-FR, Great Britain-GB, Hong Kong-HK, Indonesia-ID, India-IN, Italy-IT, Japan-JP, South Korea-KR, Kazakhstan-KZ, Lebanon-LB, Mexico-MX, Norway-NO, Philippines-PH, Poland-PL, Portugal-PT, Romania-RO, Russian Federation-RU, Sweden-SE, Slovak Republic-SK, Turkmenistan-TM, Turkey-TR, Taiwan-TW, United States-US X300 distribution includes: AE (United Arab Emirates) , AL (Albania) , AN (Netherlands Antilles), AO (Angola), AR (Argentina), AT (Austria), AU (Australia), AZ (Azerbaijan), BA (Bosnia-Herzegovina), BD (Bangladesh), BE (Belgium), BG (Bulgaria), BN (Brunei), BO (Bolivia), BR (Brazil), BY (Belarus), CA (Canada), CH (Switzerland), CL (Chile), CN (China), CO (Colombia), CV (Cape Verde), CZ (Czech Republic), DE (Germany), DK (Denmark), DZ (Algeria), EC (Ecuador), EG (Egypt), ES (Spain), ET (Ethiopia), FJ (Fiji), FR (France), GB (United Kingdom), GE (Georgia), GR (Greece), HK (Hong Kong), HN (Honduras), HR (Croatia), HU (Hungary), ID (Indonesia), IE (Ireland), IN (India), IR (Iran), IT (Italy), JO (Jordan), JP (Japan), KE (Kenya), KR (Korea, South), KW (Kuwait), KZ (Kazakhstan), LT (Lithuania), LV (Latvia), LY (Libya), MD (Moldova), ME (Montenegro), MK (Macedonia), MM (Myanmar), MU (Mauritius), MX (Mexico), MY (Malaysia), NG (Nigeria), NI (Nicaragua), NL (Netherlands), NO (Norway), NP (Nepal), NZ (New Zealand), PE (Per¿), PH (Philippines), PK (Pakistan), PL (Poland), PS (Palestine), PT (Portugal), QA (Qatar), RS (Serbia), RO (Romania), RU (Russia), RW (Rwanda), SA (Saudi Arabia), SC (Seychelles), SE (Sweden), SG (Singapore), SI (Slovenia), SK (Slovakia), SV (El Salvador), SY (Syria), TH (Thailand), TM (Turkmenistan), TN (Tunisia), TR (Turkey), TT (Trinidad/Tobago), TW (Taiwan, ROC), UA (Ukraine), US (United States of America), UY (Uruguay), UZ (Uzbekistan), VE (Venezuela), VN (Viet Nam), YE (Yemen), ZA and (South Africa) Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: |  | SoftwareChangeControl
146 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94616 | Z-1838-2011 | ACUSON X300, X300 PE and SONOVISTA X300 (all versions) with cardiovascular options installed;***ACUSON X300: 10037409, 10132987, 10038837;***SONOVISTA X300: 10133170;***ACUSON X300 PE: 10348531, 10348532, 10348533;***X300 Cardiovascular Options: 10133037, 10348777, 10349032, 10427797, 10427822, 10427816, 10563788;***Product is manufactured and distributed by ***Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA***The Siemens Acuson X300 ultrasound imaging system is intended for the following applications: General Radiology, Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Neonatal/Adult Cephalic, Cardiac, Transesophageal, Pelvic, Transcranial, OB/GYN, Urology, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides for the measurement of anatomical structures and for analysis packages that provide information that is used for clinical diagnosis purposes. ****** | ACUSON X300, X300 PE and SONOVISTA X300 | 2 | CDRH | 2011-03-28 | 2011-03-28 21:00 | 2012-06-18 | Siemens Medical Solutions USA, Inc., Mountain View, CA | There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow. | Class 2 Device Recall ACUSON X300, X300 PE and SONOVISTA X300 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ACUSON X300, X300 PE and SONOVISTA X300 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ACUSON X300, X300 PE and SONOVISTA X300 Date Initiated by Firm August 17, 2010 Date Posted March 28, 2011 Recall Status 1 Terminated 3 on June 18, 2012 Recall Number Z-1838-2011 Recall Event ID 56840 510(K)Number K072676 Product Classification System, imaging, pulsed doppler, ultrasonic - Product Code ITX Product ACUSON X300, X300 PE and SONOVISTA X300 (all versions) with cardiovascular options installed;******ACUSON X300: 10037409, 10132987, 10038837;******SONOVISTA X300: 10133170;******ACUSON X300 PE: 10348531, 10348532, 10348533;******X300 Cardiovascular Options: 10133037, 10348777, 10349032, 10427797, 10427822, 10427816, 10563788;******Product is manufactured and distributed by ******Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA******The Siemens Acuson X300 ultrasound imaging system is intended for the following applications: General Radiology, Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Neonatal/Adult Cephalic, Cardiac, Transesophageal, Pelvic, Transcranial, OB/GYN, Urology, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides for the measurement of anatomical structures and for analysis packages that provide information that is used for clinical diagnosis purposes. Code Information All units of these models. Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc., Mountain View, CA 1230 Shorebird Way P.O. Box 7393 Mountain View CA 94043 For Additional Information Contact Ms. Sheila Pickering 650-694-5993 Manufacturer Reason for Recall There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow. FDA Determined Cause 2 Software change control Action Siemens Medical Solutions USA, Inc., sent a Customer Safety Advisory Notification letter on August 17, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. ******Customers were instructed to pass the notice on to all those within their organization who need to aware of the issue until the corrective action was completed. ******For questions customers were to contact their local service support person. ******For questions regarding this recall call 650-694-5993. Distribution Worldwide Distribution - USA (nationwide)and the countries of Antares distribution includes: United Arab Emirates-AE, Austria-AT, Australia-AU, Azerbaijan-AZ, Brazil-BR, Belarus-BY, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Egypt-EG, Spain-ES, France-FR, Great Britain-GB, Guatemala-GT, Croatia-HR, Hungary-HU, Indonesia-ID, Ireland-IE, India-IN, Italy-IT, Jordan-JO, Japan-JP, South Korea-KR, Moldova-MD, Mexico-MX, Malaysia-MY, Netherlands ***NL, Norway-NO, New Zealand-NZ, Poland-PL, Portugal-PT, Romania-RO, Serbia-RS, Russian Federation-RU, Saudi Arabia-SA, Sweden-SE, Singapore-SG, Slovak Republic-SK, Suriname-SR, Turkey-TR, Taiwan-TW, Ukraine-UA, United States-US, Uruguay-UY, Uzbekistan-UZ, Venezuela-VE, Vietnam-VN, South Africa-ZA S2000 distribution includes: Austria-AT, Australia-AU, Brazil-BR, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Spain-ES, Finland-FI, France-FR, Great Britain-GB, Hong Kong-HK, Indonesia-ID, India-IN, Italy-IT, Japan-JP, South Korea-KR, Kazakhstan-KZ, Lebanon-LB, Mexico-MX, Norway-NO, Philippines-PH, Poland-PL, Portugal-PT, Romania-RO, Russian Federation-RU, Sweden-SE, Slovak Republic-SK, Turkmenistan-TM, Turkey-TR, Taiwan-TW, United States-US X300 distribution includes: AE (United Arab Emirates) , AL (Albania) , AN (Netherlands Antilles), AO (Angola), AR (Argentina), AT (Austria), AU (Australia), AZ (Azerbaijan), BA (Bosnia-Herzegovina), BD (Bangladesh), BE (Belgium), BG (Bulgaria), BN (Brunei), BO (Bolivia), BR (Brazil), BY (Belarus), CA (Canada), CH (Switzerland), CL (Chile), CN (China), CO (Colombia), CV (Cape Verde), CZ (Czech Republic), DE (Germany), DK (Denmark), DZ (Algeria), EC (Ecuador), EG (Egypt), ES (Spain), ET (Ethiopia), FJ (Fiji), FR (France), GB (United Kingdom), GE (Georgia), GR (Greece), HK (Hong Kong), HN (Honduras), HR (Croatia), HU (Hungary), ID (Indonesia), IE (Ireland), IN (India), IR (Iran), IT (Italy), JO (Jordan), JP (Japan), KE (Kenya), KR (Korea, South), KW (Kuwait), KZ (Kazakhstan), LT (Lithuania), LV (Latvia), LY (Libya), MD (Moldova), ME (Montenegro), MK (Macedonia), MM (Myanmar), MU (Mauritius), MX (Mexico), MY (Malaysia), NG (Nigeria), NI (Nicaragua), NL (Netherlands), NO (Norway), NP (Nepal), NZ (New Zealand), PE (Per¿), PH (Philippines), PK (Pakistan), PL (Poland), PS (Palestine), PT (Portugal), QA (Qatar), RS (Serbia), RO (Romania), RU (Russia), RW (Rwanda), SA (Saudi Arabia), SC (Seychelles), SE (Sweden), SG (Singapore), SI (Slovenia), SK (Slovakia), SV (El Salvador), SY (Syria), TH (Thailand), TM (Turkmenistan), TN (Tunisia), TR (Turkey), TT (Trinidad/Tobago), TW (Taiwan, ROC), UA (Ukraine), US (United States of America), UY (Uruguay), UZ (Uzbekistan), VE (Venezuela), VN (Viet Nam), YE (Yemen), ZA and (South Africa) Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = ITX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: |  | SoftwareChangeControl
147 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=86329 | Z-1495-2011 | PhD System EIA/IFAversion 2.1A software, model number 426-0246, a component of the Bio-Rad PhD System. Manufactured by Bio-Rad Laboratories, France, Distributed by Bio-Rad Laboratories, Hercules, CA.******EIA and IFA processing system that consists of a PhD fluid handling station (s) linked to a network computer via ethernet hub. The computer provides worklist generation, data management, data reduction and microplate reader control functions. | PhD System EIA/IFAversion 2.1A software | 3 | CDRH | 2011-03-02 | 2011-03-02 21:00 | 2011-03-11 | Bio-Rad Laboratories Inc | Incorrect patient results are being generated from out-of-range samples tested withc ompetitive enzyme immunoassays on the PhD system with software version 2.1A. | Class 3 Device Recall PhD System EIA/IFAversion 2.1A software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall PhD System EIA/IFAversion 2.1A software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall PhD System EIA/IFAversion 2.1A software Date Initiated by Firm September 29, 2009 Date Posted March 02, 2011 Recall Status 1 Terminated 3 on March 11, 2011 Recall Number Z-1495-2011 Recall Event ID 53690 Product Classification Station, pipeting and diluting for clinical use - Product Code JQW Product PhD System EIA/IFAversion 2.1A software, model number 426-0246, a component of the Bio-Rad PhD System. Manufactured by Bio-Rad Laboratories, France, Distributed by Bio-Rad Laboratories, Hercules, CA.************EIA and IFA processing system that consists of a PhD fluid handling station (s) linked to a network computer via ethernet hub. The computer provides worklist generation, data management, data reduction and microplate reader control functions. Code Information Model 426-0246, software version 2.1A Recalling Firm/ Manufacturer Bio-Rad Laboratories Inc 4000 Alfred Nobel Dr Hercules CA 94547-1803 For Additional Information Contact Jolene Bartilson 510-724-7000 Manufacturer Reason for Recall Incorrect patient results are being generated from out-of-range samples tested withc ompetitive enzyme immunoassays on the PhD system with software version 2.1A. FDA Determined Cause 2 Software change control Action BIO-RAD sent an URGENT: MEDICAL DEVICE CORRECTION letter dated September 29, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their customer records to identify all PhD System customers that are using version 2.1A of the EIA/IFA Software and Methods CD-ROM.******Customers who are not running competitive immunoassays on the PhD System may continue using the 2.1A software.******Customers who run competitive immunoassays must revert to software version 2.0B.******Customers were instructed to translate if necessary and add local contact information to the customer letter and distribute to all PhD System customers that are using 2.1A of the EIA/IFA Software and Methods CD-ROM. ******A Customer Medical Device Correction Response Form must be collected from each customer to ensure that they have received this important communication. Once forms have been collected from all applicable customers, they should be faxed to Bio-Rad CSD Regulatory Affairs Department along with the completed Subsidiary Medical Device Correction Response Form.******For questions regarding this recall call 510-741-3954. Quantity in Commerce 262 units Distribution Worldwide Distribution - USA including AL, CA, CO, CT, DC, FL, IL, MA, MD, MN, MO, MS, MT, NC, NY, OH, PA, SD, TN, TX, UT, and VT and the countries of Australia, New Zealand, Singapore, China, France, Sweden, and Canada Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIO-RAD sent an URGENT: MEDICAL DEVICE CORRECTION letter dated September 29, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their customer records to identify all PhD System customers that are using version 2.1A of the EIA/IFA Software and Methods CD-ROM.******Customers who are not running competitive immunoassays on the PhD System may continue using the 2.1A software.******Customers who run competitive immunoassays must revert to software version 2.0B.******Customers were instructed to translate if necessary and add local contact information to the customer letter and distribute to all PhD System customers that are using 2.1A of the EIA/IFA Software and Methods CD-ROM. ******A Customer Medical Device Correction Response Form must be collected from each customer to ensure that they have received this important communication. Once forms have been collected from all applicable customers, they should be faxed to Bio-Rad CSD Regulatory Affairs Department along with the completed Subsidiary Medical Device Correction Response Form.******For questions regarding this recall call 510-741-3954. | SoftwareChangeControl
148 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94177 | Z-0133-2011 | Computed Tomography X-Ray System MX 8000 IDT 16; Model # 728120; Serial #'s: 2513, 2518, 3147, 3217, 3014, 3050, 3068, 3188, 3091, 3051, 3099, 3126, 3183, 3303, 2539 & 3056.***Produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Computed Tomography X-Ray System MX 8000 IDT 16 | 2 | CDRH | 2010-10-22 | 2010-10-22 21:00 | 2012-09-21 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare has decided to recall the Computed Tomography X-Ray System MX 8000 IDT due to the fact that the software upgrade for the MX 8000 IDT 16 v4.2.1 may not provide the optional collimation of 4x4.5 (18mm). If the customer does not verify the increment and thickness coincide with the newly selected collimation suggested by the conversion report, there is potential for a 6mm gap to ap | Class 2 Device Recall Computed Tomography XRay System MX 8000 IDT 16 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Computed Tomography XRay System MX 8000 IDT 16 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Computed Tomography XRay System MX 8000 IDT 16 Date Initiated by Firm July 27, 2010 Date Posted October 22, 2010 Recall Status 1 Terminated 3 on September 21, 2012 Recall Number Z-0133-2011 Recall Event ID 56508 510(K)Number K012009 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Computed Tomography X-Ray System MX 8000 IDT 16; Model # 728120; Serial #'s: 2513, 2518, 3147, 3217, 3014, 3050, 3068, 3188, 3091, 3051, 3099, 3126, 3183, 3303, 2539 & 3056.******Produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. Code Information Model # 728120; Serial #'s: 2513, 2518, 3147, 3217, 3014, 3050, 3068, 3188, 3091, 3051, 3099, 3126, 3183, 3303, 2539 & 3056. Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Road Cleveland OH 44143-2131 For Additional Information Contact 440-483-7000 Manufacturer Reason for Recall Philips Healthcare has decided to recall the Computed Tomography X-Ray System MX 8000 IDT due to the fact that the software upgrade for the MX 8000 IDT 16 v4.2.1 may not provide the optional collimation of 4x4.5 (18mm). If the customer does not verify the increment and thickness coincide with the newly selected collimation suggested by the conversion report, there is potential for a 6mm gap to ap FDA Determined Cause 2 Software change control Action Philips Medical Systems (Cleveland), Inc. sent an URGENT - Field Safety Notice dated July 29, 2010 and August 23, 2010 to all customers. The notices identified the product, the problem, and the steps to be taken.******Customers were instructed to locate and correct the error. ******Field Service Engineers were to visit each site to update the software******For further information customers were to contact their local Philips Customer Care Center at 1-800-722-9377 Quantity in Commerce 16 units Distribution Product was sold in the following states: CA, DE, IN,OR, MA, NV, TN. Product was also sold in Canada and Australia Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Medical Systems (Cleveland), Inc. sent an URGENT - Field Safety Notice dated July 29, 2010 and August 23, 2010 to all customers. The notices identified the product, the problem, and the steps to be taken.******Customers were instructed to locate and correct the error. ******Field Service Engineers were to visit each site to update the software******For further information customers were to contact their local Philips Customer Care Center at 1-800-722-9377 | SoftwareChangeControl
149 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92698 | Z-2209-2010 | Medtronic Carelink Monitor Model 2490C. The Medtronic CareLink Monitor Model 2490C is designed to automatically gather information from the implanted heart device. This automatic wireless communication between the heart device and the monitor takes place at times scheduled by the doctor or clinic. The gathered heart device information is sent over a telephone line to the clinic using a toll-free number. The monitor can also send wireless heart device alerts to the clinic automatically. | Medtronic Carelink Monitor Model 2490C | 2 | CDRH | 2010-08-11 | 2010-08-11 21:00 | 2011-12-24 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic has identified that a subset of 2490C CareLink Monitors recently received an incorrect software update. Patients reported that their monitor was making a "ticking sound" and they were unable to transmit device data via CareLink. These monitors are now non-functional. ***No patient injuries have been reported as a result of this issue. | Class 2 Device Recall Medtronic Carelink Monitor Model 2490C Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Medtronic Carelink Monitor Model 2490C FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Medtronic Carelink Monitor Model 2490C Date Initiated by Firm June 11, 2010 Date Posted August 11, 2010 Recall Status 1 Terminated 3 on December 24, 2011 Recall Number Z-2209-2010 Recall Event ID 56118 PMA Number P890003 Product Classification Pulse generator, permanent, implantable - Product Code NVZ Product Medtronic Carelink Monitor Model 2490C. The Medtronic CareLink Monitor Model 2490C is designed to automatically gather information from the implanted heart device. This automatic wireless communication between the heart device and the monitor takes place at times scheduled by the doctor or clinic. The gathered heart device information is sent over a telephone line to the clinic using a toll-free number. The monitor can also send wireless heart device alerts to the clinic automatically. Code Information Lot Serial No IJX061043P IJX061063P IJX067150P IJX067158P IJX067162P IJX013664P IJX013831P IJX014701P IJX014899P IJX015305P IJX015406P IJX015768P IJX062594P IJX062607P IJX062654P IJX062655P IJX062656P IJX062661P IJX062669P IJX062670P IJX062672P IJX062673P IJX062674P IJX062683P IJX062698P IJX062701P IJX062706P IJX062716P IJX062727P IJX062734P IJX062736P IJX062738P IJX062739P IJX062741P IJX062747P IJX062763P IJX062764P Recalling Firm/ Manufacturer Medtronic Inc. Cardiac Rhythm Disease Managment 8200 Coral Sea St. N.E. Saint Paul MN 55112 For Additional Information Contact 763-514-4000 Manufacturer Reason for Recall Medtronic has identified that a subset of 2490C CareLink Monitors recently received an incorrect software update. Patients reported that their monitor was making a "ticking sound" and they were unable to transmit device data via CareLink. These monitors are now non-functional. ******No patient injuries have been reported as a result of this issue. FDA Determined Cause 2 Software change control Action Starting on 6/11/2010, Medtronic began sending Australian and Canadian consignees "Medical Device Recall" letters dated 6/10/2010. The letters describe the problem and product involved. The letters instructs consignees to order replacement monitors, notify affected patients, and instruct each patient to perform a manual transmission upon receipt of the new monitor. ******Medtronic will provide affected patients with a prepaid return kit to faciliate the return of the monitor to Medtronic.******Australian consignees should contact Scott Dunlop at 61-7 3025-3547. Canadian consignees should contact Medtronic Patient Services at 1-800-268-5346. Quantity in Commerce 37 Foreign, 0 Domestic Distribution International Distribution Only: Australia and Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = NVZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Starting on 6/11/2010, Medtronic began sending Australian and Canadian consignees "Medical Device Recall" letters dated 6/10/2010. The letters describe the problem and product involved. The letters instructs consignees to order replacement monitors, notify affected patients, and instruct each patient to perform a manual transmission upon receipt of the new monitor. ******Medtronic will provide affected patients with a prepaid return kit to faciliate the return of the monitor to Medtronic.******Australian consignees should contact Scott Dunlop at 61-7 3025-3547. Canadian consignees should contact Medtronic Patient Services at 1-800-268-5346. | SoftwareChangeControl
150 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=91258 | Z-1974-2010 | Horizon Cardiology Hemo Component Schiller Argus Pro PB-1000 Monitor unit.***Argus Pro PB 1000 firmware versions 2.41+M NIBP and 2.42.01+M.******A programmable diagnostic computer intended for acquiring, digitizing, storing, displaying and reviewing hemodynamic data for use in hospital cardiac catheterization laboratories. The system is cleared for us with a component physiological data collection unit - the Schiller Argus Pro PB 1000 Monitor unit.*** | Horizon Cardiology Hemo Component Schiller Argus Pro PB-1000 Monitor unit | 2 | CDRH | 2010-07-09 | 2010-07-09 21:00 | 2010-07-21 | McKesson Provider Technologies - Medical Imaging Group | McKesson has recently discovered that certain Horizon Cardiology Hemo systems were running specific Schiller Argus-Pro PB 1000 firmware versions that might cause a delay in the display of physiological parameters on the Horizon Cariology Hemo Real Time Monitor Screen. | Class 2 Device Recall Horizon Cardiology Hemo Component Schiller Argus Pro PB1000 Monitor unit Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Horizon Cardiology Hemo Component Schiller Argus Pro PB1000 Monitor unit FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Horizon Cardiology Hemo Component Schiller Argus Pro PB1000 Monitor unit Date Initiated by Firm April 15, 2010 Date Posted July 09, 2010 Recall Status 1 Terminated 3 on July 21, 2010 Recall Number Z-1974-2010 Recall Event ID 55598 510(K)Number K050561 Product Classification computer, diagnostic, programmable - Product Code DQK Product Horizon Cardiology Hemo Component Schiller Argus Pro PB-1000 Monitor unit.******Argus Pro PB 1000 firmware versions 2.41+M NIBP and 2.42.01+M.************A programmable diagnostic computer intended for acquiring, digitizing, storing, displaying and reviewing hemodynamic data for use in hospital cardiac catheterization laboratories. The system is cleared for us with a component physiological data collection unit - the Schiller Argus Pro PB 1000 Monitor unit. Code Information Argus Pro PB 1000 firmware versions 2.41+M NIBP and 2.42.01+M Serial numbers prefaced with PB-1000: 765.001471, 760.001566, 760.001793, 760.001882, 760.001920, 760.001934, 760.001970, 760.002043, 760.002113, 765.001168, 765.001307, 765.001520, 765.001553, 765.001599, 765.001609, 765.001510, 765.001632, 765.00165, 765.003014, 760.001320, 760,001552, 760.001555, 760,001920, 760.002107, 765.001224, 765.001306, 765.001608. Recalling Firm/ Manufacturer McKesson Provider Technologies - Medical Imaging Group 1639 State Rt 10 Ste 100 Parsippany NJ 07054-4506 For Additional Information Contact Nir Sheep 973-355-9900 Manufacturer Reason for Recall McKesson has recently discovered that certain Horizon Cardiology Hemo systems were running specific Schiller Argus-Pro PB 1000 firmware versions that might cause a delay in the display of physiological parameters on the Horizon Cariology Hemo Real Time Monitor Screen. FDA Determined Cause 2 Software change control Action The firm, McKesson, sent an "Advisory Notice" letter dated April 15, 2010, to all customers. The letter describes the problem, products affected, and actions to be taken by customers. The customers were instructed not to use the Emergency Replacement Hemo kit until they receive a replacement Schiller Argus Pro PB-1000 and they have contacted McKesson Support department at 1-866-777-0202.************If you have any questions regarding this notice, please call our Customer Support department at 1-0866-777-0202 and ask to speak to your support Manager. Quantity in Commerce 26 units; 19 in operational use, 8 as emergency replacement units. Distribution Nationwide distribution: CA, CO, MD, NH, NT, NY, TN, and TX . Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = DQK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, McKesson, sent an "Advisory Notice" letter dated April 15, 2010, to all customers. The letter describes the problem, products affected, and actions to be taken by customers. The customers were instructed not to use the Emergency Replacement Hemo kit until they receive a replacement Schiller Argus Pro PB-1000 and they have contacted McKesson Support department at 1-866-777-0202.************If you have any questions regarding this notice, please call our Customer Support department at 1-0866-777-0202 and ask to speak to your support Manager. | SoftwareChangeControl
151 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=88343 | Z-1322-2010 | Baxter 1550 Single Patient System Hemodialysis Machines; Baxter Healthcare Corporation, Deerfield, IL 60015; product codes 5M5538, 5M5538R, 5M5575, 5M5551 and 5M5551R. | Baxter 1550 Hemodialysis Machine | 2 | CDRH | 2010-04-12 | 2010-04-12 21:00 | 2011-02-03 | Baxter Healthcare Renal Div | Some of the 1550 Hemodialysis Instruments may have replacement circuit boards with incorrect versions of software. | Class 2 Device Recall Baxter 1550 Hemodialysis Machine Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Baxter 1550 Hemodialysis Machine FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Baxter 1550 Hemodialysis Machine Date Initiated by Firm January 29, 2010 Date Posted April 12, 2010 Recall Status 1 Terminated 3 on February 03, 2011 Recall Number Z-1322-2010 Recall Event ID 54373 510(K)Number K883111 Product Classification System, dialysate delivery, single patient - Product Code FKP Product Baxter 1550 Single Patient System Hemodialysis Machines; Baxter Healthcare Corporation, Deerfield, IL 60015; product codes 5M5538, 5M5538R, 5M5575, 5M5551 and 5M5551R. Code Information All Baxter 1550 instruments and/or spare circuit boards that do not have software versions 3.0A or 4.1B Recalling Firm/ Manufacturer Baxter Healthcare Renal Div 1620 Waukegan Rd Bldg R Mc Gaw Park IL 60085-6730 For Additional Information Contact Center for One Baxter 800-422-9837 Manufacturer Reason for Recall Some of the 1550 Hemodialysis Instruments may have replacement circuit boards with incorrect versions of software. FDA Determined Cause 2 Software change control Action An "Urgent Device Correction" letter dated January 29, 2010 was sent to the direct accounts via first class mail on the same date, to the attention of the Hemodialysis Administrator. The letter informed the accounts that subsequent to a prior device field correction in 2003-2004, some 1550 hemodialysis instruments may have replacement circuit boards with incorrect versions of software. The accounts were instructed to confirm the software version in their systems and replacement boards. Instructions for checking the software version were attached to the letter. A complementary software upgrade will be made available to the account if their 1550 instrument and/or spare circuit board inventory has a software version other than 3.0A or 4.1B. ******The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter and listing the serial numbers and software versions of those instruments that do nit have software 3.0A or 4.1B.************Home patients were not notified directly, the information was sent to their Dialysis Center. ************The 1550 operator's manual was revised during the 2003 device field correction. The manual correctly warns the operator to discontinue treatment if the 1550 experiences a FL04, FL07, or FLOg fault code.************To make arrangements for the software upgrade or questions regarding the communication were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 3.************For additional questions, contact the company Monday through Friday from 8:00 AM to 5:00 PM CST at (Toll Free): 800-422-9837. Quantity in Commerce 443 boards Distribution Worldwide Distribution -- United States - nationwide and Puerto Rico, Argentina, Brazil, Canada, Chile, China, Colombia, Ecuador, Guatemala, Honduras, Mexico, Panama, Philippines, Romania, Turkey and Venezuela. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = FKP - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action An "Urgent Device Correction" letter dated January 29, 2010 was sent to the direct accounts via first class mail on the same date, to the attention of the Hemodialysis Administrator. The letter informed the accounts that subsequent to a prior device field correction in 2003-2004, some 1550 hemodialysis instruments may have replacement circuit boards with incorrect versions of software. The accounts were instructed to confirm the software version in their systems and replacement boards. Instructions for checking the software version were attached to the letter. A complementary software upgrade will be made available to the account if their 1550 instrument and/or spare circuit board inventory has a software version other than 3.0A or 4.1B. ******The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter and listing the serial numbers and software versions of those instruments that do nit have software 3.0A or 4.1B.************Home patients were not notified directly, the information was sent to their Dialysis Center. ************The 1550 operator's manual was revised during the 2003 device field correction. The manual correctly warns the operator to discontinue treatment if the 1550 experiences a FL04, FL07, or FLOg fault code.************To make arrangements for the software upgrade or questions regarding the communication were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 3.************For additional questions, contact the company Monday through Friday from 8:00 AM to 5:00 PM CST at (Toll Free): 800-422-9837. | SoftwareChangeControl
152 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=88119 | Z-1143-2010 | FHC microTargeting Guideline 4000 ver 1.4 (GL4K) Software,***Catalog # MT-GL4K.******Intended to assist in functional neurological procedures. | FHC microTargeting Guideline 4000 | 2 | CDRH | 2010-03-16 | 2010-03-16 21:00 | 2010-03-18 | FHC, Inc. | The raster tab for the GL4K version 1.4 software may assign incorrect depths for waveform recordings and if used to identify the implantation target for the DBS lead, the implant may be placed in the wrong location. | Class 2 Device Recall FHC microTargeting Guideline 4000 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall FHC microTargeting Guideline 4000 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall FHC microTargeting Guideline 4000 Date Initiated by Firm December 28, 2009 Date Posted March 16, 2010 Recall Status 1 Terminated 3 on March 18, 2010 Recall Number Z-1143-2010 Recall Event ID 54278 510(K)Number K071364 Product Classification Depth Electrode - Product Code GZL Product FHC microTargeting Guideline 4000 ver 1.4 (GL4K) Software,******Catalog # MT-GL4K.************Intended to assist in functional neurological procedures. Code Information Software ver 1.4, Serial Numbers: 19454, 14423, 17349, 19215, 15573, 19424, 17583, 17589, 17588, and 17581. This version of software was released in August 2009 Recalling Firm/ Manufacturer FHC, Inc. 1201 Main Street Bowdoinham ME 04008 For Additional Information Contact same as recall contact 207-666-8190 Manufacturer Reason for Recall The raster tab for the GL4K version 1.4 software may assign incorrect depths for waveform recordings and if used to identify the implantation target for the DBS lead, the implant may be placed in the wrong location. FDA Determined Cause 2 Software change control Action FHC notified customers by telephone and by an Urgent: Medical Device Correction email, dated December 2009, beginning December 28th. Customers have been instructed to keep the instruction with the GL4K until FHC can provide a revision to the software; and, to use only "distance from zero" or "distance traveled" modes (not "distance from target" mode) during procedures using the raster tab. Customers are to notify their FHC representative to confirm receipt of the notification. A software patch has been verified and validated and will be installed on all affected units. The new version of the corrected software is Guideline 4000 version 1.41. Questions should be directed to Jason Moyer, product manager or Keri Seitz, recall coordinator at 207-666-8190. Quantity in Commerce 10 Distribution Worldwide Distribution -- USA, including states of GA, MA, NC, NJ, OH, RI, TN, TX, and WV and Bucharest, Romania. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = GZL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action FHC notified customers by telephone and by an Urgent: Medical Device Correction email, dated December 2009, beginning December 28th. Customers have been instructed to keep the instruction with the GL4K until FHC can provide a revision to the software; and, to use only "distance from zero" or "distance traveled" modes (not "distance from target" mode) during procedures using the raster tab. Customers are to notify their FHC representative to confirm receipt of the notification. A software patch has been verified and validated and will be installed on all affected units. The new version of the corrected software is Guideline 4000 version 1.41. Questions should be directed to Jason Moyer, product manager or Keri Seitz, recall coordinator at 207-666-8190. | SoftwareChangeControl
153 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=81671 | Z-2325-2009 | ARCHITECT i System Assay CD ROM - US, list 6E58-22, Version 22.01, and ARCHITECT i System Assay CD-ROM - WW (excluding US), list 6E59-23, Version 23.0, containing the ARCHITECT LH assay file; Abbott Laboratories, Diagnostic Division, Abbott Park, IL 60064 USA; Product List Number: 6E58 (US) and 6E59 (WW) Versions 23 and lower.******The Architect LH assay is a Chemiluminescent Microparticle Immunoassay for the quantitative determination of human luteinizing hormone (LH) in human serum or plasma.*** | ARCHITECT i System Assay CD ROM | 3 | CDRH | 2009-09-30 | 2009-09-30 21:00 | 2009-11-30 | Abbott Laboratories | There continues to be an elevated level of complaint activity for calibration error code 1227, when using ARCHITECT LH reagent lots (6C25-22 or 6C25-27) due to the use of assay disk versions that do not contain the updated ARCHITECT LH assay file contained on ARCHITECT i System Assay CD ROM versions 24 and higher (US LN 6E58, WW LN 6E59). | Class 3 Device Recall ARCHITECT i System Assay CD ROM Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall ARCHITECT i System Assay CD ROM FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall ARCHITECT i System Assay CD ROM Date Initiated by Firm April 24, 2009 Date Posted September 30, 2009 Recall Status 1 Terminated 3 on November 30, 2009 Recall Number Z-2325-2009 Recall Event ID 51928 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product ARCHITECT i System Assay CD ROM - US, list 6E58-22, Version 22.01, and ARCHITECT i System Assay CD-ROM - WW (excluding US), list 6E59-23, Version 23.0, containing the ARCHITECT LH assay file; Abbott Laboratories, Diagnostic Division, Abbott Park, IL 60064 USA; Product List Number: 6E58 (US) and 6E59 (WW) Versions 23 and lower.************The Architect LH assay is a Chemiluminescent Microparticle Immunoassay for the quantitative determination of human luteinizing hormone (LH) in human serum or plasma. Code Information Product List Number: 6E58 (US) Version 22.01and 6E59 (WW) Version 23.0 and lower, all serial numbers. Recalling Firm/ Manufacturer Abbott Laboratories 100 Abbott Park Rd Abbott Park IL 60064-3502 For Additional Information Contact Abbott Customer Support 877-422-2688 Manufacturer Reason for Recall There continues to be an elevated level of complaint activity for calibration error code 1227, when using ARCHITECT LH reagent lots (6C25-22 or 6C25-27) due to the use of assay disk versions that do not contain the updated ARCHITECT LH assay file contained on ARCHITECT i System Assay CD ROM versions 24 and higher (US LN 6E58, WW LN 6E59). FDA Determined Cause 2 Software change control Action Abbott issued a "Product Correction-Immediate Action Required" letter dated April 24, 2009. The letter provided users with the steps to take to determine what version of the ARCHITECT LH Assay is on their instrument and instructions for installing the assay file provided with the letter if needed. Users were requested to return a customer reply letter acknowledging receipt of the letter. ************For additional information, contact 1-877-4ABBOTT. Quantity in Commerce 1,037CD-ROMs - US, 4,676 CD-ROMs - WW Distribution Worldwide Distribution -- United States including Puerto Rico and Argentina, Australia, Brazil, Canada, Chile, Colombia, Costa Rica, Curacao, Germany, Guatemala, Hong Kong, Japan, Mexico, New Zealand, Peru, Singapore, South Korea, Taiwan, Thailand, Uruguay and Venezuela. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Abbott issued a "Product Correction-Immediate Action Required" letter dated April 24, 2009. The letter provided users with the steps to take to determine what version of the ARCHITECT LH Assay is on their instrument and instructions for installing the assay file provided with the letter if needed. Users were requested to return a customer reply letter acknowledging receipt of the letter. ************For additional information, contact 1-877-4ABBOTT. | SoftwareChangeControl
154 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=83671 | Z-1808-2009 | American Optisurgical, Horizon Phacoemulsification System, Model: HZN-2030, for ophthalmic surgery. | Horizon | 2 | CDRH | 2009-09-03 | 2009-09-04 21:00 | 2009-09-04 | American Optisurgical Inc | Device operation interrupted: A software bug causes Phaco to stop when vacuum goes above 300 mmHg while in Multi Burst mode. Use of this product at the parameters specified above could lead to a delay in surgery that might present a health risk. | Class 2 Device Recall Horizon Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Horizon FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Horizon Date Initiated by Firm March 24, 2009 Date Posted September 04, 2009 Recall Status 1 Terminated 3 on September 04, 2009 Recall Number Z-1808-2009 Recall Event ID 52697 510(K)Number K020527 Product Classification phacoemulsification system - Product Code HQC Product American Optisurgical, Horizon Phacoemulsification System, Model: HZN-2030, for ophthalmic surgery. Code Information Serial Numbers: 1114, 1116, 1118, 1120-1124, and 1126-1129. Recalling Firm/ Manufacturer American Optisurgical Inc 25501 Arctic Ocean Dr Lake Forest CA 92630 For Additional Information Contact 949-580-1266 Ext. 211 Manufacturer Reason for Recall Device operation interrupted: A software bug causes Phaco to stop when vacuum goes above 300 mmHg while in Multi Burst mode. Use of this product at the parameters specified above could lead to a delay in surgery that might present a health risk. FDA Determined Cause 2 Software change control Action American Optisurgery sent an E-mail to customers on April 2, 2009 stating the reason for recall and informed them to immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, please identify customers and identify them at once of this product recall. Their notification to customers may be enhanced by including a copy of this recall notification letter. They were then told to refer to the enclosed Recall Questionnaire for instructions on what to do with the recalled product: The recall should be carried out to the user level and return the enclosed response form as soon as possible. If there are any questions, they are advised to contact David Salzberg at (800) 576-1266. Quantity in Commerce 12 units Distribution Worldwide distribution: USA (California), Brazil, Algeria, France, South Korea, India, and Peru. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = HQC - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action American Optisurgery sent an E-mail to customers on April 2, 2009 stating the reason for recall and informed them to immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, please identify customers and identify them at once of this product recall. Their notification to customers may be enhanced by including a copy of this recall notification letter. They were then told to refer to the enclosed Recall Questionnaire for instructions on what to do with the recalled product: The recall should be carried out to the user level and return the enclosed response form as soon as possible. If there are any questions, they are advised to contact David Salzberg at (800) 576-1266. | SoftwareChangeControl
155 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=81931 | Z-1478-2009 | Philips Avalon Fetal Monitor FM20; Model Number: M2702A. | Philips Avalon Fetal Monitor | 2 | CDRH | 2009-06-26 | 2009-06-26 21:00 | 2020-09-23 09:36 | Philips Healthcare Inc. | Interruption of power-up/boot sequence may cause the paper scale and recorder speed settings of the monitor to revert to factory default may lead to misinterpretation of the fetal trace by the user and to potential incorrect clinical decisions | Class 2 Device Recall Philips Avalon Fetal Monitor Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Avalon Fetal Monitor FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Avalon Fetal Monitor Date Initiated by Firm May 12, 2009 Date Posted June 26, 2009 Recall Status 1 Terminated 3 on September 23, 2020 Recall Number Z-1478-2009 Recall Event ID 52026 510(K)Number K071800 Product Classification System, monitoring, perinatal - Product Code HGM Product Philips Avalon Fetal Monitor FM20; Model Number: M2702A. Code Information Units with serial numbers in the range of DE53001002 through DE53107383. Recalling Firm/ Manufacturer Philips Healthcare Inc. 3000 Minuteman Road Andover MA 01810 For Additional Information Contact 978-687-1501 Manufacturer Reason for Recall Interruption of power-up/boot sequence may cause the paper scale and recorder speed settings of the monitor to revert to factory default may lead to misinterpretation of the fetal trace by the user and to potential incorrect clinical decisions FDA Determined Cause 2 Software change control Action Philips Healthcare issued an "Urgent-Medical Device Correction" notification dated May 15, 2009 to Customers in the United States via UPS and notifications outside the US will be managed by Philips representatives in each affected geography. Customers were asked to follow the guidelines in the Action to be Taken by Customer/User section of the Correction Notice until they receive their software upgrade.************Questions can be made to Philips response center at 1-800-722-9377. Quantity in Commerce 5,663 units Distribution Worldwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = HGM - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare issued an "Urgent-Medical Device Correction" notification dated May 15, 2009 to Customers in the United States via UPS and notifications outside the US will be managed by Philips representatives in each affected geography. Customers were asked to follow the guidelines in the Action to be Taken by Customer/User section of the Correction Notice until they receive their software upgrade.************Questions can be made to Philips response center at 1-800-722-9377. | SoftwareChangeControl
156 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=81932 | Z-1479-2009 | Philips Avalon Fetal Monitor FM30; Model Number: M2703A. | Philips Avalon Fetal Monitor | 2 | CDRH | 2009-06-26 | 2009-06-26 21:00 | 2020-09-23 09:36 | Philips Healthcare Inc. | Interruption of power-up/boot sequence may cause the paper scale and recorder speed settings of the monitor to revert to factory default may lead to misinterpretation of the fetal trace by the user and to potential incorrect clinical decisions | Class 2 Device Recall Philips Avalon Fetal Monitor Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Avalon Fetal Monitor FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Avalon Fetal Monitor Date Initiated by Firm May 12, 2009 Date Posted June 26, 2009 Recall Status 1 Terminated 3 on September 23, 2020 Recall Number Z-1479-2009 Recall Event ID 52026 510(K)Number K071800 Product Classification System, monitoring, perinatal - Product Code HGM Product Philips Avalon Fetal Monitor FM30; Model Number: M2703A. Code Information Units with serial numbers in the range of DE53001002 through DE53107383. Recalling Firm/ Manufacturer Philips Healthcare Inc. 3000 Minuteman Road Andover MA 01810 For Additional Information Contact 978-687-1501 Manufacturer Reason for Recall Interruption of power-up/boot sequence may cause the paper scale and recorder speed settings of the monitor to revert to factory default may lead to misinterpretation of the fetal trace by the user and to potential incorrect clinical decisions FDA Determined Cause 2 Software change control Action Philips Healthcare issued an "Urgent-Medical Device Correction" notification dated May 15, 2009 to Customers in the United States via UPS and notifications outside the US will be managed by Philips representatives in each affected geography. Customers were asked to follow the guidelines in the Action to be Taken by Customer/User section of the Correction Notice until they receive their software upgrade.************Questions can be made to Philips response center at 1-800-722-9377. Quantity in Commerce 4,788 units Distribution Worldwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = HGM - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare issued an "Urgent-Medical Device Correction" notification dated May 15, 2009 to Customers in the United States via UPS and notifications outside the US will be managed by Philips representatives in each affected geography. Customers were asked to follow the guidelines in the Action to be Taken by Customer/User section of the Correction Notice until they receive their software upgrade.************Questions can be made to Philips response center at 1-800-722-9377. | SoftwareChangeControl
157 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=74895 | Z-1067-2009 | Roche/Hitachi Modular Analytics System, clinical chemistry analyzer, Modular D Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 03284522001.******Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. | Roche/Hitachi Modular Analytics System | 2 | CDRH | 2009-04-10 | 2009-04-10 21:00 | 2009-09-22 | Roche Diagnostics Corp. | If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results. | Class 2 Device Recall Roche/Hitachi Modular Analytics System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Roche/Hitachi Modular Analytics System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Roche/Hitachi Modular Analytics System Date Initiated by Firm November 12, 2008 Date Posted April 10, 2009 Recall Status 1 Terminated 3 on September 22, 2009 Recall Number Z-1067-2009 Recall Event ID 50171 510(K)Number K953239 Product Classification Chemistry Analyzer for Clinical Use - Product Code JJE Product Roche/Hitachi Modular Analytics System, clinical chemistry analyzer, Modular D Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 03284522001.************Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. Code Information All units with software version 08-03. Recalling Firm/ Manufacturer Roche Diagnostics Corp. 9115 Hague Rd Indianapolis IN 46256-1025 For Additional Information Contact 800-428-2336 Manufacturer Reason for Recall If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results. FDA Determined Cause 2 Software change control Action Consignees were notified via an Urgent Medical Device Correction letter dated 11/12/08 and instructed to only execute "module unmasking" when the unit is in "standby" mode and to consult with their physicians at the facility to determine the clinical implications and whether any clinical intervention is needed for patients treated based upon results obtained with this software. Questions should be directed to Roche Diagnostics Technical Support at 1-800-428-2336. Quantity in Commerce 22 systems in all Distribution Nationwide Distribution -- including states of California, Florida, Georgia, Hawaii, Indiana, Maryland, New York, North Carolina, Pennsylvania, and Virginia. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Consignees were notified via an Urgent Medical Device Correction letter dated 11/12/08 and instructed to only execute "module unmasking" when the unit is in "standby" mode and to consult with their physicians at the facility to determine the clinical implications and whether any clinical intervention is needed for patients treated based upon results obtained with this software. Questions should be directed to Roche Diagnostics Technical Support at 1-800-428-2336. | SoftwareChangeControl
158 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=74894 | Z-1066-2009 | Roche/Hitachi Modular Analytics System, clinical chemistry analyzer, Modular P800 Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 04998669001.******Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. | Roche/Hitachi Modular Analytics System | 2 | CDRH | 2009-04-10 | 2009-04-10 21:00 | 2009-09-22 | Roche Diagnostics Corp. | If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results. | Class 2 Device Recall Roche/Hitachi Modular Analytics System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Roche/Hitachi Modular Analytics System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Roche/Hitachi Modular Analytics System Date Initiated by Firm November 12, 2008 Date Posted April 10, 2009 Recall Status 1 Terminated 3 on September 22, 2009 Recall Number Z-1066-2009 Recall Event ID 50171 510(K)Number K953239 Product Classification Chemistry Analyzer for Clinical Use - Product Code JJE Product Roche/Hitachi Modular Analytics System, clinical chemistry analyzer, Modular P800 Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 04998669001.************Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. Code Information All units with software version 08-03. Recalling Firm/ Manufacturer Roche Diagnostics Corp. 9115 Hague Rd Indianapolis IN 46256-1025 For Additional Information Contact 800-428-2336 Manufacturer Reason for Recall If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results. FDA Determined Cause 2 Software change control Action Consignees were notified via an Urgent Medical Device Correction letter dated 11/12/08 and instructed to only execute "module unmasking" when the unit is in "standby" mode and to consult with their physicians at the facility to determine the clinical implications and whether any clinical intervention is needed for patients treated based upon results obtained with this software. Questions should be directed to Roche Diagnostics Technical Support at 1-800-428-2336. Quantity in Commerce 22 systems in all Distribution Nationwide Distribution -- including states of California, Florida, Georgia, Hawaii, Indiana, Maryland, New York, North Carolina, Pennsylvania, and Virginia. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Consignees were notified via an Urgent Medical Device Correction letter dated 11/12/08 and instructed to only execute "module unmasking" when the unit is in "standby" mode and to consult with their physicians at the facility to determine the clinical implications and whether any clinical intervention is needed for patients treated based upon results obtained with this software. Questions should be directed to Roche Diagnostics Technical Support at 1-800-428-2336. | SoftwareChangeControl
159 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=74896 | Z-1068-2009 | Roche/Hitachi Modular Analytics System, clinical chemistry analyzer, Modular P Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 03284549001.******Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. | Roche/Hitachi Modular Analytics System | 2 | CDRH | 2009-04-10 | 2009-04-10 21:00 | 2009-09-22 | Roche Diagnostics Corp. | If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results. | Class 2 Device Recall Roche/Hitachi Modular Analytics System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Roche/Hitachi Modular Analytics System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Roche/Hitachi Modular Analytics System Date Initiated by Firm November 12, 2008 Date Posted April 10, 2009 Recall Status 1 Terminated 3 on September 22, 2009 Recall Number Z-1068-2009 Recall Event ID 50171 510(K)Number K953239 Product Classification Chemistry Analyzer for Clinical Use - Product Code JJE Product Roche/Hitachi Modular Analytics System, clinical chemistry analyzer, Modular P Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 03284549001.************Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. Code Information All units with software version 08-03. Recalling Firm/ Manufacturer Roche Diagnostics Corp. 9115 Hague Rd Indianapolis IN 46256-1025 For Additional Information Contact 800-428-2336 Manufacturer Reason for Recall If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results. FDA Determined Cause 2 Software change control Action Consignees were notified via an Urgent Medical Device Correction letter dated 11/12/08 and instructed to only execute "module unmasking" when the unit is in "standby" mode and to consult with their physicians at the facility to determine the clinical implications and whether any clinical intervention is needed for patients treated based upon results obtained with this software. Questions should be directed to Roche Diagnostics Technical Support at 1-800-428-2336. Quantity in Commerce 22 systems in all Distribution Nationwide Distribution -- including states of California, Florida, Georgia, Hawaii, Indiana, Maryland, New York, North Carolina, Pennsylvania, and Virginia. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Consignees were notified via an Urgent Medical Device Correction letter dated 11/12/08 and instructed to only execute "module unmasking" when the unit is in "standby" mode and to consult with their physicians at the facility to determine the clinical implications and whether any clinical intervention is needed for patients treated based upon results obtained with this software. Questions should be directed to Roche Diagnostics Technical Support at 1-800-428-2336. | SoftwareChangeControl
160 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=74893 | Z-1065-2009 | Roche/Hitachi Modular Analytics System, clinical chemistry analyzer Modular D2400 Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 04998618001.******Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. | Roche/Hitachi Modular Analytics System | 2 | CDRH | 2009-04-10 | 2009-04-10 21:00 | 2009-09-22 | Roche Diagnostics Corp. | If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results. | Class 2 Device Recall Roche/Hitachi Modular Analytics System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Roche/Hitachi Modular Analytics System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Roche/Hitachi Modular Analytics System Date Initiated by Firm November 12, 2008 Date Posted April 10, 2009 Recall Status 1 Terminated 3 on September 22, 2009 Recall Number Z-1065-2009 Recall Event ID 50171 510(K)Number K953239 Product Classification Chemistry Analyzer for Clinical Use - Product Code JJE Product Roche/Hitachi Modular Analytics System, clinical chemistry analyzer Modular D2400 Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 04998618001.************Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. Code Information All units with software version 08-03. Recalling Firm/ Manufacturer Roche Diagnostics Corp. 9115 Hague Rd Indianapolis IN 46256-1025 For Additional Information Contact 800-428-2336 Manufacturer Reason for Recall If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results. FDA Determined Cause 2 Software change control Action Consignees were notified via an Urgent Medical Device Correction letter dated 11/12/08 and instructed to only execute "module unmasking" when the unit is in "standby" mode and to consult with their physicians at the facility to determine the clinical implications and whether any clinical intervention is needed for patients treated based upon results obtained with this software. Questions should be directed to Roche Diagnostics Technical Support at 1-800-428-2336. Quantity in Commerce 22 systems in all Distribution Nationwide Distribution -- including states of California, Florida, Georgia, Hawaii, Indiana, Maryland, New York, North Carolina, Pennsylvania, and Virginia. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Consignees were notified via an Urgent Medical Device Correction letter dated 11/12/08 and instructed to only execute "module unmasking" when the unit is in "standby" mode and to consult with their physicians at the facility to determine the clinical implications and whether any clinical intervention is needed for patients treated based upon results obtained with this software. Questions should be directed to Roche Diagnostics Technical Support at 1-800-428-2336. | SoftwareChangeControl
161 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=74885 | Z-1040-2009 | QCA and iQCA software, a video microscopy software system for quantitative estrogen receptor immunohistochemistry; Cell Analysis Inc., Evanston, IL 60201 or Highland Park, IL 60035.******Intended to measure and quantitate the percentage and intensity of positively stained nuclei in formalin-fixed, paraffin-embedded tissue specimens immunohistochemically stained for estrogen receptors. Indicated for use as an aid in the management, prognosis and prediction of therapy outcomes of breast cancer when used with reagents validated for those indications. | QCA and iQCA Software | 2 | CDRH | 2009-03-24 | 2009-03-24 21:00 | 2011-03-16 | Leica Microsystems, Inc. | A change in the software to open up the software architecture to include other analytes other than Estrogen Receptor was made without FDA review and approval. | Class 2 Device Recall QCA and iQCA Software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall QCA and iQCA Software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall QCA and iQCA Software Date Initiated by Firm November 06, 2008 Date Posted March 24, 2009 Recall Status 1 Terminated 3 on March 16, 2011 Recall Number Z-1040-2009 Recall Event ID 50163 510(K)Number K031363 Product Classification Digital Analyzer - Product Code NQN Product QCA and iQCA software, a video microscopy software system for quantitative estrogen receptor immunohistochemistry; Cell Analysis Inc., Evanston, IL 60201 or Highland Park, IL 60035.************Intended to measure and quantitate the percentage and intensity of positively stained nuclei in formalin-fixed, paraffin-embedded tissue specimens immunohistochemically stained for estrogen receptors. Indicated for use as an aid in the management, prognosis and prediction of therapy outcomes of breast cancer when used with reagents validated for those indications. Code Information all software versions 3.7 and above Recalling Firm/ Manufacturer Leica Microsystems, Inc. 2345 Waukegan Rd Bannockburn IL 60015-1515 For Additional Information Contact 847-405-6565 Manufacturer Reason for Recall A change in the software to open up the software architecture to include other analytes other than Estrogen Receptor was made without FDA review and approval. FDA Determined Cause 2 Software change control Action Leica Microsystems sent Medical Device Correction letters dated 11/6/08 to the direct accounts, informing them that all versions 3.7 and above of QCA and iQCA software are not in compliance with the original 510(k) cleared software for scoring of immunohistochemical stains in clinical applications. The QCA/iQCA software is only cleared for Estrogen Receptor staining under clinical applications. The accounts were provided with red warning stickers that state: "WARNING The clinical application of QCA software has been cleared only for the measurement of images from microscopic slides of breast cancer specimens stained for the presence of estrogen (ER) nuclear receptor protein only. Performance characteristics of the product, when used outside of this intended use, have not been established." They were instructed to apply the stickers to any existing manuals and or at each point of use, and to instruct all users about the notification. The accounts were also requested to complete and return the enclosed reply form confirming that the changes have been made. Quantity in Commerce 5 units Distribution Nationwide Distribution -- including states of CA, IL, KS, and NY. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = NQN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Leica Microsystems sent Medical Device Correction letters dated 11/6/08 to the direct accounts, informing them that all versions 3.7 and above of QCA and iQCA software are not in compliance with the original 510(k) cleared software for scoring of immunohistochemical stains in clinical applications. The QCA/iQCA software is only cleared for Estrogen Receptor staining under clinical applications. The accounts were provided with red warning stickers that state: "WARNING The clinical application of QCA software has been cleared only for the measurement of images from microscopic slides of breast cancer specimens stained for the presence of estrogen (ER) nuclear receptor protein only. Performance characteristics of the product, when used outside of this intended use, have not been established." They were instructed to apply the stickers to any existing manuals and or at each point of use, and to instruct all users about the notification. The accounts were also requested to complete and return the enclosed reply form confirming that the changes have been made. | SoftwareChangeControl
162 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=73784 | Z-1090-2009 | BD PrepStain System, Model Number: 799-14000-00; Catalog Number: 05CR000021, 05CR00021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 799-14000-00R; Catalog Number: 05CR000021, 05CR000021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 7914000-00R.******PrepStain is a liquid-based thin layer cell preparation device for preparation and staining of Prep smear slides from cervical specimens in a liquid based format (gyn use) and of cytology examination slides from non-cervical specimens such as urine and lung lavage. | BD PrepStain System | 2 | CDRH | 2009-03-19 | 2009-03-19 21:00 | 2013-01-15 | Bd Diagnostic Systems Tripath | Software Error in the non-gyn program of some devices may cause non-gyn samples to be incorrectly transferred from centrifuge tubes to slides potentially result in mismatched patient results. | Class 2 Device Recall BD PrepStain System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall BD PrepStain System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall BD PrepStain System Date Initiated by Firm September 26, 2008 Date Posted March 19, 2009 Recall Status 1 Terminated 3 on January 15, 2013 Recall Number Z-1090-2009 Recall Event ID 49548 PMA Number P970018 Product Classification Cervical Cytology Slide Processor - Product Code MKQ Product BD PrepStain System, Model Number: 799-14000-00; Catalog Number: 05CR000021, 05CR00021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 799-14000-00R; Catalog Number: 05CR000021, 05CR000021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 7914000-00R.************PrepStain is a liquid-based thin layer cell preparation device for preparation and staining of Prep smear slides from cervical specimens in a liquid based format (gyn use) and of cytology examination slides from non-cervical specimens such as urine and lung lavage. Code Information GYN software version 1.2, Non-GYN version 2.80; GYN software version 1.3.0.1, Non-GYN software version 2.80. Recalling Firm/ Manufacturer Bd Diagnostic Systems Tripath 4025 Stirrup Creek Dr Durham NC 27703-9398 For Additional Information Contact BD Diagnostics *** TriPath Technical Suppo 336-290-8271 Manufacturer Reason for Recall Software Error in the non-gyn program of some devices may cause non-gyn samples to be incorrectly transferred from centrifuge tubes to slides potentially result in mismatched patient results. FDA Determined Cause 2 Software change control Action BD Diagnostics began contacting known non-gyn customers by phone on 09/26/2008. Customers were notified by Urgent - Important Product Information letter on 09/29/2008. They were advised to stop using the system until they complete a list of actions provided in the letter. A telephone contact number for Technical Support was provided in the letter. A permanent software based solution will be provided to customers when completed. They anticipate the software will be available in October 2008. Quantity in Commerce 662 units Distribution Worldwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = MKQ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BD Diagnostics began contacting known non-gyn customers by phone on 09/26/2008. Customers were notified by Urgent - Important Product Information letter on 09/29/2008. They were advised to stop using the system until they complete a list of actions provided in the letter. A telephone contact number for Technical Support was provided in the letter. A permanent software based solution will be provided to customers when completed. They anticipate the software will be available in October 2008. | SoftwareChangeControl
163 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=71755 | Z-0297-2009 | AxSYM Digoxin II Reagent Pack, Disk Version 1. (3D53-01 or 3D53-02)******The AxSYM Digoxin II assay is a Microparticle Enzyme Immunoassay (MEIA) for the quantitative measurement of digoxin, a cardiovascular drug, in serum or plasma. The measurements obtained are used in the treatment of digoxin overdose and monitoring levels of digoxin to ensure appropriate therapy. | AxSYM Digoxin II Reagent Pack. | 3 | CDRH | 2008-11-21 | 2008-11-21 21:00 | 2011-09-09 | Abbott Diagnostic International, Ltd. | AxSYM TDM/ Transplant Assay Disks 3D53-03 and higher do not contain the AxSYM Digoxin II Assay file.******The instrument might require installation of the AxSYM Digoxin II assay file.*** | Class 3 Device Recall AxSYM Digoxin II Reagent Pack. Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall AxSYM Digoxin II Reagent Pack. FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall AxSYM Digoxin II Reagent Pack. Date Initiated by Firm May 22, 2008 Date Posted November 21, 2008 Recall Status 1 Terminated 3 on September 09, 2011 Recall Number Z-0297-2009 Recall Event ID 48628 510(K)Number K953718 Product Classification Digoxin Enzyme Immunoassay - Product Code KXT Product AxSYM Digoxin II Reagent Pack, Disk Version 1. (3D53-01 or 3D53-02)************The AxSYM Digoxin II assay is a Microparticle Enzyme Immunoassay (MEIA) for the quantitative measurement of digoxin, a cardiovascular drug, in serum or plasma. The measurements obtained are used in the treatment of digoxin overdose and monitoring levels of digoxin to ensure appropriate therapy. Code Information Lot 61275Q100, exp. date: 09/09/08; Lot 58167Q100, exp. date: 08/30/08; Lot 56039Q101, exp. date: 07/13/08; and Lot 56039Q100, exp. date: 07/13/08. Recalling Firm/ Manufacturer Abbott Diagnostic International, Ltd. Carr # 2 Km 58.0 Cruce Davila Barceloneta PR 00617-3009 For Additional Information Contact Saritza Rios 787-846-3500 Ext. 3352 Manufacturer Reason for Recall AxSYM TDM/ Transplant Assay Disks 3D53-03 and higher do not contain the AxSYM Digoxin II Assay file.************The instrument might require installation of the AxSYM Digoxin II assay file. FDA Determined Cause 2 Software change control Action A Product Information letter, dated May 22, 2008, was issued to customers. It requested that one of the following two actions be taken if the AxSYM Digoxin II assay file is not installed on your AxSYM instrument and you want to run the AxSYM Digoxin II assay. If you have one of the previously manufactured software disks listed, [AxSYM Digoxin II Assay Disk, LN 2C43-01, or higher; or AxSYM TDM/Transplant Assay Disk Version 1.0 (3D53-01) or Version 2.0 (3D53-02)], install the AxSYM Digoxin II assay file. Do not install any other assay file. If you do not have one of the assay disks listed, contact your local Customer Service representative to have this assay file installed or install the file from a previous version if you have it on hand. Quantity in Commerce 9069 kits Distribution Worldwide Distribution --- USA including states of AL, AZ, AR, CA, CT, FL, GA, IL, KS, MD, MA, MI, MN, MS, MO, NV, NJ, NY, ND, OK, OR, PA, RI, TX, VA, WA, Hawaii, and Puerto Rico, and countries of Canada, Mexico, Guatemala, Costa Rica, Venezuela, Colombia, Ecuador, Peru, Chile, Uruguay, Argentina, Puerto Rico, Germany, Japan, Singapore, Hong Kong, Taiwan, Thailand, South Korea, Australia, New Zealand, Curacao, and Panama. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = KXT - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action A Product Information letter, dated May 22, 2008, was issued to customers. It requested that one of the following two actions be taken if the AxSYM Digoxin II assay file is not installed on your AxSYM instrument and you want to run the AxSYM Digoxin II assay. If you have one of the previously manufactured software disks listed, [AxSYM Digoxin II Assay Disk, LN 2C43-01, or higher; or AxSYM TDM/Transplant Assay Disk Version 1.0 (3D53-01) or Version 2.0 (3D53-02)], install the AxSYM Digoxin II assay file. Do not install any other assay file. If you do not have one of the assay disks listed, contact your local Customer Service representative to have this assay file installed or install the file from a previous version if you have it on hand. | SoftwareChangeControl
164 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=72541 | Z-0116-2009 | Artiste MV system, Part number 8139789, medical linear accelerator equipped with COHERENCE therapist RTT4.1. Medical Linear accelerator for radiation therapy. The device utilizes a console system for selecting various treatment sequences and manages commands for all motorized movement. | Artiste MV system | 2 | CDRH | 2008-10-20 | 2008-10-20 21:00 | 2010-12-22 | Siemens Medical Solutions USA, Inc | Three issues: unexpected rotation of gantry, unexpected movement of table between beams during patient setup, and single exposure images will be overwritten by double exposure images. | Class 2 Device Recall Artiste MV system Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Artiste MV system FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Artiste MV system Date Initiated by Firm July 04, 2008 Date Posted October 20, 2008 Recall Status 1 Terminated 3 on December 22, 2010 Recall Number Z-0116-2009 Recall Event ID 49024 510(K)Number K072485 Product Classification Linear accelerator - Product Code IYE Product Artiste MV system, Part number 8139789, medical linear accelerator equipped with COHERENCE therapist RTT4.1. Medical Linear accelerator for radiation therapy. The device utilizes a console system for selecting various treatment sequences and manages commands for all motorized movement. Code Information All products with COHERENCE Therapist RTT4.1 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 4040 Nelson Ave Concord CA 94520-1200 For Additional Information Contact 925-246-8407 Manufacturer Reason for Recall Three issues: unexpected rotation of gantry, unexpected movement of table between beams during patient setup, and single exposure images will be overwritten by double exposure images. FDA Determined Cause 2 Software change control Action A letter marked "Urgent" dated July 4, 2008 was sent to customers. The letter provided a detailed description of the issues, the conditions under which these issues may occur, as well as recommendations to avoid these issues. Customer notification TH022/08/S was released on July 11, 2008. A software fix is currently being validated. All consignees notified by certified mail. Contact Siemens Medical Solutions USA, Inc. at 1-925-246-8407 for assistance. Quantity in Commerce 16 Distribution Nationwide in LA, WI, DE, PA and OH, and internationally in Netherlands, Germany, Poland, Norway, Australia, Hungary, Belgium and France Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action A letter marked "Urgent" dated July 4, 2008 was sent to customers. The letter provided a detailed description of the issues, the conditions under which these issues may occur, as well as recommendations to avoid these issues. Customer notification TH022/08/S was released on July 11, 2008. A software fix is currently being validated. All consignees notified by certified mail. Contact Siemens Medical Solutions USA, Inc. at 1-925-246-8407 for assistance. | SoftwareChangeControl
165 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=73614 | Z-0088-2009 | Envision E700 Low Airloss Therapy Surface, Made By Hill-Rom Manufacturing, Inc., Charleston, SC.******The Envision E700 Low Airloss Therapy Surface helps prevent and treat stage III and stage IV pressure ulcers in patients who weigh between 70 lb and 400 lb and are between 4'11" and 6' 4" in height. | Envision E700 Low Airloss Therapy Surface | 2 | CDRH | 2008-10-07 | 2008-10-07 21:00 | 2009-05-21 | Hill-Rom Manufacturing, Inc. | A defect in the software of the device may not allow the patient bed exam alarm or the patient movement alarm to function correctly. | Class 2 Device Recall Envision E700 Low Airloss Therapy Surface Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Envision E700 Low Airloss Therapy Surface FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Envision E700 Low Airloss Therapy Surface Date Initiated by Firm June 19, 2008 Date Posted October 07, 2008 Recall Status 1 Terminated 3 on May 21, 2009 Recall Number Z-0088-2009 Recall Event ID 49490 Product Classification Powered Flotation Therapy Bed - Product Code IOQ Product Envision E700 Low Airloss Therapy Surface, Made By Hill-Rom Manufacturing, Inc., Charleston, SC.************The Envision E700 Low Airloss Therapy Surface helps prevent and treat stage III and stage IV pressure ulcers in patients who weigh between 70 lb and 400 lb and are between 4'11" and 6' 4" in height. Code Information Serial numbers DB100527 thru DB101305. Recalling Firm/ Manufacturer Hill-Rom Manufacturing, Inc. 4349 Corporate Rd Charleston SC 29405-7445 For Additional Information Contact Joe Fogel 843-740-8586 Manufacturer Reason for Recall A defect in the software of the device may not allow the patient bed exam alarm or the patient movement alarm to function correctly. FDA Determined Cause 2 Software change control Action An internal Modification Notice dated 9/1/08 was sent via email to Service Centers (because the device is owned by the manufacturer, Hill-Rom, each Service Center in possession of the device will be the consignee). Each Service Center will perform the correction on the units affected. Each device in the affected serial number range will be tested for the software defect. If the defect is found, the circuit board that contains the defective software will be replaced. The Field Action Plan is expected to be completed on July 24, 2008. Expected completion date of the Correction will be 12 months from July, 2008. Quantity in Commerce 779 units Distribution Worldwide Distribution --- including USA and country of Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action An internal Modification Notice dated 9/1/08 was sent via email to Service Centers (because the device is owned by the manufacturer, Hill-Rom, each Service Center in possession of the device will be the consignee). Each Service Center will perform the correction on the units affected. Each device in the affected serial number range will be tested for the software defect. If the defect is found, the circuit board that contains the defective software will be replaced. The Field Action Plan is expected to be completed on July 24, 2008. Expected completion date of the Correction will be 12 months from July, 2008. | SoftwareChangeControl
166 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=71838 | Z-2203-2008 | 740 Series Multiparameter Monitor Software Version 3.0***Model Numbers: 740-1, 740M-1, 740-2T, 740M-2T, 740-2MS, 740M-2MS, 740-2NL, 740M-2NL, 740-2NN, 740M-2NN, 740-3MS, 740M-3MS, 740-3NL, 740M-3NL, 740-3NN, and 740M-3NN***CAS Medical Systems, Inc., Branford, Connecticut 06405******Indicated for use for non-invasive monitoring of blood pressure, oxygen saturation, pulse and temperature of adults, pediatric and neonatal patients, in the care of health professionals. | 740 Series Multiparameter Monitor | 2 | CDRH | 2008-09-11 | 2008-09-16 05:16 | 2010-08-05 | CAS Medical Systems, Inc. | Software: Audible alarm may be silenced if changed from the factory setting | Class 2 Device Recall 740 Series Multiparameter Monitor Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall 740 Series Multiparameter Monitor FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall 740 Series Multiparameter Monitor Date Initiated by Firm June 24, 2008 Date Posted September 16, 2008 Recall Status 1 Terminated 3 on August 05, 2010 Recall Number Z-2203-2008 Recall Event ID 48666 510(K)Number K033048 Product Classification Patient Physiological Monitor with Arrhythmia Alarms - Product Code MHX Product 740 Series Multiparameter Monitor Software Version 3.0******Model Numbers: 740-1, 740M-1, 740-2T, 740M-2T, 740-2MS, 740M-2MS, 740-2NL, 740M-2NL, 740-2NN, 740M-2NN, 740-3MS, 740M-3MS, 740-3NL, 740M-3NL, 740-3NN, and 740M-3NN******CAS Medical Systems, Inc., Branford, Connecticut 06405************Indicated for use for non-invasive monitoring of blood pressure, oxygen saturation, pulse and temperature of adults, pediatric and neonatal patients, in the care of health professionals. Code Information Serial Numbers: 0809834-0823035 Recalling Firm/ Manufacturer CAS Medical Systems, Inc. 44 E Industrial Rd Branford CT 06405-6507 For Additional Information Contact SAME 203-488-6056 Manufacturer Reason for Recall Software: Audible alarm may be silenced if changed from the factory setting FDA Determined Cause 2 Software change control Action CASMED notified customers by a Product Correction Notification letter on 6/24/08 to advise of problem of the audible alarm and a software upgrade to correct the problem. The firm requests that the alarm volume not be adjusted from its factory set value to ensure that alarms sound correctly. Also, the letter states that if the volume needs to be adjusted, that customers needs to be sure that they only select two clicks above, or two clicks below the mid-point, and that there is a corresponding beep when you make your selection. Customers are asked to review the attached list of serial numbers that will need to be updated with new software and to then email the a contact at the firm once updates have been completed. All updates must be completed by August 31, 2008. Quantity in Commerce 820 units Distribution Worldwide Distribution --- including USA and countries of Canada, Jordan, UK, France, and Taipei. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action CASMED notified customers by a Product Correction Notification letter on 6/24/08 to advise of problem of the audible alarm and a software upgrade to correct the problem. The firm requests that the alarm volume not be adjusted from its factory set value to ensure that alarms sound correctly. Also, the letter states that if the volume needs to be adjusted, that customers needs to be sure that they only select two clicks above, or two clicks below the mid-point, and that there is a corresponding beep when you make your selection. Customers are asked to review the attached list of serial numbers that will need to be updated with new software and to then email the a contact at the firm once updates have been completed. All updates must be completed by August 31, 2008. | SoftwareChangeControl
167 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=71878 | Z-2225-2008 | GE 1.5T and 3.0T Signa HDX MR System***Model Number(s): 2226300, 2377062-61, 2395001, 2395001-2, 5107849, and 5127452. This system is indicated for use as a diagnostic imaging device to produce axial sagittula, coronal and oblique images, spectroscopic images, and or spectra, dynamic images of the internal structure or organs of the entire body, | Signa HDX MR System | 2 | CDRH | 2008-09-09 | 2008-09-11 07:31 | 2011-12-17 | GE Healthcare | Possible misdiagnosis or inappropriate treatment based on the image. Signal homogeneity may be affected when BRAVO applications are set with IR Prep off. The concern is that signal inhomogeneity can result in an erroneous diagnosis. | Class 2 Device Recall Signa HDX MR System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Signa HDX MR System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Signa HDX MR System Date Initiated by Firm March 28, 2008 Date Posted September 11, 2008 Recall Status 1 Terminated 3 on December 17, 2011 Recall Number Z-2225-2008 Recall Event ID 48689 510(K)Number K052293 Product Classification Magnetic Resonance Scanner - Product Code LNH Product GE 1.5T and 3.0T Signa HDX MR System******Model Number(s): 2226300, 2377062-61, 2395001, 2395001-2, 5107849, and 5127452. This system is indicated for use as a diagnostic imaging device to produce axial sagittula, coronal and oblique images, spectroscopic images, and or spectra, dynamic images of the internal structure or organs of the entire body, Code Information 000000000R1597 000000000R5210 000000000R5871 00000007557YR2 00000007496YR3 00000010285YR5 00000008560YR5 00000009025YR8 0000000B523YR3 00000006697YR7 00000007562YR2 00000005539YR2 00000006662YR1 00000007568YR9 00000007108YR4 00000007497YR1 00000007542YR4 00000008116YR6 00000008322YR0 00000008384YR0 00000009538YR0 00000010300YR2 00000007114YR2 00000006604YR3 00000008325YR3 00000009872YR3 00000006557YR3 00000009534YR9 00000007563YR0 00000009854YR1 00000006722YR3 00000007541YR6 00000008376YR6 00000008317YR0 00000008099YR4 00000009006YR8 00000007126YR6 00000008351YR9 00000010283YR0 00000009009YR2 00000005981YR6 00000008100YR0 00000006710YR8 00000008098YR6 00000007890YR7 00000006594YR6 00000008539YR9 00000007871YR7 00000007564YR8 00000005992YR3 00000007556YR4 00000007498YR9 00000007878YR2 00000007565YR5 00000007877YR4 00000007100YR1 00000008993YR* 00000007097YR9 00000009436YR7 00000009419YR3 00000006698YR5 00000007530YR9 00000009858YR2 00000009082YR9 00000007551YR5 00000007870YR9 00000007508YR5 00000009539YR8 00000008404YR6 00000006713YR2 00000008339YR4 00000008396YR4 00000009773YR3 00000009576YR0 00000008352YR7 00000006707YR4 00000006678YR7 00000007582YR0 00000007882YR4 00000007887YR3 00000010511YR4 00000006558YR1 00000009566YR1 00000008197YR6 00000007087YR0 00000007859YR2 00000009060YR5 00000008387YR3 00000009461YR5 00000008342YR8 00000005564YRO 00000007133YR2 00000005537YR6 00000006737YR1 00000006757YR9 00000010160YR0 00000010396YR0 00000005564YR0 00000005532YR7 00000009052YR2 00000007532YR5 00000006690YR2 00000008355YR0 00000006691YR0 00000007099YR5 00000006574YR8 00000009545YR5 00000005970YR9 00000007858YR4 00000007856YR8 00000006682YR9 00000008338YR6 00000005927YR5 00000007549YR9 00000007122YR5 00000006696YR9 00000006715YR7 00000009979YR6 00000009553YR9 00000010525YR4 00000005222YR5 00000010132YR9 00000010432YR3 00000010284YR8 00000010048YR7 00000009618YR0 00000009421YR9 00000009848YR3 00000009578YR6 00000006738YR9 00000009084YR5 00000007866YR7 00000007865YR9 00000007514YR3 00000010045YR3 00000009587YR7 00000009986YR1 00000009766YR7 00000006721YR5 00000009531YR5 00000010028YR9 00000006600YR1 00000010395YR2 00000006727YR2 00000010134YR5 00000007544YR0 00000010289YR7 00000005548YR3 00000007547YR3 00000007555YR6 00000007853YR5 00000006572YR2 00000006670YR4 00000006673YR8 00000008332YR9 00000009867YR3 00000005974YR1 00000006669YR6 00000009982YR0 00000009774YR1 00000010163YR4 00000005215TR9 00000009049YR8 00000007117YR5 00000008378YR2 00000008373YR3 00000007594YR5 00000006716YR5 00000009617YR2 00000008391YR5 00000008059YR8 00000005216YR7 00000006566YR4 00000008337YR8 00000008392YR3 00000009007YR6 00000005536YR8 00000009775YR8 00000007572YR1 00000009599YR2 00000007862YR6 00000007507YR7 00000006570YR6 00000006708YR2 00000006714YR0 00000009845YR9 00000007591YR1 00000007561YR4 00000008335YR2 00000009438YR3 00000007579YR6 00000009547YR1 00000008327YR9 00000009028YR2 00000008061YR4 00000006667YR0 00000006556YR5 00000010059YR4 00000006660YR5 00000006568YR0 00000008359YR2 00000010153YR5 00000009571YR1 00000006743YR9 00000009990YR3 00000006733YR0 00000007884YR0 00000009455YR7 00000009070YR4 00000008380YR8 00000008385YR7 00000007116YR7 00000010144YR4 00000009019YR1 00000010063YR6 00000009062YR1 00000007109YR2 00000009593YR5 00000009564YR6 00000007867YR5 00000007539YR0 00000007520YR0 00000006705YR8 00000009031YR6 00000007090YR4 00000009081YR1 00000007864YR2 00000010302YR8 00000007863YR4 00000005990YR7 00000008372YR5 00000010393YR7 00000009973YR9 00000009627YR1 00000008530YR8 00000007554YR9 00000009847YR5 00000008318YR8 00000009057YR1 00000007869YR1 00000009026YR6 00000008367YR5 00000007892YR3 00000007540YR8 00000006665YR4 00000006562YR3 00000009053YR0 00000007115YR9 00000009865YR7 00000009642YR0 00000009846YR7 00000009459YR9 00000009573YR7 00000006577YR1 00000008364YR2 00000006661YR3 00000005525YR1 00000008394YR9 00000009080YR3 00000007516YR8 00000010140YR2 00000008319YR6 00000008528YR2 00000005559YR0 00000009565YR3 00000009442YR5 00000006726YR4 00000007577YR0 00000006573YR0 00000005220YR9 00000008354YR3 00000008347YR7 00000007588YR7 00000006555YR7 00000008377YR4 00000009581YR0 00000006583YR9 00000009018YR3 00000005543YR4 00000008103YR4 00000008537YR3 00000008534YR0 00000008314YR7 00000009030YR8 00000008561YR3 00000007575YR4 00000009047YR2 00000009083YR7 00000007845YR1 00000007534YR1 00000009454YR0 00000007107YR6 00000009572YR9 00000008123YR2 00000008542YR3 00000008323YR8 00000007505YR1 00000007518YR4 00000009037YR3 00000010168YR3 00000007531YR7 00000009837YR6 00000006595YR3 00000007101YR9 00000007105YR0 00000008991YR2 00000008311YR3 00000008549YR8 00000008532YR4 00000007134YR0 00000009450YR8 00000008358YR4 00000007570YR5 00000010060YR2 00000008366YR7 00000007566YR3 00000010035YR4 00000005982YR4 00000007098YR7 00000007528YR3 00000009001YR9 00000009763YR4 00000008362YR6 00000006694YR4 00000008097YR8 00000005555YR8 00000008402YR0 00000010627YR8 00000010157YR6 00000009984YR6 00000009864YR0 00000009619YR8 00000006725YR6 00000007088YR8 00000007112YR6 00000007111YR8 00000007123YR3 00000008122YR4 00000006672YR0 00000008350YR1 00000008527YR4 00000009032YR4 00000008992YR0 00000008371YR7 00000009027YR4 00000009579YR4 00000010037YR0 00000009561YR2 00000007504YR4 00000010031YR3 00000010487YR7 00000008557YR1 00000006706YR6 00000007092YR0 00000007574YR7 00000008195YR0 00000007503YR6 00000008986YR2 00000006754YR6 00000008990YR4 00000277444MR6 00000279066MR5 00000278393MR4 00000277641MR7 00000280066MR2 00000274886MR1 00000280365MR8 00000276657MR4 00000279414MR7 00000273383MRO 00000281158MR6 00000280778MR2 00000282865MR5 00000274740MR0 00000277535MR1 00000279464MR2 00000279447MR7 00000276184MR9 00000281828MR4 00000278605MR1 00000278800MR8 00000278554MR1 00000280921MR8 00000283517MR1 00000280253MR6 00000280251MR0 00000276680MR6 00000278041MR9 00000280252MR8 00000276457MR9 00000285493MR3 00000280065MR4 00000276656MR6 00000277917MR1 00000277534MR4 00000279400MR6 00000280364MR1 00000273632MR0 00000279910MR4 00000274742MR6 00000000000001 00000279067MR3 00000276245MR8 00000284034MR6 00000278391MR8 00000279166MR3 00000280775MR8 00000000000000 00000281825MR0 00000283346MR5 00000280249MR4 00000278606MR9 00000275537MR9 00000277779MR5 00000279909MR6 00000286339MR7 00000279882MR5 00000278164MR9 00000285490MR9 00000027810MR1 00000281863MR1 00000278799MR2 00000283458MR8 00000284493MR4 00000274887MR9 00000278103MR7 00000280857MR4 00000282383MR9 00000279908MR8 00000282508MR1 00000275526MR2 00000285589MR8 00000279461MR8 00000276163MR3 00000280777MR4 00000281467MR1 00000278223MR3 00000280328MR6 00000273631MR2 00000279563MR1 00000284872MR9 00000280325MR2 00000279616MR7 00000284562MR6 00000285496MR6 00000283516MR3 00000279398MR2 00000280068MR8 00000276243MR3 00000285167MR3 00000272922MR6 00000285437MR0 00000285166MR5 00000276244MR1 00000278039MR3 00000278801MR6 00000278556MR6 00000281375MR6 00000278315MR7 00000280366MR6 00000283773MR0 00000278221MR7 00000284650MR9 00000278674MR7 00000282003MR3 00000278676MR2 00000285674MR8 00000284869MR5 00000276265MR6 00000280673MR5 00000275647MR6 00000278851MR1 00000284234MR2 00000280000MR1 00000279068MR1 00000276266MR4 00000275967MR8 00000275648MR4 00000282549MR5 00000280781MR6 00000283745MR8 00000274744MR2 00000275046MR1 00000275968MR6 00000278397MR5 00000281341MR8 00000276434MR8 00000284261MR5 00000280780MR8 00000285053MR5 00000275465MR3 00000268369MR6 00000260651MR5 000000HDMR2329 000000HDMR2927 00000HDMR22498 00000HDMR10281 00000HDMR10273 00000HDMR20270 00000HDMR10462 000000HDMR2471 000000HDMR3334 000000HDMR2613 000000HDMR3288 000000HDMR2438 00000HDMR20181 000000HDMR2441 00000HDMR20228 000000HDMR2872 00000HDMR20034 00000276073MR4 00000001568422 000000HDMR2901 000000HDMR2799 00000HDMR20056 000000HDMR3407 00000DDDM10670 00000HDMR20044 00000HDMR20171 000000HDMR2444 00000HDMR10788 00000HDMR10322 000000HDMR2936 00000HDMR20294 000000HDMR2531 00000HDMR20219 00000HDMR20235 000000HDMR2829 000000HDMR2893 00000HDMR20018 000000HDMR2528 00000HDMR20061 00000HDMR10743 00000HDMR11028 000000HDMR2686 00000HDMR20206 000000HDMR2508 00000HDMR20142 00000HDMR20050 00000HDMR20151 000000HDMR2515 00000HDMR10404 000000HDMR2641 000000HDMR2888 00000HDMR20158 000000HDMR2998 00000275160MR0 000000HDMR3089 00000HDMR20276 00000HDMR20065 00000HDMR20053 000000HDMR2451 000000HDMR2463 000000HDMR2553 000000HDMR2689 000000HDMR2335 00000001516299 000000HDMR2482 00000275157MR6 00000HDMR20070 00000273807MR8 000000HDMR3373 000000HDMR1111 000000HDMR3298 00000HDMR10291 000000HDMR1005 00000HDMR10846 000000HDMR3048 000000HDMR3262 000000HDMR2591 000000HDMR2651 000000HDMR2675 00000HDMR10831 00000HDMR10334 00000HDMR10548 00000275162MR6 0000HDMR127452 000000HDMR2522 000000HDMR2344 000000HDMR2803 000000HDMR2626 00000HDMR20236 00000HDMR10739 000000HDMR1130 000000HDMR2993 00000273127MR1 00000HDMR20128 000000HDMR3088 00000HDMR20043 00000HDMR10368 00000274350MR8 00000275283MR0 00000HDMR10489 00000HDMR20084 00000HDMR10986 00000HDMR20059 000000HDMR1098 00000HDMR20068 000000HDMR3221 000000HDMR2404 00000HDMR10440 000000HDMR2560 00000HDMR10412 000000HDMR2822 00000275173MR3 000000HDMR3355 000000HDMR2448 00000275279MR8 000000HDMR2612 000000HDMR2660 00000HDMR10938 00000274359MR9 00000HDMR10764 000000HDMR2397 000000HDMR2523 00000HDMR20073 00000HDMR20289 000000HDMR2428 00000HDMR20248 000000HDMR2623 000000HDMR2756 00000276057MR7 000000HDMR2517 000000HDMR3109 000000HDMR2804 00000HDMR10351 000000HDMR3272 00000HDMR10719 00000HDMR20274 000000HDMR3299 00000HDMR20172 00000HDMR20293 000000HDMR3263 00000HDMR20047 000000HDMR2592 00000HDMR10471 00000HDMR20036 000000HDMR3050 000000HDMR2981 000000HDMR2758 000000HDMR2789 000000HDMR2524 000000HDMR2890 000000HDMR2773 000000HDMR3337 000000HDMR2873 000000HDMR2729 00000005127452 000000HDMR2904 00000HDMR10356 000030485C4228 000000HDMR2948 00000HDMR20291 000000HDMR3316 00000HDMR20269 00000HDMR10711 00000HDMR20189 000000HDMR2418 000000HDMR2755 000000HDMR3128 000000HDMR2831 00000HDMR20272 000000HDMR3199 000000HDMR3130 00000HDMR10477 000000HDMR2958 000000HDMR3459 00000HDMR20178 00000HDMR10878 00000HDMR20089 00000HDMR10787 00000HDMR20099 000000HDMR2769 000000HDMR2917 00000HDMR20290 00000HDMR19996 000000HDMR2395 000000HDMR3003 000000HDMR2483 000000HDMR2696 00000274392MR0 00000HDMR10343 000000HDMR2747 00000HDMR20246 00000HDMR10417 00000273839MR1 00000275236MR8 000000HDMR3301 000000HDMR3542 00000276031MR2 000000HDMR2313 00000273119MR8 000000HDMR2882 000000HDMR2902 00000275237MR6 00000274385MR4 00000HDMR20012 00000273832MR6 00000276038MR7 00000274383MR9 00000275168MR3 00000276042MR9 00000276102MR1 000000HDMR2572 000000HDMR2600 000000HDMR3028 000000HDMR3077 000000HDMR2835 000000HDMR2544 00000275151MR9 00000273132MR1 000000HDMR3452 00000275113MR9 00000275120MR4 00000275107MR1 00000HDMR10623 000000HDMR2521 000000HDMR2808 000000HDMR2779 000000HDMR2695 Recalling Firm/ Manufacturer GE Healthcare 3000 N Grandview Blvd Waukesha WI 53188-1615 For Additional Information Contact 262-544-3894 Manufacturer Reason for Recall Possible misdiagnosis or inappropriate treatment based on the image. Signal homogeneity may be affected when BRAVO applications are set with IR Prep off. The concern is that signal inhomogeneity can result in an erroneous diagnosis. FDA Determined Cause 2 Software change control Action Customer Letter and Operator Manual addendum were sent directly to customers in August 2008. When the software correction is issued, GE Field Engineer will visit each site to implement the correction. (Not conducted yet) Quantity in Commerce 720 units (360 US / 360 OUS) estimate Distribution AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS,MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA,WI, and WV. Countries of ******************Argentina******Australia******Austria******Belgium******Brazil******Canada******Chile******China******Czech Republic******Denmark******Finland******France******French Polynesia******Germany******Greece******Hong Kong******India******Indonesia******Ireland******Israel******Italy******Japan******Korea******Kuwait******Lebanon******Malaysia******Mexico******Netherlands******New Zealand******Norway******Peru******Philippines******Russia******Saudi Arabia******Singapore******South Africa******Spain******Sweden******Switzerland******Taiwan******Taiwan******Thailand******Turkey******United Arab Emirates and******United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LNH - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customer Letter and Operator Manual addendum were sent directly to customers in August 2008. When the software correction is issued, GE Field Engineer will visit each site to implement the correction. (Not conducted yet) | SoftwareChangeControl
168 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=67341 | Z-1243-2008 | IMPAX¿ 4.1, IMPAX ¿ 4.5, IMPAX ¿ 5.0, IMPAX ¿ 5.1 IMPAX ¿ 5.2, IMPAX ¿ 6.0, IMPAX ¿ 6.2; Picture Archiving and Communications systems (PACs), used in the acceptance, transfer, display, storage and digital processing of medical images. | IMPAX | 2 | CDRH | 2008-09-08 | 2008-09-11 07:31 | 2010-12-09 | AGFA Corp. | Java version 1.6x (JAVA 6) is not compatible with AGFA IMPAX Administration/Service Tools. Due to a Software Mismatch data can be lost. | Class 2 Device Recall IMPAX Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall IMPAX FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall IMPAX Date Initiated by Firm May 14, 2007 Date Posted September 11, 2008 Recall Status 1 Terminated 3 on December 09, 2010 Recall Number Z-1243-2008 Recall Event ID 46384 510(K)Number K022292 K040555 K050228 K050751 K053458 Product Classification Picture Archiving and Communications system - Product Code LLZ Product IMPAX 4.1, IMPAX 4.5, IMPAX 5.0, IMPAX 5.1 IMPAX 5.2, IMPAX 6.0, IMPAX 6.2; Picture Archiving and Communications systems (PACs), used in the acceptance, transfer, display, storage and digital processing of medical images. Code Information IMPAX¿ 4.1, Catalog Number: LHZQN007; IMPAX ¿ 4.5, Catalog Number: LZTJM; IMPAX ¿ 5.0, Catalog Number: L3CJE; IMPAX ¿ 5.1 Catalog Number: L4AW6; IMPAX ¿ 5.2, Catalog Number: L4EM2; IMPAX ¿ 6.0, Catalog Number: L73V4; IMPAX ¿ 6.2, Catalog Number: EWMDT. Recalling Firm/ Manufacturer AGFA Corp. 10 S Academy St Greenville SC 29601-2632 For Additional Information Contact Ray B. Myers Ph.D. 864-421-1815 Manufacturer Reason for Recall Java version 1.6x (JAVA 6) is not compatible with AGFA IMPAX Administration/Service Tools. Due to a Software Mismatch data can be lost. FDA Determined Cause 2 Software change control Action Agfa Healthcare contacted all consignees on 5/14/2007 and by registered mail on 06/07/2007. The registered Urgent Field Safety Notice letters requested the accounts take the following steps immediately: 1. Turn off automatic JAVA updates to avoid this update being applied in the background. 2. Do not manually update JAVA on their station-even for other applications. 3. If JAVA 1.6 has already been applied, do not use Service/Admin Tools. They were instructed to contact their IT department or Agfa support for assistance in backing out of this version to apply the required version for the Agfa software. A Mandatory Service Bulletin will inform and assist affected account in removing JAVA 1.6 if it has been applied and turning off the automatic updates via the JAVA Plug-in Control Panel. Quantity in Commerce 946 units (total for both types of software) Distribution U.S. Nationwide and Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Agfa Healthcare contacted all consignees on 5/14/2007 and by registered mail on 06/07/2007. The registered Urgent Field Safety Notice letters requested the accounts take the following steps immediately: 1. Turn off automatic JAVA updates to avoid this update being applied in the background. 2. Do not manually update JAVA on their station-even for other applications. 3. If JAVA 1.6 has already been applied, do not use Service/Admin Tools. They were instructed to contact their IT department or Agfa support for assistance in backing out of this version to apply the required version for the Agfa software. A Mandatory Service Bulletin will inform and assist affected account in removing JAVA 1.6 if it has been applied and turning off the automatic updates via the JAVA Plug-in Control Panel. | SoftwareChangeControl
169 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=67348 | Z-1244-2008 | WEB1000 Versions 3.1, 4.1 and 5.1, Picture Archiving and Communications systems (PACs), used in the acceptance, transfer, display, storage and digital processing of medical images. | WEB1000 | 2 | CDRH | 2008-09-08 | 2008-09-11 07:31 | 2010-12-09 | AGFA Corp. | Java version 1.6x (JAVA 6) is not compatible with AGFA IMPAX Administration/Service Tools. Due to a Software Mismatch data can be lost. | Class 2 Device Recall WEB1000 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall WEB1000 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall WEB1000 Date Initiated by Firm May 14, 2007 Date Posted September 11, 2008 Recall Status 1 Terminated 3 on December 09, 2010 Recall Number Z-1244-2008 Recall Event ID 46384 510(K)Number K022292 K040555 K050228 K050751 K053458 Product Classification Picture Archiving and Communications system - Product Code LLZ Product WEB1000 Versions 3.1, 4.1 and 5.1, Picture Archiving and Communications systems (PACs), used in the acceptance, transfer, display, storage and digital processing of medical images. Code Information WEB1000 version 3.1, Catalog Number: LD5GQ001; WEB1000 version 4.1, Catalog Number: LZS5R, LZS9Z, LZTA4 and WEB1000 version 5.1, Catalog Number: L7DOS, L7DLPU. Recalling Firm/ Manufacturer AGFA Corp. 10 S Academy St Greenville SC 29601-2632 For Additional Information Contact Ray B. Myers Ph.D. 864-421-1815 Manufacturer Reason for Recall Java version 1.6x (JAVA 6) is not compatible with AGFA IMPAX Administration/Service Tools. Due to a Software Mismatch data can be lost. FDA Determined Cause 2 Software change control Action Agfa Healthcare contacted all consignees on 5/14/2007 and by registered mail on 06/07/2007. The registered Urgent Field Safety Notice letters requested the accounts take the following steps immediately: 1. Turn off automatic JAVA updates to avoid this update being applied in the background. 2. Do not manually update JAVA on their station-even for other applications. 3. If JAVA 1.6 has already been applied, do not use Service/Admin Tools. They were instructed to contact their IT department or Agfa support for assistance in backing out of this version to apply the required version for the Agfa software. A Mandatory Service Bulletin will inform and assist affected account in removing JAVA 1.6 if it has been applied and turning off the automatic updates via the JAVA Plug-in Control Panel. Quantity in Commerce 946 units (total) Distribution U.S. Nationwide and Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Agfa Healthcare contacted all consignees on 5/14/2007 and by registered mail on 06/07/2007. The registered Urgent Field Safety Notice letters requested the accounts take the following steps immediately: 1. Turn off automatic JAVA updates to avoid this update being applied in the background. 2. Do not manually update JAVA on their station-even for other applications. 3. If JAVA 1.6 has already been applied, do not use Service/Admin Tools. They were instructed to contact their IT department or Agfa support for assistance in backing out of this version to apply the required version for the Agfa software. A Mandatory Service Bulletin will inform and assist affected account in removing JAVA 1.6 if it has been applied and turning off the automatic updates via the JAVA Plug-in Control Panel. | SoftwareChangeControl
170 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=70841 | Z-1827-2008 | MIM 4.0 and MIM 4.1 is a software package designed for use in diagnostic imaging. It is a stand alone software package which operates on Windows 2000/XP. | MM 4.0 and MIM 4.1 | 2 | CDRH | 2008-09-02 | 2008-09-03 00:30 | 2015-04-14 | MIMvista Corp | In MIM (4.0.2-4.0.5) and MIM (4.1.0-4.1.2) versions, there is an error that may occur during the calculation of statistics using the SUV Tool or Contour Statistics. It will occur only under the following specific circumstance: A fusion of a primary volume which has a Z-Axis slice separation larger than that of the secondary volume and statistics are reported on the secondary volume.*** | Class 2 Device Recall MM 4.0 and MIM 4.1 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall MM 4.0 and MIM 4.1 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall MM 4.0 and MIM 4.1 Date Initiated by Firm April 23, 2008 Date Posted September 03, 2008 Recall Status 1 Terminated 3 on April 14, 2015 Recall Number Z-1827-2008 Recall Event ID 48114 510(K)Number K060816 K071964 Product Classification Radiological Imaging Processing System - Product Code LLZ Product MIM 4.0 and MIM 4.1 is a software package designed for use in diagnostic imaging. It is a stand alone software package which operates on Windows 2000/XP. Code Information Model #'s 4.0.2-4.0.5 & 4.1.0-4.1.2 Recalling Firm/ Manufacturer MIMvista Corp 25200 Chagrin Blvd Ste 200 Cleveland OH 44122-5684 For Additional Information Contact Lynn Hanigan 216-455-0600 Manufacturer Reason for Recall In MIM (4.0.2-4.0.5) and MIM (4.1.0-4.1.2) versions, there is an error that may occur during the calculation of statistics using the SUV Tool or Contour Statistics. It will occur only under the following specific circumstance: A fusion of a primary volume which has a Z-Axis slice separation larger than that of the secondary volume and statistics are reported on the secondary volume. FDA Determined Cause 2 Software change control Action MIMvista Corp has issued this safety recall notification for MIM versions 4.0.2 through 4.1.2. An error has been discovered in the way SW and contour statistics are measured and reported under some specific circumstances. Users of these versions should immediately contact MIMvista support at (866) 421-2536 or visit httv://www.mimvista.com/recall/ to obtain complete instructions regarding this error. A new version that corrects this problem is available on the Downloads page, http://www,mimvista.com/download Quantity in Commerce 221 Distribution The product was sold in the following US states: AL, AR, AZ, CA, CO, DC, DE, FL, GA, IA, ID, IL, IN, LA, MA, MD, MI, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI & WV. It was also sold in Europe, Asia and Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action MIMvista Corp has issued this safety recall notification for MIM versions 4.0.2 through 4.1.2. An error has been discovered in the way SW and contour statistics are measured and reported under some specific circumstances. Users of these versions should immediately contact MIMvista support at (866) 421-2536 or visit httv://www.mimvista.com/recall/ to obtain complete instructions regarding this error. A new version that corrects this problem is available on the Downloads page, http://www,mimvista.com/download | SoftwareChangeControl
171 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=70261 | Z-2139-2008 | CELL-DYN Sapphire Hematology Analyzer, List number 08H00-01, LIS/ Middleware Interface Specification, Abbott Diagnostics Division, Santa Clara, CA******For in vitro diagnostic use in counting and characterizing blood cells. | CELL-DYN Sapphire Hematology Analzyer | 2 | CDRH | 2008-08-20 | 2008-08-21 00:30 | 2010-03-12 | Abbott Laboratories | Table in LIS Specification depicting the association between Record ID and Numerical Result Label is incorrect. If a record ID is used to configure the system for mapping, results will come out nonsensical. | Class 2 Device Recall CELLDYN Sapphire Hematology Analzyer Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall CELLDYN Sapphire Hematology Analzyer FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall CELLDYN Sapphire Hematology Analzyer Date Initiated by Firm April 11, 2008 Date Posted August 21, 2008 Recall Status 1 Terminated 3 on March 12, 2010 Recall Number Z-2139-2008 Recall Event ID 47801 510(K)Number K051215 Product Classification Differential Cell Counter - Product Code GKZ Product CELL-DYN Sapphire Hematology Analyzer, List number 08H00-01, LIS/ Middleware Interface Specification, Abbott Diagnostics Division, Santa Clara, CA************For in vitro diagnostic use in counting and characterizing blood cells. Code Information Revision B. Recalling Firm/ Manufacturer Abbott Laboratories 5440 Patrick Henry Dr Santa Clara CA 95054-1113 For Additional Information Contact 408-982-4800 Manufacturer Reason for Recall Table in LIS Specification depicting the association between Record ID and Numerical Result Label is incorrect. If a record ID is used to configure the system for mapping, results will come out nonsensical. FDA Determined Cause 2 Software change control Action Recall initiated on April 11, 2008. A Product Correction letter and customer reply form were sent to all currently active CELL-DYN Sapphire Customers. The letter requests that users be certain that the transmission of test results receipt by the host interface is mapped according to the Numerical Result Label, when configuring their CELL-DYN Sapphire Hematology Analyzer for results transmission to a Clinical Laboratory Information System or Middleware. Affected product was distributed to a total of 4 distributors. However, notification was made to all 154 customers that received the CELL-DYNE Sapphire product. If the customer or the health care providers that they serve have any questions regarding this information, U.S. customers should call Customer Support at 1-877-4ABBOTT. Quantity in Commerce 15 copies of the specification were sent worldwide Distribution Nationwide Distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = GKZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Recall initiated on April 11, 2008. A Product Correction letter and customer reply form were sent to all currently active CELL-DYN Sapphire Customers. The letter requests that users be certain that the transmission of test results receipt by the host interface is mapped according to the Numerical Result Label, when configuring their CELL-DYN Sapphire Hematology Analyzer for results transmission to a Clinical Laboratory Information System or Middleware. Affected product was distributed to a total of 4 distributors. However, notification was made to all 154 customers that received the CELL-DYNE Sapphire product. If the customer or the health care providers that they serve have any questions regarding this information, U.S. customers should call Customer Support at 1-877-4ABBOTT. | SoftwareChangeControl
172 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=64472 | Z-1155-2008 | GE LOGIQ 3 Expert ultrasound scanner with software versions 4.0.5, 4.1.1 and 4.1.2. GE Healthcare P.O. Box 414, Milwaukee, Wisconsin 53201 | GE LOGIQ 3 Expert | 2 | CDRH | 2008-06-11 | 2008-06-12 00:30 | 2012-05-29 | GE Healthcare Clinical Systems | Inaccurate guidelines: The BE9C is an optional, small radius Bi-planar convex probe intended for use with the LOGIQ 3 Expert system in Urology application. An attempt by the user to activate the biopsy guidelines while this probe is in use will cause incorrect guidelines to be displayed on the image. The incorrect guidelines are recognized by their orientation that is left/right reversed from t | Class 2 Device Recall GE LOGIQ 3 Expert Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GE LOGIQ 3 Expert FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GE LOGIQ 3 Expert Date Initiated by Firm September 04, 2007 Date Posted June 12, 2008 Recall Status 1 Terminated 3 on May 29, 2012 Recall Number Z-1155-2008 Recall Event ID 44884 510(K)Number K020263 Product Classification ultrasound scanner - Product Code IYN Product GE LOGIQ 3 Expert ultrasound scanner with software versions 4.0.5, 4.1.1 and 4.1.2. GE Healthcare P.O. Box 414, Milwaukee, Wisconsin 53201 Code Information Systems with software version 4.0.5, 4.1.1 and 4.1.2. Serial numbers: 26022WS4, 26023WS2, 26346WS7, 2641WS8, 31001WS1, 34381WS4, 36460WS4, 36461WS2, 43734WS3, 43735WS0, 43736WS8, 43737WS6, 43994WS3, 43995WS0, 43998WS4, 44000WS8, 44001WS6, 44002WS4, 44003WS2, 44004WS0, 44005WS7, 44011WS5, 44012WS3, 44013WS1, 44015WS6, 44016WS4, 44018WS0, 44019WS8, 44020WS6, 44100WS6, 44153WS5, 44157WS6, 44159WS2, 44270WS7, 44273WS1, 44274WS9, 44275WS6, 44276WS4, 44277WS2, 44532WS0, 44533WS8, 44817WS5, 44819WS1, 44820WS9, 44821WS7, 44822WS5, 44825WS8, 44826WS6, 44908WS2, 44909WS0, 44912WS4, 44914WS0, 44915WS7, 44989WS2, 44990WS0, 44991WS8, 44992WS6, 44994WS2, 45088WS2, 45089WS0, 5122WS7 Recalling Firm/ Manufacturer GE Healthcare Clinical Systems Manufacturer Reason for Recall Inaccurate guidelines: The BE9C is an optional, small radius Bi-planar convex probe intended for use with the LOGIQ 3 Expert system in Urology application. An attempt by the user to activate the biopsy guidelines while this probe is in use will cause incorrect guidelines to be displayed on the image. The incorrect guidelines are recognized by their orientation that is left/right reversed from t FDA Determined Cause 2 Software change control Action GE sent consignees a "GE Urgent Medical Device Correction" letter dated 8/14/07. The letter informed the customers about the problem, offer a short-term workaround and informed them that a representative will schedule a visit to upgrade the software. Quantity in Commerce 62 Distribution Worldwide Distribution - USA (NY & AL), Korea, China, India, Brazil, Mexico, Germany, Greece, Italy, Poland, Russia, Saudi Arabia, Spain, and Czech Republic. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action GE sent consignees a "GE Urgent Medical Device Correction" letter dated 8/14/07. The letter informed the customers about the problem, offer a short-term workaround and informed them that a representative will schedule a visit to upgrade the software. | SoftwareChangeControl
173 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=68116 | Z-1462-2008 | AxSYM Drugs of Abuse/Toxicology Assay Disk Version 8.0; List Number: 3D54-08; Abbott La boratories, Diagnostics Division, Abbott Park, IL 60064 | AxSYM Drugs of Abuse/Toxicology Assay Disk Version 8.0 | 3 | CDRH | 2008-05-13 | 2008-05-14 00:30 | 2009-12-01 | Abbott Laboratories | Error Message: If the Positive and Negative Interpretation Cutoff parameters (116/117) are edited after installation of the revised assay file, VRTX error #0002 in Task 40 is generated and the instrument locks up when it attempts to report an AxSYM Amphetamine/Methamphetamine II patient result. | Class 3 Device Recall AxSYM Drugs of Abuse/Toxicology Assay Disk Version 8.0 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 3 Device Recall AxSYM Drugs of Abuse/Toxicology Assay Disk Version 8.0 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 3 Device Recall AxSYM Drugs of Abuse/Toxicology Assay Disk Version 8.0 Date Initiated by Firm January 21, 2008 Date Posted May 14, 2008 Recall Status 1 Terminated 3 on December 01, 2009 Recall Number Z-1462-2008 Recall Event ID 46740 510(K)Number K953115 Product Classification Amphetamine test system - Product Code DKZ Product AxSYM Drugs of Abuse/Toxicology Assay Disk Version 8.0; List Number: 3D54-08; Abbott La boratories, Diagnostics Division, Abbott Park, IL 60064 Code Information Lot Numbers: 59311P100 and 59312P100. Recalling Firm/ Manufacturer Abbott Laboratories 100 Abbott Park Rd Abbott Park IL 60064-3502 For Additional Information Contact Abbott Customer Support 877-422-2688 Manufacturer Reason for Recall Error Message: If the Positive and Negative Interpretation Cutoff parameters (116/117) are edited after installation of the revised assay file, VRTX error #0002 in Task 40 is generated and the instrument locks up when it attempts to report an AxSYM Amphetamine/Methamphetamine II patient result. FDA Determined Cause 2 Software change control Action Consignees were sent a Product Correction-Immediate Action Required letter on 1/21/08. The letter provided users with the steps to take to temporarily correct the problem until the recalling firm is able to send a corrected assay file. Users were requested to return a customer reply letter acknowledging receipt of the letter. For additional information, contact 1-877-4ABBOTT. Quantity in Commerce 1,270 disks Distribution Worldwide Distribution including USA territory of Puerto Rico, and countries of Australia, Canada, Chile, Colombia, Germany, Mexico, New Zealand, South Korea, Taiwan, Thailand and Venezuela. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = DKZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Consignees were sent a Product Correction-Immediate Action Required letter on 1/21/08. The letter provided users with the steps to take to temporarily correct the problem until the recalling firm is able to send a corrected assay file. Users were requested to return a customer reply letter acknowledging receipt of the letter. For additional information, contact 1-877-4ABBOTT. | SoftwareChangeControl
174 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=66398 | Z-0526-2008 | BD FACS Sample Prep Assistant II, Part Number: 337170;***Product is manufactured and distributed by BD Biosciences, San Jose, CA 95131*** | BD FACS Sample Prep Assistant II | 2 | CDRH | 2008-04-09 | 2008-04-10 00:30 | 2010-12-13 | BD Biosciences | Mixing Performance: SPA II devices manufactured with the affected product motor control board may exhibit suboptimal mixing performance. | Class 2 Device Recall BD FACS Sample Prep Assistant II Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall BD FACS Sample Prep Assistant II FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall BD FACS Sample Prep Assistant II Date Initiated by Firm October 02, 2007 Date Posted April 10, 2008 Recall Status 1 Terminated 3 on December 13, 2010 Recall Number Z-0526-2008 Recall Event ID 45880 510(K)Number K050191 Product Classification Differential Cell Counter - Product Code GKI Product BD FACS Sample Prep Assistant II, Part Number: 337170;******Product is manufactured and distributed by BD Biosciences, San Jose, CA 95131 Code Information Serial Numbers: T0246 to T0290 (except for T0274,T0275,T0277 and T0279) Recalling Firm/ Manufacturer BD Biosciences 2350 Qume Dr San Jose CA 95131-1807 For Additional Information Contact Kim Ky 408-954-2610 Manufacturer Reason for Recall Mixing Performance: SPA II devices manufactured with the affected product motor control board may exhibit suboptimal mixing performance. FDA Determined Cause 2 Software change control Action On 10/1/07, all consignees were notified via Important Product Information letter and also an on-site visit by the firm's sales representative. The letter informed users of the affected product and provided instructions on the recall. For additional concerns or questions contact BD Customer Support Center at 1-877-232-8995. Distribution Worldwide Distribution --- USA including states of: NJ, IL, TX, PA, NY, MN, and AZ, and countries of Australia, Belgium, Hong Kong, Trinidad and Tobago.************************Australia, Belgium, Hong Kong, Trinidad and Tobago. 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = GKI - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: |  | SoftwareChangeControl
175 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=66172 | Z-0571-2008 | Ventana Image Analysis System (VIAS) - PATHWAY anti-HER-2/neu (4B5) primary antibody with software version 2.3.0.2, Model Number: VIAS-799-10000-00, Pathology Image Analysis System, Ventana Medical Systems, Inc., Tucson, AZ 85755 | Ventana Image Analysis System (VIAS) | 2 | CDRH | 2008-03-17 | 2008-03-18 00:30 | 2012-03-26 | Ventana Medical Systems Inc | Incomplete upgrade: The software upgrade from February 1, 2007 may not have been***completed. This software upgrade adds the algorithm (Slidetype) for the***PATHWAY anti-HER-2/neu (4B5) antibody to the VIAS system. | Class 2 Device Recall Ventana Image Analysis System (VIAS) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Ventana Image Analysis System (VIAS) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Ventana Image Analysis System (VIAS) Date Initiated by Firm June 26, 2007 Date Posted March 18, 2008 Recall Status 1 Terminated 3 on March 26, 2012 Recall Number Z-0571-2008 Recall Event ID 45798 510(K)Number K061613 Product Classification Pathology Image Analysis System - Product Code NQN Product Ventana Image Analysis System (VIAS) - PATHWAY anti-HER-2/neu (4B5) primary antibody with software version 2.3.0.2, Model Number: VIAS-799-10000-00, Pathology Image Analysis System, Ventana Medical Systems, Inc., Tucson, AZ 85755 Code Information Serial Numbers: HIS06086 HIS05042 HIS06047 HIS05018 HIS05022 HIS06046 HIS06092 HIS06059 HIS06053 HIS05053 HIS06035 HIS06038 HIS05026 HIS06100 HIS06021 HIS06074 HIS05041 HIS05017 HIS06031 HIS05051 HIS06061 HIS06028 HIS05011 HIS06083 HIS06064 HIS06057 HIS06055 HIS05047 HIS05012 HIS06054 HIS06073 HIS06088 HIS06076 HIS06002 HIS06001 HIS06004 HIS06034 HIS06015 HIS06098 HIS06069 HIS05044 HIS06033 HIS06050 HIS05046 HIS06030 HIS06095 HIS06067 HIS05028 HIS06027 HIS06056 HIS05033 HIS06049 HIS06026 HIS06008 HIS05049 HIS05013 HIS06079 HIS06072 HIS06089 HIS06084 HIS06044 HIS06007 HIS05027 HIS06051 HIS06052 HIS05025 HIS06096 HIS06066 HIS06039 HIS06065 HIS06019 HIS05052 HIS06029 HIS06022 HIS06068 HIS05048, and HIS05058 Recalling Firm/ Manufacturer Ventana Medical Systems Inc 1910 E Innovation Park Dr Tucson AZ 85755-1962 For Additional Information Contact Jeff Boone 520-877-7237 Manufacturer Reason for Recall Incomplete upgrade: The software upgrade from February 1, 2007 may not have been******completed. This software upgrade adds the algorithm (Slidetype) for the******PATHWAY anti-HER-2/neu (4B5) antibody to the VIAS system. FDA Determined Cause 2 Software change control Action Beginning in 6/07, Ventana Medical Systems, Inc., made phonecalls to labs asking to confirm software version, date of upgrade if available and confirmation of whether site uses 4B5 and Imaging together. A phone script was provided. Customers with systems not ugraded were to upgrade before further use of system. A notification letter dated 8/17/2007 was later sent to VIAS customers advising that the software upgrade from February 1 , 2007 may not have been******completed on their system. This software upgrade adds the algorithm (Slidetype) for the******PATWAY anti-HER-2/neu (4B5) antibody to their VIAS system . This Slidetype is cleared for marketing******and is required prior to imaging slides stained with PATHWAY anti-HER-2/neu (4B5).******The letter also advises that their VIAS and NEXES systems are conneded to the internet which enables this upgrade to be******performed remotely with the support of the Ventana Medical Systems Customer Support Center******(CSC). customers are instructed to call the CSC to carry out the upgrade or to schedule a******Technical Applications Specialist (TAS) to assist with the upgrade in their lab. Arrangements for the upgrade are to be made prior to imaging slides stained with PATHWAY******anti-HER-2/neu (4B5). Quantity in Commerce 77 devices Distribution Nationwide and Canada Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = NQN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Beginning in 6/07, Ventana Medical Systems, Inc., made phonecalls to labs asking to confirm software version, date of upgrade if available and confirmation of whether site uses 4B5 and Imaging together. A phone script was provided. Customers with systems not ugraded were to upgrade before further use of system. A notification letter dated 8/17/2007 was later sent to VIAS customers advising that the software upgrade from February 1 , 2007 may not have been******completed on their system. This software upgrade adds the algorithm (Slidetype) for the******PATWAY anti-HER-2/neu (4B5) antibody to their VIAS system . This Slidetype is cleared for marketing******and is required prior to imaging slides stained with PATHWAY anti-HER-2/neu (4B5).******The letter also advises that their VIAS and NEXES systems are conneded to the internet which enables this upgrade to be******performed remotely with the support of the Ventana Medical Systems Customer Support Center******(CSC). customers are instructed to call the CSC to carry out the upgrade or to schedule a******Technical Applications Specialist (TAS) to assist with the upgrade in their lab. Arrangements for the upgrade are to be made prior to imaging slides stained with PATHWAY******anti-HER-2/neu (4B5). | SoftwareChangeControl
176 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=64414 | Z-0124-2008 | Philips PageWriter Touch Cardiograph,***Model: 860284, Philips Medical Systems, Andover, MA, 01810 | Philips PageWriter Touch | 2 | CDRH | 2008-01-29 | 2008-01-30 00:30 | 2016-08-16 | Philips Medical Systems | Incorrect display: When using the 16 lead Patient Interface Module (PIM), lead tracings of V3R and V4R are reversed in the extended lead modes of Pediatric and Balanced. | Class 2 Device Recall Philips PageWriter Touch Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips PageWriter Touch FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips PageWriter Touch Date Initiated by Firm September 12, 2007 Date Posted January 30, 2008 Recall Status 1 Terminated 3 on August 16, 2016 Recall Number Z-0124-2008 Recall Event ID 44824 510(K)Number K020708 Product Classification Electrocardiograph - Product Code DPS Product Philips PageWriter Touch Cardiograph,******Model: 860284, Philips Medical Systems, Andover, MA, 01810 Code Information Serial Numbers: US10603477***US60711180***US70711234***US80711276***US80711348***US80711386***US80711459 US10603478***US60711181***US70711235***US80711283***US80711349***US80711387***US80711460 US10603491***US60711182***US70711236***US80711284***US80711350***US80711388***US80711461 US10603493***US60711183***US70711241***US80711285***US80711351***US80711389***US80711462 US10603494***US60711184***US80503039***US80711286***US80711352***US80711390***US80711463 US10603495***US60711185***US80503049***US80711287***US80711353***US80711391***US90503082 US10603496***US60711186***US80711244***US80711288***US80711354***US80711392***US90503117 US10603497***US70711197***US80711245***US80711289***US80711355***US80711393***US90503119 US10603498***US70711199***US80711246***US80711290***US80711356***US80711394***US90503120 US10603499***US70711201***US80711247***US80711291***US80711357***US80711395***US90503121 US10603530***US70711205***US80711248***US80711292***US80711358***US80711396***US90503126 US10710463***US70711207***US80711249***US80711293***US80711359***US80711398***US90503127 US10710464***US70711208***US80711250***US80711294***US80711360***US80711399***US90503128 US10710465***US70711209***US80711251***US80711295***US80711361***US80711400***US90503129 US50502760***US70711210***US80711252***US80711296***US80711362***US80711402***US90503143 US50710968***US70711211***US80711253***US80711297***US80711363***US80711403***US90503148 US60502824***US70711212***US80711254***US80711298***US80711364***US80711404***US90503149 US60502825***US70711213***US80711255***US80711299***US80711365***US80711405***US90503150 US60502826***US70711214***US80711256***US80711300***US80711366***US80711406***US90503151 US60711064***US70711215***US80711257***US80711301***US80711367***US80711407***US90503152 US60711065***US70711216***US80711258***US80711302***US80711368***US80711408***US90503153 US60711067***US70711217***US80711259***US80711303***US80711369***US80711409***US90503154 US60711068***US70711218***US80711260***US80711304***US80711370***US80711411***US90503155 US60711070***US70711219***US80711261***US80711305***US80711371***US80711412***US90503156 US60711071***US70711220***US80711262***US80711306***US80711372***US80711413***US90503157 US60711073***US70711221***US80711263***US80711307***US80711373***US80711416***US90503158 US60711074***US70711222***US80711264***US80711308***US80711374***US80711417***US90503159 US60711075***US70711223***US80711265***US80711309***US80711375***US80711418***USD0301315 US60711076***US70711224***US80711266***US80711310***US80711376***US80711419***USD0610313 US60711077***US70711225***US80711267***US80711339***US80711377***US80711420***USN0301134 US60711078***US70711226***US80711268***US80711340***US80711378***US80711421***USN0400155 US60711079***US70711227***US80711269***US80711341***US80711379***US80711422***USN0402121 US60711080***US70711228***US80711270***US80711342***US80711380***US80711431***USN0402126 US60711081***US70711229***US80711271***US80711343***US80711381***US80711439***USN0402127 US60711176***US70711230***US80711272***US80711344***US80711382***US80711440***USN0402128 US60711177***US70711231***US80711273***US80711345***US80711383***US80711455***USN0402129 US60711178***US70711232***US80711274***US80711346***US80711384***US80711457***USN0610295 US60711179***US70711233***US80711275***US80711347***US80711385***US80711458***and USO0610104 Recalling Firm/ Manufacturer Philips Medical Systems 3000 Minuteman Road Andover MA 01810 For Additional Information Contact Steve Smyth 978-687-1501 Manufacturer Reason for Recall Incorrect display: When using the 16 lead Patient Interface Module (PIM), lead tracings of V3R and V4R are reversed in the extended lead modes of Pediatric and Balanced. FDA Determined Cause 2 Software change control Action Philips Medical Systems notified affected customers by Customer Letter and Urgent - Device Correction Notice dated 9/12/07. The correction will be a software upgrade. In the US, the Philips Response Center will contact US customers and arrange their upgrades, either via a software download, or by sending a Field Service Engineer to the customer site to perform the upgrade. Until new software is installed, customers are being advised to use the 12 lead module, and to disable or disconnect the 16 lead connection. Quantity in Commerce 266 units Distribution Worldwide Distribution - USA, Canada, Australia, Hungary, New Zealand, Nigeria, Norway, Qatar, Taiwan, and UAE. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = DPS - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Medical Systems notified affected customers by Customer Letter and Urgent - Device Correction Notice dated 9/12/07. The correction will be a software upgrade. In the US, the Philips Response Center will contact US customers and arrange their upgrades, either via a software download, or by sending a Field Service Engineer to the customer site to perform the upgrade. Until new software is installed, customers are being advised to use the 12 lead module, and to disable or disconnect the 16 lead connection. | SoftwareChangeControl
177 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=66252 | Z-0451-2008 | Philips Medical Systems, M3290A IntelliVue Information Center Software for M3170 Intellivue Patient Link, Catalog Number: 865007 and 865015; Running Software Versions J.00.23, J.00.24, J.00.25, and J.00.26, Manufactured and Distributed by; Philips Medical Systems, Andover, MA 01810*** | Philips Medical Systems IntelliVue Information Center | 2 | CDRH | 2007-12-10 | 2007-12-11 00:30 | 2016-08-15 | Philips Medical Systems | Alarm failure : If changes are made to the Caregroup Editor configuration, Care Group alarm status, including alarm text and prompt tones, may fail to be displayed or annunciated at bedside monitors within the Care Group | Class 2 Device Recall Philips Medical Systems IntelliVue Information Center Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Medical Systems IntelliVue Information Center FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Medical Systems IntelliVue Information Center Date Initiated by Firm November 13, 2007 Date Posted December 11, 2007 Recall Status 1 Terminated 3 on August 15, 2016 Recall Number Z-0451-2008 Recall Event ID 45843 510(K)Number K062271 Product Classification Arrhythmia Detector and Alarm - Product Code MHX Product Philips Medical Systems, M3290A IntelliVue Information Center Software for M3170 Intellivue Patient Link, Catalog Number: 865007 and 865015; Running Software Versions J.00.23, J.00.24, J.00.25, and J.00.26, Manufactured and Distributed by; Philips Medical Systems, Andover, MA 01810 Code Information Serial Numbers: 4711A00132 4725A02724 4711A00354***4725A02791 4716A00562***4725A02897 4717A00702***4725A02927 4717A00703***4725A03255 4717A00705***4732A03719 4717A00706***4732A03725 4717A00724***4732A03764 4717A00725***4732A03765 4717A00799***4732A03766 4717A00857***4732A04028 4718A00957***4732A04031 4718A01077***4732A04134 4718A01078***4732A04178 4718A01120***4732A04400 4718A01121***4732A04623 4718A01454***4732A04671 4721A01670***4732A04675 4721A01711***4732A04746 4721A01731***4732A04868 4721A01773***4732A04869 4721A01797***4732A04870 4721A01828***4732A04871 4721A01919***4732A04872 4721A01993***4732A04873 4721A01994***4732A04874 4721A01997***4732A04875 4721A02164***4732A04876 4725A02354***4732A04877 4732A04890 4732A04896 Recalling Firm/ Manufacturer Philips Medical Systems 3000 Minuteman Road Andover MA 01810 For Additional Information Contact SAME 978-687-1501 Manufacturer Reason for Recall Alarm failure : If changes are made to the Caregroup Editor configuration, Care Group alarm status, including alarm text and prompt tones, may fail to be displayed or annunciated at bedside monitors within the Care Group FDA Determined Cause 2 Software change control Action Philips Medical notified customers by letter dated November 2007 titled : Urgent Medical Device Correction Notice. Customers will be instructed to follow the Procedure to Mitigate Risk section * of the notice while they await the correction of their devices with a mandatory software upgrade to be performed by service representatives. * (no changes are to be made to the Caregroup Editor configuration on any system. If changes have already been made to the configuration prior to receipt of this notification, turn each bedside monitor "off " and then "on" in order to reset the system-Customer Care Contact telephone 1-800-722-9377) Quantity in Commerce 60 units Distribution Worldwide Distribution -USA including states of CA, MA, TX, IN, MI, FL, MN, NC, MS, HI, RI, NH, and PA, and countries of Canada, Australia, Belgium, France, Germany, Ireland, Italy, Japan, Switzerland, United Arab Emirates, and The United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Medical notified customers by letter dated November 2007 titled : Urgent Medical Device Correction Notice. Customers will be instructed to follow the Procedure to Mitigate Risk section * of the notice while they await the correction of their devices with a mandatory software upgrade to be performed by service representatives. * (no changes are to be made to the Caregroup Editor configuration on any system. If changes have already been made to the configuration prior to receipt of this notification, turn each bedside monitor "off " and then "on" in order to reset the system-Customer Care Contact telephone 1-800-722-9377) | SoftwareChangeControl
178 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=65617 | Z-0294-2008 | Medtronic 0-arm Imaging System, Product Catalog Number: B1-700-00027, Medtronic Navigation, Inc., Littleton, MA 01460***(software version 3.0) | Medtronic 0-arm Imaging System | 2 | CDRH | 2007-11-26 | 2007-11-27 00:30 | 2007-11-27 | Medtronic Navigation, Inc | Reversed Instrument Position : a defect in the software version 3.0 causes an error in the position of surgical instruments and relative motion are reversed 180' on the StealthStation image. (If the StealthStation is not used, or if the patient orientation is prone or supine, the problem does not occur) | Class 2 Device Recall Medtronic 0arm Imaging System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Medtronic 0arm Imaging System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Medtronic 0arm Imaging System Date Initiated by Firm October 17, 2007 Date Posted November 27, 2007 Recall Status 1 Terminated 3 on November 27, 2007 Recall Number Z-0294-2008 Recall Event ID 45492 510(K)Number K050996 Product Classification Imaging System - Product Code JZL Product Medtronic 0-arm Imaging System, Product Catalog Number: B1-700-00027, Medtronic Navigation, Inc., Littleton, MA 01460******(software version 3.0) Code Information System Serial Numbers: 103, 104, 109, 117, 119 (Germany) 125, and 144 (Canada) Recalling Firm/ Manufacturer Medtronic Navigation, Inc 826 Coal Creek Circle Louisville CO 80027-9710 For Additional Information Contact 720-890-3200 Ext. 3303 Manufacturer Reason for Recall Reversed Instrument Position : a defect in the software version 3.0 causes an error in the position of surgical instruments and relative motion are reversed 180' on the StealthStation image. (If the StealthStation is not used, or if the patient orientation is prone or supine, the problem does not occur) FDA Determined Cause 2 Software change control Action Medtronic Navigation notified customers by letter on October 17, 2007. Medtronic Field service will perform upgrade (software version 3.0.1) at user site. Quantity in Commerce 7 units Distribution Worldwide: USA including states of CO, IN, NE, and MN and countries of Canada and Germany 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JZL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Medtronic Navigation notified customers by letter on October 17, 2007. Medtronic Field service will perform upgrade (software version 3.0.1) at user site. | SoftwareChangeControl
179 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=64656 | Z-0146-2008 | ARCHITECT c8000 Processing Module, for In Vitro Diagnostics; List Number 106-01; Affected software: v2.10, v2.11, v2.12, v2.20, v2.20DB, v2.30, v2.60, v3.00 and v3.10; Distributed by Abbott Laboratories, Irving, TX. | ARCHITECT cSystem | 2 | CDRH | 2007-11-06 | 2007-11-07 00:30 | 2010-10-04 | Abbott Laboratories, Inc | Incorrect Sample/Test Identification: Due to a software timing defect with certain versions of ARCHITECT cSystem Software, test results may be incorrectly associated to a Sample Identification during specific sample-processing conditions of the Robotic Sample Handler. | Class 2 Device Recall ARCHITECT cSystem Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ARCHITECT cSystem FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ARCHITECT cSystem Date Initiated by Firm September 05, 2007 Date Posted November 07, 2007 Recall Status 1 Terminated 3 on October 04, 2010 Recall Number Z-0146-2008 Recall Event ID 44961 510(K)Number K980367 Product Classification Processing Module, for In Vitro Diagnostics; - Product Code JJE Product ARCHITECT c8000 Processing Module, for In Vitro Diagnostics; List Number 106-01; Affected software: v2.10, v2.11, v2.12, v2.20, v2.20DB, v2.30, v2.60, v3.00 and v3.10; Distributed by Abbott Laboratories, Irving, TX. Code Information All Serial Numbers: ARCHITECT cSystems using ARCHITECT System Software Versions: ---Software Version--List Number--Part Number-- 2.10-- 05F48-11-- 7-99077-01 2.11-- 05F48-12 --7-99077-02 2.12 --None --------7-99077-03 2.20 --05F48-13 ---7-201738-01 2.20DB --05F48-14 ---7-201738-02 2.30-- None --202246-01 2.60-- 05F48-17-- 7-203715-01 3.00-- 05F48-15-- 7-202246-01 3.10-- 05F48-18-- 7-203715-02 Recalling Firm/ Manufacturer Abbott Laboratories, Inc 1921 Hurd Drive PO Box 152020 Irving TX 75038 For Additional Information Contact 972-518-6154 Manufacturer Reason for Recall Incorrect Sample/Test Identification: Due to a software timing defect with certain versions of ARCHITECT cSystem Software, test results may be incorrectly associated to a Sample Identification during specific sample-processing conditions of the Robotic Sample Handler. FDA Determined Cause 2 Software change control Action Notification letters sent to consignees on 09/17/07. Consignees informed of the software defect and the conditions that would cause the software error to occur. Consignees told they would be contacted by an Abbott Field Service Rep who would install ARCHITECT System Software v3.11 and were asked to respond via fax using Customer Reply form acknowledging receipt and understanding of correction letter. Firm recalling to hospital/user level. Quantity in Commerce 2070 units. Distribution Worldwide Distribution: USA (including Puerto Rico), Argentina, Australia, Bahamas, Bermuda, Brazil, Cayman Islands, Canada, Chile, China, Colombia, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Hong Kong, Jamaica, Mexico, New Zealand, Panama, Singapore, Thailand, Trinidad and Tobago, Uruguay, and Venezuela. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Notification letters sent to consignees on 09/17/07. Consignees informed of the software defect and the conditions that would cause the software error to occur. Consignees told they would be contacted by an Abbott Field Service Rep who would install ARCHITECT System Software v3.11 and were asked to respond via fax using Customer Reply form acknowledging receipt and understanding of correction letter. Firm recalling to hospital/user level. | SoftwareChangeControl
180 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=64657 | Z-0147-2008 | ARCHITECT c16000 Processing Module, for In Vitro Diagnostics; List Number 3L77-01; Affected software: v2.10, v2.11, v2.12, v2.20, v2.20DB, v2.30, v2.60, v3.00 and v3.10; Distributed by Abbott Laboratories, Irving, TX. | ARCHITECT cSystem | 2 | CDRH | 2007-11-06 | 2007-11-07 00:30 | 2010-10-04 | Abbott Laboratories, Inc | Incorrect Sample/Test Identification: Due to a software timing defect with certain versions of ARCHITECT cSystem Software, test results may be incorrectly associated to a Sample Identification during specific sample-processing conditions of the Robotic Sample Handler. | Class 2 Device Recall ARCHITECT cSystem Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ARCHITECT cSystem FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ARCHITECT cSystem Date Initiated by Firm September 05, 2007 Date Posted November 07, 2007 Recall Status 1 Terminated 3 on October 04, 2010 Recall Number Z-0147-2008 Recall Event ID 44961 Product Classification Processing Module, for In Vitro Diagnostics - Product Code JJE Product ARCHITECT c16000 Processing Module, for In Vitro Diagnostics; List Number 3L77-01; Affected software: v2.10, v2.11, v2.12, v2.20, v2.20DB, v2.30, v2.60, v3.00 and v3.10; Distributed by Abbott Laboratories, Irving, TX. Code Information All Serial Numbers--ARCHITECT cSystems using ARCHITECT System Software Versions: ---Software Version--List Number--Part Number-- 2.10-- 05F48-11-- 7-99077-01 2.11-- 05F48-12 --7-99077-02 2.12 --None --------7-99077-03 2.20 --05F48-13 ---7-201738-01 2.20DB --05F48-14 ---7-201738-02 2.30-- None --202246-01 2.60-- 05F48-17-- 7-203715-01 3.00-- 05F48-15-- 7-202246-01 3.10-- 05F48-18-- 7-203715-02 Recalling Firm/ Manufacturer Abbott Laboratories, Inc 1921 Hurd Drive PO Box 152020 Irving TX 75038 For Additional Information Contact 972-518-6154 Manufacturer Reason for Recall Incorrect Sample/Test Identification: Due to a software timing defect with certain versions of ARCHITECT cSystem Software, test results may be incorrectly associated to a Sample Identification during specific sample-processing conditions of the Robotic Sample Handler. FDA Determined Cause 2 Software change control Action Notification letters sent to consignees on 09/17/07. Consignees informed of the software defect and the conditions that would cause the software error to occur. Consignees told they would be contacted by an Abbott Field Service Rep who would install ARCHITECT System Software v3.11 and were asked to respond via fax using Customer Reply form acknowledging receipt and understanding of correction letter. Firm recalling to hospital/user level. Quantity in Commerce 89 units. Distribution Worldwide Distribution: USA (including Puerto Rico), Argentina, Australia, Bahamas, Bermuda, Brazil, Cayman Islands, Canada, Chile, China, Colombia, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Hong Kong, Jamaica, Mexico, New Zealand, Panama, Singapore, Thailand, Trinidad and Tobago, Uruguay, and Venezuela. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Notification letters sent to consignees on 09/17/07. Consignees informed of the software defect and the conditions that would cause the software error to occur. Consignees told they would be contacted by an Abbott Field Service Rep who would install ARCHITECT System Software v3.11 and were asked to respond via fax using Customer Reply form acknowledging receipt and understanding of correction letter. Firm recalling to hospital/user level. | SoftwareChangeControl
181 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=49265 | Z-0223-2007 | Roche Modular Drive (rchmdlri), analytical laboratory data interface. | Roche Modular Drive | 2 | CDRH | 2006-11-28 09:59 | 2006-11-29 00:30 | 2015-06-11 | Data Innovations, Inc. | Software of modular driver may incorrectly report patient results as Quality Control Results | Class 2 Device Recall Roche Modular Drive Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Roche Modular Drive FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Roche Modular Drive Date Initiated by Firm November 03, 2006 Date Posted November 29, 2006 Recall Status 1 Terminated 3 on June 11, 2015 Recall Number Z-0223-2007 Recall Event ID 36752 Product Classification analytical laboratory data interface. - Product Code JQP Product Roche Modular Drive (rchmdlri), analytical laboratory data interface. Code Information Software Version: v8.00.0016, v8.00.0017, v7.00.0023 and v7.00.0024 Recalling Firm/ Manufacturer Data Innovations, Inc. 120 Kimball Ave Ste 100 South Burlington VT 05403-6837 For Additional Information Contact Katharine Laddison 802-264-3420 Manufacturer Reason for Recall Software of modular driver may incorrectly report patient results as Quality Control Results FDA Determined Cause 2 Software change control Action Data Innovations notified users by email listserv and by letter dated 10/25/06. Quantity in Commerce 56 Distribution Nationwide and Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Data Innovations notified users by email listserv and by letter dated 10/25/06. | SoftwareChangeControl
182 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=49266 | Z-0224-2007 | Roche Cobas driver (rchcob6i), analytical laboratory data interface. | Roche Cobas driver | 2 | CDRH | 2006-11-28 09:59 | 2006-11-29 00:30 | 2015-06-11 | Data Innovations, Inc. | Software of modular driver may incorrectly report patient results as Quality Control Results | Class 2 Device Recall Roche Cobas driver Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Roche Cobas driver FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Roche Cobas driver Date Initiated by Firm November 03, 2006 Date Posted November 29, 2006 Recall Status 1 Terminated 3 on June 11, 2015 Recall Number Z-0224-2007 Recall Event ID 36752 Product Classification analytical laboratory data interface. - Product Code JQP Product Roche Cobas driver (rchcob6i), analytical laboratory data interface. Code Information Software version: v8.00.0001 and v7.00.0001 Recalling Firm/ Manufacturer Data Innovations, Inc. 120 Kimball Ave Ste 100 South Burlington VT 05403-6837 For Additional Information Contact Katharine Laddison 802-264-3420 Manufacturer Reason for Recall Software of modular driver may incorrectly report patient results as Quality Control Results FDA Determined Cause 2 Software change control Action Data Innovations notified users by email listserv and by letter dated 10/25/06. Quantity in Commerce 56 Distribution Nationwide and Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Data Innovations notified users by email listserv and by letter dated 10/25/06. | SoftwareChangeControl
183 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=43332 | Z-0380-06 | Cytomics FC500 with***CXP Software*** | Cytomics FC500 with CXP Software | 2 | CDRH | 2006-01-10 12:05 | 2006-01-11 00:30 | 2012-01-26 | Beckman Coulter Inc | Under certain conditions an incorrect Tube ID (Barcode) can be displayed and printed on the Runtime Panel Report in CXP Cytometer Software Version 2.0 | Class 2 Device Recall Cytomics FC500 with CXP Software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Cytomics FC500 with CXP Software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Cytomics FC500 with CXP Software Date Initiated by Firm November 23, 2005 Date Posted January 11, 2006 Recall Status 1 Terminated 3 on January 26, 2012 Recall Number Z-0380-06 Recall Event ID 34247 Product Classification Analyzer, Chemistry, Micro, For Clinical Use - Product Code JJF Product Cytomics FC500 with******CXP Software Code Information Part Numbers 629636 629637 Recalling Firm/ Manufacturer Beckman Coulter Inc 200 S Kraemer Blvd Brea CA 92822-6208 For Additional Information Contact Nora Zerounian 714-961-3634 Manufacturer Reason for Recall Under certain conditions an incorrect Tube ID (Barcode) can be displayed and printed on the Runtime Panel Report in CXP Cytometer Software Version 2.0 FDA Determined Cause 2 Software change control Action A Product Corrective Action letter will be mailed the week of Nov 28, 2005 to all FC500 wi CXP Software customers. Each user will be contacted initially and then once more (if needed) in order to determine the rate of response from the Response form provided. Quantity in Commerce 179 Distribution Nationwide and Canada Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action A Product Corrective Action letter will be mailed the week of Nov 28, 2005 to all FC500 wi CXP Software customers. Each user will be contacted initially and then once more (if needed) in order to determine the rate of response from the Response form provided. | SoftwareChangeControl