## Sheet1
Unnamed: 0 | WEB_ADDRESS | RECALL_NUMBER | PRODUCT_DESCRIPTION | TRADE_NAME | RECALL_CLASS | CENTER | CENTER_CLASSIFICATION_DT | POSTED_INTERNET_DT | TERMINATION_DT | FIRM_NAME | MANUFACTURER_RECALL_REASON | PAGE_TEXT | ACTION | Recall Type
0 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=212390 | Z-1250-2025 | Brand Name: TBS Osteo***Product Name: TBS iNsight***Model/Catalog Number: TBS iNsight v3.x / pn: OS-TBSi***Software Version: V3.x: (1) V3.1.2, (2) V3.1.1, (3) V3.1.0, (4)V3.0.0***Product Description: TBS iNsight is a class II Medical Device software that is installed on bone densitometers for analysis of bone microarchitecture and bone health management***Component: No*** | TBS iNsight | 2 | CDRH | 2025-02-28 | 2025-02-28 20:07 |  | Medimaps Group | Potential variability in calculations from fast array scans compared to array scans when operating on Hologic Horizon machines. | Class 2 Device Recall TBS iNsight Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TBS iNsight FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TBS iNsight Date Initiated by Firm February 03, 2025 Date Posted February 28, 2025 Recall Status 1 Open 3 , Classified Recall Number Z-1250-2025 Recall Event ID 96233 510(K)Number K152299 Product Classification Densitometer, bone - Product Code KGI Product Brand Name: TBS Osteo******Product Name: TBS iNsight******Model/Catalog Number: TBS iNsight v3.x / pn: OS-TBSi******Software Version: V3.x: (1) V3.1.2, (2) V3.1.1, (3) V3.1.0, (4)V3.0.0******Product Description: TBS iNsight is a class II Medical Device software that is installed on bone densitometers for analysis of bone microarchitecture and bone health management******Component: No Code Information UDI-DI: (1) B214OSTBSI3120, (2) B214OSTBSI3110, (3) B214OSTBSIV310, (4) B214OSTBSIV300. Recalling Firm/ Manufacturer Medimaps Group Fongit Chemin des Aulx 18 Plan-les-Ouates Switzerland Manufacturer Reason for Recall Potential variability in calculations from fast array scans compared to array scans when operating on Hologic Horizon machines. FDA Determined Cause 2 Software Design Change Action On February 5, 2025 URGENT MEDICAL DEVICE CORRECTION letters were sent to customers. Actions to be taken by the user: Follow patient management recommendations. This notice needs to be passed on all organisation where the potentially affected devices have been transferred: the DXA manufacturer (Hologic), direct customers, distributors and field partners.***Awareness on this notice and resulting action will be maintained through customer communications for an appropriate period of time until the end of the FSCA to ensure effectiveness of the corrective action.***Customers are requested to report problem-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback.****Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA s MedWatch Adverse Event Reporting program either online, by regular mail or by fax.***Action Being Taken by the Manufacturer: Software upgrade***Medimaps will issue a maintenance release of the software by the end of February in the worst case and deploy it progressively in the field. The exact plan will be further communicated to customers. The associated FSCA for strongly recommended software update should be completed in 6 months starting from the release of the update version at the end of February 2025. Quantity in Commerce 2036 units (724 US, 1312 OUS) Distribution Domestic: AL, AZ, CA, CO, DE, FL, IL, LA, MA, MD, MI, MN, MO, MS, NC, NM, NY, OH, PA, SC, TX, UT, WA.******International: ALGERIA, ANDORRA, AUSTRALIA, AUSTRIA, BELGIUM, BOSNIA AND HERZEGOVINA, BRAZIL, CHILE, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONDURAS, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, SOUTH KOREA, LEBANON, MALAYSIA, MEXICO, MONTENEGRO, MOROCCO, NEPAL, NETHERLANDS, NEW ZEALAND, NORWAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, ROMANIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = KGI - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On February 5, 2025 URGENT MEDICAL DEVICE CORRECTION letters were sent to customers. Actions to be taken by the user: Follow patient management recommendations. This notice needs to be passed on all organisation where the potentially affected devices have been transferred: the DXA manufacturer (Hologic), direct customers, distributors and field partners.***Awareness on this notice and resulting action will be maintained through customer communications for an appropriate period of time until the end of the FSCA to ensure effectiveness of the corrective action.***Customers are requested to report problem-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback.****Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA s MedWatch Adverse Event Reporting program either online, by regular mail or by fax.***Action Being Taken by the Manufacturer: Software upgrade***Medimaps will issue a maintenance release of the software by the end of February in the worst case and deploy it progressively in the field. The exact plan will be further communicated to customers. The associated FSCA for strongly recommended software update should be completed in 6 months starting from the release of the update version at the end of February 2025. | SoftwareDesignChange
1 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211967 | Z-1019-2025 | Ivenix Infusion System (IIS), LVP Software LVP-SW-0005.******Indicated for use in a hospital and in outpatient care environments for the controlled administration of fluids through clinically accepted routes of administration. | Ivenix Infusion System (IIS), LVP Software; LVP-SW-0005 | 1 | CDRH | 2025-02-04 10:58 | 2025-02-04 20:07 |  | Fresenius Kabi USA, LLC | Large Volume Pump Software, version 5.9.2 and earlier has potential for the following anomalies: ***1. The pump may become nonfunctional if during an alarm condition the Pause Audio option is repeated 70 time or more. ***2. If a secondary infusion is started at the exact moment a primary infusion completes and VTBI reaches 0, it will switch to primary. The primary infusion will infuse at the previously programed primary rate and continue until the infusion is stopped or the bag is empty. | Class 1 Device Recall Ivenix Infusion System (IIS), LVP Software; LVPSW0005 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall Ivenix Infusion System (IIS), LVP Software; LVPSW0005 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall Ivenix Infusion System (IIS), LVP Software; LVPSW0005 Date Initiated by Firm January 10, 2025 Date Posted February 04, 2025 Recall Status 1 Open 3 , Classified Recall Number Z-1019-2025 Recall Event ID 96085 510(K)Number K183311 Product Classification Pump, infusion - Product Code FRN Product Ivenix Infusion System (IIS), LVP Software LVP-SW-0005.************Indicated for use in a hospital and in outpatient care environments for the controlled administration of fluids through clinically accepted routes of administration. Code Information UDI-DI: 00811505030122. Software 5.9.2 and prior Recalling Firm/ Manufacturer Fresenius Kabi USA, LLC 50 High St Ste 50 North Andover MA 01845-2620 For Additional Information Contact Rebecca McCandless 847-393-5451 Manufacturer Reason for Recall Large Volume Pump Software, version 5.9.2 and earlier has potential for the following anomalies: ******1. The pump may become nonfunctional if during an alarm condition the Pause Audio option is repeated 70 time or more. ******2. If a secondary infusion is started at the exact moment a primary infusion completes and VTBI reaches 0, it will switch to primary. The primary infusion will infuse at the previously programed primary rate and continue until the infusion is stopped or the bag is empty. FDA Determined Cause 2 Software Design Change Action Fresenius Kabi began notifying consignees on 01/10/2024 via emailed letter. Consignees were instructed to check if your device contains the affected software, review and become familiar with the temporary Risk Mitigation actions for each anomaly outlined in the letter and notify customers/personnel if the product was transferred or further distributed. Additionally, to facilitate the installation of the pump software, LVP SW 5.10, your facility must install the new Ivenix Infusion Management System (IMS) software version (5.2) on your IMS server and install the new software version (5.10) for the Ivenix Large Volume Pump (LVP). Please post this letter in areas that affected devices are stored and used such that users of this device are aware of the issues and recommended risk mitigation measures. Lastly, consignees were requested to complete and return the Customer Reply Form. Quantity in Commerce 23 systems Distribution US distribution to CA, CO, GA, ID, MI, MS, NE, NJ, NV, OK, TX, UT, VA, WA, WI. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = FRN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Fresenius Kabi began notifying consignees on 01/10/2024 via emailed letter. Consignees were instructed to check if your device contains the affected software, review and become familiar with the temporary Risk Mitigation actions for each anomaly outlined in the letter and notify customers/personnel if the product was transferred or further distributed. Additionally, to facilitate the installation of the pump software, LVP SW 5.10, your facility must install the new Ivenix Infusion Management System (IMS) software version (5.2) on your IMS server and install the new software version (5.10) for the Ivenix Large Volume Pump (LVP). Please post this letter in areas that affected devices are stored and used such that users of this device are aware of the issues and recommended risk mitigation measures. Lastly, consignees were requested to complete and return the Customer Reply Form. | SoftwareDesignChange
2 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211818 | Z-1006-2025 | Brand Name: VITROS***Product Name: VITROS XT 7600 Integrated System***Model/Catalog Number: 6844461 (New), 6272222 (Certified), 6904003 (China)***Software Version: version 3.8.0 or higher***Product Description: In vitro diagnostic chemistry analyzer.******FDA 3 letter product code: JJE***Component: No, not a component*** | VITROS XT 7600 Integrated System | 2 | CDRH | 2025-01-21 | 2025-01-21 20:06 |  | Ortho-Clinical Diagnostics, INc. | The VITROS XT 7600 Integrated System software versions 3.8.0 may not correctly notify users if the Signal Reagent Pack exceeds the on-board and/or shelf-life expiration date. This could result in the use of expired products, which would increase the risk of no results or erroneous results from the instrument.*** | Class 2 Device Recall VITROS XT 7600 Integrated System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS XT 7600 Integrated System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS XT 7600 Integrated System Date Initiated by Firm December 19, 2024 Date Posted January 21, 2025 Recall Status 1 Open 3 , Classified Recall Number Z-1006-2025 Recall Event ID 96031 510(K)Number K182063 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product Brand Name: VITROS******Product Name: VITROS XT 7600 Integrated System******Model/Catalog Number: 6844461 (New), 6272222 (Certified), 6904003 (China)******Software Version: version 3.8.0 or higher******Product Description: In vitro diagnostic chemistry analyzer.************FDA 3 letter product code: JJE******Component: No, not a component Code Information Lot Code: VITROS XT 7600 Integrated System***Product Code: 6844461 (New)***UDI: 10758750012343******VITROS XT 7600 System (Certified)***Product Code: 6272222***UDI: 10758750012343*********Lot Code: VITROS XT 7600 Integrated System***Product Code: 6904003 (China)***UDI: 10758750035809******Range of installed serial numbers:76000001-76002199***Manufacture date range: July 2018 to present***Distribution date range; July 2018 to present******Total systems currently installed in the US: 994***Total systems currently installed outside the US: 899******Note1: The catalogue number 6844461 is used to denote the VITROS XT 7600 Integrated System. Catalogue number 6272222 is used to denote VITROS XT 7600 System Certified. A VITROS XT 7600 Integrated System maintains its original serial number through the certification process. Therefore, the manufacturing and distribution information for the two catalogue numbers will be combined.******Note2: When manufactured the system s serial numbers are sequential, but not all analyzers are released for distribution and therefore not installed. The range of serial numbers provided includes the identified minimum and maximum in serial numbers associated with all currently installed systems. Therefore, the amount of each system manufactured may not equal the amount of each system installed. Additionally, not all serial numbers listed are associated with an active installed analyzer in use today. ******Note3: The value given for "Product Quantity Distributed" is the total number of active installed systems that are in use today. This number is known to fluctuate as individual systems are installed and/or decommissioned. Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, INc. 513 Technology Blvd Rochester NY 14626-3601 For Additional Information Contact Joe Falvo 585-4533452 Manufacturer Reason for Recall The VITROS XT 7600 Integrated System software versions 3.8.0 may not correctly notify users if the Signal Reagent Pack exceeds the on-board and/or shelf-life expiration date. This could result in the use of expired products, which would increase the risk of no results or erroneous results from the instrument. FDA Determined Cause 2 Software Design Change Action On December 19, 2024 URGENT PRODUCT CORRECTION NOTIFICATION letters were sent to customers. REQUIRED ACTION***" Do not use expired Signal Reagent packs. See the Questions & Answers section for more details.***" Complete the enclosed Confirmation of Receipt form no later than December 31, 2024.***" Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System in your laboratory until the issue has been resolved.***Contact Information***We apologize for the inconvenience this will cause your laboratory. If you have further questions, please contact Global Services Organization at 1-800-421-3311.******To report adverse events, contact your local Global Services Organization. Adverse reactions or quality problems experienced with the use of this product may also be reported to the FDA's MedWatch Adverse Event Reporting program: www.fda.gov/medwatch/report.htm, 1-800-332-1088. Quantity in Commerce 1,893 units (994 US, 899 OUS) Distribution Worldwide distribution - US Nationwide and the countries of Australia, Belgium, Bermuda, Brazil, Canada, Chile, China, Colombia, Denmark, France, Germany, India, Italy, Japan, Mexico, Netherlands, Norway, Portugal, Russia, Singapore, Spain, Sweden, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On December 19, 2024 URGENT PRODUCT CORRECTION NOTIFICATION letters were sent to customers. REQUIRED ACTION***" Do not use expired Signal Reagent packs. See the Questions & Answers section for more details.***" Complete the enclosed Confirmation of Receipt form no later than December 31, 2024.***" Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System in your laboratory until the issue has been resolved.***Contact Information***We apologize for the inconvenience this will cause your laboratory. If you have further questions, please contact Global Services Organization at 1-800-421-3311.******To report adverse events, contact your local Global Services Organization. Adverse reactions or quality problems experienced with the use of this product may also be reported to the FDA's MedWatch Adverse Event Reporting program: www.fda.gov/medwatch/report.htm, 1-800-332-1088. | SoftwareDesignChange
3 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211828 | Z-1007-2025 | Brand Name: VITROS***Product Name: VITROS 3600 Immunodiagnostic System***Model/Catalog Number: 6802783 (New), 6802914 (Certified/Refurbished), 6904000 (China)***Software Version: version 3.8.0 or higher***Product Description: In vitro diagnostic chemistry analyzer.******FDA 3 letter product code: JJE***Component: no, not a component*** | VITROS 3600 Immunodiagnostic System | 2 | CDRH | 2025-01-21 | 2025-01-21 20:06 |  | Ortho-Clinical Diagnostics, INc. | The VITROS 3600 Immunodiagnostic System software versions 3.8.0 may not correctly notify users if the Signal Reagent Pack exceeds the on-board and/or shelf-life expiration date. This could result in the use of expired products, which would increase the risk of no results or erroneous results from the instrument.*** | Class 2 Device Recall VITROS 3600 Immunodiagnostic System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS 3600 Immunodiagnostic System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS 3600 Immunodiagnostic System Date Initiated by Firm December 19, 2024 Date Posted January 21, 2025 Recall Status 1 Open 3 , Classified Recall Number Z-1007-2025 Recall Event ID 96031 510(K)Number K083173 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product Brand Name: VITROS******Product Name: VITROS 3600 Immunodiagnostic System******Model/Catalog Number: 6802783 (New), 6802914 (Certified/Refurbished), 6904000 (China)******Software Version: version 3.8.0 or higher******Product Description: In vitro diagnostic chemistry analyzer.************FDA 3 letter product code: JJE******Component: no, not a component Code Information Lot Code: VITROS 3600 Immunodiagnostic System***Product Code: 6802783 (New)***UDI: 10758750002979******VITROS 3600 Immunodiagnostic System (Certified/Refurbished)***Product Code: 6802914***UDI: 10758750002979******VITROS 3600 Immunodiagnostic System (China)***Product Code: 6904000***UDI: 10758750035779******Range of installed serial numbers:36000004-36001776***Manufacture date range: June 2008 to present***Distribution date range; June 2008 to present******Total systems currently installed in the US: 99***Total systems currently installed outside the US: 882******Note1: The catalogue number 6802783 is used to denote the VITROS 3600 Immunodiagnostic System. Catalogue number 6802914 is used to denote the VITROS 3600 Immunodiagnostic System that have been refurbished, these systems can also be referred to as certified. A VITROS 3600 Immunodiagnostic System Certified/Refurbished maintains its original serial number through the refurbishment/certification process. Therefore, the manufacturing and distribution information for the two catalogue numbers will be combined.******Note2: When manufactured the system s serial numbers are sequential, but not all analyzers are released for distribution and therefore not installed. The range of serial numbers provided includes the identified minimum and maximum in serial numbers associated with all currently installed systems. Therefore, the amount of each system manufactured may not equal the amount of each system installed. Additionally, not all serial numbers listed are associated with an active installed analyzer in use today. ******Note3: The value given for "Product Quantity Distributed" is the total number of active installed systems that are in use today. This number is known to fluctuate as individual systems are installed and/or decommissioned. Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, INc. 513 Technology Blvd Rochester NY 14626-3601 For Additional Information Contact Joe Falvo 585-4533452 Manufacturer Reason for Recall The VITROS 3600 Immunodiagnostic System software versions 3.8.0 may not correctly notify users if the Signal Reagent Pack exceeds the on-board and/or shelf-life expiration date. This could result in the use of expired products, which would increase the risk of no results or erroneous results from the instrument. FDA Determined Cause 2 Software Design Change Action On December 19, 2024 URGENT PRODUCT CORRECTION NOTIFICATION letters were sent to customers. REQUIRED ACTION***" Do not use expired Signal Reagent packs. See the Questions & Answers section for more details.***" Complete the enclosed Confirmation of Receipt form no later than December 31, 2024.***" Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System in your laboratory until the issue has been resolved.***Contact Information***We apologize for the inconvenience this will cause your laboratory. If you have further questions, please contact Global Services Organization at 1-800-421-3311.******To report adverse events, contact your local Global Services Organization. Adverse reactions or quality problems experienced with the use of this product may also be reported to the FDA's MedWatch Adverse Event Reporting program: www.fda.gov/medwatch/report.htm, 1-800-332-1088. Quantity in Commerce 981 units (99 US, 882 OUS) Distribution Worldwide distribution - US Nationwide and the countries of Australia, Belgium, Bermuda, Brazil, Canada, Chile, China, Colombia, Denmark, France, Germany, India, Italy, Japan, Mexico, Netherlands, Norway, Portugal, Russia, Singapore, Spain, Sweden, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On December 19, 2024 URGENT PRODUCT CORRECTION NOTIFICATION letters were sent to customers. REQUIRED ACTION***" Do not use expired Signal Reagent packs. See the Questions & Answers section for more details.***" Complete the enclosed Confirmation of Receipt form no later than December 31, 2024.***" Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System in your laboratory until the issue has been resolved.***Contact Information***We apologize for the inconvenience this will cause your laboratory. If you have further questions, please contact Global Services Organization at 1-800-421-3311.******To report adverse events, contact your local Global Services Organization. Adverse reactions or quality problems experienced with the use of this product may also be reported to the FDA's MedWatch Adverse Event Reporting program: www.fda.gov/medwatch/report.htm, 1-800-332-1088. | SoftwareDesignChange
4 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211816 | Z-1005-2025 | Brand Name: VITROS***Product Name: VITROS 5600 Integrated System***Model/Catalog Number: 6802413 (New), 6802415 (Certified/Refurbished), 6904001 (China)***Software Version: version 3.8.0 or higher***Product Description: In vitro diagnostic chemistry analyzer.******FDA 3 letter product code: JJE***Component: No, not a component*** | VITROS 5600 Integrated System | 2 | CDRH | 2025-01-21 | 2025-01-21 20:06 |  | Ortho-Clinical Diagnostics, INc. | The VITROS 5600 Integrated System software versions 3.8.0 may not correctly notify users if the Signal Reagent Pack exceeds the on-board and/or shelf-life expiration date. This could result in the use of expired products, which would increase the risk of no results or erroneous results from the instrument.*** | Class 2 Device Recall VITROS 5600 Integrated System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS 5600 Integrated System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS 5600 Integrated System Date Initiated by Firm December 19, 2024 Date Posted January 21, 2025 Recall Status 1 Open 3 , Classified Recall Number Z-1005-2025 Recall Event ID 96031 510(K)Number K081543 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product Brand Name: VITROS******Product Name: VITROS 5600 Integrated System******Model/Catalog Number: 6802413 (New), 6802415 (Certified/Refurbished), 6904001 (China)******Software Version: version 3.8.0 or higher******Product Description: In vitro diagnostic chemistry analyzer.************FDA 3 letter product code: JJE******Component: No, not a component Code Information Lot Code: VITROS 5600 Integrated System (New)***Product Code: 6802413***UDI: 10758750002740******VITROS 5600 Integrated System (Certified/Refurbished)***Product Code: 6802915***UDI: 10758750002740******Lot Code: VITROS 5600 Integrated System (China)***Product Code: 6904001***UDI: 10758750035786******Range of installed serial numbers: 56000015-56004962***Manufacture date range: August 2008 to present***Distribution date range; August 2008 to present******Total systems currently installed in the US: 905***Total systems currently installed outside the US:2344******Note1: The catalogue number 6802413 is used to denote the VITROS 5600 Integrated System. Catalogue number 6802415 is used to denote those VITROS 5600 Integrated Systems that have been refurbished, these systems can also be referred to as certified. A VITROS 5600 Integrated System maintains its original serial number through the refurbishment/certification process. Therefore, the manufacturing and distribution information for the two catalogue numbers will be combined.******Note2: When manufactured the system s serial numbers are sequential, but not all analyzers are released for distribution and therefore not installed. The range of serial numbers provided includes the identified minimum and maximum in serial numbers associated with all currently installed systems. Therefore, the amount of each system manufactured may not equal the amount of each system installed. Additionally, not all serial numbers listed are associated with an active installed analyzer in use today. ******Note3: The value given for "Product Quantity Distributed" is the total number of active installed systems that are in use today. This number is known to fluctuate as individual systems are installed and/or decommissioned. Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, INc. 513 Technology Blvd Rochester NY 14626-3601 For Additional Information Contact Joe Falvo 585-4533452 Manufacturer Reason for Recall The VITROS 5600 Integrated System software versions 3.8.0 may not correctly notify users if the Signal Reagent Pack exceeds the on-board and/or shelf-life expiration date. This could result in the use of expired products, which would increase the risk of no results or erroneous results from the instrument. FDA Determined Cause 2 Software Design Change Action On December 19, 2024 URGENT PRODUCT CORRECTION NOTIFICATION letters were sent to customers. REQUIRED ACTION***" Do not use expired Signal Reagent packs. See the Questions & Answers section for more details.***" Complete the enclosed Confirmation of Receipt form no later than December 31, 2024.***" Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System in your laboratory until the issue has been resolved.***Contact Information***We apologize for the inconvenience this will cause your laboratory. If you have further questions, please contact Global Services Organization at 1-800-421-3311.******To report adverse events, contact your local Global Services Organization. Adverse reactions or quality problems experienced with the use of this product may also be reported to the FDA's MedWatch Adverse Event Reporting program: www.fda.gov/medwatch/report.htm, 1-800-332-1088. Quantity in Commerce 3,249 units (905 US, 2,344 OUS) Distribution Worldwide distribution - US Nationwide and the countries of Australia, Belgium, Bermuda, Brazil, Canada, Chile, China, Colombia, Denmark, France, Germany, India, Italy, Japan, Mexico, Netherlands, Norway, Portugal, Russia, Singapore, Spain, Sweden, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On December 19, 2024 URGENT PRODUCT CORRECTION NOTIFICATION letters were sent to customers. REQUIRED ACTION***" Do not use expired Signal Reagent packs. See the Questions & Answers section for more details.***" Complete the enclosed Confirmation of Receipt form no later than December 31, 2024.***" Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System in your laboratory until the issue has been resolved.***Contact Information***We apologize for the inconvenience this will cause your laboratory. If you have further questions, please contact Global Services Organization at 1-800-421-3311.******To report adverse events, contact your local Global Services Organization. Adverse reactions or quality problems experienced with the use of this product may also be reported to the FDA's MedWatch Adverse Event Reporting program: www.fda.gov/medwatch/report.htm, 1-800-332-1088. | SoftwareDesignChange
5 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210892 | Z-0570-2025 | Arial Server Software Versions 11.1.4 through 11.1.6, SKU 54305, in conjunction with the Arial Mobile Application or Arial Care Giver Application | Arial Server Software | 2 | CDRH | 2024-11-27 | 2024-11-27 20:07 |  | Securitas Healthcare LLC | Certain software versions will not notify end users of calls initiated by call points with the Arial Emergency and Nurse Call System. | Class 2 Device Recall Arial Server Software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Arial Server Software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Arial Server Software Date Initiated by Firm November 01, 2024 Date Posted November 27, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-0570-2025 Recall Event ID 95652 Product Classification System, environmental control, powered - Product Code IQA Product Arial Server Software Versions 11.1.4 through 11.1.6, SKU 54305, in conjunction with the Arial Mobile Application or Arial Care Giver Application Code Information All Systems using Arial Server Software versions 11.1.4 through 11.1.6 in conjunction with the Arial Mobile Application or Arial Care Giver Application Recalling Firm/ Manufacturer Securitas Healthcare LLC 4600 Vine St Lincoln NE 68503-2823 For Additional Information Contact Ameya Oke 402-475-4002 Manufacturer Reason for Recall Certain software versions will not notify end users of calls initiated by call points with the Arial Emergency and Nurse Call System. FDA Determined Cause 2 Software Design Change Action An URGENT: MEDICAL DEVICE CORRECTION letter dated 10/31/24 was sent to customers.******Necessary Actions to be Taken by Clients***Instructions for determining which versions of the Arial Server Software you are using are attached hereto as Exhibit C. Clients using Arial Server Software versions 11.1.4.9, 11.1.5.5, or 11.1.6.2 should:***1. Not decline previously cleared calls through the Arial App;***2. Restart the Arial Server Software if an end user has declined a previously completed ***call using the Arial App. After restarting the Arial Server Software, calls from the affected***call point(s) will begin to be received by the Arial Server Software and the Arial Server***Software will notify end users as intended. Instructions for restarting the Arial Server***Software are included on Exhibit B.***3. Contact Securitas Healthcare Technical Support at 1 (800) 380-8883 to upgrade the Arial***Server Software to version 11.1.7 or later.******For general questions concerning this communication please contact Securitas Healthcare***Technical Support at 1 (800) 380-8883.******We ask that you please complete and return a copy of the acknowledgement form attached hereto as Exhibit A. Quantity in Commerce 111 systems Distribution Worldwide distribution - US Nationwide and the countries of Canada, Switzerland. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action An URGENT: MEDICAL DEVICE CORRECTION letter dated 10/31/24 was sent to customers.******Necessary Actions to be Taken by Clients***Instructions for determining which versions of the Arial Server Software you are using are attached hereto as Exhibit C. Clients using Arial Server Software versions 11.1.4.9, 11.1.5.5, or 11.1.6.2 should:***1. Not decline previously cleared calls through the Arial App;***2. Restart the Arial Server Software if an end user has declined a previously completed ***call using the Arial App. After restarting the Arial Server Software, calls from the affected***call point(s) will begin to be received by the Arial Server Software and the Arial Server***Software will notify end users as intended. Instructions for restarting the Arial Server***Software are included on Exhibit B.***3. Contact Securitas Healthcare Technical Support at 1 (800) 380-8883 to upgrade the Arial***Server Software to version 11.1.7 or later.******For general questions concerning this communication please contact Securitas Healthcare***Technical Support at 1 (800) 380-8883.******We ask that you please complete and return a copy of the acknowledgement form attached hereto as Exhibit A. | SoftwareDesignChange
6 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210132 | Z-0046-2025 | Welch Allyn Life2000 Ventilation System, REF MS-01-0118, within exterior carton codes: 1) BT-20-0002, 2) BT-20-0002A, 3) BT-20-0002AP, 4) BT-20-0007, 5) BT-20-0011, and 6) RMS010118CP | Welch Allyn Life2000 Ventilation System | 1 | CDRH | 2024-10-22 14:23 | 2024-10-22 21:06 |  | Baxter Healthcare Corporation | The Life2000 ventilator may fail to initiate the Low Gas Pressure alarm if the pressure gas source (Life2000 compressor, oxygen cylinder or wall source) is not supplied to the ventilator before initiating therapy. | Class 1 Device Recall Welch Allyn Life2000 Ventilation System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall Welch Allyn Life2000 Ventilation System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall Welch Allyn Life2000 Ventilation System Date Initiated by Firm September 12, 2024 Date Posted October 22, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-0046-2025 Recall Event ID 95398 510(K)Number K170037 Product Classification Ventilator, continuous, facility use - Product Code CBK Product Welch Allyn Life2000 Ventilation System, REF MS-01-0118, within exterior carton codes: 1) BT-20-0002, 2) BT-20-0002A, 3) BT-20-0002AP, 4) BT-20-0007, 5) BT-20-0011, and 6) RMS010118CP Code Information UDI/DI 00887761978089 or 00815410020537, All Life2000 Ventilators with software version 06.08.00.00 Recalling Firm/ Manufacturer Baxter Healthcare Corporation 1 Baxter Pkwy Deerfield IL 60015-4625 For Additional Information Contact Center for One Baxter 800-422-9837 Manufacturer Reason for Recall The Life2000 ventilator may fail to initiate the Low Gas Pressure alarm if the pressure gas source (Life2000 compressor, oxygen cylinder or wall source) is not supplied to the ventilator before initiating therapy. FDA Determined Cause 2 Software Design Change Action Baxter issued an URGENT MEDICAL DEVICE CORRECTION notice to its consignees 09/12/2024 via USPS first class mail. The notice explained the issue with the device, the hazard involved, and requested the following: "Actions to be Taken by Patients***1. Prior to starting therapy, patients should ensure that the Life2000 ventilator is connected to a pressure gas source (Life2000 compressor, oxygen cylinder, or wall source) first and then turned on. The system setup instructions can be found on the Quick Reference Guide, page 11-13.***2. Refer to Attachment A for the website link and QR code for the Quick Reference Guide, Instructions for Use and additional information regarding Troubleshooting Guidance. ***3. If you received this communication directly from Baxter, acknowledge receipt of this letter using one of the two methods detailed on the enclosed Home Patient Reply Form Instruction Sheet. This step is required per FDA guidelines. Acknowledging receipt promptly will prevent you from receiving repeat notifications. If you do not complete the acknowledgement, you will receive a phone call from OnProcess Technology on behalf of Baxter to confirm your receipt of this notification.***4. If you purchased this product from a distributor or wholesaler, please note that responding on the Baxter customer portal is not applicable. If a response is requested by your distributor or wholesaler, please respond to the supplier according to their instructions. ***Further Information and Support***For general questions regarding this Urgent Medical Device Correction, please contact Advanced Respiratory, Inc. Home Care Customer Service at 800-426-4224, option 3, between the hours of 7:30 am and 6:00 pm Central Time or the Advanced Respiratory, Inc. Clinical Support team at 800-397-9071." Quantity in Commerce 10,957 units Distribution US nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = CBK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Baxter issued an URGENT MEDICAL DEVICE CORRECTION notice to its consignees 09/12/2024 via USPS first class mail. The notice explained the issue with the device, the hazard involved, and requested the following: "Actions to be Taken by Patients***1. Prior to starting therapy, patients should ensure that the Life2000 ventilator is connected to a pressure gas source (Life2000 compressor, oxygen cylinder, or wall source) first and then turned on. The system setup instructions can be found on the Quick Reference Guide, page 11-13.***2. Refer to Attachment A for the website link and QR code for the Quick Reference Guide, Instructions for Use and additional information regarding Troubleshooting Guidance. ***3. If you received this communication directly from Baxter, acknowledge receipt of this letter using one of the two methods detailed on the enclosed Home Patient Reply Form Instruction Sheet. This step is required per FDA guidelines. Acknowledging receipt promptly will prevent you from receiving repeat notifications. If you do not complete the acknowledgement, you will receive a phone call from OnProcess Technology on behalf of Baxter to confirm your receipt of this notification.***4. If you purchased this product from a distributor or wholesaler, please note that responding on the Baxter customer portal is not applicable. If a response is requested by your distributor or wholesaler, please respond to the supplier according to their instructions. ***Further Information and Support***For general questions regarding this Urgent Medical Device Correction, please contact Advanced Respiratory, Inc. Home Care Customer Service at 800-426-4224, option 3, between the hours of 7:30 am and 6:00 pm Central Time or the Advanced Respiratory, Inc. Clinical Support team at 800-397-9071." | SoftwareDesignChange
7 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210142 | Z-0139-2025 | Trinias unity, Digital Angiography System with Catheterization Table, Model: KS-100 This system consists of the following main components: C arm, catherization table, image processing equipment, x-ray high voltage generator, x-ray tube. There are several types of each component. | TRINIAS | 2 | CDRH | 2024-10-22 | 2024-10-22 21:07 |  | SHIMADZU CORPORATION MEDICAL SYSTEMS DIVISION | If digital angiography system is used with a specific catheterization table and the operator directs simultaneous movement of the table and the C-arm using direct memory, then direct memory button is released, the table may continue to move, which the operator can stop by pressing the stop button. | Class 2 Device Recall TRINIAS Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TRINIAS FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TRINIAS Date Initiated by Firm September 06, 2024 Date Posted October 22, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-0139-2025 Recall Event ID 95406 510(K)Number K123508 Product Classification Interventional fluoroscopic x-ray system - Product Code OWB Product Trinias unity, Digital Angiography System with Catheterization Table, Model: KS-100 This system consists of the following main components: C arm, catherization table, image processing equipment, x-ray high voltage generator, x-ray tube. There are several types of each component. Code Information UDI-DI: 04540217058327, 04540217049066, 04540217058143******Serial Numbers: 41EFCD7AC001, 41E58C773001, 41EFCD79C001, 41EFCD78B001, 41EFC73A9001, 41EFCD7D4001, 41E58C76B001, 41EFC73D9001, 41EFCD77C001, 41EFC73A5001, 41EFC73C5001, 41EFCD79A001, 41EFCD7A6001. Recalling Firm/ Manufacturer SHIMADZU CORPORATION MEDICAL SYSTEMS DIVISION Nishinokyo Kuwabara -Cho Nakagyo-Ku; 1 Kyoto Japan For Additional Information Contact 8175823-1305 Manufacturer Reason for Recall If digital angiography system is used with a specific catheterization table and the operator directs simultaneous movement of the table and the C-arm using direct memory, then direct memory button is released, the table may continue to move, which the operator can stop by pressing the stop button. FDA Determined Cause 2 Software Design Change Action On 9/6/24, recall notices were sent to customers who were asked to do the following:******1) Firmware for the system controller requires updating. The work to correct this issue will***be performed by your Authorized Service provider.******2) Schedule time with your Distributer to have your system updated.******3) Complete and return the response form to drodriguez@shimadzu-usa.com******Questions can be directed to the firm's National Technical Support:***Toll Free: (844)487-2767 Available 8:00 to 5:00 PM, Monday - Friday Quantity in Commerce 13 Distribution US Nationwide distribution in the states of MA, IL, SC, KY, SC, OH, TN. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = OWB - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 9/6/24, recall notices were sent to customers who were asked to do the following:******1) Firmware for the system controller requires updating. The work to correct this issue will***be performed by your Authorized Service provider.******2) Schedule time with your Distributer to have your system updated.******3) Complete and return the response form to drodriguez@shimadzu-usa.com******Questions can be directed to the firm's National Technical Support:***Toll Free: (844)487-2767 Available 8:00 to 5:00 PM, Monday - Friday | SoftwareDesignChange
8 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209798 | Z-3160-2024 | GE Healthcare Centricity Universal Viewer Zero Footprint Client, Model Numbers: ***a) 2110344-008; ***b) 2110344-009; ***c) 2110344-010; ***d) 2110344-011; ***e) 2110344-013; ***f) 2110344-019; ***g) 2110344-025; ***h) 2110344-029; ***i) 2110344-030; ***j) 2110344-032******; Radiological image processing system | GE Healthcare Centricity | 2 | CDRH | 2024-09-17 | 2024-09-17 21:07 |  | GE Healthcare | Footprint Client (ZFP) versions v6.0 SP9.x and SP10.x where the latest addended report is not shown by default to the user. The issue occurs when the addendum is created on the same day as the original report but at a different time, and only when launched on the following browsers.***" Internet Explorer 10 (IE10)***" Internet Explorer 11 (IE11)***" Firefox. | Class 2 Device Recall GE Healthcare Centricity Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GE Healthcare Centricity FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GE Healthcare Centricity Date Initiated by Firm July 23, 2024 Date Posted September 17, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-3160-2024 Recall Event ID 95238 510(K)Number K131977 Product Classification System, image processing, radiological - Product Code LLZ Product GE Healthcare Centricity Universal Viewer Zero Footprint Client, Model Numbers: ******a) 2110344-008; ******b) 2110344-009; ******c) 2110344-010; ******d) 2110344-011; ******e) 2110344-013; ******f) 2110344-019; ******g) 2110344-025; ******h) 2110344-029; ******i) 2110344-030; ******j) 2110344-032************; Radiological image processing system Code Information UDI/DI: 00840682102988: ***a) 2110344-008, Installed Product ID Numbers: B54834-Centricity Universal Viewer Zero Footprint / ZFP-02394744, KMCH-Centricity Universal Viewer Zero Footprint / ZFP-01819848, 355920-Centricity Universal Viewer Zero Footprint / ZFP-20996307, 225717-Centricity Universal Viewer Zero Footprint / ZFP-19666806, RF0827-Centricity Universal Viewer Zero Footprint / ZFP-19460712; ***b) 2110344-009, Installed Product ID Numbers: 1472-15-Centricity Universal Viewer Zero Footprint / ZFP-01372918, 1495-3-Centricity Universal Viewer Zero Footprint / ZFP-02348000, 1495-3-Centricity Universal Viewer Zero Footprint / ZFP-02348009, 1495-3-Centricity Universal Viewer Zero Footprint / ZFP-02348010, 1495-3-Centricity Universal Viewer Zero Footprint / ZFP-02348035, 1407636-Centricity Universal Viewer Zero Footprint / ZFP-19632983, SA1008-Centricity Universal Viewer Zero Footprint / ZFP-20620864; ***c) 2110344-010, Installed Product ID Numbers: 3922-18-Centricity Universal Viewer Zero Footprint / ZFP-00098227, 3922-18-Centricity Universal Viewer Zero Footprint / ZFP-00098228, 3922-18-Centricity Universal Viewer Zero Footprint / ZFP-14194760, 3922-18-Centricity Universal Viewer Zero Footprint / ZFP-16023936, 3922-18-Centricity Universal Viewer Zero Footprint / ZFP-16023938, 3922-18-Centricity Universal Viewer Zero Footprint / ZFP-16023939, 3922-18-Centricity Universal Viewer Zero Footprint / ZFP-16023940, 3922-18-Centricity Universal Viewer Zero Footprint / ZFP-16023941, 3922-18-Centricity Universal Viewer Zero Footprint / ZFP-16023942, 3922-18-Centricity Universal Viewer Zero Footprint / ZFP-17192653, 1472-19-Centricity Universal Viewer Zero Footprint / ZFP-22125883, 946739-Centricity Universal Viewer Zero Footprint / ZFP-14767595, 753891-Centricity Universal Viewer Zero Footprint / ZFP-21739897, 753891-Centricity Universal Viewer Zero Footprint / ZFP-19737302, JP-340499-Centricity Universal Viewer Zero Footprint / ZFP-01275434, RU8902-Centricity Universal Viewer Zero Footprint / ZFP-17775055, RU9383-Centricity Universal Viewer Zero Footprint / ZFP-16868184; ***d) 2110344-011, Installed Product ID Numbers: 3805-14-Centricity Universal Viewer Zero Footprint / ZFP-01564901, 3805-14-Centricity Universal Viewer Zero Footprint / ZFP-02098755, 3805-14-Centricity Universal Viewer Zero Footprint / ZFP-02098757, 3805-14-Centricity Universal Viewer Zero Footprint / ZFP-02098758, 4084-1-Centricity Universal Viewer Zero Footprint / ZFP-02301039, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20036722; ***e) 2110344-013, Installed Product ID Numbers: RU1473-Centricity Universal Viewer Zero Footprint / ZFP-02331227, RU3152-Centricity Universal Viewer Zero Footprint / ZFP-02314018, RF1194-Centricity Universal Viewer Zero Footprint / ZFP-19716329, RU6265-Centricity Universal Viewer Zero Footprint / ZFP-20080267, RU1107-Centricity Universal Viewer Zero Footprint / ZFP-01620728, RU9227-Centricity Universal Viewer Zero Footprint / ZFP-20080351, RU8412-Centricity Universal Viewer Zero Footprint / ZFP-19727241, RU7361-Centricity Universal Viewer Zero Footprint / ZFP-20084480; ***f) 2110344-019, Installed Product ID Numbers: 104586-2-Centricity Universal Viewer Zero Footprint / ZFP-01299076, 8121-6-Centricity Universal Viewer Zero Footprint / ZFP-01808438, 1472-15-Centricity Universal Viewer Zero Footprint / ZFP-01367875, 1472-15-Centricity Universal Viewer Zero Footprint / ZFP-01372919, 1472-15-Centricity Universal Viewer Zero Footprint / ZFP-01372920, 1472-15-Centricity Universal Viewer Zero Footprint / ZFP-01372921, 1472-15-Centricity Universal Viewer Zero Footprint / ZFP-01372922, 1495-1-Centricity Universal Viewer Zero Footprint / ZFP-00556036, 1495-1-Centricity Universal Viewer Zero Footprint / ZFP-00556037, 1495-1-Centricity Universal Viewer Zero Footprint / ZFP-02347655, 444402-Centricity Universal Viewer Zero Footprint / ZFP-20091853, JP-340206-Centricity Universal Viewer Zero Footprint / ZFP-01275380, JP-690277-Centricity Universal Viewer Zero Footprint / ZFP-19665249; ***g) 2110344-025, Installed Product ID Numbers: 104440-1-Centricity Universal Viewer Zero Footprint / ZFP-23264377, RU1152-Centricity Universal Viewer Zero Footprint / ZFP-20579574; ***h) 2110344-029, Installed Product ID Numbers: 105906-2-Centricity Universal Viewer Zero Footprint / ZFP-02312614, 4653-1-Centricity Universal Viewer Zero Footprint / ZFP-00161865, 10414-1-Centricity Universal Viewer Zero Footprint / ZFP-02371135, 10414-1-Centricity Universal Viewer Zero Footprint / ZFP-02371136, 10414-1-Centricity Universal Viewer Zero Footprint / ZFP-02371137, 10414-1-Centricity Universal Viewer Zero Footprint / ZFP-02371138, 10414-1-Centricity Universal Viewer Zero Footprint / ZFP-02371139, 10414-1-Centricity Universal Viewer Zero Footprint / ZFP-01983569, 10414-1-Centricity Universal Viewer Zero Footprint / ZFP-02257545, 4415-2-Centricity Universal Viewer Zero Footprint / ZFP-19609719, 4415-2-Centricity Universal Viewer Zero Footprint / ZFP-19609721, 4415-2-Centricity Universal Viewer Zero Footprint / ZFP-19609722, 4415-2-Centricity Universal Viewer Zero Footprint / ZFP-19609724, 4415-2-Centricity Universal Viewer Zero Footprint / ZFP-19609725, 4415-2-Centricity Universal Viewer Zero Footprint / ZFP-19609727, 4415-2-Centricity Universal Viewer Zero Footprint / ZFP-19609728, 4049-1-Centricity Universal Viewer Zero Footprint / ZFP-01805026, 4049-1-Centricity Universal Viewer Zero Footprint / ZFP-02302924, 103944-3-Centricity Universal Viewer Zero Footprint / ZFP-01477454, 103944-3-Centricity Universal Viewer Zero Footprint / ZFP-02286832, 3805-14-Centricity Universal Viewer Zero Footprint / ZFP-02098759, 1275-6-Centricity Universal Viewer Zero Footprint / ZFP-00161288, 1275-6-Centricity Universal Viewer Zero Footprint / ZFP-00161686, 1275-6-Centricity Universal Viewer Zero Footprint / ZFP-00161687, 1275-6-Centricity Universal Viewer Zero Footprint / ZFP-18469046, 1275-6-Centricity Universal Viewer Zero Footprint / ZFP-20653788, 106710-1-Centricity Universal Viewer Zero Footprint / ZFP-01810611, 106710-1-Centricity Universal Viewer Zero Footprint / ZFP-14768105, 106710-1-Centricity Universal Viewer Zero Footprint / ZFP-14769712, 102753-1-Centricity Universal Viewer Zero Footprint / ZFP-01810676, 102753-1-Centricity Universal Viewer Zero Footprint / ZFP-01978730, 102753-1-Centricity Universal Viewer Zero Footprint / ZFP-01978731, 102753-1-Centricity Universal Viewer Zero Footprint / ZFP-01978732, 105525-1-Centricity Universal Viewer Zero Footprint / ZFP-02315655, 105525-1-Centricity Universal Viewer Zero Footprint / ZFP-21306260, 104015-1-Centricity Universal Viewer Zero Footprint / ZFP-02157305, 3212-5-Centricity Universal Viewer Zero Footprint / ZFP-01740249, 3212-8-Centricity Universal Viewer Zero Footprint / ZFP-14761752, 3212-5-Centricity Universal Viewer Zero Footprint / ZFP-02157261, 10547-2-Centricity Universal Viewer Zero Footprint / ZFP-01294055, 10547-2-Centricity Universal Viewer Zero Footprint / ZFP-02337580, 9915-2-Centricity Universal Viewer Zero Footprint / ZFP-01459010, 9915-2-Centricity Universal Viewer Zero Footprint / ZFP-02257259, 9866-2-Centricity Universal Viewer Zero Footprint / ZFP-00101400, 9866-2-Centricity Universal Viewer Zero Footprint / ZFP-00101401, 10832-3-Centricity Universal Viewer Zero Footprint / ZFP-21207910, 105468-1-Centricity Universal Viewer Zero Footprint / ZFP-02254840, 4289-3-Centricity Universal Viewer Zero Footprint / ZFP-01976550, 102193-2-Centricity Universal Viewer Zero Footprint / ZFP-14765252, 1609-1-Centricity Universal Viewer Zero Footprint / ZFP-02344894, 10510-3-Centricity Universal Viewer Zero Footprint / ZFP-22188687, 3719-9-Centricity Universal Viewer Zero Footprint / ZFP-00161362, 3719-9-Centricity Universal Viewer Zero Footprint / ZFP-00161368, 3066-1-Centricity Universal Viewer Zero Footprint / ZFP-00101769, 3066-1-Centricity Universal Viewer Zero Footprint / ZFP-00101770, 3719-1-Centricity Universal Viewer Zero Footprint / ZFP-01987326, 1331-24-Centricity Universal Viewer Zero Footprint / ZFP-02041554, 1331-24-Centricity Universal Viewer Zero Footprint / ZFP-02190655, 11673-4-Centricity Universal Viewer Zero Footprint / ZFP-01814396, 101647-1-Centricity Universal Viewer Zero Footprint / ZFP-02205362, 106558-2-Centricity Universal Viewer Zero Footprint / ZFP-01987441, 4204-1-Centricity Universal Viewer Zero Footprint / ZFP-01459008, 6901-11-Centricity Universal Viewer Zero Footprint / ZFP-02341586, 6901-11-Centricity Universal Viewer Zero Footprint / ZFP-02354814, 6901-11-Centricity Universal Viewer Zero Footprint / ZFP-02354815, 6901-11-Centricity Universal Viewer Zero Footprint / ZFP-02354816, 6901-11-Centricity Universal Viewer Zero Footprint / ZFP-15291321, 6901-11-Centricity Universal Viewer Zero Footprint / ZFP-15291322, 6901-11-Centricity Universal Viewer Zero Footprint / ZFP-15291332, 6901-11-Centricity Universal Viewer Zero Footprint / ZFP-15291333, 6901-11-Centricity Universal Viewer Zero Footprint / ZFP-15291334, 6901-11-Centricity Universal Viewer Zero Footprint / ZFP-15291335, 6901-11-Centricity Universal Viewer Zero Footprint / ZFP-16355110, 105461-1-Centricity Universal Viewer Zero Footprint / ZFP-02035195, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-01974088, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-02354817, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-02354818, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-02354819, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-02354820, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-15291319, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-15291320, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-17594364, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-19020773, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-19020774, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-20621990, 6901-12-Centricity Universal Viewer Zero Footprint / ZFP-20621993, 1847-3-Centricity Universal Viewer Zero Footprint / ZFP-02097811, 1847-3-Centricity Universal Viewer Zero Footprint / ZFP-02255598, 1847-2-Centricity Universal Viewer Zero Footprint / ZFP-00169967, 1847-2-Centricity Universal Viewer Zero Footprint / ZFP-00169968, 1847-2-Centricity Universal Viewer Zero Footprint / ZFP-00169969, 1696-1-Centricity Universal Viewer Zero Footprint / ZFP-20694814, 1696-1-Centricity Universal Viewer Zero Footprint / ZFP-20699361, 1696-1-Centricity Universal Viewer Zero Footprint / ZFP-02186950, 1696-1-Centricity Universal Viewer Zero Footprint / ZFP-02269625, 100729-1-Centricity Universal Viewer Zero Footprint / ZFP-01721608, 6051-Centricity Universal Viewer Zero Footprint / ZFP-02277661, 2892781_Stvincentssyd-Centricity Universal Viewer Zero Footprint / ZFP-02186467, 1698348_Alfred-Centricity Universal Viewer Zero Footprint / ZFP-02032817, 1586578_Nedlands-Centricity Universal Viewer Zero Footprint / ZFP-20034444, 123456-Centricity Universal Viewer Zero Footprint / ZFP-14310910, 423834-Centricity Universal Viewer Zero Footprint / ZFP-24370371, 423834-Centricity Universal Viewer Zero Footprint / ZFP-24370372, 423834-Centricity Universal Viewer Zero Footprint / ZFP-24370373, 333266-Centricity Universal Viewer Zero Footprint / ZFP-16482049, 13628-Centricity Universal Viewer Zero Footprint / ZFP-02336211, 104295-2-Centricity Universal Viewer Zero Footprint / ZFP-01286791, 631789-Centricity Universal Viewer Zero Footprint / ZFP-01286792, 631789-Centricity Universal Viewer Zero Footprint / ZFP-21268829, B0201-12-Centricity Universal Viewer Zero Footprint / ZFP-02031960, 2227709-Centricity Universal Viewer Zero Footprint / ZFP-20080947, 1301843-Centricity Universal Viewer Zero Footprint / ZFP-14430019, S540080-Centricity Universal Viewer Zero Footprint / ZFP-00043934, M44850-Centricity Universal Viewer Zero Footprint / ZFP-01817813, M41884-Centricity Universal Viewer Zero Footprint / ZFP-02345947, PF1005-Centricity Universal Viewer Zero Footprint / ZFP-02301061, PF1005-Centricity Universal Viewer Zero Footprint / ZFP-02301134, U1982906 / NC1076-Centricity Universal Viewer Zero Footprint / ZFP-01987930, 4625913X - HC0565-Centricity Universal Viewer Zero Footprint / ZFP-00545594, X9000399 - HC0399-Centricity Universal Viewer Zero Footprint / ZFP-01282975, HC1451-Centricity Universal Viewer Zero Footprint / ZFP-01288243, 1657844X - 518086-Centricity Universal Viewer Zero Footprint / ZFP-00158637, X9000245 - HC0245-Centricity Universal Viewer Zero Footprint / ZFP-01462020, 362144-Centricity Universal Viewer Zero Footprint / ZFP-02263754, PC0070-Centricity Universal Viewer Zero Footprint / ZFP-02042361, UHAM02-Centricity Universal Viewer Zero Footprint / ZFP-00059329, X9003387 - HHEI01-Centricity Universal Viewer Zero Footprint / ZFP-21310713, 434487-Centricity Universal Viewer Zero Footprint / ZFP-22169581, 145774-Centricity Universal Viewer Zero Footprint / ZFP-01814045, 1257845X - 434487-Centricity Universal Viewer Zero Footprint / ZFP-22203508, 0828550068-Centricity Universal Viewer Zero Footprint / ZFP-02094418, 828550053-Centricity Universal Viewer Zero Footprint / ZFP-02253600, 434519-Centricity Universal Viewer Zero Footprint / ZFP-20677457, A1892038-Centricity Universal Viewer Zero Footprint / ZFP-02093306, A56306-Centricity Universal Viewer Zero Footprint / ZFP-02180921, A56600-Centricity Universal Viewer Zero Footprint / ZFP-01454929, A6700602-Centricity Universal Viewer Zero Footprint / ZFP-02178761, A60858-Centricity Universal Viewer Zero Footprint / ZFP-02286910, JP-380108-Centricity Universal Viewer Zero Footprint / ZFP-02255381, JP-340255-Centricity Universal Viewer Zero Footprint / ZFP-01974303, JP-280245-Centricity Universal Viewer Zero Footprint / ZFP-02255463, JP-280250-Centricity Universal Viewer Zero Footprint / ZFP-02360544, JP-140140-Centricity Universal Viewer Zero Footprint / ZFP-02154910, JP-140051-Centricity Universal Viewer Zero Footprint / ZFP-02099295, JP-290033-Centricity Universal Viewer Zero Footprint / ZFP-01275435, JP-290033-Centricity Universal Viewer Zero Footprint / ZFP-01275437, JP-290033-Centricity Universal Viewer Zero Footprint / ZFP-01275438, JP-250378-Centricity Universal Viewer Zero Footprint / ZFP-02186431, JP-090905-Centricity Universal Viewer Zero Footprint / ZFP-02041046, JP-130422-Centricity Universal Viewer Zero Footprint / ZFP-01980318, JP-130527-Centricity Universal Viewer Zero Footprint / ZFP-02093919, JP-130626-Centricity Universal Viewer Zero Footprint / ZFP-02043333, JP-130761-Centricity Universal Viewer Zero Footprint / ZFP-02248293, KW1102-Centricity Universal Viewer Zero Footprint / ZFP-02191178, 59219_ERPerror-Centricity Universal Viewer Zero Footprint / ZFP-02156484, 60020_ERPerror-Centricity Universal Viewer Zero Footprint / ZFP-02307677, 60020_ERPerror-Centricity Universal Viewer Zero Footprint / ZFP-02291060, 709434-Centricity Universal Viewer Zero Footprint / ZFP-14765518, SA1075-Centricity Universal Viewer Zero Footprint / ZFP-01457247, UP285A-Centricity Universal Viewer Zero Footprint / ZFP-01815172, GPX458-Centricity Universal Viewer Zero Footprint / ZFP-01980394, GPX459-Centricity Universal Viewer Zero Footprint / ZFP-01980471, GPQ131-Centricity Universal Viewer Zero Footprint / ZFP-02254265, 407237-Centricity Universal Viewer Zero Footprint / ZFP-02290167, 529010-Centricity Universal Viewer Zero Footprint / ZFP-01824677, 34433-Centricity Universal Viewer Zero Footprint / ZFP-01294788, 16798599-Centricity Universal Viewer Zero Footprint / ZFP-01475999, 34435-Centricity Universal Viewer Zero Footprint / ZFP-00545336, 34449-Centricity Universal Viewer Zero Footprint / ZFP-22226438, 80674-Centricity Universal Viewer Zero Footprint / ZFP-01425411, 80674-Centricity Universal Viewer Zero Footprint / ZFP-21318253; ***i) 2110344-030, Installed Product ID Numbers: 4084-1-Centricity Universal Viewer Zero Footprint / ZFP-19600711, 9862-1-Centricity Universal Viewer Zero Footprint / ZFP-00100263, 9862-1-Centricity Universal Viewer Zero Footprint / ZFP-14765956, 1495-3-Centricity Universal Viewer Zero Footprint / ZFP-00096091, 1495-3-Centricity Universal Viewer Zero Footprint / ZFP-02348002, 1495-3-Centricity Universal Viewer Zero Footprint / ZFP-00096090, 1503588-Centricity Universal Viewer Zero Footprint / ZFP-14549617, 6891796-Centricity Universal Viewer Zero Footprint / ZFP-19631000, 306184-Centricity Universal Viewer Zero Footprint / ZFP-19620282, 324227-Centricity Universal Viewer Zero Footprint / ZFP-02040656, 320092-Centricity Universal Viewer Zero Footprint / ZFP-19675475, 329997-Centricity Universal Viewer Zero Footprint / ZFP-19732482, 3440585-Centricity Universal Viewer Zero Footprint / ZFP-02286138, 3440585-Centricity Universal Viewer Zero Footprint / ZFP-20092101, 327750-Centricity Universal Viewer Zero Footprint / ZFP-20904827, JP-380123-Centricity Universal Viewer Zero Footprint / ZFP-20101283, JP-400021-Centricity Universal Viewer Zero Footprint / ZFP-19649615, JP-030135-Centricity Universal Viewer Zero Footprint / ZFP-01275425, JP-420052-Centricity Universal Viewer Zero Footprint / ZFP-14761038, JP-420052-Centricity Universal Viewer Zero Footprint / ZFP-15111121, JP-420052-Centricity Universal Viewer Zero Footprint / ZFP-15111126, JP-272900-Centricity Universal Viewer Zero Footprint / ZFP-02041843, JP-270263-Centricity Universal Viewer Zero Footprint / ZFP-01275411, 59219_ERPerror-Centricity Universal Viewer Zero Footprint / ZFP-01299106, 651240-Centricity Universal Viewer Zero Footprint / ZFP-02173181, 328816-Centricity Universal Viewer Zero Footprint / ZFP-17805275, 4879575-Centricity Universal Viewer Zero Footprint / ZFP-15743070, U-4J5U54-Centricity Universal Viewer Zero Footprint / ZFP-01455930, RU7579-Centricity Universal Viewer Zero Footprint / ZFP-18957266, 41400-Centricity Universal Viewer Zero Footprint / ZFP-02291025; ***j) 2110344-032, Installed Product ID Numbers: 3915-10-Centricity Universal Viewer Zero Footprint / ZFP-02360681, 3915-10-Centricity Universal Viewer Zero Footprint / ZFP-18867313, 4653-1-Centricity Universal Viewer Zero Footprint / ZFP-00161866, 4653-1-Centricity Universal Viewer Zero Footprint / ZFP-19618935, 4653-1-Centricity Universal Viewer Zero Footprint / ZFP-21780800, 103710-1-Centricity Universal Viewer Zero Footprint / ZFP-02308031, 104586-2-Centricity Universal Viewer Zero Footprint / ZFP-23066609, 4415-6-Centricity Universal Viewer Zero Footprint / ZFP-01810693, 2195-15-Centricity Universal Viewer Zero Footprint / ZFP-01291411, 2195-15-Centricity Universal Viewer Zero Footprint / ZFP-00165673, 2195-15-Centricity Universal Viewer Zero Footprint / ZFP-00168142, 2195-15-Centricity Universal Viewer Zero Footprint / ZFP-00168143, 2195-15-Centricity Universal Viewer Zero Footprint / ZFP-01448429, 2195-10-Centricity Universal Viewer Zero Footprint / ZFP-20999747, 105281-1-Centricity Universal Viewer Zero Footprint / ZFP-01820425, 100256-3-Centricity Universal Viewer Zero Footprint / ZFP-01810795, 100256-3-Centricity Universal Viewer Zero Footprint / ZFP-19604803, 100256-3-Centricity Universal Viewer Zero Footprint / ZFP-19604804, 100256-3-Centricity Universal Viewer Zero Footprint / ZFP-19604805, 100256-3-Centricity Universal Viewer Zero Footprint / ZFP-19604806, 4415-8-Centricity Universal Viewer Zero Footprint / ZFP-20274847, 4415-2-Centricity Universal Viewer Zero Footprint / ZFP-19609729, 4415-7-Centricity Universal Viewer Zero Footprint / ZFP-20274843, 8142-4-Centricity Universal Viewer Zero Footprint / ZFP-02397428, 8142-4-Centricity Universal Viewer Zero Footprint / ZFP-19715369, 8142-4-Centricity Universal Viewer Zero Footprint / ZFP-19715370, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-00099694, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-01294051, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-02350740, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-16265092, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-20019523, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-20019524, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-20019525, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-20019526, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-20019527, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-20036734, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-20036735, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-20036736, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-20036737, 105953-2-Centricity Universal Viewer Zero Footprint / ZFP-02345969, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-22268979, 100620-23-Centricity Universal Viewer Zero Footprint / ZFP-22268125, 103696-8-Centricity Universal Viewer Zero Footprint / ZFP-16868149, 4420-6-Centricity Universal Viewer Zero Footprint / ZFP-01810727, 4420-6-Centricity Universal Viewer Zero Footprint / ZFP-20032680, 4420-6-Centricity Universal Viewer Zero Footprint / ZFP-20032681, 4420-6-Centricity Universal Viewer Zero Footprint / ZFP-20032690, 2204-1-Centricity Universal Viewer Zero Footprint / ZFP-01893572, 2317-3-Centricity Universal Viewer Zero Footprint / ZFP-00165278, 2317-3-Centricity Universal Viewer Zero Footprint / ZFP-00165279, 2317-3-Centricity Universal Viewer Zero Footprint / ZFP-02337066, 2317-3-Centricity Universal Viewer Zero Footprint / ZFP-02337067, 105441-1-Centricity Universal Viewer Zero Footprint / ZFP-01987335, 105950-1-Centricity Universal Viewer Zero Footprint / ZFP-02344194, 105950-1-Centricity Universal Viewer Zero Footprint / ZFP-02369612, 103835-2-Centricity Universal Viewer Zero Footprint / ZFP-02314205, 103835-2-Centricity Universal Viewer Zero Footprint / ZFP-02336078, 103835-2-Centricity Universal Viewer Zero Footprint / ZFP-15050663, 103835-2-Centricity Universal Viewer Zero Footprint / ZFP-15050664, 103835-2-Centricity Universal Viewer Zero Footprint / ZFP-15050665, 103944-3-Centricity Universal Viewer Zero Footprint / ZFP-02286833, 103944-3-Centricity Universal Viewer Zero Footprint / ZFP-02286834, 103944-3-Centricity Universal Viewer Zero Footprint / ZFP-02286835, 103944-3-Centricity Universal Viewer Zero Footprint / ZFP-02286836, 10345-1-Centricity Universal Viewer Zero Footprint / ZFP-01419025, 10345-1-Centricity Universal Viewer Zero Footprint / ZFP-02189674, 10345-9-Centricity Universal Viewer Zero Footprint / ZFP-01425634, 10363-2-Centricity Universal Viewer Zero Footprint / ZFP-00555476, 10363-2-Centricity Universal Viewer Zero Footprint / ZFP-21992896, 104473-1-Centricity Universal Viewer Zero Footprint / ZFP-01813754, 10388-1-Centricity Universal Viewer Zero Footprint / ZFP-14762271, 10388-1-Centricity Universal Viewer Zero Footprint / ZFP-19714873, 10388-1-Centricity Universal Viewer Zero Footprint / ZFP-19714877, 10388-1-Centricity Universal Viewer Zero Footprint / ZFP-19714878, 10536-1-Centricity Universal Viewer Zero Footprint / ZFP-00161373, 10536-1-Centricity Universal Viewer Zero Footprint / ZFP-00161372, 10536-1-Centricity Universal Viewer Zero Footprint / ZFP-21325653, 1635-8-Centricity Universal Viewer Zero Footprint / ZFP-18227238, 9880-6-Centricity Universal Viewer Zero Footprint / ZFP-21775960, 102719-1-Centricity Universal Viewer Zero Footprint / ZFP-00098294, 102719-1-Centricity Universal Viewer Zero Footprint / ZFP-00098295, 102719-6-Centricity Universal Viewer Zero Footprint / ZFP-21318832, 100171-15-Centricity Universal Viewer Zero Footprint / ZFP-20869504, 100171-15-Centricity Universal Viewer Zero Footprint / ZFP-00098296, B0201-64-Centricity Universal Viewer Zero Footprint / ZFP-22255230, 446860-Centricity Universal Viewer Zero Footprint / ZFP-22255231, 102768-1-Centricity Universal Viewer Zero Footprint / ZFP-02302156, 10283-1-Centricity Universal Viewer Zero Footprint / ZFP-00097951, 3228-10-Centricity Universal Viewer Zero Footprint / ZFP-01442252, 3228-12-Centricity Universal Viewer Zero Footprint / ZFP-01451967, 3228-12-Centricity Universal Viewer Zero Footprint / ZFP-02003801, 3228-12-Centricity Universal Viewer Zero Footprint / ZFP-02256569, 10725-1-Centricity Universal Viewer Zero Footprint / ZFP-00099453, 10725-1-Centricity Universal Viewer Zero Footprint / ZFP-00099454, 10725-1-Centricity Universal Viewer Zero Footprint / ZFP-15291356, 10725-1-Centricity Universal Viewer Zero Footprint / ZFP-19609380, 10725-7-Centricity Universal Viewer Zero Footprint / ZFP-15291372, 4362-1-Centricity Universal Viewer Zero Footprint / ZFP-14640680, 4362-1-Centricity Universal Viewer Zero Footprint / ZFP-19732060, 3909-2-Centricity Universal Viewer Zero Footprint / ZFP-14640650, 3909-2-Centricity Universal Viewer Zero Footprint / ZFP-20095312, 3909-2-Centricity Universal Viewer Zero Footprint / ZFP-20095313, 1542-5-Centricity Universal Viewer Zero Footprint / ZFP-00116852, 1542-5-Centricity Universal Viewer Zero Footprint / ZFP-01970562, 1542-5-Centricity Universal Viewer Zero Footprint / ZFP-02345660, 105482-1-Centricity Universal Viewer Zero Footprint / ZFP-02093928, 104073-1-Centricity Universal Viewer Zero Footprint / ZFP-19660132, 104073-1-Centricity Universal Viewer Zero Footprint / ZFP-21281793, 1609-1-Centricity Universal Viewer Zero Footprint / ZFP-00104758, 1609-1-Centricity Universal Viewer Zero Footprint / ZFP-00104772, 1609-1-Centricity Universal Viewer Zero Footprint / ZFP-00104773, 1609-1-Centricity Universal Viewer Zero Footprint / ZFP-02344892, 4418-3-Centricity Universal Viewer Zero Footprint / ZFP-15291378, 4418-3-Centricity Universal Viewer Zero Footprint / ZFP-19604646, 3922-25-Centricity Universal Viewer Zero Footprint / ZFP-01290685, 100254-7-Centricity Universal Viewer Zero Footprint / ZFP-00096557, 100254-7-Centricity Universal Viewer Zero Footprint / ZFP-14776418, 100254-7-Centricity Universal Viewer Zero Footprint / ZFP-14776426, 100254-1-Centricity Universal Viewer Zero Footprint / ZFP-21752638, 3922-17-Centricity Universal Viewer Zero Footprint / ZFP-00098229, B0201-64-Centricity Universal Viewer Zero Footprint / ZFP-22255230, 11673-4-Centricity Universal Viewer Zero Footprint / ZFP-02349623, 11673-4-Centricity Universal Viewer Zero Footprint / ZFP-02349624, 9894-1-Centricity Universal Viewer Zero Footprint / ZFP-02173723, 4157-3-Centricity Universal Viewer Zero Footprint / ZFP-00100067, 4157-3-Centricity Universal Viewer Zero Footprint / ZFP-00108112, 4157-3-Centricity Universal Viewer Zero Footprint / ZFP-02363496, 4157-3-Centricity Universal Viewer Zero Footprint / ZFP-02363498, 4157-6-Centricity Universal Viewer Zero Footprint / ZFP-02290179, 4157-3-Centricity Universal Viewer Zero Footprint / ZFP-20943326, 4157-11-Centricity Universal Viewer Zero Footprint / ZFP-21282125, 4541-1-Centricity Universal Viewer Zero Footprint / ZFP-00105159, 4541-1-Centricity Universal Viewer Zero Footprint / ZFP-00105160, 102359-1-Centricity Universal Viewer Zero Footprint / ZFP-02186080, 103918-1-Centricity Universal Viewer Zero Footprint / ZFP-00100288, 103696-4-Centricity Universal Viewer Zero Footprint / ZFP-20698641, 103696-4-Centricity Universal Viewer Zero Footprint / ZFP-00100649, 103696-33-Centricity Universal Viewer Zero Footprint / ZFP-20098431, 103696-30-Centricity Universal Viewer Zero Footprint / ZFP-21237479, 103696-4-Centricity Universal Viewer Zero Footprint / ZFP-20019515, 103696-4-Centricity Universal Viewer Zero Footprint / ZFP-22284069, 103696-4-Centricity Universal Viewer Zero Footprint / ZFP-20036716, 4325-52-Centricity Universal Viewer Zero Footprint / ZFP-01367195, 4325-52-Centricity Universal Viewer Zero Footprint / ZFP-02345640, 4325-52-Centricity Universal Viewer Zero Footprint / ZFP-02345641, 4325-51-Centricity Universal Viewer Zero Footprint / ZFP-01299096, 103696-3-Centricity Universal Viewer Zero Footprint / ZFP-16868150, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-00103256, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-02351584, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20019518, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20019519, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20019521, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20019522, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20036724, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20036725, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20036726, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20036727, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20036728, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20036731, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20036732, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-20036733, 100085-5-Centricity Universal Viewer Zero Footprint / ZFP-22284088, 105348-1-Centricity Universal Viewer Zero Footprint / ZFP-01893555, 3726-1-Centricity Universal Viewer Zero Footprint / ZFP-00599126, 2440-6-Centricity Universal Viewer Zero Footprint / ZFP-00160891, 2440-6-Centricity Universal Viewer Zero Footprint / ZFP-00160911, 2440-6-Centricity Universal Viewer Zero Footprint / ZFP-20903230, 100505-1-Centricity Universal Viewer Zero Footprint / ZFP-02375853, 102667-4-Centricity Universal Viewer Zero Footprint / ZFP-02355027, 102667-4-Centricity Universal Viewer Zero Footprint / ZFP-02355028, 102667-4-Centricity Universal Viewer Zero Footprint / ZFP-02183317, 102667-4-Centricity Universal Viewer Zero Footprint / ZFP-02348702, 102667-4-Centricity Universal Viewer Zero Footprint / ZFP-02348704, 103246-1-Centricity Universal Viewer Zero Footprint / ZFP-01973593, DZ27??-Centricity Universal Viewer Zero Footprint / ZFP-02176131, 1503586-Centricity Universal Viewer Zero Footprint / ZFP-21016603, 000009335-Centricity Universal Viewer Zero Footprint / ZFP-19631916, AH5045-Centricity Universal Viewer Zero Footprint / ZFP-17594407, 313084-Centricity Universal Viewer Zero Footprint / ZFP-14776791, 268882-Centricity Universal Viewer Zero Footprint / ZFP-19714224, 293834-Centricity Universal Viewer Zero Footprint / ZFP-21244289, 293834-Centricity Universal Viewer Zero Footprint / ZFP-21244291, 293834-Centricity Universal Viewer Zero Footprint / ZFP-21244292, 293834-Centricity Universal Viewer Zero Footprint / ZFP-02004797, 283039-Centricity Universal Viewer Zero Footprint / ZFP-01819807, 283039-Centricity Universal Viewer Zero Footprint / ZFP-02170574, 283039-Centricity Universal Viewer Zero Footprint / ZFP-02283242, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605119, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605120, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605122, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605124, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605125, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605126, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605127, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605128, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605129, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605133, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605247, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605248, 105800-Centricity Universal Viewer Zero Footprint / ZFP-00605250, 105800-Centricity Universal Vie |  | SoftwareDesignChange
9 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209118 | Z-2580-2024 | VITROS Immunodiagnostic VITROS Folate Reagent Pack 1/2 | VITROS Immunodiagnostic VITROS Folate Reagent Pack 1/2 | 2 | CDRH | 2024-08-12 | 2024-08-12 21:09 |  | Ortho-Clinical Diagnostics, Inc. | It was confirmed that fibrinogen in patient plasma samples precipitates out of solution upon the addition of folate stabilizer reagent as part of the pre-treatment process causing the increase in "TM5-4MB" condition codes which lead to a delay of results. | Class 2 Device Recall VITROS Immunodiagnostic VITROS Folate Reagent Pack 1/2 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS Immunodiagnostic VITROS Folate Reagent Pack 1/2 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS Immunodiagnostic VITROS Folate Reagent Pack 1/2 Date Initiated by Firm July 03, 2024 Date Posted August 12, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-2580-2024 Recall Event ID 95020 510(K)Number K001266 K984166 Product Classification Acid, folic, radioimmunoassay - Product Code CGN Product VITROS Immunodiagnostic VITROS Folate Reagent Pack 1/2 Code Information Model No 1513266; UDI/GTIN# 110758750009237; Lots within expiry: 3380, 3390, 3400, 3410, 3422, 3430, 3440, 3450, 3461, 3471, 3480. Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, Inc. 100 Indigo Creek Dr Rochester NY 14626-5101 For Additional Information Contact Joe Falvo 585-4533452 Manufacturer Reason for Recall It was confirmed that fibrinogen in patient plasma samples precipitates out of solution upon the addition of folate stabilizer reagent as part of the pre-treatment process causing the increase in "TM5-4MB" condition codes which lead to a delay of results. FDA Determined Cause 2 Software Design Change Action On July 15, 2024, IMPORTANT PRODUCT CORRECTION NOTIFICATION letters were sent to customers. REQUIRED ACTION***1. Send the enclosed Customer Letter and Confirmation of Receipt form to all customers who were shipped VITROS Immunodiagnostic Products Folate Reagent Pack 1/2 from your facility within the past 12 months.***2. Complete the enclosed Confirmation of Receipt form no later than July 23, 2024. If you have further questions, please contact the Global Services Organization at 1-800-421-3311. Quantity in Commerce 18,467 units Distribution Worldwide distribution - US Nationwide and the countries of Australia, Belgium, Bermuda, Brazil, Canada, Chile, China, Columbia, Denmark, France, Germany, India, Italy, Japan, Mexico, Netherlands, Norway, Portugal, Russia, Singapore, Sweden, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = CGN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On July 15, 2024, IMPORTANT PRODUCT CORRECTION NOTIFICATION letters were sent to customers. REQUIRED ACTION***1. Send the enclosed Customer Letter and Confirmation of Receipt form to all customers who were shipped VITROS Immunodiagnostic Products Folate Reagent Pack 1/2 from your facility within the past 12 months.***2. Complete the enclosed Confirmation of Receipt form no later than July 23, 2024. If you have further questions, please contact the Global Services Organization at 1-800-421-3311. | SoftwareDesignChange
10 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=208061 | Z-2325-2024 | Kosmos on iOS impacting iOS, Model Numbers: P007639, P007944, software is part of a general-purpose diagnostic ultrasound system. | Kosmos | 2 | CDRH | 2024-07-10 18:07 | 2024-07-10 21:07 |  | Echonous Inc | Ultrasound system has a bug in affected iOS software, which after DICOM export, can cause inaccuracy in postprocessed measurements done in 3rd party software, such as PACS or other post-processing software, which may lead to inaccurate diagnoses or treatment. | Class 2 Device Recall Kosmos Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Kosmos FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Kosmos Date Initiated by Firm May 09, 2024 Date Posted July 10, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-2325-2024 Recall Event ID 94709 510(K)Number K212100 Product Classification System, imaging, pulsed doppler, ultrasonic - Product Code IYN Product Kosmos on iOS impacting iOS, Model Numbers: P007639, P007944, software is part of a general-purpose diagnostic ultrasound system. Code Information Model/UDI-DI/Software versions:***P007639/09353278000717/1.0,***P007944/09353278000762/ 2.1, 2.2. Recalling Firm/ Manufacturer Echonous Inc 8310 154th Ave Ne Ste 200 Bldg B Redmond WA 98052-6180 Manufacturer Reason for Recall Ultrasound system has a bug in affected iOS software, which after DICOM export, can cause inaccuracy in postprocessed measurements done in 3rd party software, such as PACS or other post-processing software, which may lead to inaccurate diagnoses or treatment. FDA Determined Cause 2 Software Design Change Action On 5/9/24, notices were sent to customers informing them to refrain from performing measurements on Kosmos iOS images with third-party PACS or post-processing software, after DICOM export. Also, they were informed, that a software patch would be ready to upload soon.******On 5/22/24, notices were sent to customers informing them that software patch 2.1.4.53 was available and that users should download this software performing measurements on Kosmos iOS images with third-party PACS or post-processing software.******Follow the steps below and download this new version as soon as possible:***1. Open your currently installed Kosmos Ultrasound app and from the home screen go***to "Settings" > "DICOM". Write down your MWL and PACS profiles. You will need to reference these profiles after you download the new app. Please note: your MWL and PACS profiles will not be retained when you download the new app. If you***don't use DICOM or MWL then you can skip this step.***2. Export any images or clips you wish to keep via USB or DICOM. Please note:***existing images or clips will not be retained when you download the new app.***3. Delete the Kosmos Ultrasound app that is currently on your iPad(s)***4. Download Kosmos Ultrasound 2.1.4.53 from the Apple App store***5. Connect your iPad(s) to Wi-Fi and open the Kosmos Ultrasound 2.1.4.53 app to***download any purchased licenses and when prompted enable drivers***6. From the home screen go to "Settings" > "DICOM" and reconfigure your MWL***and PACS profiles with the profile information you wrote down in step 1. If you***don't use DICOM or MWL then you can skip this step also.******Fill out and return the acknowledgement form.******Questions or concerns can be directed to Customer Service at (844) 854-0800 service@echonous.com Quantity in Commerce 272 Distribution US: ID, TX, WI, KS, IL, GA, MI, CA, OH, FL, AK, UT, MD, NY, NV, CO, ND, WA, DC, LA, KY, MN, HI, NM, TN, NJ, ID, MI, VT, AZ, MO, AL, MA, NC******OUS: Denmark, Switzerland, Austria, Italy, France, United Kingdom, Germany, Spain, Romania, Finland, Romania, Australia, India, Hong Kong, Thailand Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = IYN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 5/9/24, notices were sent to customers informing them to refrain from performing measurements on Kosmos iOS images with third-party PACS or post-processing software, after DICOM export. Also, they were informed, that a software patch would be ready to upload soon.******On 5/22/24, notices were sent to customers informing them that software patch 2.1.4.53 was available and that users should download this software performing measurements on Kosmos iOS images with third-party PACS or post-processing software.******Follow the steps below and download this new version as soon as possible:***1. Open your currently installed Kosmos Ultrasound app and from the home screen go***to "Settings" > "DICOM". Write down your MWL and PACS profiles. You will need to reference these profiles after you download the new app. Please note: your MWL and PACS profiles will not be retained when you download the new app. If you***don't use DICOM or MWL then you can skip this step.***2. Export any images or clips you wish to keep via USB or DICOM. Please note:***existing images or clips will not be retained when you download the new app.***3. Delete the Kosmos Ultrasound app that is currently on your iPad(s)***4. Download Kosmos Ultrasound 2.1.4.53 from the Apple App store***5. Connect your iPad(s) to Wi-Fi and open the Kosmos Ultrasound 2.1.4.53 app to***download any purchased licenses and when prompted enable drivers***6. From the home screen go to "Settings" > "DICOM" and reconfigure your MWL***and PACS profiles with the profile information you wrote down in step 1. If you***don't use DICOM or MWL then you can skip this step also.******Fill out and return the acknowledgement form.******Questions or concerns can be directed to Customer Service at (844) 854-0800 service@echonous.com | SoftwareDesignChange
11 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=208053 | Z-2214-2024 | Software version 5.3 in: ***RAPIDPoint 500e Blood Gas System (USA)***Siemens Material Number: 11416751******RAPIDPoint 500e Blood Gas System (China) 11416752;***RAPIDPoint 500e Blood Gas System (Japan) 11416754;***RAPIDPoint 500e Blood Gas System (ROW) 11416755 ;*** | RAPIDPoint 500e Blood Gas System | 2 | CDRH | 2024-06-26 | 2024-06-26 21:06 |  | Siemens Healthcare Diagnostics Inc | Siemens Healthineers has identified an issue with the RAPIDPoint500e system software version 5.3. This issue affects how the sample source is identified at the LIS when using the capillary mode. With this software version, samples run in the capillary mode are incorrectly labeled as arterial when displayed on the LIS, leading to the potential for results to be flagged according to an established arterial range. Capillary sample results are correctly identified as capillary on both the RAPIDPoint 500e system display and the instrument printouts. The issue only affects the new software version 5.3 and only capillary sample mode is impacted. The inability to distinguish specimen sources for blood gas may lead to incorrect diagnoses and management decisions. | Class 2 Device Recall RAPIDPoint 500e Blood Gas System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall RAPIDPoint 500e Blood Gas System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall RAPIDPoint 500e Blood Gas System Date Initiated by Firm May 17, 2024 Date Posted June 26, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-2214-2024 Recall Event ID 94704 510(K)Number K192240 Product Classification Electrode measurement, blood-gases (pco2, po2) and blood ph - Product Code CHL Product Software version 5.3 in: ******RAPIDPoint 500e Blood Gas System (USA)******Siemens Material Number: 11416751************RAPIDPoint 500e Blood Gas System (China) 11416752;******RAPIDPoint 500e Blood Gas System (Japan) 11416754;******RAPIDPoint 500e Blood Gas System (ROW) 11416755 ; Code Information UDI-DI:***(1)RAPIDPoint 500e Blood Gas System (USA) 00630414286143;***(2)RAPIDPoint 500e Blood Gas System (China) 00630414286150;***(3) RAPIDPoint 500e Blood Gas System (Japan) 00630414286167;***(4) RAPIDPoint 500e Blood Gas System (ROW) 00630414286174******Software Version 5.3. Recalling Firm/ Manufacturer Siemens Healthcare Diagnostics Inc 2 Edgewater Dr Norwood MA 02062-4637 For Additional Information Contact SAME 781-269-3000 Manufacturer Reason for Recall Siemens Healthineers has identified an issue with the RAPIDPoint500e system software version 5.3. This issue affects how the sample source is identified at the LIS when using the capillary mode. With this software version, samples run in the capillary mode are incorrectly labeled as arterial when displayed on the LIS, leading to the potential for results to be flagged according to an established arterial range. Capillary sample results are correctly identified as capillary on both the RAPIDPoint 500e system display and the instrument printouts. The issue only affects the new software version 5.3 and only capillary sample mode is impacted. The inability to distinguish specimen sources for blood gas may lead to incorrect diagnoses and management decisions. FDA Determined Cause 2 Software Design Change Action Siemens Healthineers issued Urgent Medical Device Correction Letter (POC 24-015.A.US)*** on 5/1/7/24 to US customers via Fed'x. An Urgent Field Safety Notice (UFSN) POC 24-015.A.OUS was issued to all OUS Siemens reporting sites with affected customers for distribution and regional***reporting to their local competent authorities.*** Letter states reason for recall, health risk and action to take:***" Please review this letter with your Medical Director to determine the appropriate course of action.***" Siemens will convert all RAPIDPoint 500e systems at software v5.3 to v5.2 regardless of whether the capillary sample mode is used. The software conversion will be performed by Siemens Service group. Please contact Siemens customer support for the system software conversion.***" Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days.***" Adverse reactions or quality problems experienced with the use of the affected products listed in the above table may be reported to the FDA s MedWatch Adverse Event Reporting program either online, by regular mail or by fax.***" Please retain this letter with your laboratory records and forward this letter to those who may have received this product.***" For any questions, you may call our 24/7 Customer Service USA line at 877-229-3711. Quantity in Commerce 9 units US Distribution Worldwide - US Nationwide distribution in the states of FL, GA, LA, MI, NC and the countries of Chile, Costa Rica, France, Germany, Japan, Lebanon, Mexico, Netherlands, New Zealand, Poland, Spain, Turkey, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = CHL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens Healthineers issued Urgent Medical Device Correction Letter (POC 24-015.A.US)*** on 5/1/7/24 to US customers via Fed'x. An Urgent Field Safety Notice (UFSN) POC 24-015.A.OUS was issued to all OUS Siemens reporting sites with affected customers for distribution and regional***reporting to their local competent authorities.*** Letter states reason for recall, health risk and action to take:***" Please review this letter with your Medical Director to determine the appropriate course of action.***" Siemens will convert all RAPIDPoint 500e systems at software v5.3 to v5.2 regardless of whether the capillary sample mode is used. The software conversion will be performed by Siemens Service group. Please contact Siemens customer support for the system software conversion.***" Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days.***" Adverse reactions or quality problems experienced with the use of the affected products listed in the above table may be reported to the FDA s MedWatch Adverse Event Reporting program either online, by regular mail or by fax.***" Please retain this letter with your laboratory records and forward this letter to those who may have received this product.***" For any questions, you may call our 24/7 Customer Service USA line at 877-229-3711. | SoftwareDesignChange
12 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=206914 | Z-1609-2024 | t:connect mobile app used in conjunction with t:slim X2 insulin pump with Control-IQ technology | t:connect mobile app | 1 | CDRH | 2024-05-06 15:02 | 2024-05-06 21:06 |  | Tandem Diabetes Care, Inc. | During normal use, the mobile app version 2.7 may crash and be automatically relaunched by the iOS operating system. This cycle intermittently repeats, which leads to excessive Bluetooth communication that may result in pump battery drain and may lead to the pump shutting down sooner than typically expected. Pump shutdown will cause insulin delivery to suspend, which could lead to an under-delivery of insulin and may result in hyperglycemia, including severe hyperglycemia. | Class 1 Device Recall t:connect mobile app Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall t:connect mobile app FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall t:connect mobile app Date Initiated by Firm March 05, 2024 Date Posted May 06, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1609-2024 Recall Event ID 94312 510(K)Number K203234 Product Classification Alternate controller enabled insulin infusion pump - Product Code QFG Product t:connect mobile app used in conjunction with t:slim X2 insulin pump with Control-IQ technology Code Information version 2.7 or later /UDI: 00850006613274 Recalling Firm/ Manufacturer Tandem Diabetes Care, Inc. 11075 Roselle St San Diego CA 92121-1204 For Additional Information Contact Ms. Susan Morrison 877-801-6901 Manufacturer Reason for Recall During normal use, the mobile app version 2.7 may crash and be automatically relaunched by the iOS operating system. This cycle intermittently repeats, which leads to excessive Bluetooth communication that may result in pump battery drain and may lead to the pump shutting down sooner than typically expected. Pump shutdown will cause insulin delivery to suspend, which could lead to an under-delivery of insulin and may result in hyperglycemia, including severe hyperglycemia. FDA Determined Cause 2 Software Design Change Action On 10/09/2024, Tandem send via email an "URGENT MEDICAL DEVICE CORRECTION UPDATE" informing customers of additional actions required (uninstall and reinstall mobile app version 2.8.2) to address the faster than expected battery depletion issue.******Customers are instructed to:***1.Touch and hold the app icon.***2.Tap Remove App, then tap Delete App to confirm.***3.Go to Apple App Store.***4.Search t:connect mobile to locate the t:connect mobile app.***5.Tap Download, or the cloud icon with an arrow.***6.Login with your t:connect account credentials and re-pair with your Tandem t:slim X2 pump. For complete re-pair instructions, please visit our website here.******Between the dates of August 9-12, 2024, ******Tandem issued a Press Release and sent an "URGENT MEDICAL DEVICE CORRECTION UPDATE" Letter due to becoming aware that users may still be experiencing pump battery depletion after updating the mobile app to version 2.7.1.******On August 19, 2024, Tandem issued a press release announcing the future release of the next version of the t:connect mobile app version 2.8.2. ******On August 20, 2024, Tandem sent a follow-up email and in-app push notification to all affected users who Tandem has a valid email announcing the release of T:connect mobile app version 2.8.2.****** - The customer email communication will include a URL and QR code to provide a method for users to acknowledge receipt. ******Between the dates of August 21, 2024 - September 4, 2024, Tandem will send an "Urgent Medical Device Correction Update" to all affected app users via First-Class Postal mail. This notification will outline the release of T:connect mobile app version 2.8.2. Quantity in Commerce 85,863 applications Distribution US nationwide distribution including Puerto Rico. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = QFG - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 10/09/2024, Tandem send via email an "URGENT MEDICAL DEVICE CORRECTION UPDATE" informing customers of additional actions required (uninstall and reinstall mobile app version 2.8.2) to address the faster than expected battery depletion issue.******Customers are instructed to:***1.Touch and hold the app icon.***2.Tap Remove App, then tap Delete App to confirm.***3.Go to Apple App Store.***4.Search t:connect mobile to locate the t:connect mobile app.***5.Tap Download, or the cloud icon with an arrow.***6.Login with your t:connect account credentials and re-pair with your Tandem t:slim X2 pump. For complete re-pair instructions, please visit our website here.******Between the dates of August 9-12, 2024, ******Tandem issued a Press Release and sent an "URGENT MEDICAL DEVICE CORRECTION UPDATE" Letter due to becoming aware that users may still be experiencing pump battery depletion after updating the mobile app to version 2.7.1.******On August 19, 2024, Tandem issued a press release announcing the future release of the next version of the t:connect mobile app version 2.8.2. ******On August 20, 2024, Tandem sent a follow-up email and in-app push notification to all affected users who Tandem has a valid email announcing the release of T:connect mobile app version 2.8.2.****** - The customer email communication will include a URL and QR code to provide a method for users to acknowledge receipt. ******Between the dates of August 21, 2024 - September 4, 2024, Tandem will send an "Urgent Medical Device Correction Update" to all affected app users via First-Class Postal mail. This notification will outline the release of T:connect mobile app version 2.8.2. | SoftwareDesignChange
13 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=206605 | Z-1484-2024 | LVP Software of the Ivenix Infusion System (IIS), Version 5.8.0. | Ivenix Infusion Pump LVP Software | 1 | CDRH | 2024-04-11 13:41 | 2024-04-11 21:06 |  | Fresenius Kabi USA, LLC | Software has anomalies that have the potential to cause underdose, overdose, or delay in therapy which could lead to serious patient harm or death. | Class 1 Device Recall Ivenix Infusion Pump LVP Software Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall Ivenix Infusion Pump LVP Software FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall Ivenix Infusion Pump LVP Software Date Initiated by Firm March 07, 2024 Date Posted April 11, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1484-2024 Recall Event ID 94228 510(K)Number K183311 Product Classification Pump, infusion - Product Code FRN Product LVP Software of the Ivenix Infusion System (IIS), Version 5.8.0. Code Information Product LVP-SW-0004; UDI-DI 00811505030122. Recalling Firm/ Manufacturer Fresenius Kabi USA, LLC 50 High St Ste 50 North Andover MA 01845-2620 For Additional Information Contact Rebecca McCandless 847-550-2913 Manufacturer Reason for Recall Software has anomalies that have the potential to cause underdose, overdose, or delay in therapy which could lead to serious patient harm or death. FDA Determined Cause 2 Software Design Change Action A letter titled URGENT - Medical Device Field Correction, dated 3/7/24, was delivered by hand to consignees or by mail notifying them of this recall event. This notice details eleven software anomalies that could pose a potential risk of serious patient harm or death and the mitigations consignees should perform until software is updated. Consignees are to update their software to LVP software version 5.9.1 by reaching out to their Fresenius Kabi representative at 1-978-775-8100 to schedule installation. Consignees are to complete and return the Customer Reply Form via fax or email. Questions about the recall notice can be directed to Fresenius Kabi by email at Ivenix_support@fresenius-kabi.com or by phone at 1-978-775-8100 (Monday through Friday from 8:30 am to 5:00 pm). Quantity in Commerce 17 units Distribution US Distribution: CA, CO, ID, MI, NE, NJ, NV, OK, TX, UT, VA, WA, & WI. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = FRN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action A letter titled URGENT - Medical Device Field Correction, dated 3/7/24, was delivered by hand to consignees or by mail notifying them of this recall event. This notice details eleven software anomalies that could pose a potential risk of serious patient harm or death and the mitigations consignees should perform until software is updated. Consignees are to update their software to LVP software version 5.9.1 by reaching out to their Fresenius Kabi representative at 1-978-775-8100 to schedule installation. Consignees are to complete and return the Customer Reply Form via fax or email. Questions about the recall notice can be directed to Fresenius Kabi by email at Ivenix_support@fresenius-kabi.com or by phone at 1-978-775-8100 (Monday through Friday from 8:30 am to 5:00 pm). | SoftwareDesignChange
14 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205928 | Z-1230-2024 | VITROS 4600 Chemistry System Catalog Nos. 6802445 (regular) and 6900440 (refurbished) | VITROS 4600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:06 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. | Class 2 Device Recall VITROS 4600 Immunodiagnostic System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS 4600 Immunodiagnostic System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS 4600 Immunodiagnostic System Date Initiated by Firm January 26, 2024 Date Posted March 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1230-2024 Recall Event ID 93927 510(K)Number K031924 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product VITROS 4600 Chemistry System Catalog Nos. 6802445 (regular) and 6900440 (refurbished) Code Information VITROS System Software Versions 3.8.0 or 3.8.1***Regular UDI-DI 10758750012343***Refurbished UD-DI 10758750033201******Serial Numbers:******US: ***J46000208***J46000209***J46000240***J46000497***J46000864***J46000865***J46000898***J46001030***J46001064***J46001125***J46001132***J46001293***J46001462***J46001703***J46001704*********OUS: ***J46000150***J46000933***J46000162***J46001029***J46000190***J46001041***J46000247***J46001042***J46000296***J46001087***J46000393***J46001119***J46000410***J46001186***J46000478***J46001187***J46000499***J46001210***J46000531***J46001217***J46000564***J46001289***J46000617***J46001541***J46000772***J46001565***J46000788***J46001571***J46000800***J46001609***J46000801***J46001629***J46000812***J46001639***J46000821***J46001645***J46000855***J46001680***J46000859***J46000948***J46000861***J46000983***J46000910 Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, Inc. 100 Indigo Creek Dr Rochester NY 14626-5101 For Additional Information Contact Mr. Joe Falvo 585-453-3000 Manufacturer Reason for Recall A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. FDA Determined Cause 2 Software Design Change Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. Quantity in Commerce 15 US; 43 OUS Distribution Worldwide distribution - US Nationwide and the countries of Australia, Belgium, Brazil, ***Canada, Chile, China, Colombia, Denmark, France, Germany, India, Italy, Japan, Mexico,***Norway, Portugal, Russia, Singapore, Spain, Sweden, The Netherlands, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange
15 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205924 | Z-1228-2024 | VITROS XT 7600 Integrated System, Catalog No. 6844461 | VITROS XT 7600 Integrated System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:07 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. | Class 2 Device Recall VITROS XT 7600 Integrated System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS XT 7600 Integrated System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS XT 7600 Integrated System Date Initiated by Firm January 26, 2024 Date Posted March 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1228-2024 Recall Event ID 93925 510(K)Number K182063 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product VITROS XT 7600 Integrated System, Catalog No. 6844461 Code Information VITROS System Software Version 3.8.1***UDI-DI 10758750031610******Serial Numbers ***US: ***J76000199***J76001383***J76000207***J76001385***J76000353***J76001392***J76000456***J76001404***J76000459***J76001406***J76000460***J76001516***J76000468***J76001518***J76000495***J76001519***J76000502***J76001520***J76000561***J76001581***J76001207***J76001583***J76001230***J76001611***J76001231***J76001620***J76001232***J76001644***J76001235***J76001647***J76001239***J76001660***J76001260***J76001668***J76001263***J76001671***J76001270***J76001674***J76001274***J76001692***J76001331***J76001700***J76001351***J76001748************OUS:***J76000118***J76001070***J76000253***J76001074***J76000265***J76001095***J76000274***J76001121***J76000281***J76001130***J76000370***J76001152***J76000447***J76001174***J76000475***J76001328***J76000620***J76001377***J76000855***J76001378***J76000856***J76001396***J76000898***J76001399***J76000901***J76001407***J76000973***J76001408***J76000976***J76001484***J76001001***J76001673***J76001027***J76001761***J76001035 Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, Inc. 100 Indigo Creek Dr Rochester NY 14626-5101 For Additional Information Contact Mr. Joe Falvo 585-453-3000 Manufacturer Reason for Recall A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. FDA Determined Cause 2 Software Design Change Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. Quantity in Commerce 44 US; 35 OUS Distribution Worldwide - US Nationwide distribution in the states of AZ, CA, CT, FL, IA, ID, IL, IN, KY, ***LA, MO, NE, NY, OH, OR, SC, SD, TX, WA, WI, WV and the countries of Australia, Brazil, Canada, China, Colombia, India, Mexico, Singapore, United Kingdom, Italy, Spain, Russia, France, and Belgium. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange
16 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205930 | Z-1232-2024 | VITROS XT 3400 Chemistry System, Catalog No. 6844458 | VITROS XT 3400 Chemistry System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:06 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. | Class 2 Device Recall VITROS XT 3400 Chemistry System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS XT 3400 Chemistry System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS XT 3400 Chemistry System Date Initiated by Firm January 26, 2024 Date Posted March 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1232-2024 Recall Event ID 93927 510(K)Number K182063 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product VITROS XT 3400 Chemistry System, Catalog No. 6844458 Code Information VITROS System Software Versions 3.8.0 or 3.8.1***UDI-DI 10758750031986******Serial Numbers:******US: ***J34500140***J34500180***J34500257***J34500289***J34500319***J34500386***J34500422***J34500447***J34500473***J34500474***J34500510***J34500539***J34500547***J34500579***J34500588***J34500622***J34500656***J34500777***J34500818***J34500861***J34500887***J34500893***J34500896***J34500911***J34500931******OUS: ***J34500123***J34500128***J34500129***J34500130***J34500190***J34500261***J34500269***J34500271***J34500292***J34500301***J34500337***J34500356***J34500363***J34500515***J34500520***J34500536***J34500582***J34500609***J34500628***J34500658***J34500711***J34500793***J34500802***J34500940***J34500960***J34500962 Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, Inc. 100 Indigo Creek Dr Rochester NY 14626-5101 For Additional Information Contact Mr. Joe Falvo 585-453-3000 Manufacturer Reason for Recall A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. FDA Determined Cause 2 Software Design Change Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. Quantity in Commerce 25 US; 26 OUS Distribution Worldwide distribution - US Nationwide and the countries of Australia, Belgium, Brazil, ***Canada, Chile, China, Colombia, Denmark, France, Germany, India, Italy, Japan, Mexico,***Norway, Portugal, Russia, Singapore, Spain, Sweden, The Netherlands, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange
17 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205931 | Z-1233-2024 | VITROS XT 7600 Integrated System, Catalog No. 6844461 | VITROS XT 7600 Integrated System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:06 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. | Class 2 Device Recall VITROS XT 7600 Integrated System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS XT 7600 Integrated System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS XT 7600 Integrated System Date Initiated by Firm January 26, 2024 Date Posted March 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1233-2024 Recall Event ID 93927 510(K)Number K182063 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product VITROS XT 7600 Integrated System, Catalog No. 6844461 Code Information VITROS System Software Versions 3.8.0 or 3.8.1***UDI-DI 10758750031610******Serial Numbers:******US: ***J76000050***J76000878***J76001555***J76000108***J76000905***J76001558***J76000144***J76000909***J76001559***J76000148***J76000922***J76001565***J76000179***J76000930***J76001570***J76000192***J76000968***J76001577***J76000195***J76001106***J76001581***J76000199***J76001171***J76001583***J76000200***J76001172***J76001589***J76000206***J76001206***J76001592***J76000207***J76001207***J76001595***J76000208***J76001218***J76001596***J76000241***J76001230***J76001597***J76000298***J76001231***J76001599***J76000313***J76001232***J76001600***J76000317***J76001234***J76001611***J76000324***J76001235***J76001616***J76000333***J76001239***J76001617***J76000342***J76001260***J76001620***J76000352***J76001263***J76001638***J76000353***J76001270***J76001639***J76000420***J76001274***J76001643***J76000422***J76001292***J76001644***J76000456***J76001293***J76001647***J76000459***J76001294***J76001648***J76000460***J76001331***J76001650***J76000468***J76001340***J76001652***J76000495***J76001351***J76001659***J76000502***J76001357***J76001660***J76000524***J76001370***J76001666***J76000525***J76001376***J76001667***J76000537***J76001383***J76001668***J76000561***J76001384***J76001669***J76000609***J76001385***J76001671***J76000709***J76001392***J76001674***J76000743***J76001404***J76001675***J76000747***J76001406***J76001678***J76000783***J76001516***J76001679***J76000790***J76001517***J76001687***J76000803***J76001518***J76001692***J76000805***J76001519***J76001700***J76000806***J76001520***J76001703***J76000817***J76001535***J76001705***J76000862***J76001540***J76001711***J76000866***J76001549***J76001714***J76000868***J76001552***J76001748***J76000871***J76001553***J76001879***J76000873************OUS:******J76000106***J76000976***J76000118***J76000983***J76000211***J76001001***J76000253***J76001027***J76000257***J76001035***J76000265***J76001070***J76000274***J76001074***J76000276***J76001095***J76000278***J76001111***J76000280***J76001121***J76000281***J76001130***J76000284***J76001141***J76000288***J76001142***J76000289***J76001144***J76000290***J76001152***J76000292***J76001155***J76000293***J76001159***J76000370***J76001174***J76000436***J76001183***J76000446***J76001189***J76000447***J76001272***J76000475***J76001273***J76000620***J76001285***J76000628***J76001309***J76000760***J76001314***J76000782***J76001328***J76000807***J76001368***J76000813***J76001377***J76000838***J76001378***J76000845***J76001379***J76000855***J76001396***J76000856***J76001399***J76000888***J76001407***J76000898***J76001408***J76000901***J76001452***J76000973***J76001457******J76001484******J76001547******J76001562******J76001563******J76001673******J76001761 Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, Inc. 100 Indigo Creek Dr Rochester NY 14626-5101 For Additional Information Contact Mr. Joe Falvo 585-453-3000 Manufacturer Reason for Recall A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. FDA Determined Cause 2 Software Design Change Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. Quantity in Commerce 142 US; 78 OUS Distribution Worldwide distribution - US Nationwide and the countries of Australia, Belgium, Brazil, ***Canada, Chile, China, Colombia, Denmark, France, Germany, India, Italy, Japan, Mexico,***Norway, Portugal, Russia, Singapore, Spain, Sweden, The Netherlands, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange
18 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205912 | Z-1224-2024 | VITROS 3600 Immunodiagnostic System Catalog Nos. 6802783 (regular) and 6802914 (refurbished) | VITROS 3600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:07 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. | Class 2 Device Recall VITROS 3600 Immunodiagnostic System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS 3600 Immunodiagnostic System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS 3600 Immunodiagnostic System Date Initiated by Firm January 26, 2024 Date Posted March 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1224-2024 Recall Event ID 93925 510(K)Number K083173 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product VITROS 3600 Immunodiagnostic System Catalog Nos. 6802783 (regular) and 6802914 (refurbished) Code Information VITROS System Software Version 3.8.1***Regular UDI-DI 10758750002979***Refurbished UD-DI 10758750007103******Serial Numbers ***US: J36001167***OUS: J36000427***J36001117***J36001558***J36000580***J36001154***J36001570***J36000626***J36001229***J36001573***J36000688***J36001427***J36001575***J36000877***J36001451***J36001579***J36000920***J36001488***J36001582***J36000929***J36001489***J36001583***J36000942***J36001495***J36001595***J36001012***J36001519***J36001602***J36001051***J36001528***J36001608***J36001108***J36001530***J36001632***J36001111***J36001535***J36001640 Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, Inc. 100 Indigo Creek Dr Rochester NY 14626-5101 For Additional Information Contact Mr. Joe Falvo 585-453-3000 Manufacturer Reason for Recall A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. FDA Determined Cause 2 Software Design Change Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. Quantity in Commerce 1 US; 36 OUS Distribution Worldwide - US Nationwide distribution in the states of AZ, CA, CT, FL, IA, ID, IL, IN, KY, ***LA, MO, NE, NY, OH, OR, SC, SD, TX, WA, WI, WV and the countries of Australia, Brazil, Canada, China, Colombia, India, Mexico, Singapore, United Kingdom, Italy, Spain, Russia, France, and Belgium. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange
19 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205927 | Z-1229-2024 | VITROS 3600 Immunodiagnostic System Catalog Nos. 6802783 (regular) and 6802914 (refurbished) | VITROS 3600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:06 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. | Class 2 Device Recall VITROS 3600 Immunodiagnostic System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS 3600 Immunodiagnostic System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS 3600 Immunodiagnostic System Date Initiated by Firm January 26, 2024 Date Posted March 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1229-2024 Recall Event ID 93927 510(K)Number K083173 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product VITROS 3600 Immunodiagnostic System Catalog Nos. 6802783 (regular) and 6802914 (refurbished) Code Information VITROS System Software Versions 3.8.0 or 3.8.1***Regular UDI-DI 10758750002979***Refurbished UD-DI 10758750007103******Serial Numbers:******US: J36000109***J36000292***J36000489***J36000955***J36001167***J36001222***J36001283***J36001330***J36001348******OUS: ***J36000148***J36001000***J36001495***J36000340***J36001012***J36001513***J36000348***J36001024***J36001515***J36000406***J36001026***J36001519***J36000407***J36001027***J36001528***J36000427***J36001051***J36001530***J36000446***J36001088***J36001535***J36000511***J36001108***J36001558***J36000517***J36001111***J36001570***J36000580***J36001112***J36001573***J36000612***J36001117***J36001575***J36000626***J36001154***J36001579***J36000688***J36001229***J36001581***J36000753***J36001237***J36001582***J36000861***J36001245***J36001583***J36000877***J36001246***J36001595***J36000920***J36001401***J36001600***J36000923***J36001403***J36001602***J36000929***J36001427***J36001608***J36000933***J36001451***J36001618***J36000941***J36001488***J36001632***J36000942***J36001489***J36001640 Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, Inc. 100 Indigo Creek Dr Rochester NY 14626-5101 For Additional Information Contact Mr. Joe Falvo 585-453-3000 Manufacturer Reason for Recall A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. FDA Determined Cause 2 Software Design Change Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. Quantity in Commerce 9 US; 66 OUS Distribution Worldwide distribution - US Nationwide and the countries of Australia, Belgium, Brazil, ***Canada, Chile, China, Colombia, Denmark, France, Germany, India, Italy, Japan, Mexico,***Norway, Portugal, Russia, Singapore, Spain, Sweden, The Netherlands, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange
20 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205923 | Z-1227-2024 | VITROS XT 3400 Chemistry System, Catalog No. 6844458 | VITROS XT 3400 Chemistry System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:07 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. | Class 2 Device Recall VITROS XT 3400 Chemistry System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS XT 3400 Chemistry System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS XT 3400 Chemistry System Date Initiated by Firm January 26, 2024 Date Posted March 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1227-2024 Recall Event ID 93925 510(K)Number K182063 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product VITROS XT 3400 Chemistry System, Catalog No. 6844458 Code Information VITROS System Software Version 3.8.1***UDI-DI 10758750031986******Serial Numbers ***US: ***J34500180***J34500289***J34500422***J34500473***J34500622***J34500861***J34500887*********OUS:***J34500515***J34500609***J34500802***J34500940 Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, Inc. 100 Indigo Creek Dr Rochester NY 14626-5101 For Additional Information Contact Mr. Joe Falvo 585-453-3000 Manufacturer Reason for Recall A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. FDA Determined Cause 2 Software Design Change Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. Quantity in Commerce 7 US; 4 OUS Distribution Worldwide - US Nationwide distribution in the states of AZ, CA, CT, FL, IA, ID, IL, IN, KY, ***LA, MO, NE, NY, OH, OR, SC, SD, TX, WA, WI, WV and the countries of Australia, Brazil, Canada, China, Colombia, India, Mexico, Singapore, United Kingdom, Italy, Spain, Russia, France, and Belgium. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange
21 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205921 | Z-1225-2024 | VITROS 4600 Chemistry System Catalog Nos. 6802445 (regular) and 6900440 (refurbished) | VITROS 4600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:07 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. | Class 2 Device Recall VITROS 4600 Immunodiagnostic System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS 4600 Immunodiagnostic System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS 4600 Immunodiagnostic System Date Initiated by Firm January 26, 2024 Date Posted March 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1225-2024 Recall Event ID 93925 510(K)Number K031924 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product VITROS 4600 Chemistry System Catalog Nos. 6802445 (regular) and 6900440 (refurbished) Code Information VITROS System Software Version 3.8.1***Regular UDI-DI 10758750012343***Refurbished UD-DI 10758750033201******Serial Numbers ***US: J46000240***J46000898***J46001030***J46001293******OUS: J46000150***J46001041***J46000190***J46001042***J46000499***J46001186***J46000564***J46001187***J46000801***J46001210***J46000821***J46001217***J46000910***J46001289***J46000933***J46001639***J46001029***J46001645 Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, Inc. 100 Indigo Creek Dr Rochester NY 14626-5101 For Additional Information Contact Mr. Joe Falvo 585-453-3000 Manufacturer Reason for Recall A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. FDA Determined Cause 2 Software Design Change Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. Quantity in Commerce 4 US; 18 OUS Distribution Worldwide - US Nationwide distribution in the states of AZ, CA, CT, FL, IA, ID, IL, IN, KY, ***LA, MO, NE, NY, OH, OR, SC, SD, TX, WA, WI, WV and the countries of Australia, Brazil, Canada, China, Colombia, India, Mexico, Singapore, United Kingdom, Italy, Spain, Russia, France, and Belgium. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange
22 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205929 | Z-1231-2024 | VITROS 5600 Integrated System Catalog Nos. 6802413 (regular) and 6802915 (refurbished) | VITROS 5600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:06 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. | Class 2 Device Recall VITROS 5600 Immunodiagnostic System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS 5600 Immunodiagnostic System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS 5600 Immunodiagnostic System Date Initiated by Firm January 26, 2024 Date Posted March 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1231-2024 Recall Event ID 93927 510(K)Number K081543 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product VITROS 5600 Integrated System Catalog Nos. 6802413 (regular) and 6802915 (refurbished) Code Information VITROS System Software Versions 3.8.0 or 3.8.1***Regular UDI-DI 10758750002740***Refurbished UD-DI 10758750007110******Serial Numbers:******US: ***J56000116***J56003369***J56000317***J56003373***J56001292***J56003377***J56001523***J56003433***J56001629***J56003477***J56001885***J56003478***J56001886***J56003484***J56001910***J56003496***J56002041***J56003541***J56002159***J56003570***J56002198***J56003578***J56002199***J56003581***J56002201***J56003583***J56002249***J56003632***J56002301***J56003646***J56002308***J56003655***J56002461***J56003668***J56002463***J56003669***J56002466***J56003670***J56002474***J56003672***J56002477***J56003725***J56002500***J56003730***J56002591***J56003752***J56002603***J56003759***J56002607***J56003774***J56002608***J56003911***J56002615***J56004068***J56002698***J56004137***J56002703***J56004174***J56002704***J56004240***J56002722***J56004325***J56002769***J56004326***J56002895***J56004485***J56003035***J56004491***J56003045***J56004498***J56003076***J56004506***J56003113***J56004512***J56003199***J56004549***J56003220***J56004551***J56003287***J56004571***J56003299***J56004577***J56003317***J56004587***J56003330***J56004589***J56003334***J56004604***J56003342***J56004606***J56003356***************OUS: ***J56000190***J56001929***J56000246***J56001944***J56000418***J56001979***J56000425***J56002028***J56000468***J56002064***J56000512***J56002085***J56000533***J56002099***J56000665***J56002116***J56000671***J56002218***J56000674***J56002258***J56000680***J56002300***J56000719***J56002410***J56000809***J56002460***J56000831***J56002518***J56000836***J56002615***J56000910***J56002625***J56000929***J56002637***J56000993***J56002784***J56000994***J56002821***J56001015***J56002862***J56001051***J56002913***J56001097***J56002926***J56001111***J56002929***J56001114***J56002966***J56001140***J56003061***J56001208***J56003162***J56001210***J56003179***J56001237***J56003260***J56001260***J56003285***J56001320***J56003401***J56001355***J56003451***J56001378***J56003464***J56001389***J56003508***J56001469***J56003523***J56001509***J56003538***J56001591***J56003919***J56001631***J56003951***J56001689***J56004078***J56001731***J56004082***J56001733***J56004308***J56001765***J56004309***J56001780***J56004346***J56001793***J56004355***J56001826***J56004414***J56001844***J56004415***J56001853***J56004531***J56001925***J56004695 Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, Inc. 100 Indigo Creek Dr Rochester NY 14626-5101 For Additional Information Contact Mr. Joe Falvo 585-453-3000 Manufacturer Reason for Recall A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. FDA Determined Cause 2 Software Design Change Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. Quantity in Commerce 91 US; 94 OUS Distribution Worldwide distribution - US Nationwide and the countries of Australia, Belgium, Brazil, ***Canada, Chile, China, Colombia, Denmark, France, Germany, India, Italy, Japan, Mexico,***Norway, Portugal, Russia, Singapore, Spain, Sweden, The Netherlands, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange
23 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205922 | Z-1226-2024 | VITROS 5600 Integrated System Catalog Nos. 6802413 (regular) and 6802915 (refurbished) | VITROS 5600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:07 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. | Class 2 Device Recall VITROS 5600 Immunodiagnostic System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITROS 5600 Immunodiagnostic System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITROS 5600 Immunodiagnostic System Date Initiated by Firm January 26, 2024 Date Posted March 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-1226-2024 Recall Event ID 93925 510(K)Number K081543 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product VITROS 5600 Integrated System Catalog Nos. 6802413 (regular) and 6802915 (refurbished) Code Information VITROS System Software Version 3.8.1***Regular UDI-DI 10758750002740***Refurbished UD-DI 10758750007110******Serial Numbers ***US: J56001885***J56003484***J56001886***J56003541***J56001910***J56003570***J56002308***J56003583***J56002461***J56003646***J56002474***J56003669***J56002608***J56003670***J56002698***J56004068***J56002704***J56004137***J56003342***J56004174***J56003356***J56004491***J56003377***J56004506***J56003433***J56004549***J56003477***J56004551***J56003478***J56004577*********OUS: J56000425***J56002518***J56000512***J56002615***J56000533***J56002625***J56000674***J56002637***J56000680***J56002926***J56000719***J56002929***J56000809***J56003061***J56000831***J56003260***J56000836***J56003285***J56001051***J56003401***J56001111***J56003464***J56001114***J56003508***J56001208***J56003523***J56001731***J56003538***J56001765***J56004346***J56001826***J56004355***J56001853***J56004415***J56002085***J56004531***J56002116***J56004695 Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics, Inc. 100 Indigo Creek Dr Rochester NY 14626-5101 For Additional Information Contact Mr. Joe Falvo 585-453-3000 Manufacturer Reason for Recall A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. FDA Determined Cause 2 Software Design Change Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. Quantity in Commerce 30 US: 38 OUS Distribution Worldwide - US Nationwide distribution in the states of AZ, CA, CT, FL, IA, ID, IL, IN, KY, ***LA, MO, NE, NY, OH, OR, SC, SD, TX, WA, WI, WV and the countries of Australia, Brazil, Canada, China, Colombia, India, Mexico, Singapore, United Kingdom, Italy, Spain, Russia, France, and Belgium. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange
24 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205373 | Z-0864-2024 | smiths medical Medfusion Model 4000 syringe pump | smiths medical Medfusion | 1 | CDRH | 2024-02-01 14:28 | 2024-02-01 20:06 |  | Smiths Medical ASD Inc. | Medfusion Model 4000 syringe infusion pumps, with the following software versions, may have the following issues: 1. Delivery During Motor Not Running High Priority Alarm (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 2. Re-administered Loading Dose (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 3. Incorrect Critical Data Failure Alarm, 4. Interruption of Bolus or Loading Dose Delivery (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 5. Incorrect Total Bolus/Loading Dose Displayed (v1.1.0, v1.1.1, v1.1.2), 6. Volume Limit Before Bolus/Loading Dose Complete (v1.1.2), 7. Drug Library Lower Limit Displayed Incorrectly (v1.6.0, v1.6.1), 8. Depleted Battery Alarm (All versions previous to v1.6.5), 9. Loss of Wireless Connectivity (v1.5.0, v1.5.1, v1.6.0, v1.6.1, v1.6.4), 10. PharmGuard Server Password (v2.3, v2.4, v2.5). ***Smiths Medical corrected all issues included in this notification in previous software updates and the corrections were carried forward into all subsequent software releases.***Please ensure you have the most recent Medfusion software installed on your pumps.*** | Class 1 Device Recall smiths medical Medfusion Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall smiths medical Medfusion FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall smiths medical Medfusion Date Initiated by Firm December 19, 2023 Date Posted February 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-0864-2024 Recall Event ID 93757 510(K)Number K111386 Product Classification Pump, infusion - Product Code FRN Product smiths medical Medfusion Model 4000 syringe pump Code Information All Medfusion Model 4000 with software versions: v1.0.0, v1.1.0, v1.1.1, v1.1.2, v1.5.0, v1.5.1, 1.6.0, v1.6.1, v1.6.4, v2.3, v2.4, v2.5, and all versions prior to v1.6.5. Recalling Firm/ Manufacturer Smiths Medical ASD Inc. 6000 Nathan Ln N Minneapolis MN 55442-1690 For Additional Information Contact icumed.custhelp.com/app/market-action 763-383-3072 Manufacturer Reason for Recall Medfusion Model 4000 syringe infusion pumps, with the following software versions, may have the following issues: 1. Delivery During Motor Not Running High Priority Alarm (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 2. Re-administered Loading Dose (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 3. Incorrect Critical Data Failure Alarm, 4. Interruption of Bolus or Loading Dose Delivery (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 5. Incorrect Total Bolus/Loading Dose Displayed (v1.1.0, v1.1.1, v1.1.2), 6. Volume Limit Before Bolus/Loading Dose Complete (v1.1.2), 7. Drug Library Lower Limit Displayed Incorrectly (v1.6.0, v1.6.1), 8. Depleted Battery Alarm (All versions previous to v1.6.5), 9. Loss of Wireless Connectivity (v1.5.0, v1.5.1, v1.6.0, v1.6.1, v1.6.4), 10. PharmGuard Server Password (v2.3, v2.4, v2.5). ***Smiths Medical corrected all issues included in this notification in previous software updates and the corrections were carried forward into all subsequent software releases.***Please ensure you have the most recent Medfusion software installed on your pumps. FDA Determined Cause 2 Software Design Change Action Smiths Medical issued an URGENT MEDICAL DEVICE CORRRECTION notice to its consignees on 12/19/2023 by letter delivered via traceable means. This notification details the issues and the affected software versions. Please note that the pump displays the installed software version on the startup screen after the pump is powered on. The notice explained the issues with the pumps, the potential risks involved with their use, and requested the following actions be taken: ***1. Locate all affected pumps in your possession and ensure all users or potential users***of these devices are immediately made aware of this notification and proposed***mitigations. Please ensure you have the most recent Medfusion software installed on***your pumps. ***2. Complete and return the attached Response Form to smithsmedical8045@sedgwick.com within ten days of receipt to acknowledge your understanding of this notification. ***3. DISTRIBUTORS: If you have distributed potentially affected products to your customers, please immediately forward this notice to them. Request that they complete the response form and return it to smithsmedical8045@sedgwick.com. *** ***To report adverse events or product complaints: Global Complaint Management, globalcomplaints@icumed.com, 1-(866)-216-8806******Additional information or technical assistance: Technical Support, icumed.custhelp.com/app/market-action******Questions about this Field Correction Notice: Field Corrections, icumed.custhelp.com/app/market-action Quantity in Commerce 60146 pumps Distribution US nationwide, Canada, Great Britain, Bermuda, Australia, Germany, Singapore, Saudi Arabia, Lebanon, and Malaysia Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = FRN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Smiths Medical issued an URGENT MEDICAL DEVICE CORRRECTION notice to its consignees on 12/19/2023 by letter delivered via traceable means. This notification details the issues and the affected software versions. Please note that the pump displays the installed software version on the startup screen after the pump is powered on. The notice explained the issues with the pumps, the potential risks involved with their use, and requested the following actions be taken: ***1. Locate all affected pumps in your possession and ensure all users or potential users***of these devices are immediately made aware of this notification and proposed***mitigations. Please ensure you have the most recent Medfusion software installed on***your pumps. ***2. Complete and return the attached Response Form to smithsmedical8045@sedgwick.com within ten days of receipt to acknowledge your understanding of this notification. ***3. DISTRIBUTORS: If you have distributed potentially affected products to your customers, please immediately forward this notice to them. Request that they complete the response form and return it to smithsmedical8045@sedgwick.com. *** ***To report adverse events or product complaints: Global Complaint Management, globalcomplaints@icumed.com, 1-(866)-216-8806******Additional information or technical assistance: Technical Support, icumed.custhelp.com/app/market-action******Questions about this Field Correction Notice: Field Corrections, icumed.custhelp.com/app/market-action | SoftwareDesignChange
25 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205366 | Z-0863-2024 | smiths medical Medfusion Model 3500 Syringe pump | smiths medical Medfusion | 1 | CDRH | 2024-02-01 13:44 | 2024-02-01 20:07 |  | Smiths Medical ASD Inc. | Medfusion syringe pumps, with software versions before v6.0.0, may have the following issues:1. Delivery During Motor Not Running High Priority Alarm, 2. Infusion Restarted with Incorrect Parameters, 3. Screen Lock, 4. Interruption of Bolus or Loading Dose Delivery, 5. Pump Displays Incorrect Bolus/Loading Dose, 6. Loading/Bolus Dose Below the Minimum Recommended Rate, 7. Motor Rate Error, 8. Incorrect Recall Last Settings, 9. Corrupt Configuration, 10. Auto Lock, 11. Toolbox Configuration Loading Dose Time Values.******Smiths Medical corrected all issues included in this notification in previous software updates and the corrections were carried forward into all subsequent software releases. Please ensure you have the most recent Medfusion software (v6.0.0) installed on your pumps. | Class 1 Device Recall smiths medical Medfusion Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall smiths medical Medfusion FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall smiths medical Medfusion Date Initiated by Firm December 19, 2023 Date Posted February 01, 2024 Recall Status 1 Open 3 , Classified Recall Number Z-0863-2024 Recall Event ID 93754 510(K)Number K040899 Product Classification Pump, infusion - Product Code FRN Product smiths medical Medfusion Model 3500 Syringe pump Code Information All Medfusion Model 3500 pumps with software v3.X, v4.X, and v5.0.0 and was***addressed in software v6.0.0 (2012) and all subsequent versions. Recalling Firm/ Manufacturer Smiths Medical ASD Inc. 6000 Nathan Ln N Minneapolis MN 55442-1690 For Additional Information Contact Field Corrections 763-383-3072 Manufacturer Reason for Recall Medfusion syringe pumps, with software versions before v6.0.0, may have the following issues:1. Delivery During Motor Not Running High Priority Alarm, 2. Infusion Restarted with Incorrect Parameters, 3. Screen Lock, 4. Interruption of Bolus or Loading Dose Delivery, 5. Pump Displays Incorrect Bolus/Loading Dose, 6. Loading/Bolus Dose Below the Minimum Recommended Rate, 7. Motor Rate Error, 8. Incorrect Recall Last Settings, 9. Corrupt Configuration, 10. Auto Lock, 11. Toolbox Configuration Loading Dose Time Values.************Smiths Medical corrected all issues included in this notification in previous software updates and the corrections were carried forward into all subsequent software releases. Please ensure you have the most recent Medfusion software (v6.0.0) installed on your pumps. FDA Determined Cause 2 Software Design Change Action Smiths Medical issued an URGENT MEDICAL DEVICE CORRRECTION notice to its consignees on 12/16/2023 by letter delivered via traceable means. The notice explained the issues with the pumps, the potential risks involved with their use, and requested the following actions be taken: ***1. Locate all affected pumps in your possession and ensure all users or potential users***of these devices are immediately made aware of this notification and proposed***mitigations. Please ensure you have the most recent Medfusion software installed on***your pumps.***2. Complete and return the attached Response Form to smithsmedical8045@sedgwick.com***within ten days of receipt to acknowledge your understanding of this notification.***3. DISTRIBUTORS: If you have distributed potentially affected products to your customers,***please immediately forward this notice to them. Request that they complete the response form and return it to smithsmedical8045@sedgwick.com. ***Inquires: ***To report adverse events or product complaints: Global Complaint Management, globalcomplaints@icumed.com, 1-(866)-216-8806******Additional information or technical assistance: Technical Support, icumed.custhelp.com/app/market-action******Questions about this Field Correction Notice: Field Corrections, icumed.custhelp.com/app/market-action Quantity in Commerce 93177 pumps Distribution worldwide Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = FRN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Smiths Medical issued an URGENT MEDICAL DEVICE CORRRECTION notice to its consignees on 12/16/2023 by letter delivered via traceable means. The notice explained the issues with the pumps, the potential risks involved with their use, and requested the following actions be taken: ***1. Locate all affected pumps in your possession and ensure all users or potential users***of these devices are immediately made aware of this notification and proposed***mitigations. Please ensure you have the most recent Medfusion software installed on***your pumps.***2. Complete and return the attached Response Form to smithsmedical8045@sedgwick.com***within ten days of receipt to acknowledge your understanding of this notification.***3. DISTRIBUTORS: If you have distributed potentially affected products to your customers,***please immediately forward this notice to them. Request that they complete the response form and return it to smithsmedical8045@sedgwick.com. ***Inquires: ***To report adverse events or product complaints: Global Complaint Management, globalcomplaints@icumed.com, 1-(866)-216-8806******Additional information or technical assistance: Technical Support, icumed.custhelp.com/app/market-action******Questions about this Field Correction Notice: Field Corrections, icumed.custhelp.com/app/market-action | SoftwareDesignChange
26 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204312 | Z-0560-2024 | Roche Cobas 8000 and Cobas PRO Interface Module (CO8 IM)- IVD FlexLab Automation System -a modular system designed to automate Pre-Analytical and Post-Analytical processing, sample handling in order to automate sample processing in the Laboratory.******Hardware versions: FLX-247-00***Firmware versions: COBASa_2.2.0, COBASa_2.3.0 | Inpeco FlexLab | 2 | CDRH | 2023-12-15 | 2023-12-15 20:07 |  | Inpeco S.A. | Firmware of the Interface Module with Roche Cobas 8000 and Cobas PRO (CO8 IM) may lead to a delayed sample tube processing. | Class 2 Device Recall Inpeco FlexLab Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Inpeco FlexLab FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Inpeco FlexLab Date Initiated by Firm November 03, 2023 Date Posted December 15, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-0560-2024 Recall Event ID 93423 510(K)Number K121012 Product Classification Electrode, ion specific, potassium - Product Code CEM Product Roche Cobas 8000 and Cobas PRO Interface Module (CO8 IM)- IVD FlexLab Automation System -a modular system designed to automate Pre-Analytical and Post-Analytical processing, sample handling in order to automate sample processing in the Laboratory.************Hardware versions: FLX-247-00******Firmware versions: COBASa_2.2.0, COBASa_2.3.0 Code Information FlexLab UDI-DI:07640172340004***Hardware versions: FLX-247-00***Firmware versions: COBASa_2.2.0, COBASa_2.3.0***Serial Number US: ***FLX.0208******Serial Numbers OUS:***Serial Number***FLX.0116***FLX.0165***FLX.0178***FLX.067***FLX.0127***FLX.0128***FLX.0150***FLX.0154***FLX.012***FLX.0121***FLX.0132***FLX.031***FLX.048***FLX.049 Recalling Firm/ Manufacturer Inpeco S.A. Via San Gottardo 10 Lugano Switzerland Manufacturer Reason for Recall Firmware of the Interface Module with Roche Cobas 8000 and Cobas PRO (CO8 IM) may lead to a delayed sample tube processing. FDA Determined Cause 2 Software Design Change Action Inpeco issued Urgent Medical Device Correction Letter (Recall-identifier: 3010825766 11/06/23 002 C) via email on 11/3/23. Letter states reason for recall, health risk and action to take:***Please apply the following instructions to detect the sample tubes potentially impacted by the issue (Step 1) and to manage them according to the specific occurred scenario (Step 2). If Step 1 leads to determine that the issue has not occurred, Step 2 is not to be performed.***Note: The test orders status displayed on the IUI is not overwritten neither changed in any way due to the issue: it would be possible to monitor the turn-around time of sample tubes processed by the CO8 IM in order to identify any overdue sample tube with tests still to be executed.***Contact your local technical support provider if you need assistance in implementing these actions. Your service provider will contact you to schedule the firmware upgrade.***Until the service visit please maintain awareness on this notice and apply the actions***recommended above.***Transmission of this Recall Letter:***Please transfer this notice to whom it might concern.***Please complete and return the Customer Letter Receipt Confirmation and Implementation Check form attached to this letter within 30 days directly to the email address specified in the email communication.***Contact reference person:***For any clarification you may need, do not hesitate to contact:***Eva Balzarotti - Regulatory Affairs Manager***E-mail: regulatory.affairs@inpeco.com***Phone: (+41) 91 9118 224 Quantity in Commerce 1 unit US: 14 units OUS Distribution Worldwide - US Nationwide distribution in the state of MA and the countries of AUSTRIA, BELGIUM, DENMARK, ITALY, NORWAY, SPAIN. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = CEM - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Inpeco issued Urgent Medical Device Correction Letter (Recall-identifier: 3010825766 11/06/23 002 C) via email on 11/3/23. Letter states reason for recall, health risk and action to take:***Please apply the following instructions to detect the sample tubes potentially impacted by the issue (Step 1) and to manage them according to the specific occurred scenario (Step 2). If Step 1 leads to determine that the issue has not occurred, Step 2 is not to be performed.***Note: The test orders status displayed on the IUI is not overwritten neither changed in any way due to the issue: it would be possible to monitor the turn-around time of sample tubes processed by the CO8 IM in order to identify any overdue sample tube with tests still to be executed.***Contact your local technical support provider if you need assistance in implementing these actions. Your service provider will contact you to schedule the firmware upgrade.***Until the service visit please maintain awareness on this notice and apply the actions***recommended above.***Transmission of this Recall Letter:***Please transfer this notice to whom it might concern.***Please complete and return the Customer Letter Receipt Confirmation and Implementation Check form attached to this letter within 30 days directly to the email address specified in the email communication.***Contact reference person:***For any clarification you may need, do not hesitate to contact:***Eva Balzarotti - Regulatory Affairs Manager***E-mail: regulatory.affairs@inpeco.com***Phone: (+41) 91 9118 224 | SoftwareDesignChange
27 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204258 | Z-0522-2024 | Medtronic SynchroMed, Model A10 | Medtronic | 2 | CDRH | 2023-12-09 | 2023-12-09 20:06 |  | Medtronic Inc. | The previous software application version (1.1.300) is missing a decimal separator (a comma) for parameter range guidance values displayed on some of the programming screens: Catheter, Reservoir, Infusion, Bolus, myPTM, and Alarm. | Class 2 Device Recall Medtronic Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Medtronic FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Medtronic Date Initiated by Firm July 13, 2020 Date Posted December 09, 2023 Recall Status 1 Completed Recall Number Z-0522-2024 Recall Event ID 93401 PMA Number P860004 Product Classification Pump, infusion, implanted, programmable - Product Code LKK Product Medtronic SynchroMed, Model A10 Code Information GTIN 00643169771031, All serial numbers with software version 1.1.300. Recalling Firm/ Manufacturer Medtronic Inc. 8200 Coral Sea St Ne Mounds View MN 55112-4391 For Additional Information Contact Medtronic Customer Service 800-962-9888 Manufacturer Reason for Recall The previous software application version (1.1.300) is missing a decimal separator (a comma) for parameter range guidance values displayed on some of the programming screens: Catheter, Reservoir, Infusion, Bolus, myPTM, and Alarm. FDA Determined Cause 2 Software Design Change Action Medtronic issued a SAFETY NOTIFICATION to its consignees on 07/13/2020 by letter. The notice explained the problem with the device and requested the software be upgraded to version 1.1.342. Quantity in Commerce 2543 units Distribution Foreign Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. PMA Database PMAs with Product Code = LKK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Medtronic issued a SAFETY NOTIFICATION to its consignees on 07/13/2020 by letter. The notice explained the problem with the device and requested the software be upgraded to version 1.1.342. | SoftwareDesignChange
28 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204260 | Z-0517-2024 | SagiPlan 2.2, CE0344, Rx Only , Brachyterapy Treatment Planning System | SagiPlan 2.2 | 2 | CDRH | 2023-12-08 15:55 | 2023-12-08 20:08 |  | Bebig Isotopentechnik Gmbh | Due to software malfunction, numerical values may be rounded resulting in inaccurate measurements. | Class 2 Device Recall SagiPlan 2.2 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SagiPlan 2.2 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SagiPlan 2.2 Date Initiated by Firm October 17, 2023 Date Posted December 08, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-0517-2024 Recall Event ID 93399 510(K)Number K141900 Product Classification System, planning, radiation therapy treatment - Product Code MUJ Product SagiPlan 2.2, CE0344, Rx Only , Brachyterapy Treatment Planning System Code Information SagiPlan, article # 1374-0600, version 2.2 Recalling Firm/ Manufacturer Bebig Isotopentechnik Gmbh Robert-Rossle-Str. 10 Berlin Germany Manufacturer Reason for Recall Due to software malfunction, numerical values may be rounded resulting in inaccurate measurements. FDA Determined Cause 2 Software Design Change Action On October 17, 2023, Eckert & Ziegler BEBIG GmbH issued a "Urgent Field Safety Notice" via email. Eckert & Ziegler asked consignees to take the following actions: ******1. If the Length Unit is set to millimeters, user must carefully review and, if necessary, adjust the Image Sequence Plane Separation value during import when working with slice thicknesses that have two decimals, as shown in Figure 1. Users must consider that this approach is not intrinsically safe, therefore solution 1 or 2 is highly recommended ******2. If you have used mm as length unit and may have used two decimal numbers in mm for slice thickness, carefully review the patient cases and act accordingly.******3. This notice needs to be passed on to all those who need to be aware within your organization or to any organization where the devices have been transferred to.***Please reply to this email by 20th of October 2023 confirming that you have received and understood this information and that you have forwarded it to the people which already have received the named products or will receive them. Quantity in Commerce 2 software licenses in US, 255 software licenses in OUS Distribution US: MD***OUS: Germany Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = MUJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On October 17, 2023, Eckert & Ziegler BEBIG GmbH issued a "Urgent Field Safety Notice" via email. Eckert & Ziegler asked consignees to take the following actions: ******1. If the Length Unit is set to millimeters, user must carefully review and, if necessary, adjust the Image Sequence Plane Separation value during import when working with slice thicknesses that have two decimals, as shown in Figure 1. Users must consider that this approach is not intrinsically safe, therefore solution 1 or 2 is highly recommended ******2. If you have used mm as length unit and may have used two decimal numbers in mm for slice thickness, carefully review the patient cases and act accordingly.******3. This notice needs to be passed on to all those who need to be aware within your organization or to any organization where the devices have been transferred to.***Please reply to this email by 20th of October 2023 confirming that you have received and understood this information and that you have forwarded it to the people which already have received the named products or will receive them. | SoftwareDesignChange
29 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203408 | Z-0127-2024 | TruDi Navigation System, Model: FG-2000-00, Catalog: ENS022B, when used with TruDi Curette, Models: TDC0005Z and TDC0005 | TruDi Navigation System | 2 | CDRH | 2023-10-20 13:19 | 2023-10-20 21:07 |  | Acclarent, Inc. | When using affected curette and software, there is a discrepancy between the actual curette tip location and the location displayed on navigation systems intended for use during surgical procedures of the Ear Nose and Throat (ENT) and ENT skull base surgery, which may cause delayed/prolonged surgery, cerebrospinal fluid leak, visual impairment, or skull base structural damage. | Class 2 Device Recall TruDi Navigation System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TruDi Navigation System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TruDi Navigation System Date Initiated by Firm September 21, 2023 Date Posted October 20, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-0127-2024 Recall Event ID 93087 510(K)Number K192397 Product Classification Ear, Nose, and Throat Stereotaxic Instrument - Product Code PGW Product TruDi Navigation System, Model: FG-2000-00, Catalog: ENS022B, when used with TruDi Curette, Models: TDC0005Z and TDC0005 Code Information TruDi Navigation System, UDI-DI: 10846835018639, software version: V2.3.1 Update (2.3.1.144 and 2.3.1.166) Recalling Firm/ Manufacturer Acclarent, Inc. 31 Technology Dr Suite 200 Irvine CA 92618-2346 For Additional Information Contact 949-784-9350 Manufacturer Reason for Recall When using affected curette and software, there is a discrepancy between the actual curette tip location and the location displayed on navigation systems intended for use during surgical procedures of the Ear Nose and Throat (ENT) and ENT skull base surgery, which may cause delayed/prolonged surgery, cerebrospinal fluid leak, visual impairment, or skull base structural damage. FDA Determined Cause 2 Software Design Change Action On 9/21/23, correction notices were mailed to customers who were asked to do the following:******Please cease using the affected curette device with the affected navigation system until the new version of navigation system software version 2.3.2 is installed.***2. Please do not return any of the curettes within your inventory.***3. Complete and return the business response form via email to OneMD-Field-Actions@its.jnj.com***4. Forward this notice to any personnel in your facility that needs to be informed.***5. If any of the affected products have been forwarded to another facility, contact that facility and provide a copy of this to the relevant personnel.***6. Post a copy of this notice in a visible area for awareness******If any adverse reactions or quality problems are experienced during the use of the affected navigation system with the curette, please report any issues to the firm via the product complaint phone number:1-877-775-2789, Option # 5. Quantity in Commerce 141 Distribution US: CA, IL, WI, MA, TX, CO, FL, ID, KS, OH, OR, NE, MO, AZ, HI, TN, NY, ME, SC, NC, CT, DC, MN, LA, OK, PA, AL, IA, MI, SD Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = PGW - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 9/21/23, correction notices were mailed to customers who were asked to do the following:******Please cease using the affected curette device with the affected navigation system until the new version of navigation system software version 2.3.2 is installed.***2. Please do not return any of the curettes within your inventory.***3. Complete and return the business response form via email to OneMD-Field-Actions@its.jnj.com***4. Forward this notice to any personnel in your facility that needs to be informed.***5. If any of the affected products have been forwarded to another facility, contact that facility and provide a copy of this to the relevant personnel.***6. Post a copy of this notice in a visible area for awareness******If any adverse reactions or quality problems are experienced during the use of the affected navigation system with the curette, please report any issues to the firm via the product complaint phone number:1-877-775-2789, Option # 5. | SoftwareDesignChange
30 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203459 | Z-0094-2024 | Sapphire Infusion Pumps, Models: Multi-Therapy (REF 15031-000-0028), Epidural (REF 15032-000-0027) and SapphirePlus (REF 15038-000-0001) | Sapphire MT Infusion Pump, Sapphire Epidural Infusion Pump, SapphirePlus | 1 | CDRH | 2023-10-19 20:34 | 2023-10-20 21:07 |  | EITAN MEDICAL LTD | Infusion Pumps with affected software revision may fail to detect air in line, which may lead to air embolism. | Class 1 Device Recall Sapphire MT Infusion Pump, Sapphire Epidural Infusion Pump, SapphirePlus Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall Sapphire MT Infusion Pump, Sapphire Epidural Infusion Pump, SapphirePlus FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall Sapphire MT Infusion Pump, Sapphire Epidural Infusion Pump, SapphirePlus Date Initiated by Firm September 11, 2023 Date Posted October 20, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-0094-2024 Recall Event ID 93105 510(K)Number K192860 Product Classification Pump, infusion - Product Code FRN Product Sapphire Infusion Pumps, Models: Multi-Therapy (REF 15031-000-0028), Epidural (REF 15032-000-0027) and SapphirePlus (REF 15038-000-0001) Code Information Software Revision: 16.10.1 or 16.10.2******REF/UDI-DI: 15031-000-0028/7290109150109, 15032-000-0027/7290109150147, 15038-000-0001/7290109150161 Recalling Firm/ Manufacturer EITAN MEDICAL LTD 29, Yad Kharutsim Netanya Israel For Additional Information Contact 972.73.2388888 Manufacturer Reason for Recall Infusion Pumps with affected software revision may fail to detect air in line, which may lead to air embolism. FDA Determined Cause 2 Software Design Change Action On 9/11/23, recall notices were emailed to customers informing them of the following:******The Affected Pumps may fail to identify air in line when the following conditions are met: ***1. The pump is operating on battery power (not connected to a power supply) ***2. Treatments are water-like solutions (TPN and other opaque solutions are not affected) ******Until the Affected Pumps in your facility have been updated with new software, users are advised to only operate the effected devices as follows: ***i. Connect to continuous AC power during treatment (i.e., connect to a power supply) and/or; ***ii. Use with air eliminating filters sets: ***- Available Eitan Medical sets with filters: AP206-01, AP240-01, AP 214-01, AP 204-01, AP223-01, or; ***- Connect a 3rd party air eliminating filter set extension to the end of the Sapphire set. ******Alternatively, use software version - Rev15.10. ******Complete and return the Customer Response form via email to compliance@eitanmedical.com ******To report adverse events or product complaints, contact Global Complaint Management, complaints@eitanmedical.com ******If additional information or technical assistance, contact Customer Service and Technical Assistance, customerservice@eitanmedical.com 1-877-541-9944 Quantity in Commerce 1383 Distribution US nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = FRN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 9/11/23, recall notices were emailed to customers informing them of the following:******The Affected Pumps may fail to identify air in line when the following conditions are met: ***1. The pump is operating on battery power (not connected to a power supply) ***2. Treatments are water-like solutions (TPN and other opaque solutions are not affected) ******Until the Affected Pumps in your facility have been updated with new software, users are advised to only operate the effected devices as follows: ***i. Connect to continuous AC power during treatment (i.e., connect to a power supply) and/or; ***ii. Use with air eliminating filters sets: ***- Available Eitan Medical sets with filters: AP206-01, AP240-01, AP 214-01, AP 204-01, AP223-01, or; ***- Connect a 3rd party air eliminating filter set extension to the end of the Sapphire set. ******Alternatively, use software version - Rev15.10. ******Complete and return the Customer Response form via email to compliance@eitanmedical.com ******To report adverse events or product complaints, contact Global Complaint Management, complaints@eitanmedical.com ******If additional information or technical assistance, contact Customer Service and Technical Assistance, customerservice@eitanmedical.com 1-877-541-9944 | SoftwareDesignChange
31 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203324 | Z-0041-2024 | Real Intelligence CORI RI.KNEE v2.0 TKA (Total Knee Arthroplasty) software version 2.0 (2.0.0.5)***Product Number: ROB10299*** | REAL INTELLIGENCE CORI Welcome Box USA | 2 | CDRH | 2023-10-05 | 2023-10-05 21:06 |  | Blue Belt Technologies, Inc | CORI software was missing a planning stage that appears in cases of significant knee deformation (outside of the range of 3 degrees valgus to 7 degrees varus) when the surgeon chooses to define the femoral rotational reference via the posterior condylar axis | Class 2 Device Recall REAL INTELLIGENCE CORI Welcome Box USA Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall REAL INTELLIGENCE CORI Welcome Box USA FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall REAL INTELLIGENCE CORI Welcome Box USA Date Initiated by Firm September 07, 2023 Date Posted October 05, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-0041-2024 Recall Event ID 93046 510(K)Number K231963 Product Classification Orthopedic stereotaxic instrument - Product Code OLO Product Real Intelligence CORI RI.KNEE v2.0 TKA (Total Knee Arthroplasty) software version 2.0 (2.0.0.5)******Product Number: ROB10299 Code Information UDI: (01) 00885556773338***Lot Numbers: 51141585, 51108475, 51108689 Recalling Firm/ Manufacturer Blue Belt Technologies, Inc 2875 Railroad St Pittsburgh PA 15201 Manufacturer Reason for Recall CORI software was missing a planning stage that appears in cases of significant knee deformation (outside of the range of 3 degrees valgus to 7 degrees varus) when the surgeon chooses to define the femoral rotational reference via the posterior condylar axis FDA Determined Cause 2 Software Design Change Action Smith & Nephew, Inc. (Blue Belt Technologies) issued Urgent Medical Device Correction Notice (C-2023-08) on September 7, 2023 via overnight mail. Letter states reason for recall, health risk and action to take:******Required Actions for user/hospital:***1. Complete Correction Notice Acknowledgement Form and submit to FieldActions@smith-nephew.com.***2. Prior to use of software, contact your sales representative to update settings.***3. Upon completion of settings update, you may use the CORI Surgical System until the software update is deployed.***4. A Smith & Nephew Robotics Representative will be deployed to your site to implement a new software version to correct the issue once it is available.***Email: FieldActions@smith-nephew.com******Actions for Smith & Nephew Representatives***1. Work with customer to update CORI settings to prevent incorrect configuration.***2. Complete the attached Correction Notice Acknowledgement Form.***3. Upon completion of software upgrade, submit CORI In-Field Service Record to Field Action Team. Quantity in Commerce 56 units Distribution US Nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = OLO - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Smith & Nephew, Inc. (Blue Belt Technologies) issued Urgent Medical Device Correction Notice (C-2023-08) on September 7, 2023 via overnight mail. Letter states reason for recall, health risk and action to take:******Required Actions for user/hospital:***1. Complete Correction Notice Acknowledgement Form and submit to FieldActions@smith-nephew.com.***2. Prior to use of software, contact your sales representative to update settings.***3. Upon completion of settings update, you may use the CORI Surgical System until the software update is deployed.***4. A Smith & Nephew Robotics Representative will be deployed to your site to implement a new software version to correct the issue once it is available.***Email: FieldActions@smith-nephew.com******Actions for Smith & Nephew Representatives***1. Work with customer to update CORI settings to prevent incorrect configuration.***2. Complete the attached Correction Notice Acknowledgement Form.***3. Upon completion of software upgrade, submit CORI In-Field Service Record to Field Action Team. | SoftwareDesignChange
32 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=202870 | Z-2618-2023 | Omnipod 5 Automated Insulin Delivery System | Omnipod 5 Automated Insulin Delivery System | 2 | CDRH | 2023-09-20 | 2023-09-20 21:06 |  | Insulet Corporation | An error message was received when using the Omnipod 5 App on compatible smartphones that prevented phone control users from controlling the Omnipod 5 Automated Insulin Delivery System from their Omnipod 5 App which may cause a delay in therapy. To reduce the volume of inquiries Insulet was receiving from customers, an email was sent to all phone control users and the cause of the error message was resolved by reverting to prior certificates on the Insulet cloud to secure communications between the Omnipod 5 App sotware and the Insulet Cloud. This issue was limited to only certain compatible Android smartphones and did not affect the Omnipod 5 Pods or the dedicated locked-down Controller provided by Insulet. All other Omnipod 5 users were still able to manage their insulin with the device without use of the smartphone app and affected users were able to switch to the locked-down control provided to all users when they initiated the device during the time the App was unavailable to them. Further, the Omnipod itself continued to deliver insulin as per its pre-programmed settings so although affected users who did not have the locked-down controller readily available may not have been able to command insulin boluses during this time, they continued to receive basal insulin without interruption. Lastly, users of the Omnipod 5 were advised during training to always have backup supplies ready in the event of a device malfunction which would include alternative means to deliver insulin subcutaneously until the issue is resolved to minimize disruption to their insulin regimen. There were 2,168 complaints received out of 7,838 users with the Android compatible smartphones affected and 9 MDRs were submitted, however, there were no serious injuries reported or reports of erroneous results as the primary issue was a delay in therapy. This defect cannot cause false results that could negatively impact patients. | Class 2 Device Recall Omnipod 5 Automated Insulin Delivery System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Omnipod 5 Automated Insulin Delivery System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Omnipod 5 Automated Insulin Delivery System Date Initiated by Firm February 28, 2023 Date Posted September 20, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-2618-2023 Recall Event ID 92909 510(K)Number K203768 Product Classification Alternate controller enabled insulin infusion pump - Product Code QFG Product Omnipod 5 Automated Insulin Delivery System Code Information PT-000559 Omnipod 5 App; all versions of software installed******UDI-DI: 10385081120302 Recalling Firm/ Manufacturer Insulet Corporation 50 & 100 Nagog Park Acton MA 01720-3440 For Additional Information Contact 978-600-7451 Manufacturer Reason for Recall An error message was received when using the Omnipod 5 App on compatible smartphones that prevented phone control users from controlling the Omnipod 5 Automated Insulin Delivery System from their Omnipod 5 App which may cause a delay in therapy. To reduce the volume of inquiries Insulet was receiving from customers, an email was sent to all phone control users and the cause of the error message was resolved by reverting to prior certificates on the Insulet cloud to secure communications between the Omnipod 5 App sotware and the Insulet Cloud. This issue was limited to only certain compatible Android smartphones and did not affect the Omnipod 5 Pods or the dedicated locked-down Controller provided by Insulet. All other Omnipod 5 users were still able to manage their insulin with the device without use of the smartphone app and affected users were able to switch to the locked-down control provided to all users when they initiated the device during the time the App was unavailable to them. Further, the Omnipod itself continued to deliver insulin as per its pre-programmed settings so although affected users who did not have the locked-down controller readily available may not have been able to command insulin boluses during this time, they continued to receive basal insulin without interruption. Lastly, users of the Omnipod 5 were advised during training to always have backup supplies ready in the event of a device malfunction which would include alternative means to deliver insulin subcutaneously until the issue is resolved to minimize disruption to their insulin regimen. There were 2,168 complaints received out of 7,838 users with the Android compatible smartphones affected and 9 MDRs were submitted, however, there were no serious injuries reported or reports of erroneous results as the primary issue was a delay in therapy. This defect cannot cause false results that could negatively impact patients. FDA Determined Cause 2 Software Design Change Action Insulet emailed all phone control users on February 28, 2023 that an investigation into the issue was underway. Insulet determined root cause and resolved the issue early on March 1, 2023. Insulet sent a follow up email to all phone control users on March 1 to notify them that the issue had been resolved. Quantity in Commerce 11,016 users Distribution Distribution throughout the US Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = QFG - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Insulet emailed all phone control users on February 28, 2023 that an investigation into the issue was underway. Insulet determined root cause and resolved the issue early on March 1, 2023. Insulet sent a follow up email to all phone control users on March 1 to notify them that the issue had been resolved. | SoftwareDesignChange
33 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=202222 | Z-2414-2023 | FreeStyle Libre 2 Readers, REF: 71951-01 and 71953-01, used with FreeStyle Libre 2 System |  | 2 | CDRH | 2023-08-17 | 2023-08-17 21:07 |  | Abbott Diabetes Care, Inc. | If blood glucose monitoring system users attempt to start a new sensor, when the old sensor wasn't used for the full 14-day wear time, then the reader may display "Incompatible Sensor" message. If users don't have the previous sensor, or FreeStyle Libre 2 app, or the sensor's full wear time hasn't ended, users may be unable to start a new sensor, which could result in no or delayed glucose results | Class 2 Device Recall Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Date Initiated by Firm July 12, 2023 Date Posted August 17, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-2414-2023 Recall Event ID 92708 510(K)Number K193371 Product Classification Integrated Continuous Glucose Monitoring System, Factory Calibrated, Not For Use With Automated Insulin Delivery Systems - Product Code QLG Product FreeStyle Libre 2 Readers, REF: 71951-01 and 71953-01, used with FreeStyle Libre 2 System Code Information REF/UDI-DI/Lots:***71951-01/357599805005/02A289D, 03B082D, 02B082D, 02B109D, 03B251D, 04B251D, 01C073D, 01C084D, 02C087D, 01C143D, 03C152D, 03C166D, 01C174D, 03C222D, 01C256D, 02C313D, 01D018D, 02D093D, 02D129D;******71953-01/357599803001/01A198D, 02A198D, 03A198D, 04A198D, 05A198D, 01A202D, 04A202D, 01A204D, 02A203D, 02A205D, 03A205D, 01A211D, 02A211D, 03A210D, 02A212D, 04A218D, 05A218D, 06A218D, 01A219D, 04A225D, 05A225D, 01A231D, 03A231D, 02A232D, 06A232D, 02A240D, 02A248D, 01A259D, 01A260D, 04A262D, 03A266D, 01A267D, 02A268D, 03A268D, 02A280D, 01A281D, 01A283D, 02A283D, 01A290D, 02A297D, 02A298D, 03A296D, 02A303D, 02A304D, 04A325D, 01A328D, 03A325D, 02A340D, 03A344D, 02A351D, 03A350D, 01A352D, 03A352D, 02A358D, 01A364D, 02A363D, 01B007F, 02B007F, 01B008F, 01B011F, 03B012F, 01B012F, 01B014F, 05B013F, 01B015F, 04B013F, 02B013F, 11B021F, 02B034F, 02B035F, 04B034F, 04B035F, 06B034F, 07B034F, 02B036F, 05B035F, 01B036F, 01B039F, 03B039F, 03B036F, 03B041F, 05B041F, 02B042F, 03B043F, 02B046F, 05B043F, 06B046F, 09B046F, 01B071F, 05B071F, 06B074F, 04B074F, 02B076F, 05B075F, 01B077F, 05B076F, 02B078F, 05B078F, 02B081F, 02B082F, 04B082F, 02B083D, 05B082D, 03B109D, 02B203D, 03B203D, 01B205D, 01B207D, 01B208D, 02B208D, 02B211D, 02B216D, 02B229D, 01B288D, 02B289D, 01B289D, 02B305D, 01B306F, 02B306F, 01B306D, 02B308F, 03B306F, 01B308F, 01B309F, 01B309D, 01B310D, 02B312D, 01B313D, 03B312D, 03B313D, 05B322F, 07B322F, 01B323F, 01B326F, 06B322F, 01B305D, 01B327F, 03B326F, 04B326F, 02B328F, 01B308D, 01B333D, 01B334D, 01B335F, 03B335F, 02B333D, 01B336F, 02B336F, 01B337F, 02B337F, 01B302F, 01B340F, 02B340F, 03B302F, 04B301F, 01B337D, 01B341F, 02B341F, 01B341D, 02B341D, 01B342F, 02B342F, 02B343F, 01B343F, 04B302F, 02B344F, 06B349F, 03B351D, 04B352D, 04B354D, 02B356D, 01B356D, 04B356D, 05B356D, 01B362D, 02B362D, 02B363D, 01C006D, 03C005D, 02C005D, 02C010D, 01C011D, 01C012D, 03C013D, 02C013D, 02C025F, 03C028F, 01C028F, 02C031F, 03C031F, 01C034F, 02C032F, 01C074D, 03C074D, 02C073D, 02C074D, 03C075D, 01C075D, 02C080D, 02C112D, 03C117D, 03C123D, 03C126D, 02C130D, 03C129D, 02C133D, 03C133D, 01C133D, 01C136D, 02C136D, 01C137D, 01C138D, 04C138D, 01C140D, 02C140D, 03C140D, 02C143D, 01C144D, 02C145D, 05C146D, 01C146D, 01C159D, 03C159D, 01C161D, 03C165D, 05C165D, 02C172D, 01C181D, 02C181D, 01C182D, 01C186D, 01C188D, 02C188D, 01C194D, 02C193D, 03C196D, 03C199D, 05C199D, 01C201D, 02C199D, 03C207D, 01C207D, 01C209D, 02C208D, 03C213D, 02C223D, 01C223D, 01C224D, 02C224D, 01C227D, 01C228D, 02C229D, 03C224D, 02C236D, 01C237D, 03C236D, 02C237D, 03C238D, 04C241D, 01C242D, 01C243D, 03C242D, 02C242D, 02C243D, 04C243D, 01C245D, 02C245D, 01C249D, 02C250D, 03C251D, 02C255D, 03C256D, 01C257D, 02C256D, 02C257D, 03C257D, 02C262D, 02C265D, 01C266D, 03C270D, 08C271D, 01C272D, 01C276D, 01C273D, 04C276D, 02C278D, 03C279D, 01C283D, 02C287D, 04C292D, 02C294D, 03C294D, 06C292D, 01C298D, 01C299D, 02C298D, 01C305D, 02C305D, 03C305D, 01C307D, 01C308D, 01C313D, 02C312D, 01C317D, 01C318D, 03C320D, 02C321D, 01C322D, 01C325D, 02C325D, 01C326D, 01C332D, 01C327D, 01C334D, 02C334D, 04C334D, 01C340D, 02C340D, 01C341D, 03C340D, 01C346D, 02C355D, 01D010D, 01D011D, 02D011D, 01D012D, 02D018D, 02D017D, 02D023D, 02D028D, 02D019D, 03D024D, 01D025D, 02D025D, 01D027D, 03D027D, 01D033D, 03D034D, 04D034D, 05D033D, 01D040D, 02D040D, 02D038D, 06D055D, 03D041D, 02D045D, 02D053D, 03D058D, 02D056D, 02D060D, 03D067D, 04D068D, 03D065D, 02D065D, 02D066D, 02D069D, 01D072D, 03D069D, 02D072D, 01D076D, 02D075D, 03D075D, 01D095D, 03D083D, 01D081D, 02D095D, 02D081D, 02D082D, 03D080D, 02D087D, 06D103D, 01D102D, 02D107D, 01D107D, 01D104D, 03D107D, 02D104D, 01D111D, 01D115D, 02D115D, 02D116D, 03D115D, 01D128D, 01D129D, 05D130D, 04D125D, 02D128D, 03D136D, 02D136D, 03D142D, 01D139D, 03D144D, 02D125D, 01D150D, 03D151D, 02D151D, 03D152D, 01D156D, 02D157D, 02D160D, 02D156D, 01D159D, 01D177D, 02D177D Recalling Firm/ Manufacturer Abbott Diabetes Care, Inc. 1360 S Loop Rd Alameda CA 94502-7000 For Additional Information Contact Lindsy Delco 510-749-5000 Manufacturer Reason for Recall If blood glucose monitoring system users attempt to start a new sensor, when the old sensor wasn't used for the full 14-day wear time, then the reader may display "Incompatible Sensor" message. If users don't have the previous sensor, or FreeStyle Libre 2 app, or the sensor's full wear time hasn't ended, users may be unable to start a new sensor, which could result in no or delayed glucose results FDA Determined Cause 2 Software Design Change Action On 7/12/23, correction notices were e-mailed or mailed to customers and distributors who were asked to do the following:******If the "Incompatible Sensor" message is received when applying a new sensor and attempting it to start it with the Reader, follow these steps below.******If previous Sensor retained:***a) Wait 3 hours before scanning the previous sensor. While waiting, please use a blood glucose meter to check glucose.***b) "Replace Sensor" should appear.***c) Scan the new Sensor and "New Sensor Starting Up" screen should appear.******If previous Sensor not retained:***a) Users can use the free FreeStyle Libre 2 app to activate the new Sensor and get glucose data on the user's compatible smartphone.***-The new Sensor that generated the "Incompatible Sensor" error message on the***Reader can still be used with the app.***-App can be download from the Apple App Store or Google Play Store.***-To check smartphone compatibility, visit: https://www.freestyle.abbott/usen//products/freestyle-libre-app/compatibility.html***b) Open the FreeStyle Libre 2 app***c) Follow the instructions on the app to start the new Sensor.******-Users can start the next Sensor with the Reader once a "Sensor Ended" is received.***-Users can wait until the 14-day wear time for the previous Sensor has ended to activate the new Sensor with the Reader. The Reader Home Screen will indicate sensor wear days remaining.******The Reader can be used as a blood glucose meter; this functionality isn't impacted by the Incompatible Sensor message.******If not able to scan the previous Sensor and a compatible smartphone is not available, contact Customer Service at 1-855-632-8658 to receive further assistance. Additionally, contact Customer Service to report any adverse reactions or quality problems. ******Distributors should acknowledge receipt of the notice and provide the customer notice to all customers who were supplied Sensors since April 20, 2023. Quantity in Commerce 1,722,645 Distribution US Nationwide distribution in the states of CA, FL, MI, NY, OH, TX, UT, VA. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = QLG - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 7/12/23, correction notices were e-mailed or mailed to customers and distributors who were asked to do the following:******If the "Incompatible Sensor" message is received when applying a new sensor and attempting it to start it with the Reader, follow these steps below.******If previous Sensor retained:***a) Wait 3 hours before scanning the previous sensor. While waiting, please use a blood glucose meter to check glucose.***b) "Replace Sensor" should appear.***c) Scan the new Sensor and "New Sensor Starting Up" screen should appear.******If previous Sensor not retained:***a) Users can use the free FreeStyle Libre 2 app to activate the new Sensor and get glucose data on the user's compatible smartphone.***-The new Sensor that generated the "Incompatible Sensor" error message on the***Reader can still be used with the app.***-App can be download from the Apple App Store or Google Play Store.***-To check smartphone compatibility, visit: https://www.freestyle.abbott/usen//products/freestyle-libre-app/compatibility.html***b) Open the FreeStyle Libre 2 app***c) Follow the instructions on the app to start the new Sensor.******-Users can start the next Sensor with the Reader once a "Sensor Ended" is received.***-Users can wait until the 14-day wear time for the previous Sensor has ended to activate the new Sensor with the Reader. The Reader Home Screen will indicate sensor wear days remaining.******The Reader can be used as a blood glucose meter; this functionality isn't impacted by the Incompatible Sensor message.******If not able to scan the previous Sensor and a compatible smartphone is not available, contact Customer Service at 1-855-632-8658 to receive further assistance. Additionally, contact Customer Service to report any adverse reactions or quality problems. ******Distributors should acknowledge receipt of the notice and provide the customer notice to all customers who were supplied Sensors since April 20, 2023. | SoftwareDesignChange
34 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199871 | Z-1676-2023 | The 4Kscore Test is an in vitro serum or plasma test that combines the results of four immunoassays (Roche Elecsys total PSA (prostate specific antigen), Roche Elecsys free PSA, intact PSA, and human kallikrein 2) into a single numerical score that also incorporates the following information: a patients age, previous biopsy, and digital rectal exam (DRE). Patients file (sample type, prior biopsy status, DRE status and age) are uploaded to the Specimen Processing Module (SPM). The results from the four assays are exported automatically to Lab Manager Application (B2 LIS) which automatically triggers the 4Kscore calculation when values of the four analytes are available. The 4Kscore test results are determined automatically by the validated laboratory information system algorithm calculation software. | The 4Kscore Test | 2 | CDRH | 2023-06-02 | 2023-06-02 21:07 |  | BioReference Health, LLC | SPM software anomalies that may lead to the generation of erroneous 4Kscore Test results | Class 2 Device Recall The 4Kscore Test Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall The 4Kscore Test FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall The 4Kscore Test Date Initiated by Firm November 18, 2022 Date Posted June 02, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-1676-2023 Recall Event ID 92205 PMA Number P190022 Product Classification Multi-analyte test system with algorithmic analysis for detection of prostate cancer - Product Code QRF Product The 4Kscore Test is an in vitro serum or plasma test that combines the results of four immunoassays (Roche Elecsys total PSA (prostate specific antigen), Roche Elecsys free PSA, intact PSA, and human kallikrein 2) into a single numerical score that also incorporates the following information: a patient***s age, previous biopsy, and digital rectal exam (DRE). Patient***s file (sample type, prior biopsy status, DRE status and age) are uploaded to the Specimen Processing Module (SPM). The results from the four assays are exported automatically to Lab Manager Application (B2 LIS) which automatically triggers the 4Kscore calculation when values of the four analytes are available. The 4Kscore test results are determined automatically by the validated laboratory information system algorithm calculation software. Code Information UDI: 00850038860004 Recalling Firm/ Manufacturer BioReference Health, LLC 481 Edward H Ross Dr Elmwood Park NJ 07407-3118 For Additional Information Contact 800-229-5227 Manufacturer Reason for Recall SPM software anomalies that may lead to the generation of erroneous 4Kscore Test results FDA Determined Cause 2 Software Design Change Action Customers were notified via telephone between 11/18/22 and 12/31/22.******In addition, the following information was provided within the corrected test results sent to customers:****** The 4Kscore result has been corrected to reflect the history of a prostatic nodule and/or the prior negative biopsy result, which was inaccurately posted into the algorithm, due to a bioinformatics error. Quantity in Commerce 1 system (662 Patients test results affected) Distribution US Nationwide distribution in the state of New Jersey. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = QRF - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were notified via telephone between 11/18/22 and 12/31/22.******In addition, the following information was provided within the corrected test results sent to customers:****** The 4Kscore result has been corrected to reflect the history of a prostatic nodule and/or the prior negative biopsy result, which was inaccurately posted into the algorithm, due to a bioinformatics error. | SoftwareDesignChange
35 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199540 | Z-1579-2023 | BD Pyxis MedStation ES (Med ES Main), REF: 323; BD Pyxis MedStation ES (Med ES Main Tower), REF: 352; BD Pyxis Anesthesia Station ES, REF: 327; BD Pyxis CII Safe ES Tower Main, REF: 1116-00; BD Pyxis CII Safe ES, Desktop PC, W-LESS SCNR, BIO, REF: 107-255-01 ********* | BD Pyxis MedStation ES (Med ES Main) | 2 | CDRH | 2023-05-11 | 2023-05-11 21:07 |  | CareFusion 303, Inc. | Automated dispensing cabinet software is experiencing: 1) ES device download failure resulting in partial patient information loss that leads to 2) partial loss of patient transaction data, seen as a server upload processing error. Outdated/incorrect information display may contribute to the removal of medication to which the patient is allergic, incorrect dose amount, or discontinued medication. | Class 2 Device Recall BD Pyxis MedStation ES (Med ES Main) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall BD Pyxis MedStation ES (Med ES Main) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall BD Pyxis MedStation ES (Med ES Main) Date Initiated by Firm April 05, 2023 Date Posted May 11, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-1579-2023 Recall Event ID 92061 Product Classification Cabinet, table and tray, anesthesia - Product Code BRY Product BD Pyxis MedStation ES (Med ES Main), REF: 323; BD Pyxis MedStation ES (Med ES Main Tower), REF: 352; BD Pyxis Anesthesia Station ES, REF: 327; BD Pyxis CII Safe ES Tower Main, REF: 1116-00; BD Pyxis CII Safe ES, Desktop PC, W-LESS SCNR, BIO, REF: 107-255-01 Code Information REF/UDI-DI/Software: 323/10885403512667/1.7.0, 1.7.1, 1.7.2, 1.7.3, 1.7.4;***352/10885403512674/1.7.0, 1.7.1, 1.7.2, 1.7.3, 1.7.4;***327/10885403477836/1.7.0, 1.7.1, 1.7.2, 1.7.3, 1.7.4;***1116-00/10885403512605/1.7.0, 1.7.1, 1.7.2, 1.7.3, 1.7.4;***107-255-01/1.7.0, 1.7.1, 1.7.2, 1.7.3, 1.7.4 Recalling Firm/ Manufacturer CareFusion 303, Inc. 10020 Pacific Mesa Blvd San Diego CA 92121-4386 For Additional Information Contact 888-876-7128 Manufacturer Reason for Recall Automated dispensing cabinet software is experiencing: 1) ES device download failure resulting in partial patient information loss that leads to 2) partial loss of patient transaction data, seen as a server upload processing error. Outdated/incorrect information display may contribute to the removal of medication to which the patient is allergic, incorrect dose amount, or discontinued medication. FDA Determined Cause 2 Software Design Change Action On 4/5/23, correction notices were distributed to Biomedical Engineering, Risk Management, Nursing, Pharmacy Directors who were informed of the following:******1) Remind users that automated dispensing cabinet devices are not a replacement for the medical record (e.g., medication orders, medication administration record (MAR), and patient allergy information) and are not to be relied on as the source of truth for medication administration. Consult the hospital medical record before removing any medication.***2. Please ensure this letter is forwarded to applicable facilities if devices have been transferred and that all appropriate personnel have been made aware of this issue.***3. Complete and return the response form to BDRC24@bd.com******Recalling Firm Actions:***1. For customers who have reported an upload processing failure, the support team can provide a copy of affected data that is unable to be moved to the ES Server data repository.***2. A software update is available. The recalling firm will contact you to schedule a time to install the software update.******For further assistance contact Customer Support at 1-866-583-8783, 5:00am PST to 4:00pm PST, Monday - Friday; or Technical Support, bd.com/self-service, 1-800-727-6102, 24 hours, 7 days a week Quantity in Commerce 1,474 Distribution US Nationwide distribution in the states of KS, NY, NC, LA, MA, MI. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 4/5/23, correction notices were distributed to Biomedical Engineering, Risk Management, Nursing, Pharmacy Directors who were informed of the following:******1) Remind users that automated dispensing cabinet devices are not a replacement for the medical record (e.g., medication orders, medication administration record (MAR), and patient allergy information) and are not to be relied on as the source of truth for medication administration. Consult the hospital medical record before removing any medication.***2. Please ensure this letter is forwarded to applicable facilities if devices have been transferred and that all appropriate personnel have been made aware of this issue.***3. Complete and return the response form to BDRC24@bd.com******Recalling Firm Actions:***1. For customers who have reported an upload processing failure, the support team can provide a copy of affected data that is unable to be moved to the ES Server data repository.***2. A software update is available. The recalling firm will contact you to schedule a time to install the software update.******For further assistance contact Customer Support at 1-866-583-8783, 5:00am PST to 4:00pm PST, Monday - Friday; or Technical Support, bd.com/self-service, 1-800-727-6102, 24 hours, 7 days a week | SoftwareDesignChange
36 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=198114 | Z-1280-2023 | Scan Monitor/ScanWatch, model number hwa09******Wearable, Bluetooth-connected wrist watch that records heart activity (ECG) and oxygen saturation (SpO2)*** | Scan Monitor/ScanWatch | 2 | CDRH | 2023-03-22 | 2023-03-22 21:06 |  | Withings | A software bug eliminated the initial ECG activation and review of the ECG under a physician. However, the device is only cleared for use under the care of a physician. | Class 2 Device Recall Scan Monitor/ScanWatch Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Scan Monitor/ScanWatch FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Scan Monitor/ScanWatch Date Initiated by Firm December 19, 2022 Date Posted March 22, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-1280-2023 Recall Event ID 91523 510(K)Number K201456 Product Classification Electrocardiograph - Product Code DPS Product Scan Monitor/ScanWatch, model number hwa09************Wearable, Bluetooth-connected wrist watch that records heart activity (ECG) and oxygen saturation (SpO2) Code Information UDI-DI 03700546706400******Serial numbers***00:24:e4:e6:94:a0***00:24:e4:ee:d6:a3***00:24:e4:ef:3d:52***00:24:e4:ef:2b:ca***00:24:e4:e6:d2:fe***00:24:e4:ef:22:d4***00:24:e4:c0:9e:e8***00:24:e4:c0:90:ff***00:24:e4:ee:b6:9a***00:24:e4:d3:ee:74***a4:7e:fa:02:33:88***a4:7e:fa:02:36:4d***00:24:e4:ef:46:09***00:24:e4:c0:8e:6c***00:24:e4:ef:0c:07***a4:7e:fa:02:34:3d***00:24:e4:e7:5f:ce***00:24:e4:c0:8a:df***00:24:e4:e6:da:83***00:24:e4:d3:ff:c2***00:24:e4:c0:86:f2***00:24:e4:ef:45:29***a4:7e:fa:02:58:e1***00:24:e4:c0:a3:f4***00:24:e4:e6:fb:5c***00:24:e4:c0:88:cc***00:24:e4:ee:d8:3b***a4:7e:fa:02:37:55***00:24:e4:e7:4a:76***00:24:e4:e6:d6:f3***00:24:e4:c0:4b:08***00:24:e4:c0:ba:1a***00:24:e4:d3:fb:2c***a4:7e:fa:02:39:6d***00:24:e4:d4:11:37***00:24:e4:e6:d0:1a***a4:7e:fa:02:4a:71***00:24:e4:e7:41:dd***a4:7e:fa:02:49:5e***00:24:e4:e6:c4:ca***00:24:e4:d3:fe:44***00:24:e4:ef:0a:db***00:24:e4:c0:90:31***00:24:e4:e6:8c:c4***00:24:e4:e7:32:eb***00:24:e4:a3:d4:2a***00:24:e4:e6:c9:ae***00:24:e4:d3:f9:ef***00:24:e4:c0:8a:7b***00:24:e4:d4:13:51***00:24:e4:e6:7d:f9***00:24:e4:e6:f5:65***a4:7e:fa:02:37:a5***00:24:e4:e6:f1:9c***00:24:e4:e6:ca:b7***00:24:e4:e7:58:76***00:24:e4:e6:f4:eb***a4:7e:fa:02:93:af***00:24:e4:c0:aa:f1***00:24:e4:ee:b2:ac***00:24:e4:c0:a8:41***00:24:e4:e7:5f:30***00:24:e4:c0:52:6c***00:24:e4:e6:f5:9a***00:24:e4:ee:b7:c8***00:24:e4:e6:c8:69***00:24:e4:ef:37:2c***00:24:e4:d3:fc:73***00:24:e4:d3:f4:b3***00:24:e4:e6:d9:bc***a4:7e:fa:02:2f:bc***00:24:e4:e7:2d:0b***00:24:e4:e7:60:7b***00:24:e4:c0:b4:aa***00:24:e4:c0:a5:31***00:24:e4:d4:04:85***00:24:e4:ef:2b:ab***00:24:e4:d3:fc:65***a4:7e:fa:02:4a:3e***00:24:e4:e7:5c:4d***a4:7e:fa:02:9c:47***00:24:e4:d3:f9:80***00:24:e4:e7:2c:f3***00:24:e4:e7:57:3d***00:24:e4:c0:9f:b0***a4:7e:fa:02:39:4e***a4:7e:fa:02:3f:40***00:24:e4:e6:f1:98***00:24:e4:ee:c2:14***00:24:e4:e7:50:48***00:24:e4:e7:3d:7d***00:24:e4:ef:3a:bc***00:24:e4:d4:18:6d***00:24:e4:e7:43:39***00:24:e4:e7:45:4c***00:24:e4:c0:2b:bb***00:24:e4:ef:55:6d***00:24:e4:d3:fc:92***00:24:e4:e6:c9:7f***00:24:e4:d4:1b:71***00:24:e4:d4:23:56***00:24:e4:ef:2f:5f***00:24:e4:d4:24:a7***00:24:e4:ef:2b:03***00:24:e4:c0:8f:52***00:24:e4:d3:ed:4b***00:24:e4:d3:f5:2c***00:24:e4:d4:15:a5***00:24:e4:c0:88:8d***00:24:e4:e6:92:bc***00:24:e4:c0:be:23***00:24:e4:e6:c6:b1***00:24:e4:d3:f3:c2***00:24:e4:c0:8f:a8***00:24:e4:e7:2f:a1***00:24:e4:d4:15:4a***00:24:e4:e7:3a:13***00:24:e4:d3:fe:4a***00:24:e4:c0:a0:ba***00:24:e4:c0:ba:eb***00:24:e4:e7:5f:b5***00:24:e4:e6:88:66***00:24:e4:c0:84:77***00:24:e4:c0:74:d9***00:24:e4:ee:cd:27***00:24:e4:ee:b8:22***00:24:e4:c0:a5:f0***00:24:e4:c0:95:79***00:24:e4:e6:a9:b9***00:24:e4:e6:89:eb***00:24:e4:d3:f6:b0***00:24:e4:d4:07:65***00:24:e4:d4:04:d1***00:24:e4:e7:61:31***00:24:e4:d4:16:92***00:24:e4:c0:9b:87***00:24:e4:c0:78:f6***00:24:e4:e7:4b:1c***00:24:e4:d3:fd:7a***00:24:e4:e6:cd:17***00:24:e4:c0:ae:8c***00:24:e4:d4:1d:bd***00:24:e4:e7:0d:db***00:24:e4:ef:4f:c1***00:24:e4:ef:05:cb***00:24:e4:c0:91:c0***a4:7e:fa:02:31:dc***00:24:e4:d4:09:2b***00:24:e4:e7:2f:fa***00:24:e4:c0:9f:e4***00:24:e4:c0:9f:3e***00:24:e4:d3:f0:5a***00:24:e4:e6:d5:d1***00:24:e4:e6:de:c5***a4:7e:fa:02:24:72***00:24:e4:d4:09:5f***00:24:e4:c0:97:57***00:24:e4:e7:3f:a6***00:24:e4:e7:40:91***00:24:e4:c0:be:16***00:24:e4:d4:18:4f***00:24:e4:e7:55:42***00:24:e4:e6:b0:78***00:24:e4:ee:ae:76***00:24:e4:ee:b3:b5***00:24:e4:e7:0d:03***00:24:e4:ef:31:16***00:24:e4:ef:15:87***00:24:e4:ee:cd:1a***00:24:e4:e6:dd:6b***00:24:e4:ef:29:33***00:24:e4:ef:09:4d***00:24:e4:d3:f8:dd***00:24:e4:e7:25:51***00:24:e4:e7:5a:ed***00:24:e4:e6:91:78***00:24:e4:c0:8b:8c***00:24:e4:e7:29:1a***00:24:e4:e6:f0:ff***00:24:e4:c0:b3:ee***00:24:e4:c0:b8:50***00:24:e4:e6:8b:fc***00:24:e4:e7:62:06***00:24:e4:c0:b7:f3***00:24:e4:ef:46:51***00:24:e4:d5:15:3b***00:24:e4:e6:82:8d***00:24:e4:e7:5c:bb***00:24:e4:ef:5b:37***00:24:e4:e6:d0:73***00:24:e4:e6:cf:ad***00:24:e4:d4:14:47***00:24:e4:e6:f9:0b***00:24:e4:e6:c9:7c***00:24:e4:e6:c8:19***00:24:e4:e6:84:0d***00:24:e4:e6:f3:3f***00:24:e4:e6:f7:fb***00:24:e4:c0:83:9c***00:24:e4:ee:db:11***00:24:e4:e6:d6:c5***00:24:e4:c0:95:11***00:24:e4:e6:ca:a6***00:24:e4:ef:45:04***a4:7e:fa:02:36:52***00:24:e4:e7:60:7e***00:24:e4:ef:0d:bd***00:24:e4:e6:eb:cb***00:24:e4:e6:c8:39***00:24:e4:c0:b3:76***00:24:e4:e6:d4:09***00:24:e4:e6:f4:f9***00:24:e4:c0:52:f5***00:24:e4:d3:f8:71***00:24:e4:c0:2c:4a***00:24:e4:e6:94:2a***a4:7e:fa:02:34:9d***00:24:e4:e6:ee:50***00:24:e4:ef:08:c9***00:24:e4:ef:2f:cf***00:24:e4:e6:e8:62***a4:7e:fa:02:34:42***00:24:e4:e6:a6:a5***00:24:e4:ef:37:f3***00:24:e4:ee:bd:06***00:24:e4:e7:4c:37***00:24:e4:ef:45:a2***00:24:e4:e6:8a:e3***00:24:e4:ef:41:fc***00:24:e4:e6:e7:77***00:24:e4:d4:16:2a***00:24:e4:ef:47:c2***00:24:e4:e7:3d:ae***00:24:e4:e6:bc:a3***00:24:e4:ef:3b:c5***00:24:e4:d4:03:8c***00:24:e4:ef:0d:26***00:24:e4:ef:5b:4f***00:24:e4:e6:c7:03***a4:7e:fa:02:57:19***00:24:e4:e6:c6:be***00:24:e4:ef:31:69***00:24:e4:ee:bd:e4***00:24:e4:e6:93:0e***00:24:e4:e6:bb:13***00:24:e4:ef:3c:28***00:24:e4:c0:63:b7***a4:7e:fa:02:33:50***00:24:e4:ee:cf:73***00:24:e4:e6:b8:9f***00:24:e4:ee:dc:01***00:24:e4:d4:10:d5***00:24:e4:e6:cf:c6***00:24:e4:e7:43:08***00:24:e4:e6:c8:92***a4:7e:fa:02:49:77***00:24:e4:ef:30:10***00:24:e4:e6:d4:d3***a4:7e:fa:02:4b:a5***00:24:e4:e6:ee:6a***00:24:e4:d4:18:56***00:24:e4:e6:ed:e2***00:24:e4:e6:ca:87***00:24:e4:d4:14:8a***a4:7e:fa:02:3a:bb***00:24:e4:ef:38:dd***00:24:e4:e6:cf:f2***00:24:e4:d4:0a:08***00:24:e4:d5:0a:de***00:24:e4:e7:44:16***00:24:e4:e6:d8:ee***00:24:e4:e6:c8:d6***00:24:e4:ef:39:e9***00:24:e4:ef:46:38***00:24:e4:e6:e9:8d***00:24:e4:d3:f5:3d***00:24:e4:e6:80:1a***00:24:e4:e6:d8:17***00:24:e4:ef:50:f5***00:24:e4:e6:af:fa***00:24:e4:c0:8c:23***a4:7e:fa:02:31:6f***a4:7e:fa:02:3b:77***00:24:e4:d3:f0:ca***00:24:e4:e6:d5:97***00:24:e4:e6:8d:e3***a4:7e:fa:02:3f:fa***00:24:e4:ef:3f:78***00:24:e4:e6:c7:3c***00:24:e4:e7:60:88***a4:7e:fa:02:40:ed***00:24:e4:e7:2c:c3***00:24:e4:d5:04:c1***00:24:e4:e6:87:0a***00:24:e4:d4:12:16***00:24:e4:d4:14:27***00:24:e4:d4:25:0e***00:24:e4:d3:ed:44***00:24:e4:e6:d0:f9***00:24:e4:e7:41:d3***00:24:e4:d3:f8:87***00:24:e4:e6:f8:e4***00:24:e4:e7:54:e7***00:24:e4:ef:2a:e7***00:24:e4:c0:9a:88***00:24:e4:e6:fc:c6***00:24:e4:d4:1b:b1***00:24:e4:e6:d9:4a***00:24:e4:e6:d4:f1***00:24:e4:ef:0b:20***00:24:e4:e7:3a:01***00:24:e4:e7:66:29***00:24:e4:e7:2d:26***00:24:e4:c0:af:9d***00:24:e4:e6:ca:a9***00:24:e4:e6:cd:1c***00:24:e4:e6:a6:2a***00:24:e4:ef:0b:5c***00:24:e4:d4:16:12***a4:7e:fa:02:2f:79***00:24:e4:e6:96:02***00:24:e4:e7:30:79***00:24:e4:ef:07:da***00:24:e4:d4:1e:3a***00:24:e4:ee:b5:39***00:24:e4:e6:92:9d***00:24:e4:ee:b3:bf***00:24:e4:ef:17:69***00:24:e4:e6:ce:f6***00:24:e4:d3:f0:05***00:24:e4:ef:0a:eb***00:24:e4:e7:50:e1***a4:7e:fa:02:31:9d***00:24:e4:e6:ca:11***00:24:e4:e6:d8:60***00:24:e4:e6:c7:38***00:24:e4:e7:3f:85***00:24:e4:e7:64:d1***a4:7e:fa:02:35:53***00:24:e4:d5:01:9d***00:24:e4:e7:57:c9***00:24:e4:ef:0a:8c***00:24:e4:e6:ef:14***00:24:e4:e7:62:10***a4:7e:fa:02:3a:a7***00:24:e4:d4:01:93***00:24:e4:ee:cb:d9***a4:7e:fa:02:3c:11***00:24:e4:e6:7f:62***00:24:e4:e6:c6:2d***a4:7e:fa:02:33:b7***00:24:e4:d3:f8:f7***00:24:e4:ef:0c:b4***00:24:e4:e7:2c:dc***00:24:e4:e6:7f:11***00:24:e4:d3:f4:14***00:24:e4:e7:03:63***00:24:e4:d4:17:c2***00:24:e4:e6:7d:8d***00:24:e4:e7:28:84***00:24:e4:ef:22:ce***a4:7e:fa:02:30:57***00:24:e4:d3:e9:cd***00:24:e4:e6:bd:69***00:24:e4:e7:5e:4b***00:24:e4:c0:b7:d6***00:24:e4:e7:64:de***00:24:e4:ef:45:ff***00:24:e4:ef:39:5a***a4:7e:fa:02:34:be***00:24:e4:e6:88:04***00:24:e4:ef:5e:35***00:24:e4:d3:f4:2b***00:24:e4:e6:c6:c0***00:24:e4:ef:29:73***00:24:e4:ef:2f:36***00:24:e4:d3:fd:00***a4:7e:fa:02:27:c2***00:24:e4:ef:2b:92***00:24:e4:e6:ca:d8***00:24:e4:d4:24:c9***00:24:e4:d4:1e:d0***00:24:e4:ef:30:69***00:24:e4:ef:30:7c***00:24:e4:c0:22:e1***00:24:e4:ee:dc:da***00:24:e4:ef:19:76***a4:7e:fa:02:30:a3***00:24:e4:d4:07:6d***00:24:e4:ef:36:59***00:24:e4:ef:13:d6***00:24:e4:ef:45:bf***00:24:e4:e6:8d:30***a4:7e:fa:02:96:7d***00:24:e4:e6:b7:b7***00:24:e4:e6:ae:8f***00:24:e4:e6:d1:9f***00:24:e4:c0:ae:a0***00:24:e4:ef:0e:10***00:24:e4:d4:23:6a***a4:7e:fa:02:9b:5f***00:24:e4:e7:23:4c***00:24:e4:e6:d8:e3***00:24:e4:e6:c6:c9***00:24:e4:e6:eb:46***a4:7e:fa:02:34:6f***00:24:e4:e7:5f:42***00:24:e4:e6:f7:e6***00:24:e4:d5:05:2d***00:24:e4:ef:17:22***00:24:e4:e7:28:d7***00:24:e4:e6:cc:cb***00:24:e4:ef:59:a0***00:24:e4:d4:1f:54***a4:7e:fa:02:96:f7***00:24:e4:ee:bd:aa***00:24:e4:d3:f0:82***00:24:e4:e6:e9:1a***00:24:e4:c0:9f:40***00:24:e4:e6:cf:70***00:24:e4:e6:b3:93***00:24:e4:e6:a6:96***00:24:e4:e6:cf:24***a4:7e:fa:02:5a:14***00:24:e4:ef:05:2d***00:24:e4:e6:d0:9b***00:24:e4:e7:3c:0f***00:24:e4:ef:33:35***00:24:e4:ef:45:cf***00:24:e4:e6:b5:28***00:24:e4:e6:cb:02***00:24:e4:e6:c6:5e***00:24:e4:c0:b9:88***00:24:e4:e7:40:01***00:24:e4:d5:0a:18***00:24:e4:ee:cd:cc***00:24:e4:e6:d4:55***00:24:e4:a3:91:94***00:24:e4:ef:07:67***00:24:e4:e6:a0:81***00:24:e4:c0:a3:61***00:24:e4:ef:56:4c***00:24:e4:e6:96:c4***a4:7e:fa:02:39:3a***00:24:e4:ee:b3:d8***00:24:e4:e7:12:3b***00:24:e4:ef:36:00***00:24:e4:ee:c7:80***00:24:e4:d3:ef:61***00:24:e4:ef:2e:0f***00:24:e4:e6:81:dc***00:24:e4:c0:9d:db***00:24:e4:e7:3b:21***00:24:e4:d3:fc:04***00:24:e4:e6:e6:f9***00:24:e4:c0:4b:d0***00:24:e4:d4:03:aa***00:24:e4:e7:2f:fe***00:24:e4:d4:20:b5***00:24:e4:ee:b8:1a***00:24:e4:c0:a4:20***00:24:e4:ee:ca:1e***00:24:e4:d4:0f:22***00:24:e4:e6:96:df***00:24:e4:ef:29:3e***00:24:e4:c0:ba:71***00:24:e4:c0:83:4f***00:24:e4:c0:80:e3***00:24:e4:e6:cb:39***a4:7e:fa:02:4a:33***00:24:e4:e7:2d:f3***00:24:e4:e6:87:f6***00:24:e4:ee:c7:96***00:24:e4:ef:0c:f7***00:24:e4:d4:18:3d***00:24:e4:e6:85:75***00:24:e4:ee:b4:71***00:24:e4:e6:cd:16***00:24:e4:c0:be:49***00:24:e4:ee:cd:24***00:24:e4:c0:b1:02***00:24:e4:ef:44:88***00:24:e4:e6:f8:df***00:24:e4:ef:43:59***00:24:e4:e7:39:0a***00:24:e4:c0:21:74***00:24:e4:ee:c9:e0***00:24:e4:d5:0a:b4***00:24:e4:ee:b4:7e***00:24:e4:d3:f9:1f***00:24:e4:e7:0d:12***00:24:e4:c0:96:03***00:24:e4:c0:a3:ac***00:24:e4:c0:8e:6d***00:24:e4:ef:19:1e***00:24:e4:c0:29:55***00:24:e4:d3:f8:1b***00:24:e4:e7:60:c2***00:24:e4:ee:d0:30***00:24:e4:e7:49:b3***00:24:e4:e7:3d:5d***00:24:e4:ef:56:6a***00:24:e4:e6:c3:97***a4:7e:fa:02:a8:49***00:24:e4:e7:4a:04***00:24:e4:c0:50:4b***00:24:e4:ee:b7:ef***00:24:e4:e6:ca:8d***00:24:e4:e7:63:f0***00:24:e4:d4:1a:73***00:24:e4:e7:27:61***00:24:e4:ee:c8:fb***00:24:e4:e6:c9:f8***00:24:e4:d5:07:9c***00:24:e4:ef:3b:41***00:24:e4:e7:3d:7b***00:24:e4:e7:41:c8***00:24:e4:c0:9c:39***00:24:e4:d3:f3:0c***00:24:e4:e6:ff:58***00:24:e4:ef:0e:42***00:24:e4:c0:a3:2a***00:24:e4:c0:a7:21***a4:7e:fa:02:a0:23***00:24:e4:d3:ea:b2***00:24:e4:e6:f0:92***a4:7e:fa:02:7f:85***00:24:e4:d5:0a:53***00:24:e4:d4:0d:84***00:24:e4:e6:87:77***00:24:e4:ee:c0:c1***00:24:e4:e6:92:c4***00:24:e4:e7:12:23***00:24:e4:c0:ab:42***00:24:e4:ee:cb:0d***00:24:e4:e6:bd:d3***00:24:e4:e6:b0:23***00:24:e4:d3:ea:36***00:24:e4:d4:05:f2***00:24:e4:ef:45:f0***00:24:e4:b1:60:1a***00:24:e4:ef:36:7d***00:24:e4:e7:51:e6***00:24:e4:d4:1f:6f***00:24:e4:d4:17:a5***00:24:e4:c0:90:45***00:24:e4:ee:ce:c7***00:24:e4:e6:c9:c0***00:24:e4:ee:af:59***00:24:e4:d3:f1:ee***00:24:e4:e7:59:4b***00:24:e4:e6:d8:49***00:24:e4:e7:32:f6***00:24:e4:d4:1c:19***00:24:e4:c0:83:50***00:24:e4:e7:3c:e1***00:24:e4:d4:02:dc***00:24:e4:ef:46:4d***00:24:e4:e7:57:83***a4:7e:fa:02:94:bd***00:24:e4:e6:d9:aa***a4:7e:fa:02:49:d2***00:24:e4:ee:bd:e7***00:24:e4:c0:24:c8***a4:7e:fa:02:37:54***00:24:e4:e7:43:45***00:24:e4:c0:4c:bc***00:24:e4:d3:f1:ff***00:24:e4:ee:cc:03***00:24:e4:e6:da:0d***00:24:e4:e7:47:5f***00:24:e4:e6:ac:1d***00:24:e4:c0:91:4e***00:24:e4:e7:2b:cb***00:24:e4:ef:37:cd***a4:7e:fa:02:3a:c4***00:24:e4:b1:2f:8e***00:24:e4:ee:c8:fc***00:24:e4:ef:28:29***00:24:e4:e7:3e:24***00:24:e4:e7:53:fc***00:24:e4:c0:b3:83***00:24:e4:e7:28:b4***00:24:e4:e6:7d:b7***00:24:e4:e6:cb:6f***00:24:e4:ee:db:6e***00:24:e4:d4:0d:a1***00:24:e4:c0:b0:78***00:24:e4:ee:c7:70***00:24:e4:d4:15:54***00:24:e4:ef:58:96***00:24:e4:c0:27:54***00:24:e4:ef:0d:a9***00:24:e4:d3:f4:00***00:24:e4:ee:c9:e2***00:24:e4:ef:2e:81***00:24:e4:d4:23:67***00:24:e4:d5:07:c4***00:24:e4:d4:0e:40***00:24:e4:ef:0b:91***00:24:e4:e7:4b:b6***00:24:e4:d4:17:4f***00:24:e4:e6:c6:cd***00:24:e4:e6:f8:4a***00:24:e4:ef:2d:5e***00:24:e4:e6:a1:7c***a4:7e:fa:02:31:9b***00:24:e4:d4:16:6f***00:24:e4:e6:d0:95***00:24:e4:ef:2c:c6***00:24:e4:e6:96:b2***00:24:e4:d4:0b:87***00:24:e4:ee:c9:e9***00:24:e4:e7:3a:c0***00:24:e4:d3:fb:52***a4:7e:fa:02:a1:8a***a4:7e:fa:02:30:1a***00:24:e4:c0:2f:6d***00:24:e4:ef:53:7c***00:24:e4:e6:86:5a***00:24:e4:ef:59:74***00:24:e4:ef:39:dc***a4:7e:fa:02:4b:93***00:24:e4:ee:d6:87***00:24:e4:ee:ce:f3***a4:7e:fa:02:7f:bf***00:24:e4:d4:0c:f8***00:24:e4:ef:3c:1f***00:24:e4:d4:1d:a4***00:24:e4:ee:bf:ba***00:24:e4:ee:be:ec***00:24:e4:ef:2a:3f***00:24:e4:d4:17:9e***00:24:e4:ef:3f:32***00:24:e4:e6:f1:4a***00:24:e4:ef:2a:d7***00:24:e4:c0:98:8c***00:24:e4:e7:63:d5***00:24:e4:e6:f1:03***00:24:e4:ee:c6:e6***00:24:e4:ee:ae:7b***00:24:e4:d4:04:63***00:24:e4:ef:2b:5e***00:24:e4:e6:ed:f4***00:24:e4:e7:5c:b7***00:24:e4:e6:ca:58***00:24:e4:ef:28:a8***00:24:e4:e7:3f:f4***00:24:e4:c0:4d:7d***00:24:e4:d4:13:ec***00:24:e4:e7:33:24***00:24:e4:e6:81:28***00:24:e4:ef:59:42***00:24:e4:e6:ca:a3***a4:7e:fa:02:39:2e***00:24:e4:e6:85:b8***00:24:e4:d4:10:1d***00:24:e4:ef:5d:a8***00:24:e4:e7:0d:55***00:24:e4:c0:b8:c7***a4:7e:fa:02:a9:d6***00:24:e4:d5:08:a7***00:24:e4:ee:bc:f7***a4:7e:fa:02:ff:77***00:24:e4:d4:1a:83***00:24:e4:bf:f6:28***00:24:e4:c0:70:27***00:24:e4:ef:20:2c***00:24:e4:e7:48:ea***00:24:e4:e7:43:cb***a4:7e:fa:02:3a:ee***00:24:e4:d4:0d:a9***00:24:e4:ee:cc:0b***00:24:e4:c0:b2:a1***00:24:e4:ee:c4:0e***a4:7e:fa:02:30:97***00:24:e4:ef:2e:82***00:24:e4:c0:b7:81***00:24:e4:e6:bb:69***00:24:e4:c0:a2:8f***00:24:e4:ef:5a:bf***00:24:e4:e6:da:27***00:24:e4:ef:5f:0f***00:24:e4:c0:bd:86***00:24:e4:e6:89:2c***00:24:e4:e7:2c:f1***00:24:e4:ee:ed:19***00:24:e4:d5:08:b0***00:24:e4:c0:21:57***a4:7e:fa:02:38:ae***a4:7e:fa:02:3c:12***00:24:e4:e7:31:2f***00:24:e4:c0:3d:9d***00:24:e4:ef:40:b7***00:24:e4:d5:06:14***00:24:e4:c0:a4:a2***00:24:e4:e7:3d:42***00:24:e4:d4:19:1e***00:24:e4:ee:c5:6d***00:24:e4:ef:2b:b2***00:24:e4:e6:f3:29***00:24:e4:e6:f8:b6***a4:7e:fa:03:01:9e***00:24:e4:c0:b9:5b***00:24:e4:e7:44:b2***00:24:e4:e7:63:8c***00:24:e4:e7:10:7e***00:24:e4:c0:4d:69***00:24:e4:d4:06:60***a4:7e:fa:02:4c:1e***00:24:e4:e7:50:6d***a4:7e:fa:02:31:2b***00:24:e4:ee:c9:8c***00:24:e4:ef:5e:29***00:24:e4:e7:31:f4***00:24:e4:e6:ca:ad***00:24:e4:e6:8c:2f***00:24:e4:c0:5b:4e***a4:7e:fa:02:25:5a***00:24:e4:d3:ec:16***00:24:e4:d3:f3:54***00:24:e4:ef:38:9c***00:24:e4:ee:bd:cc***00:24:e4:ee:b5:3a***00:24:e4:e7:50:ed***00:24:e4:d4:14:cb***00:24:e4:c0:8c:07***00:24:e4:e7:12:04***00:24:e4:d4:0c:e2***00:24:e4:c0:4a:86***a4:7e:fa:02:41:23***00:24:e4:d3:f3:2f***00:24:e4:e7:3e:bb***00:24:e4:e7:4d:36***00:24:e4:e6:fa:3e***00:24:e4:d4:10:19***a4:7e:fa:02:2f:da***00:24:e4:d4:04:8b***00:24:e4:c0:be:17***00:24:e4:ee:ce:b8***00:24:e4:d4:0f:00***00:24:e4:d3:f3:fe***00:24:e4:e7:42:46***00:24:e4:d4:23:38***00:24:e4:e7:60:4f***00:24:e4:ef:45:71***00:24:e4:ee:cd:49***00:24:e4:d4:0b:11***00:24:e4:e6:c7:cc***00:24:e4:e6:c6:bd***00:24:e4:ef:3d:61***00:24:e4:c0:8a:6b***00:24:e4:c0:a8:05***00:24:e4:c0:a6:e5***00:24:e4:d3:f1:4e***a4:7e:fa:02:96:e2***00:24:e4:e6:d0:18***00:24:e4:c0:a7:05***00:24:e4:d5:08:c5***00:24:e4:e6:f9:0a***00:24:e4:c0:6f:95***00:24:e4:e6:d2:26***00:24:e4:c0:87:4a***00:24:e4:e7:47:3a***00:24:e4:d5:02:a0***00:24:e4:e7:3a:14***00:24:e4:ee:b8:26***00:24:e4:e7:43:9d***00:24:e4:c0:94:b9***00:24:e4:e7:4a:57***00:24:e4:ef:2e:f4***00:24:e4:ee:d6:77***00:24:e4:c0:ae:be***a4:7e:fa:02:30:99***00:24:e4:d4:09:69***00:24:e4:c0:9d:fc***00:24:e4:c0:84:bd***a4:7e:fa:02:59:5c***a4:7e:fa:02:31:a7***00:24:e4:d5:06:61***00:24:e4:e6:d3:6d***00:24:e4:e6:88:fe***00:24:e4:d4:1e:fc***00:24:e4:d5:03:54***00:24:e4:d5:08:22***00:24:e4:d3:f3:47***00:24:e4:d4:0c:ba***00:24:e4:d4:0e:47***00:24:e4:bf:a7:cd***00:24:e4:ef:08:cf***00:24:e4:e6:88:44***00:24:e4:d5:03:bc***00:24:e4:ef:49:5c***00:24:e4:e6:ea:21***00:24:e4:d4:1a:7c***00:24:e4:ef:47:d0***00:24:e4:e6:ca:02***a4:7e:fa:02:9c:ed***00:24:e4:ef:38:67***00:24:e4:e7:42:42***00:24:e4:c0:c1:73***00:24:e4:d4:15:1f***00:24:e4:e7:3f:b2***00:24:e4:c0:27:85***00:24:e4:e7:29:4c***a4:7e:fa:02:2a:66***a4:7e:fa:02:89:38***00:24:e4:bf:f4:dd***00:24:e4:c0:74:21***00:24:e4:ef:26:8f***00:24:e4:c0:b0:f6***00:24:e4:c0:31:d9***00:24:e4:e6:7d:fa***00:24:e4:ee:b3:a0***00:24:e4:ef:05:39***00:24:e4:e7:44:59***00:24:e4:e6:ee:49***00:24:e4:c0:bf:38***a4:7e:fa:02:26:e6***00:24:e4:e7:2c:d4***00:24:e4:d5:06:a9***00:24:e4:e7:63:d7***00:24:e4:ef:40:df***00:24:e4:ef:46:6c***00:24:e4:d4:12:96***00:24:e4:ef:0c:15***00:24:e4:e7:2a:7f***00:24:e4:e6:d3:e1***00:24:e4:e7:3c:12***00:24:e4:c0:97:e0***a4:7e:fa:02:3b:a9***00:24:e4:e6:a1:60***00:24:e4:d3:f3:82***00:24:e4:e7:4a:e8***00:24:e4:c0:ab:ff***00:24:e4:c0:4b:de***00:24:e4:c0:bf:23***a4:7e:fa:02:5e:63***a4:7e:fa:02:39:1c***00:24:e4:c0:26:5a***00:24:e4:e6:e7:64***00:24:e4:e6:8b:8e***00:24:e4:e6:8e:2d***00:24:e4:c0:a7:dc***00:24:e4:ef:2a:6e***00:24:e4:c0:c2:ec***00:24:e4:e6:81:d1***00:24:e4:e7:43:79***00:24:e4:ef:59:3f***00:24:e4:c0:4a:60***00:24:e4:e7:4b:23***00:24:e4:d3:f0:dc***a4:7e:fa:02:35:55***00:24:e4:ee:dd:2b***00:24:e4:e6:e9:71***00:24:e4:ef:46:80***00:24:e4:e6:a5:d0***00:24:e4:e6:d7:7e***00:24:e4:d4:1e:17***00:24:e4:e7:08:b8***00:24:e4:c0:93:a2***00:24:e4:c0:50:33***00:24:e4:d3:ec:47***00:24:e4:ee:b7:de***00:24:e4:e6:c2:db***00:24:e4:e6:e6:9d***00:24:e4:e6:c9:59***00:24:e4:d3:f2:a8***00:24:e4:e7:0f:5c***00:24:e4:c0:24:3d***00:24:e4:ee:c1:bb***00:24:e4:d3:f9:59***00:24:e4:d4:1f:4f***00:24:e4:e7:63:c3***00:24:e4:e7:57:9a***00:24:e4:c0:af:00***00:24:e4:e7:39:ac***00:24:e4:e6:e0:9b***00:24:e4:e6:cc:87***00:24:e4:e6:ed:4d***00:24:e4:ef:3f:77***00:24:e4:e6:cf:55***00:24:e4:e7:5f:d5***00:24:e4:e6:f3:36***00:24:e4:ef:2f:02***a4:7e:fa:02:31:81***00:24:e4:e6:83:00***00:24:e4:ef:41:2f***00:24:e4:e7:63:e4***00:24:e4:e6:90:fb***00:24:e4:e6:ec:b3***00:24:e4:d4:14:3f***00:24:e4:d3:ed:e0***00:24:e4:c0:99:62***00:24:e4:c0:9f:dc***a4:7e:fa:02:52:3e***00:24:e4:e6:b6:be***00:24:e4:e7:50:7b***00:24:e4:e6:c8:ca***00:24:e4:d4:10:e0***00:24:e4:c0:8d:61***00:24:e4:e6:f5:58***00:24:e4:c0:9f:59***00:24:e4:ee:c0:94***00:24:e4:e6:eb:bf***00:24:e4:ef:0d:a0***00:24:e4:e7:50:ef***00:24:e4:e7:4e:02***00:24:e4:ef:45:24***00:24:e4:e6:c1:bb***00:24:e4:ee:af:5f***a4:7e:fa:02:88:4c***00:24:e4:c0:b5:37***00:24:e4:d4:15:e1***00:24:e4:ee:cc:6e***00:24:e4:ee:ba:5d***00:24:e4:e6:7d:7f***00:24:e4:c0:b7:00***00:24:e4:e7:04:69***00:24:e4:ef:3f:04***00:24:e4:c0:78:bc***00:24:e4:c0:2e:b7***00:24:e4:d4:02:f9***00:24:e4:d3:f5:34***00:24:e4:c0:94:f7***00:24:e4:c0:b8:3e***00:24:e4:d4:1e:38***00:24:e4:ef:2d:ca***a4:7e:fa:02:f8:33***00:24:e4:d3:f2:c4***00:24:e4:ef:06:7f***00:24:e4:c0:8e:4d***00:24:e4:c0:a5:30***00:24:e4:c0:b5:31***00:24:e4:c0:b8:8e***00:24:e4:d3:f8:cb***00:24:e4:c0:85:b1***a4:7e:fa:02:39:74***00:24:e4:ef:2d:e0***00:24:e4:e7:32:ff***00:24:e4:e6:c8:31***00:24:e4:e7:5f:63***00:24:e4:d4:22:29***00:24:e4:d4:22:2f***00:24:e4:d4:05:df***00:24:e4:e7:4d:44***a4:7e:fa:02:99:9c***00:24:e4:c0:96:ce***00:24:e4:e7:4c:2c***00:24:e4:e7:4b:c6***00:24:e4:e7:28:c6***00:24:e4:d3:fd:8e***00:24:e4:e7:41:71***00:24:e4:c0:ae:83***00:24:e4:c0:2f:9f***00:24:e4:e6:d0:6f***00:24:e4:e7:28:b6***00:24:e4:e6:cb:18***00:24:e4:e7:4a:49***00:24:e4:ef:44:fe***00:24:e4:c0:37:9d***00:24:e4:e6:f1:28***a4:7e:fa:02:29:75***00:24:e4:ef:29:9f***00:24:e4:c0:74:07***00:24:e4:d3:fa:18***00:24:e4:d3:f5:90***00:24:e4:e7:42:30***00:24:e4:d3:ec:31***00:24:e4:e6:e8:22***00:24:e4:c0:af:40***00:24:e4:e7:4d:f6***00:24:e4:d4:0e:34***00:24:e4:d4:1f:17***00:24:e4:c0:98:7b***00:24:e4:d3:ef:4c***00:24:e4:c0:7a:e9***a4:7e:fa:02:9d:3e***00:24:e4:e7:4d:cf***00:24:e4:d3:fd:b7***00:24:e4:c0:84:56***00:24:e4:d4:04:0e***00:24:e4:d3:f0:ac***00:24:e4:d4:1f:c5***00:24:e4:ef:44:9c***00:24:e4:e6:88:f4***00:24:e4:d5:09:b9***00:24:e4:ee:ce:2a***00:24:e4:e6:f6:cc***00:24:e4:e7:48:db***00:24:e4:e7:65:2a***00:24:e4:d3:eb:d6***00:24:e4:ee:cd:b8***00:24:e4:e7:5b:83***00:24:e4:c0:88:3d***00:24:e4:ee:c2:24***00:24:e4:e7:46:93***00:24:e4:e7:3a:11***00:24:e4:e7:4b:18***00:24:e4:c0:c0:cf***00:24:e4:e7:31:f9***00:24:e4:c0:4f:02***00:24:e4:e7:56:7b***00:24:e4:ef:0a:e9***00:24:e4:d3:fe:5f***00:24:e4:c0:b1:9a***00:24:e4:d5:03:ed***00:24:e4:ef:4a:50***00:24:e4:ee:b8:bd***00:24:e4:d4:21:6a***00:24:e4:e7:3a:b2***00:24:e4:ee:cf:a8***00:24:e4:d4:06:c2***00:24:e4:e7:50:07***00:24:e4:d4:0b:43***00:24:e4:e7:38:fc***00:24:e4:e6:e9:a4***00:24:e4:e7:5b:90***00:24:e4:c0:9c:51***00:24:e4:c0:c2:8d***00:24:e4:ef:46:1a***00:24:e4:e6:ca:7a***00:24:e4:ee:cc:76***00:24:e4:d4:18:17***00:24:e4:d4:17:b9***00:24:e4:e6:80:97***00:24:e4:d3:f1:ef***00:24:e4:d4:19:39***00:24:e4:c0:8b:ae***00:24:e4:e6:cc:a9***00:24:e4:d3:fa:ea***00:24:e4:ef:41:15***00:24:e4:e7:2a:8b***00:24:e4:e6:89:4f***00:24:e4:ee:b1:9c***00:24:e4:e6:d6:d0***00:24:e4:ee:c4:38***00:24:e4:ef:15:3d***00:24:e4:ee:c4:0c***00:24:e4:c0:a4:01***00:24:e4:e7:3a:20***00:24:e4:e6:d0:03***a4:7e:fa:02:58:7e***00:24:e4:ef:04:f9***00:24:e4:e7:5e:01***a4:7e:fa:02:33:42***00:24:e4:e6:f2:21***00:24:e4:c0:53:a2***00:24:e4:c0:8c:c8***00:24:e4:e7:3a:4e***00:24:e4:c0:93:9c***00:24:e4:c0:b8:c0***00:24:e4:d4:24:17***00:24:e4:e6:85:fc***00:24:e4:ef:5b:51***00:24:e4:e7:5f:9c***00:24:e4:d4:1c:d8***00:24:e4:d3:f8:2f***00:24:e4:e7:59:55***00:24:e4:d4:14:e0***00:24:e4:c0:8f:15***00:24:e4:e7:5b:bb***00:24:e4:ee:b3:d9***00:24:e4:d4:23:18***00:24:e4:ef:17:62***00:24:e4:d3:ee:5b***00:24:e4:e7:32:37***00:24:e4:e6:80:98***00:24:e4:c0:bb:1b***00:24:e4:e7:3a:32***00:24:e4:bf:f4:65***00:24:e4:c0:96:88***00:24:e4:e6:f5:4f***00:24:e4:e7:44:58***00:24:e4:e7:33:56***00:24:e4:e7:43:09***00:24:e4:e6:91:a6***00:24:e4:ef:2f:a6***00:24:e4:e7:66:32***00:24:e4:e7:4d:19***00:24:e4:d4:04:4f***a4:7e:fa:02:34:d2***00:24:e4:d3:f2:ad***00:24:e4:e7:26:83***00:24:e4:e7:3f:eb***00:24:e4:ef:5a:5c***00:24:e4:e6:7f:6b***00:24:e4:e6:d0:cb***00:24:e4:e7:49:f6***00:24:e4:ef:3d:26***00:24:e4:e7:29:aa***00:24:e4:ef:39:d7***00:24:e4:c0:29:76***00:24:e4:d5:01:46***00:24:e4:c0:90:71***00:24:e4:d4:1c:ca***00:24:e4:e6:86:66***00:24:e4:e6:d8:8f***00:24:e4:d4:1a:d7***00:24:e4:e6:bb:d3***00:24:e4:c0:20:ef***a4:7e:fa:02:62:77***00:24:e4:ef:08:69***00:24:e4:e6:fa:21***00:24:e4:ef:49:a2***00:24:e4:c0:a2:73***00:24:e4:ef:29:4c***00:24:e4:ef:45:11***00:24:e4:ef:0d:42***00:24:e4:e6:e8:9c***00:24:e4:ef:59:cc***00:24:e4:d4:04:20***00:24:e4:e7:52:70***00:24:e4:ef:17:46***a4:7e:fa:02:45:38***00:24:e4:e6:cf:a1***00:24:e4:e6:8b:3f***00:24:e4:e6:d0:a1***00:24:e4:d4:13:0a***00:24:e4:c0:1f:fc***00:24:e4:e6:cf:05***00:24:e4:ef:2c:9c***00:24:e4:d4:04:a8***00:24:e4:d4:0d:a6***00:24:e4:e6:cc:af***00:24:e4:e6:d7:24***00:24:e4:d4:0b:c7***00:24:e4:e6:d8:61***00:24:e4:ee:dc:a0***00:24:e4:e6:b0:ad***a4:7e:fa:02:55:50***00:24:e4:e7:11:64***00:24:e4:e7:65:3c***00:24:e4:c0:9c:5c***00:24:e4:d5:09:0f***00:24:e4:ef:2d:59***00:24:e4:e6:ba:ab***00:24:e4:e6:c9:b6***00:24:e4:e6:cc:80***00:24:e4:ef:14:05***00:24:e4:ef:34:65***00:24:e4:e6:eb:1d***00:24:e4:d4:1b:07***00:24:e4:e7:57:23***00:24:e4:e7:46:d6***00:24:e4:e6:ca:8f***00:24:e4:ef:39:82***00:24:e4:e6:cd:8c***00:24:e4:e7:3a:39***00:24:e4:c0:53:29***00:24:e4:ef:43:ef***00:24:e4:d4:17:38***00:24:e4:e6:80:7c***a4:7e:fa:02:3c:44***00:24:e4:c0:b2:8f***00:24:e4:e6:cd:0b***00:24:e4:e6:f9:55***a4:7e:fa:02:31:83***00:24:e4:e6:89:d7***00:24:e4:e7:61:00***00:24:e4:e6:cd:00***00:24:e4:e7:67:4a***00:24:e4:ee:db:5e***00:24:e4:e6:ea:59***00:24:e4:e7:06:46***00:24:e4:e6:c6:83***00:24:e4:d5:07:9e***00:24:e4:c0:93:61***00:24:e4:ef:39:8a***00:24:e4:d4:07:ad***00:24:e4:ef:2d:25***00:24:e4:e6:e7:f2***00:24:e4:e6:8d:31***00:24:e4:e7:4c:9f***00:24:e4:e7:66:35***a4:7e:fa:02:47:b3***00:24:e4:e6:cc:d4***00:24:e4:e6:d6:a6***00:24:e4:d4:16:bc***00:24:e4:e6:7f:1e***a4:7e:fa:02:3d:ca***00:24:e4:d4:15:a1***00:24:e4:e6:f3:0a***00:24:e4:d5:08:c4***00:24:e4:e6:95:38***00:24:e4:d5:06:a7***00:24:e4:e6:cc:e3***a4:7e:fa:02:3f:34***00:24:e4:ef:2f:89***00:24:e4:e7:5b:1e***00:24:e4:e7:5b:10***00:24:e4:e7:5b:97***00:24:e4:e6:f3:ee***00:24:e4:e6:b5:cc***00:24:e4:e7:3f:04***00:24:e4:e6:cb:6c***00:24:e4:e6:cf:ee***00:24:e4:c0:ac:5b***00:24:e4:d5:04:56***00:24:e4:e6:96:82***00:24:e4:e6:f8:fd***00:24:e4:d3:f7:38***a4:7e:fa:02:3b:a7***00:24:e4:e6:d3:10***00:24:e4:ee:db:72***00:24:e4:e7:39:c6***00:24:e4:e6:f8:f0***00:24:e4:c0:4c:b8***00:24:e4:e6:8e:41***a4:7e:fa:02:33:ec***a4:7e:fa:02:45:be***00:24:e4:ef:06:c9***00:24:e4:e7:2e:4a***00:24:e4:e7:4c:52***00:24:e4:e6:b1:9b***00:24:e4:e6:e9:58***00:24:e4:ee:c7:ba***00:24:e4:e6:ce:62***a4:7e:fa:02:54:af***00:24:e4:d4:17:00***00:24:e4:e7:0a:c6***00:24:e4:e6:a5:a3***00:24:e4:d4:17:3a***00:24:e4:c0:bc:50***00:24:e4:e6:cf:42***00:24:e4:ef:0a:0e***00:24:e4:e7:3e:79***00:24:e4:e6:e8:29***00:24:e4:e7:4b:b4***a4:7e:fa:02:59:32***00:24:e4:d4:17:fd***00:24:e4:d4:19:db***00:24:e4:c0:bc:13***00:24:e4:c0:53:7c***00:24:e4:d4:0b:e3***00:24:e4:e7:06:f1***00:24:e4:e6:84:d7***00:24:e4:e6:c0:bb***a4:7e:fa:02:4a:a1***00:24:e4:ef:34:5b***00:24:e4:c0:ab:68***00:24:e4:e7:2a:b8***a4:7e:fa:02:44:4e***a4:7e:fa:02:48:f6***00:24:e4:e6:cc:e6***00:24:e4:e6:82:22***00:24:e4:ef:38:82***00:24:e4:d5:08:92***00:24:e4:c0:86:53***00:24:e4:d5:06:63***00:24:e4:e6:bc:01***00:24:e4:ef:40:13***00:24:e4:e7:43:b8***a4:7e:fa:02:39:53***00:24:e4:e6:ad:48***00:24:e4:e6:d0:0a***a4:7e:fa:02:35:67***00:24:e4:e6:85:55***00:24:e4:d4:09:90***00:24:e4:e6:f7:78***00:24:e4:d4:05:8e***00:24:e4:e6:cf:8d***00:24:e4:c0:89:85***00:24:e4:ef:0b:39***a4:7e:fa:02:92:11***a4:7e:fa:02:33:90***00:24:e4:e7:66:e1***a4:7e:fa:02:31:e3***00:24:e4:d4:15:af***00:24:e4:e6:d8:56***00:24:e4:ee:b6:8e***00:24:e4:e6:d7:a4***00:24:e4:e7:60:a1***00:24:e4:e6:d3:73***00:24:e4:ef:2f:e1***00:24:e4:e6:be:7d***00:24:e4:e6:87:9f***00:24:e4:e7:3d:17***00:24:e4:e6:f3:1b***00:24:e4:e7:5e:05***00:24:e4:d3:fa:9c***a4:7e:fa:02:31:bc***00:24:e4:e6:f6:de***00:24:e4:e7:5f:b2***00:24:e4:e6:e8:57***00:24:e4:e6:d8:0f***00:24:e4:e7:5c:da***00:24:e4:e7:0a:a6***00:24:e4:e6:93:60***00:24:e4:ef:0e:3e***a4:7e:fa:02:27:84***00:24:e4:e6:cc:85***00:24:e4:e7:5d:65***00:24:e4:e7:66:8a***00:24:e4:e7:41:ea***a4:7e:fa:02:64:b1***00:24:e4:e6:d7:b4***a4:7e:fa:02:3e:aa***a4:7e:fa:02:2c:33***a4:7e:fa:02:27:a5***00:24:e4:e7:59:35***00:24:e4:c0:21:91***00:24:e4:e7:06:e5***00:24:e4:ef:5f:f8***a4:7e:fa:02:38:11***00:24:e4:e6:ac:d6***a4:7e:fa:02:40:85***00:24:e4:ef:3b:36***00:24:e4:d4:10:71***00:24:e4:ef:09:e0***00:24:e4:e7:2d:1d***00:24:e4:ef:48:10***00:24:e4:ef:1e:65***00:24:e4:e7:62:d3***00:24:e4:e6:f0:f9***00:24:e4:e6:83:04***00:24:e4:c0:b1:ee***00:24:e4:c0:8b:b1***00:24:e4:e7:5d:ce***a4:7e:fa:02:63:76***00:24:e4:ef:3a:f8***00:24:e4:c0:8f:09***00:24:e4:c0:91:cc***00:24:e4:ef:06:0b***00:24:e4:e6:cb:5a***00:24:e4:e6:d3:5c***00:24:e4:e6:b2:6c***00:24:e4:c0:6f:63***00:24:e4:c0:b4:af***00:24:e4:e7:2f:c1***00:24:e4:ef:08:cb***00:24:e4:d4:05:01***00:24:e4:ee:c0:98***00:24:e4:d4:22:7b***00:24:e4:e7:4c:32***00:24:e4:e6:c0:90***00:24:e4:c0:96:7f***00:24:e4:d5:5d:23***a4:7e:fa:02:83:c3***00:24:e4:e6:8d:67***00:24:e4:e6:aa:c4***00:24:e4:e6:dd:d6***a4:7e:fa:02:31:95***00:24:e4:e7:39:1a***00:24:e4:e7:3e:ed***00:24:e4:e6:c7:51***00:24:e4:e7:31:d1***a4:7e:fa:02:29:98***00:24:e4:e6:e8:99***00:24:e4:d4:11:ea***00:24:e4:c0:72:1a***a4:7e:fa:02:2d:9b***00:24:e4:e6:a4:7e***00:24:e4:e7:09:cd***00:24:e4:c0:8e:81***00:24:e4:d3:ef:05***00:24:e4:e6:f8:4c***00:24:e4:e7:58:44***00:24:e4:ef:07:fe***00:24:e4:ef:0a:45***00:24:e4:e6:e0:af***00:24:e4:e6:e8:76***a4:7e:fa:02:3c:0b***00:24:e4:e7:28:d1***00:24:e4:ef:3f:ba***00:24:e4:d4:1a:51***00:24:e4:ef:3d:64***00:24:e4:e6:b9:04***00:24:e4:e7:5b:f3***a4:7e:fa:02:33:97***00:24:e4:e6:c6:91***00:24:e4:ef:0a:3c***00:24:e4:ef:29:28***00:24:e4:e6:ca:d5***00:24:e4:e6:f3:6d***00:24:e4:e6:ce:54***00:24:e4:e6:d1:5f***00:24:e4:c0:88:f2***00:24:e4:d3:fa:83***00:24:e4:e6:e5:79***00:24:e4:d4:25:10***00:24:e4:e7:2f:84***a4:7e:fa:02:89:4a***00:24:e4:c0:b6:ce***00:24:e4:ef:30:59***00:24:e4:e6:ca:8e***00:24:e4:e6:ca:c0***00:24:e4:ef:48:f8***00:24:e4:d3:f4:eb***00:24:e4:d5:15:b6***a4:7e:fa:02:64:bd***00:24:e4:ef:28:43***00:24:e4:c0:a4:62***00:24:e4:d5:01:52***00:24:e4:e6:83:ea***00:24:e4:ef:2e:e4***00:24:e4:d5:00:4f***00:24:e4:e6:c9:ed***00:24:e4:e7:27:66***00:24:e4:e6:c9:e7***a4:7e:fa:02:2d:15***00:24:e4:c0:b4:6c***a4:7e:fa:02:45:ad***00:24:e4:c0:a3:d0***00:24:e4:e6:f4:d5***00:24:e4:ef:30:6c***00:24:e4:ef:12:1c***00:24:e4:ef:5a:19***a4:7e:fa:02:37:57***00:24:e4:ef:2b:07***00:24:e4:e6:d8:a8***00:24:e4:e6:fa:2e***00:24:e4:e7:42:12***00:24:e4:e6:9d:13***00:24:e4:e7:2b:2a***00:24:e4:e6:eb:09***a4:7e:fa:02:a8:4b***a4:7e:fa:02:48:ec***00:24:e4:e6:fc:db***00:24:e4:e6:d1:f6***00:24:e4:e6:d7:ac***00:24:e4:e6:ca:a8***00:24:e4:e6:d0:97***00:24:e4:ef:2a:97***00:24:e4:e6:c7:1e***00:24:e4:ef:4e:a3***00:24:e4:ef:43:39***00:24:e4:c0:b1:61***00:24:e4:e7:46:1f***00:24:e4:e7:5e:0c***00:24:e4:ef:1f:27***00:24:e4:e7:44:07***00:24:e4:d4:0b:f9***00:24:e4:e6:f7:44***00:24:e4:ef:0d:0b***00:24:e4:e6:d5:21***00:24:e4:e6:ca:b2***00:24:e4:e7:2f:1d***00:24:e4:d3:f2:b6***00:24:e4:e7:58:04***00:24:e4:d5:04:52***00:24:e4:e6:d7:fa***a4:7e:fa:02:35:ab***00:24:e4:ef:2b:66***a4:7e:fa:02:67:1c***00:24:e4:c0:a3:25***00:24:e4:c0:b3:6c***00:24:e4:ef:1c:c3***00:24:e4:ef:34:43***00:24:e4:e7:50:7c***00:24:e4:e6:c9:7a***00:24:e4:c0:48:48***00:24:e4:e6:d8:28***00:24:e4:e6:cd:d9***a4:7e:fa:02:a7:a7***00:24:e4:d5:06:a5***00:24:e4:e7:3a:04***00:24:e4:e6:d8:4e***a4:7e:fa:02:82:f3***00:24:e4:ef:2f:d9***00:24:e4:e6:82:44***00:24:e4:ef:53:3a***00:24:e4:ef:12:01***00:24:e4:e6:d5:29***00:24:e4:e6:8c:ec***00:24:e4:ef:59:d0***00:24:e4:e6:fa:1c***00:24:e4:e7:5d:3f***00:24:e4:e6:af:a0***00:24:e4:e6:d4:21***a4:7e:fa:02:4f:ca***00:24:e4:e7:5d:32***00:24:e4:ef:30:b9***00:24:e4:ef:35:8e***a4:7e:fa:02:32:47***00:24:e4:e6:d8:a0***00:24:e4:e6:cd:c5***00:24:e4:e6:94:79***00:24:e4:e6:96:c0***00:24:e4:e6:c6:39***00:24:e4:e6:f0:c3***00:24:e4:ef:2b:3f***00:24:e4:d4:0e:7c***00:24:e4:e6:d5:53***00:24:e4:e6:b2:71***00:24:e4:d5:4d:c2***00:24:e4:e6:d5:f3***00:24:e4:e7:06:41***00:24:e4:d3:f3:5b***00:24:e4:ef:31:eb***00:24:e4:ef:28:6a***00:24:e4:ef:29:f1***00:24:e4:e6:93:f4***00:24:e4:ef:3f:fe***00:24:e4:e6:d7:c4***00:24:e4:e6:97:7d***00:24:e4:e6:b4:3f***00:24:e4:ef:2d:e6***00:24:e4:ef:45:d7***00:24:e4:ef:37:57***00:24:e4:ef:4b:50***00:24:e4:ef:4a:42***a4:7e:fa:02:52:b1***00:24:e4:d5:05:98***00:24:e4:e7:62:16***a4:7e:fa:02:31:a5***00:24:e4:ef:19:c7***00:24:e4:e6:ae:b3***00:24:e4:e6:ed:22***00:24:e4:ee:d9:0e***00:24:e4:e6:e5:ee***00:24:e4:ef:2e:04***00:24:e4:ef:10:83***00:24:e4:e6:87:c9***00:24:e4:e6:ec:eb***00:24:e4:ef:30:e6***a4:7e:fa:02:a1:9d***00:24:e4:c0:a3:a1***a4:7e:fa:02:29:b7***00:24:e4:e6:f6:5d***00:24:e4:c0:89:1c***00:24:e4:e6:d7:b6***a4:7e:fa:02:5e:ee***00:24:e4:c0:97:75***00:24:e4:e7:04:be***a4:7e:fa:02:50:79 Recalling Firm/ Manufacturer Withings 37 Bis Rue Du General Leclerc Issy les Moulineaux CEDEX France For Additional Information Contact Customer Service 888-2294116 Manufacturer Reason for Recall A software bug eliminated the initial ECG activation and review of the ECG under a physician. However, the device is only cleared for use under the care of a physician. FDA Determined Cause 2 Software Design Change Action Users were initially notified via push notifications in the companion app on December 19, 2022, which instructed users to update the app and activate the ECG feature under a physician.******On February 10, 2023, the firm followed up with customers via Medical Device Recall letters. Customers were instructed to ensure their Health Mate app is updated to iOS 5.13.0 or higher, and to ensure their device is synchronized with the Health Mate app. Customers should follow the instructions on their Health Mate app to activate their ECG feature.******Contact Information for questions:***Email: deals-support@withings.com***Phone number: 1-888-229-4116 (Availability |  | SoftwareDesignChange
37 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=198299 | Z-1192-2023 | Alaris PC Unit 8015 utilizing 802.11 b/g or 802.11 a/b/g wireless network cards, with IP addresses from x.x.x.224 to x.x.x.255 | Alaris PC Unit | 2 | CDRH | 2023-03-01 11:30 | 2023-03-01 20:07 |  | CareFusion 303, Inc. | Infusion pump PCs with specific software/network cards/IP addresses can have network/wireless connection failure. Could cause inability to receive new datasets and inability to transfer retrospective analytics data to connected systems. Users with interfaces to electronic medication record or admin discharge transfer, may be unable to receive/send data; infusions may need to be manually programmed | Class 2 Device Recall Alaris PC Unit Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Alaris PC Unit FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Alaris PC Unit Date Initiated by Firm January 24, 2023 Date Posted March 01, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-1192-2023 Recall Event ID 91582 510(K)Number K072105 Product Classification Pump, infusion - Product Code FRN Product Alaris PC Unit 8015 utilizing 802.11 b/g or 802.11 a/b/g wireless network cards, with IP addresses from x.x.x.224 to x.x.x.255 Code Information All PC Units, UDI-DI: 10885403801549, 10885403801532, with software versions 12.1.0, 12.1.2, 12.1.3, utilizing affected network cards and IP addresses Recalling Firm/ Manufacturer CareFusion 303, Inc. 10020 Pacific Mesa Blvd San Diego CA 92121-4386 Manufacturer Reason for Recall Infusion pump PCs with specific software/network cards/IP addresses can have network/wireless connection failure. Could cause inability to receive new datasets and inability to transfer retrospective analytics data to connected systems. Users with interfaces to electronic medication record or admin discharge transfer, may be unable to receive/send data; infusions may need to be manually programmed FDA Determined Cause 2 Software Design Change Action On 01/24/23, notices were mailed to customers, who were told to do the following:******1) Inspect the back of the PC Unit and observe the wireless card. Wireless cards that protrude out of the PC Unit are impacted.***2) Software version can be seen on the PCU using the following steps: a) Press Options Key, b) Press Page Down soft key, c) Press Software Versions soft key to display Software Versions menu, d) Press View soft key for the PCU.******To prevent this issue from occurring, do one of the following:***1. Exclude IP address ranges ending in x.x.x.224 to x.x.x.255 from all the DHCP scopes where the affected PC Units are being utilized.***2. Assign the affected PC Units IP addresses utilizing DHCP reservations to prevent the PC Unit from obtaining an IP address in the aforementioned range.***3. Assign the affected PC Units a static IP address outside of the aforementioned range.******Actions for Clinical Users: Check if wireless icon is illuminated. If the PC Unit is not connected to the wireless network when using Interoperability, the clinical user can input all infusion parameters manually. In addition, the clinical end user can manually document the infusion in the EMR. Data sets can also manually be programmed or uploaded to the PC unit using a serial cable. Continue to rely on hospital established protocols for wireless downtime procedures.******Customers who require additional assistance can contact: Customer Order Management, 1-800-482-4822, GMB-CTS-CustCareInfusion@bd.com; Technical Support, 1-866-488-1408, DL-US-INF-TechSupport@bd.com; Clinical & Pharmacy Support, 1-858-617-1316, GMB-AlarisMedSafetyProgram@bd.com; BD Designated Complaint Handling Unit, 1-888-812-3266, productcomplaints@bd.com******Complete and return the customer response form to BDRC39@bd.com Quantity in Commerce 83,185 Distribution US: FL, CA, OR, IL, WI, NJ, AZ, NV, CO, WY, NE, TX, MI, VA, OH, NC, LA, NM, AL, TN, AR, PA, CT, IN, KS, GA, SC, MD, UT, MA, DC, OK, IA, NY, WA, SD, ND, MN, KY, MO, WV, MS, DE, ME, MT, ID, NH, HI, AK, VT******OUS: Canada, Qatar, Saudi Arabia,, Kuwait, United Arab Emirates, Israel, South Africa, Zimbabwe, Botswana, Australia, New Zealand Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = FRN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 01/24/23, notices were mailed to customers, who were told to do the following:******1) Inspect the back of the PC Unit and observe the wireless card. Wireless cards that protrude out of the PC Unit are impacted.***2) Software version can be seen on the PCU using the following steps: a) Press Options Key, b) Press Page Down soft key, c) Press Software Versions soft key to display Software Versions menu, d) Press View soft key for the PCU.******To prevent this issue from occurring, do one of the following:***1. Exclude IP address ranges ending in x.x.x.224 to x.x.x.255 from all the DHCP scopes where the affected PC Units are being utilized.***2. Assign the affected PC Units IP addresses utilizing DHCP reservations to prevent the PC Unit from obtaining an IP address in the aforementioned range.***3. Assign the affected PC Units a static IP address outside of the aforementioned range.******Actions for Clinical Users: Check if wireless icon is illuminated. If the PC Unit is not connected to the wireless network when using Interoperability, the clinical user can input all infusion parameters manually. In addition, the clinical end user can manually document the infusion in the EMR. Data sets can also manually be programmed or uploaded to the PC unit using a serial cable. Continue to rely on hospital established protocols for wireless downtime procedures.******Customers who require additional assistance can contact: Customer Order Management, 1-800-482-4822, GMB-CTS-CustCareInfusion@bd.com; Technical Support, 1-866-488-1408, DL-US-INF-TechSupport@bd.com; Clinical & Pharmacy Support, 1-858-617-1316, GMB-AlarisMedSafetyProgram@bd.com; BD Designated Complaint Handling Unit, 1-888-812-3266, productcomplaints@bd.com******Complete and return the customer response form to BDRC39@bd.com | SoftwareDesignChange
38 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197513 | Z-0959-2023 | Getinge Flow-i Anesthesia Systems***Flow-i C20, Part no. 6677200***Flow-i C30, Part no. 6677300***Flow-i C40, Part no. 6677400****** | Getinge Flow-i | 1 | CDRH | 2023-01-26 | 2023-01-26 20:07 |  | Getinge Usa Sales Inc | Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia. | Class 1 Device Recall Getinge Flowi Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall Getinge Flowi FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall Getinge Flowi Date Initiated by Firm December 22, 2022 Date Posted January 26, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-0959-2023 Recall Event ID 91339 510(K)Number K191027 Product Classification Gas-machine, anesthesia - Product Code BSZ Product Getinge Flow-i Anesthesia Systems******Flow-i C20, Part no. 6677200******Flow-i C30, Part no. 6677300******Flow-i C40, Part no. 6677400 Code Information Part #*** UDI***6677200***7325710001349***Serial Numbers 1202-8027******Part #*** UDI***6677300***7325710001363***Serial Numbers 1058-1201******Part #*** UDI***6677400***7325710001318***Serial Numbers 2034-8446******Software versions 4.8.0, 4.8.1, 4.8.2, and 4.8.3. Recalling Firm/ Manufacturer Getinge Usa Sales Inc 1 Geoffrey Way Wayne NJ 07470-2035 For Additional Information Contact Stephanie Moretti 973-709-7170 Manufacturer Reason for Recall Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia. FDA Determined Cause 2 Software Design Change Action On December 22, 2022, the firm notified affected customers through Urgent Medical Device Correction letters.******Customers were instructed to examine their inventory immediately to determine if they have any of the affected serial numbers. If you have an affected system running on software version 4.8.0 through 4.8.3, you may continue use of the device until the software is updated with the following considerations:******1. If during use, the device is switched from Automatic to Manual and it is found that the manual bag is empty the user can activate the Emergency ventilation or turn off/on the device to reset the disabled valves function.******2. Users may also use the resuscitator that is available in the Operating Room.******Getinge will update the software to correct the issue free of charge. You can arrange a visit by a Getinge trained or authorized service technician by contacting Getinge Customer Technical Support at (888) 9GETUSA / (888) 943 8872 (select option 1, then option 2), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. (Eastern Time Zone). Quantity in Commerce 8,016 Total (7,766 OUS; 250 US) Distribution US nationwide and Worldwide Distribution:******Domestic distribution to AL, AR, AZ, CA, CO, DC, FL, GA, IN, MA, MI, MO, NY, OH, OR, PA, TX, VA, WA, WV. Foreign distribution worldwide. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = BSZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On December 22, 2022, the firm notified affected customers through Urgent Medical Device Correction letters.******Customers were instructed to examine their inventory immediately to determine if they have any of the affected serial numbers. If you have an affected system running on software version 4.8.0 through 4.8.3, you may continue use of the device until the software is updated with the following considerations:******1. If during use, the device is switched from Automatic to Manual and it is found that the manual bag is empty the user can activate the Emergency ventilation or turn off/on the device to reset the disabled valves function.******2. Users may also use the resuscitator that is available in the Operating Room.******Getinge will update the software to correct the issue free of charge. You can arrange a visit by a Getinge trained or authorized service technician by contacting Getinge Customer Technical Support at (888) 9GETUSA / (888) 943 8872 (select option 1, then option 2), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. (Eastern Time Zone). | SoftwareDesignChange
39 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197515 | Z-0961-2023 | Getinge Flow-e Anesthesia System, Part No. 6887900*** | Getinge Flow-e | 1 | CDRH | 2023-01-26 | 2023-01-26 20:07 |  | Getinge Usa Sales Inc | Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia. | Class 1 Device Recall Getinge Flowe Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall Getinge Flowe FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall Getinge Flowe Date Initiated by Firm December 22, 2022 Date Posted January 26, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-0961-2023 Recall Event ID 91339 510(K)Number K191027 Product Classification Gas-machine, anesthesia - Product Code BSZ Product Getinge Flow-e Anesthesia System, Part No. 6887900 Code Information Part #*** UDI***6887900 07325710010457***Serial Numbers 50012-50197******Software versions 4.8.0, 4.8.1, 4.8.2, and 4.8.3. Recalling Firm/ Manufacturer Getinge Usa Sales Inc 1 Geoffrey Way Wayne NJ 07470-2035 For Additional Information Contact Stephanie Moretti 973-709-7170 Manufacturer Reason for Recall Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia. FDA Determined Cause 2 Software Design Change Action On December 22, 2022, the firm notified affected customers through Urgent Medical Device Correction letters.******Customers were instructed to examine their inventory immediately to determine if they have any of the affected serial numbers. If you have an affected system running on software version 4.8.0 through 4.8.3, you may continue use of the device until the software is updated with the following considerations:******1. If during use, the device is switched from Automatic to Manual and it is found that the manual bag is empty the user can activate the Emergency ventilation or turn off/on the device to reset the disabled valves function.******2. Users may also use the resuscitator that is available in the Operating Room.******Getinge will update the software to correct the issue free of charge. You can arrange a visit by a Getinge trained or authorized service technician by contacting Getinge Customer Technical Support at (888) 9GETUSA / (888) 943 8872 (select option 1, then option 2), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. (Eastern Time Zone). Quantity in Commerce 8.016 Total (7.766 OUS; 250 US) Distribution US nationwide and Worldwide Distribution:******Domestic distribution to AL, AR, AZ, CA, CO, DC, FL, GA, IN, MA, MI, MO, NY, OH, OR, PA, TX, VA, WA, WV. Foreign distribution worldwide. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = BSZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On December 22, 2022, the firm notified affected customers through Urgent Medical Device Correction letters.******Customers were instructed to examine their inventory immediately to determine if they have any of the affected serial numbers. If you have an affected system running on software version 4.8.0 through 4.8.3, you may continue use of the device until the software is updated with the following considerations:******1. If during use, the device is switched from Automatic to Manual and it is found that the manual bag is empty the user can activate the Emergency ventilation or turn off/on the device to reset the disabled valves function.******2. Users may also use the resuscitator that is available in the Operating Room.******Getinge will update the software to correct the issue free of charge. You can arrange a visit by a Getinge trained or authorized service technician by contacting Getinge Customer Technical Support at (888) 9GETUSA / (888) 943 8872 (select option 1, then option 2), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. (Eastern Time Zone). | SoftwareDesignChange
40 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197514 | Z-0960-2023 | Getinge Flow-c Anesthesia System, Part No. 6887700*** | Getinge Flow-c | 1 | CDRH | 2023-01-26 | 2023-01-26 20:07 |  | Getinge Usa Sales Inc | Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia. | Class 1 Device Recall Getinge Flowc Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall Getinge Flowc FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall Getinge Flowc Date Initiated by Firm December 22, 2022 Date Posted January 26, 2023 Recall Status 1 Open 3 , Classified Recall Number Z-0960-2023 Recall Event ID 91339 510(K)Number K191027 Product Classification Gas-machine, anesthesia - Product Code BSZ Product Getinge Flow-c Anesthesia System, Part No. 6887700 Code Information Part #*** UDI***6887700 07325710009765***Serial Numbers 4003-5078******Software versions 4.8.0, 4.8.1, 4.8.2, and 4.8.3. Recalling Firm/ Manufacturer Getinge Usa Sales Inc 1 Geoffrey Way Wayne NJ 07470-2035 For Additional Information Contact Stephanie Moretti 973-709-7170 Manufacturer Reason for Recall Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia. FDA Determined Cause 2 Software Design Change Action On December 22, 2022, the firm notified affected customers through Urgent Medical Device Correction letters.******Customers were instructed to examine their inventory immediately to determine if they have any of the affected serial numbers. If you have an affected system running on software version 4.8.0 through 4.8.3, you may continue use of the device until the software is updated with the following considerations:******1. If during use, the device is switched from Automatic to Manual and it is found that the manual bag is empty the user can activate the Emergency ventilation or turn off/on the device to reset the disabled valves function.******2. Users may also use the resuscitator that is available in the Operating Room.******Getinge will update the software to correct the issue free of charge. You can arrange a visit by a Getinge trained or authorized service technician by contacting Getinge Customer Technical Support at (888) 9GETUSA / (888) 943 8872 (select option 1, then option 2), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. (Eastern Time Zone). Quantity in Commerce 8.016 Total (7.766 OUS; 250 US) Distribution US nationwide and Worldwide Distribution:******Domestic distribution to AL, AR, AZ, CA, CO, DC, FL, GA, IN, MA, MI, MO, NY, OH, OR, PA, TX, VA, WA, WV. Foreign distribution worldwide. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = BSZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On December 22, 2022, the firm notified affected customers through Urgent Medical Device Correction letters.******Customers were instructed to examine their inventory immediately to determine if they have any of the affected serial numbers. If you have an affected system running on software version 4.8.0 through 4.8.3, you may continue use of the device until the software is updated with the following considerations:******1. If during use, the device is switched from Automatic to Manual and it is found that the manual bag is empty the user can activate the Emergency ventilation or turn off/on the device to reset the disabled valves function.******2. Users may also use the resuscitator that is available in the Operating Room.******Getinge will update the software to correct the issue free of charge. You can arrange a visit by a Getinge trained or authorized service technician by contacting Getinge Customer Technical Support at (888) 9GETUSA / (888) 943 8872 (select option 1, then option 2), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. (Eastern Time Zone). | SoftwareDesignChange
41 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=196379 | Z-0198-2023 | Canon Aplio ultrasound system, Model: Aplio i900 TUS-AI900, Aplio i800 TUS-AI800, Aplio i700 TUS-AI700, Aplio a550 CUS-AA550, Aplio a450 CUS-AA450*** | Canon Aplio ultrasound system | 2 | CDRH | 2022-11-07 | 2022-11-08 20:06 |  | Canon Medical System, USA, INC. | After completion, ultrasound system's Stress Echo may not terminate if user does not manually terminate the protocol. During next patient exam, an abnormal confirmation message will be displayed "Same view was saved in this phase. Do you use this image with StressEcho Review?" If "no", acquired images not saved, and if "yes" images assigned to the previous patient; could cause treatment delay. | Class 2 Device Recall Canon Aplio ultrasound system Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Canon Aplio ultrasound system FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Canon Aplio ultrasound system Date Initiated by Firm September 29, 2022 Date Posted November 08, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-0198-2023 Recall Event ID 90980 510(K)Number K212333 K212960 Product Classification System, imaging, pulsed doppler, ultrasonic - Product Code IYN Product Canon Aplio ultrasound system, Model: Aplio i900 TUS-AI900, Aplio i800 TUS-AI800, Aplio i700 TUS-AI700, Aplio a550 CUS-AA550, Aplio a450 CUS-AA450 Code Information Model/UDI-DI, with software: Aplio i900 TUS-AI900/ 04987670102523,***Aplio i800 TUS-AI800/ 04987670102516,***Aplio i700 TUS-AI700/ 04987670102509,******With System Version:V4.6/V5.0/V5.1, SP0005 to SP1014, or V6.0/V6.5 SP0000 to SP1009;******Aplio a550 CUS-AA550/ 04987670103353,***Aplio a450 CUS-AA450/ 04987670103346,******With System Version: V3.0, SP0005 to SP1014, or V4.0/V4.5, SP0000 to SP1009 Recalling Firm/ Manufacturer Canon Medical System, USA, INC. 2441 Michelle Dr Tustin CA 92780-7047 For Additional Information Contact 714-730-5000 Manufacturer Reason for Recall After completion, ultrasound system's Stress Echo may not terminate if user does not manually terminate the protocol. During next patient exam, an abnormal confirmation message will be displayed "Same view was saved in this phase. Do you use this image with StressEcho Review?" If "no", acquired images not saved, and if "yes" images assigned to the previous patient; could cause treatment delay. FDA Determined Cause 2 Software Design Change Action On 9/29/22, correction notices were mailed to customers who were asked to do the following until the software corrective action had been implemented: Terminate Stress Echo by pressing "X" or "Quit" on the touch panel when completing the exam. If the exam is completed without terminating Stress Echo, and the confirmation message shown in section 2 is displayed when saving the first image in the next Stress Echo, terminate Stress Echo by pressing "X" or "Quit", and then start Stress Echo again. However, if the confirmation message is not displayed when saving the first image in Stress Echo, the function can be used normally. ******Share the contents of this letter with all users and reviewing radiologist as well as clinical engineering or biomedical group at your facility. Complete and return the response form. ***For software scheduling questions, please contact your respective Area Service Manager or our InTouch team at intouchdispatch@us.medical.canon. Quantity in Commerce 65 Distribution US: MT, MO, NC, SC, MN, WA, KY, TN, AR, IL, IA, MI, WV, KS, CA, HA, CO, LA, OK, ME, KY, NY, PA, GA, FL Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = IYN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 9/29/22, correction notices were mailed to customers who were asked to do the following until the software corrective action had been implemented: Terminate Stress Echo by pressing "X" or "Quit" on the touch panel when completing the exam. If the exam is completed without terminating Stress Echo, and the confirmation message shown in section 2 is displayed when saving the first image in the next Stress Echo, terminate Stress Echo by pressing "X" or "Quit", and then start Stress Echo again. However, if the confirmation message is not displayed when saving the first image in Stress Echo, the function can be used normally. ******Share the contents of this letter with all users and reviewing radiologist as well as clinical engineering or biomedical group at your facility. Complete and return the response form. ***For software scheduling questions, please contact your respective Area Service Manager or our InTouch team at intouchdispatch@us.medical.canon. | SoftwareDesignChange
42 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=195973 | Z-0133-2023 | da Vinci X (IS4200) and Xi (IS4000) systems, which consists of finished devices Surgeon Console (SSC), Patient Cart (PSCART), Vision System (VSS) | da Vinci X (IS4200) and Xi (IS4000) systems | 2 | CDRH | 2022-10-25 10:26 | 2022-10-25 21:06 |  | Intuitive Surgical, Inc. | Inadvertent energy delivery from surgical system instrument if 1) Force bipolar and bipolar instruments installed on system 2) Force bipolar connected to Force Triad generator, bipolar not connected to generator 3) Yellow pedal associated with force bipolar is pressed, released with head in surgeon console 4) Blue pedal associated with bipolar is pressed, resulting in force bipolar energy delivery | Class 2 Device Recall da Vinci X (IS4200) and Xi (IS4000) systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall da Vinci X (IS4200) and Xi (IS4000) systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall da Vinci X (IS4200) and Xi (IS4000) systems Date Initiated by Firm September 20, 2022 Date Posted October 25, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-0133-2023 Recall Event ID 90826 510(K)Number K131861 K171294 Product Classification System, surgical, computer controlled instrument - Product Code NAY Product da Vinci X (IS4200) and Xi (IS4000) systems, which consists of finished devices Surgeon Console (SSC), Patient Cart (PSCART), Vision System (VSS) Code Information Software version P10 (OS4 v10.0.0/A70_P10_B738) ******SOFTWARE,EMBEDDED RLS,IS4000 AND IS4200,A70_P10_B738 (Model: 610092-738) on Model Name/Model Number/UDI-DI:***ASSY,PSCART,IS4000,4-ARM,P10/380652-55/00886874110720,***ASSY,SSC,IS4000,P10/380677-28/00886874110744,***ASSY,VSS VISION SYSTEM,IS4000,P10/381121-45/00886874110898,***ASSY,PSCART,IS4200,P10/380620-47/00886874115404,***ASSY,VSC,IS4000,RECONDITIONED,P10/380721-15/00886874115343,***ASSY,SSC,IS4000,RECONDITIONED,P10/380723-16/00886874115374 Recalling Firm/ Manufacturer Intuitive Surgical, Inc. 1266 Kifer Rd Bldg 101 Sunnyvale CA 94086-5304 For Additional Information Contact Alex Mathews 408-523-2100 Manufacturer Reason for Recall Inadvertent energy delivery from surgical system instrument if 1) Force bipolar and bipolar instruments installed on system 2) Force bipolar connected to Force Triad generator, bipolar not connected to generator 3) Yellow pedal associated with force bipolar is pressed, released with head in surgeon console 4) Blue pedal associated with bipolar is pressed, resulting in force bipolar energy delivery FDA Determined Cause 2 Software Design Change Action On 09/20/22, correction notices were mailed to customers.***Customers were advised of the following:***1) The system software version can be located on the Vision Cart Touchscreen in the Troubleshooting Tab.***2) Ensure all device users read and understand this notice.***3) Notice should be passed on to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred.***4) Ensure proper instrument connections to the generator(s) and continue to follow all existing warnings and cautions found in user manuals 551400-15, USER MANUAL,SYS,P10,US,IS4000, 553873-07, USER MANUAL,I-A,MDR P10,IS4000-IS4200 and 554070-04, ADD,I-A,FORCE BIPOLAR,IS4000,IS4200.***5) Complete and return the acknowledgement form.***6) The recalling firm will follow-up with customers and schedule a software upgrade. ***7) After the upgrade, inform all personnel that the software update/upgrade was completed.***8) If you have experienced this issue or require additional assistance, contact Technical Support: 800-876-1310, Option 1 (24/7) or customerservice@intusurg.com Quantity in Commerce 162 Distribution US: MS, PA, WI, KY, TN, MI, FL, GA, NC, NV, LA, OH, IL, MO, IN, CA, AK, MD, TX, IA, MA, NE, AZ, WA, CO, AL, SC Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = NAY - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 09/20/22, correction notices were mailed to customers.***Customers were advised of the following:***1) The system software version can be located on the Vision Cart Touchscreen in the Troubleshooting Tab.***2) Ensure all device users read and understand this notice.***3) Notice should be passed on to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred.***4) Ensure proper instrument connections to the generator(s) and continue to follow all existing warnings and cautions found in user manuals 551400-15, USER MANUAL,SYS,P10,US,IS4000, 553873-07, USER MANUAL,I-A,MDR P10,IS4000-IS4200 and 554070-04, ADD,I-A,FORCE BIPOLAR,IS4000,IS4200.***5) Complete and return the acknowledgement form.***6) The recalling firm will follow-up with customers and schedule a software upgrade. ***7) After the upgrade, inform all personnel that the software update/upgrade was completed.***8) If you have experienced this issue or require additional assistance, contact Technical Support: 800-876-1310, Option 1 (24/7) or customerservice@intusurg.com | SoftwareDesignChange
43 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=195576 | Z-1563-2022 | DigitalDiagnost C50 1.1 | DigitalDiagnost C50 | 2 | CDRH | 2022-08-31 | 2022-09-01 21:06 |  | Philips North America Llc | Following a system restart where the Table Height 2 (TH2) is the default setting and then switched to Wallstand VS2, operators will see an incorrect orientation of the image on the first examination due to an issue in the firmware of the Wallstand VS2 board. The system will rotate the amplimat field selection by 90 degrees. The wrong amplimat field selection may cause an incorrect dose of radiation to occur. Additionally, the anatomic position markers may become mispositioned and could potentially be associated with the opposite side of the anatomy. ***(Added 1/3/2023), Expansion of the root cause to include incorrect amplimat field selection by 90 degrees following the first exposure using Wallstand VS2 and a subsequent system restart. Previously, this issue only occurred following a system restart where the Table Height 2 (TH2) is the default setting. | Class 2 Device Recall DigitalDiagnost C50 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall DigitalDiagnost C50 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall DigitalDiagnost C50 Date Initiated by Firm February 25, 2022 Date Posted September 01, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-1563-2022 Recall Event ID 90728 510(K)Number K163410 K201725 Product Classification System, x-ray, stationary - Product Code KPR Product DigitalDiagnost C50 1.1 Code Information Model Number 712204 Serial Number 210182 210187 220002 210049 210051 210097 210052 210053 210055 220020 Recalling Firm/ Manufacturer Philips North America Llc 222 Jacobs St Cambridge MA 02141-2289 For Additional Information Contact Dusty Leppert, CHP 978-228-0190 Manufacturer Reason for Recall Following a system restart where the Table Height 2 (TH2) is the default setting and then switched to Wallstand VS2, operators will see an incorrect orientation of the image on the first examination due to an issue in the firmware of the Wallstand VS2 board. The system will rotate the amplimat field selection by 90 degrees. The wrong amplimat field selection may cause an incorrect dose of radiation to occur. Additionally, the anatomic position markers may become mispositioned and could potentially be associated with the opposite side of the anatomy. ***(Added 1/3/2023), Expansion of the root cause to include incorrect amplimat field selection by 90 degrees following the first exposure using Wallstand VS2 and a subsequent system restart. Previously, this issue only occurred following a system restart where the Table Height 2 (TH2) is the default setting. FDA Determined Cause 2 Software Design Change Action Philips notified customers of the issue via an Electronic Product Radiation Defect notification letter dated July 12, 2022. Philips will schedule an appointment with customers to install the software update at no cost to customers. For questions contact Philips Customer Care Solutions Center 1-800-722-9377. Quantity in Commerce 11 systems (updated 1/3/2023) Distribution Nationwide Distribution including Puerto Rico Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = KPR 510(K)s with Product Code = KPR - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips notified customers of the issue via an Electronic Product Radiation Defect notification letter dated July 12, 2022. Philips will schedule an appointment with customers to install the software update at no cost to customers. For questions contact Philips Customer Care Solutions Center 1-800-722-9377. | SoftwareDesignChange
44 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194648 | Z-1513-2022 | Xhibit Telemetry Receiver, Model: 96280 | Xhibit Telemetry Receiver, Model: 96280 | 2 | CDRH | 2022-08-11 |  |  | Spacelabs Healthcare, Inc. | Software related alarm escalation defect occurs after approximately 25 days of continuous use. Technical alarm escalation for the following alarms, signal loss, all leads off, low battery, SpO2 sensor off, and signal interference, does not occur in the telemetry receiver as specified. Caregiver may not respond to the initial technical alarm and may not be aware of the continuing technical failure. | Class 2 Device Recall Xhibit Telemetry Receiver, Model: 96280 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Xhibit Telemetry Receiver, Model: 96280 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Xhibit Telemetry Receiver, Model: 96280 Date Initiated by Firm May 31, 2022 Create Date August 11, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-1513-2022 Recall Event ID 90529 510(K)Number K141156 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product Xhibit Telemetry Receiver, Model: 96280 Code Information UDI-DI: 10841522100246, Software version 1.4.0.***Serial Numbers:***6280-001538,***6280-000590,***6280-000589,***6280-000339,***9280-000711,***6280-001103,***6280-001102,***6280-001469,***6280-001467,***6280-001686,***6280-001399,***6280-001323,***6280-000457,***6280-000456,***6280-000455,***6280-000440,***6280-000439,***6280-000438,***6280-000429,***6280-000476,***6280-000475,***6280-001693,***6280-000703,***6280-000702,***6280-000701,***6280-000700,***6280-000314,***6280-000258,***6280-000157,***6280-001223,***6280-001222,***6280-001221,***6280-001220,***6280-001219,***6280-001218,***6280-001217,***6280-001216,***6280-001215,***6280-001214,***6280-001207,***6280-001127,***6280-001125,***6280-001124,***6280-001123,***6280-001056,***6280-001055,***6280-001054,***6280-001053,***6280-001052,***6280-001051,***6280-001050,***6280-001288,***6280-001287,***6280-001490,***6280-001692,***6280-000588,***6280-000517,***6280-000516,***6280-000515,***6280-001668,***6280-000206,***6280-000162,***6280-000210,***6280-000663,***6280-000377,***6280-001589,***6280-001385,***6280-001384,***6280-001382,***6280-001405,***6280-000594,***6280-000593,***6280-000592,***6280-000596,***6280-000274,***6280-001844,***6280-001843,***6280-001721,***6280-001720,***6280-000136,***6280-001784,***6280-001783,***6280-000359,***6280-000340,***6280-000336,***6280-001718,***6280-001706,***6280-001705,***6280-001704,***6280-001703,***6280-001702,***6280-001699,***6280-001698,***6280-001672,***6280-001669,***6280-000622,***6280-001797,***6280-000097,***6280-001558,***6280-001551,***6280-001649,***6280-001648,***6280-001647,***6280-001646,***6280-001645,***6280-001594,***6280-000093,***6280-001168,***6280-001531,***6280-001670,***6280-000219,***6280-000550,***6280-000878,***6280-000877,***6280-000872,***6280-000862,***6280-000858,***6280-000857,***6280-000221,***6280-000089,***6280-000204,***6280-000212,***6280-000510,***6280-000509,***6280-000209,***6280-000320,***6280-000538,***6280-000528,***6280-000205,***6280-000361,***6280-000360,***6280-000223,***6280-000963,***6280-000529,***6280-001575,***6280-001574,***6280-001573,***6280-001572,***6280-001571,***6280-001570,***6280-000467,***6280-000466,***6280-000465,***6280-000464,***6280-000463,***6280-000462,***6280-000461,***6280-000460,***6280-000459,***6280-000458,***6280-000222,***6280-000264,***6280-000275,***6280-000354,***6280-001677,***6280-001048,***6280-001043,***6280-001041,***6280-001038,***6280-001037,***6280-001036,***6280-001023,***6280-001022,***6280-001016,***6280-000787,***6280-001418,***6280-001419,***6280-001416,***6280-000894,***6280-000866,***6280-000865,***6280-000863,***6280-000902,***6280-000898,***6280-000724,***6280-000420,***6280-000419,***6280-000798,***6280-000325,***6280-000532,***6280-000521,***6280-000213,***6280-000254,***6280-000856,***6280-000855,***6280-000854,***6280-000853,***6280-000852,***6280-000851,***6280-001143,***6280-001142,***6280-001141,***6280-000472,***6280-000375,***6280-000368,***6280-000421,***6280-000519,***6280-001213,***6280-000836,***6280-000835,***6280-000834,***6280-000833,***6280-000832,***6280-000831,***6280-000586,***6280-001473,***6280-000319,***6280-000263,***6280-000312,***6280-000579,***6280-000092,***6280-000090,***6280-000084,***6280-000083,***6280-000079,***6280-001380,***6280-001466,***6280-001458,***6280-001611,***6280-001592,***6280-001591,***6280-000388,***6280-000161,***6280-000160,***6280-000159,***6280-000158,***6280-000066,***6280-000053,***6280-000174,***6280-000173,***6280-000172,***6280-000170,***6280-000168,***6280-000167,***6280-000077,***6280-000598,***6280-000355,***6280-000344,***6280-000341,***6280-000338,***6280-000468,***6280-000423,***6280-000478,***6280-000477,***6280-000454,***6280-000428,***6280-000261,***6280-001614,***6280-001642,***6280-001641,***6280-001640,***6280-001639,***6280-001638,***6280-001637,***6280-001636,***6280-001635,***6280-001634,***6280-001633,***6280-001632,***6280-001631,***6280-001630,***6280-001629,***6280-001913,***6280-000138,***6280-000437,***6280-000436,***6280-000177,***6280-000176,***6280-000337,***6280-001857,***6280-001853,***6280-001749,***6280-001359,***6280-001160,***6280-001159,***6280-001158,***6280-001157,***6280-001156,***6280-001154,***6280-001132,***6280-001112,***6280-001109,***6280-000991,***6280-000983,***6280-001189,***6280-001188,***6280-001205,***6280-001204,***6280-001203,***6280-001202,***6280-001200,***6280-001194,***6280-001193,***6280-001184,***6280-001166,***6280-000358,***6280-000357,***6280-000345,***6280-000432,***6280-000431,***6280-000811,***6280-000693,***6280-000690,***6280-000929,***6280-001063,***6280-001062,***6280-001061,***6280-001060,***6280-000391,***6280-001273,***6280-001554,***6280-001553,***6280-001261,***6280-000426,***6280-001525,***6280-000229,***6280-000876,***6280-000875,***6280-000745,***6280-000208,***6280-000574,***6280-000218,***6280-001087,***6280-000045,***6280-000041,***6280-000374,***6280-000373,***6280-001257,***6280-001256,***6280-001255,***6280-001254,***6280-001253,***6280-001252,***6280-001251,***6280-001250,***6280-001249,***6280-000840,***6280-000486,***6280-001825,***6280-000543,***6280-000542,***6280-000839,***6280-000879,***6280-001325,***6280-001071,***6280-001070,***6280-001069,***6280-000900,***6280-000897,***6280-001822,***6280-001874,***6280-001872,***6280-001871,***6280-000695,***6280-000540,***6280-000541,***6280-000539,***6280-000166,***6280-000624,***6280-000614,***6280-000613,***6280-000612,***6280-000611,***6280-000610,***6280-000609,***6280-000608,***6280-000607,***6280-000630,***6280-000629,***6280-000628,***6280-000627,***6280-000626,***6280-001644,***6280-001619,***6280-001617,***6280-000587,***6280-000490,***6280-000489,***6280-000488,***6280-000487,***6280-000974,***6280-000973,***6280-000972,***6280-000971,***6280-000970,***6280-000969,***6280-000968,***6280-000967,***6280-000966,***6280-000965,***6280-000329,***6280-000326,***6280-000318,***6280-001621,***6280-001609,***6280-001596,***6280-001536,***6280-001381,***6280-001340,***6280-001339,***6280-001338,***6280-001337,***6280-001336,***6280-001335,***6280-001334,***6280-001315,***6280-001312,***6280-001311,***6280-000547,***6280-000546,***6280-000545,***6280-000544,***6280-000518,***6280-001764,***6280-001752,***6280-001751,***6280-001746,***6280-001548,***6280-001408,***6280-001431,***6280-001729,***6280-001046,***6280-000601,***6280-001350,***6280-000689,***6280-001487,***6280-001475,***6280-000928,***6280-000927,***6280-000926,***6280-000925,***6280-000901,***6280-001329,***6280-001331,***6280-001330,***6280-001342,***6280-001341,***6280-000514,***6280-000513,***6280-000512,***6280-000511,***6280-000530,***6280-000108,***6280-000107,***6280-000106,***6280-000105,***6280-000104,***6280-000103,***6280-000102,***6280-000101,***6280-000100,***6280-000600,***6280-000276,***6280-001374,***6280-000195,***6280-000192,***6280-000189,***6280-000187,***6280-000178,***6280-000194,***6280-001820,***6280-001805,***6280-001444,***6280-001225,***6280-000850,***6280-000725,***6280-000694,***6280-000553,***6280-000549,***6280-000548,***6280-000330,***6280-000317,***6280-000188,***6280-000182,***6280-000290,***6280-001818,***6280-001765,***6280-001602,***6280-001601,***6280-001450,***6280-001449,***6280-001377,***6280-001376,***6280-001178,***6280-001049,***6280-000796,***6280-000743,***6280-000491,***6280-000356,***6280-000295,***6280-000293,***6280-000292,***6280-000291,***6280-000231,***6280-000196,***6280-000061,***6280-000060,***6280-000473,***6280-000232,***6280-000062,***6280-001319,***6280-001313,***6280-000384,***6280-000199,***6280-000198,***6280-000065,***6280-000064,***6280-000237,***6280-000406,***6280-000401,***6280-000362,***6280-000270,***6280-000400,***6280-001245,***6280-001101,***6280-000873,***6280-000871,***6280-000870,***6280-000861,***6280-000860,***6280-000815,***6280-000554,***6280-000527,***6280-000499,***6280-000416,***6280-000371,***6280-000197,***6280-000523,***6280-000063 Recalling Firm/ Manufacturer Spacelabs Healthcare, Inc. 35301 Se Center St Snoqualmie WA 98065-9216 For Additional Information Contact Wendy Machleid 425-396-3300 Manufacturer Reason for Recall Software related alarm escalation defect occurs after approximately 25 days of continuous use. Technical alarm escalation for the following alarms, signal loss, all leads off, low battery, SpO2 sensor off, and signal interference, does not occur in the telemetry receiver as specified. Caregiver may not respond to the initial technical alarm and may not be aware of the continuing technical failure. FDA Determined Cause 2 Software Design Change Action On 05/31/22, correction notices were mailed to customers who were asked to take the following actions that can be taken to minimize risk to patients or users:***Weigh the benefits versus the risks when deciding whether to continue to use the alarm escalation feature within XTR 1.4. ******Suggestions from Spacelabs to mitigate the risk:***- Set the priority to the same level as the escalation max on the product. This will ensure that the highest priority alarm is triggered for the technical alarm.***- Do not run the XTR 1.4 for more than 24 days consecutively. Restarting the XTR after 24 days will mitigate the issue for another 24 days.***- The recalling firm is working on a 1.4.1 release of XTR which will address this defect.******For additional information or technical assistance, please contact: Technical Support, 1-800-522-7025 and select 2 Quantity in Commerce 529 Distribution Worldwide distribution - US Nationwide distribution in the states of NY, AR, CO, FL, NJ, AL, KY, TN, SC, WA, OK, MN, PA, GA, VA, CA, OH, IA, KS, MI, IN, TX, MT, MO, ME, LA, IL, SD, MA, OR, ID, MD, AZ, AK and the countries of MB, ON, FR, QC, BC, IS, CT, RM. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 05/31/22, correction notices were mailed to customers who were asked to take the following actions that can be taken to minimize risk to patients or users:***Weigh the benefits versus the risks when deciding whether to continue to use the alarm escalation feature within XTR 1.4. ******Suggestions from Spacelabs to mitigate the risk:***- Set the priority to the same level as the escalation max on the product. This will ensure that the highest priority alarm is triggered for the technical alarm.***- Do not run the XTR 1.4 for more than 24 days consecutively. Restarting the XTR after 24 days will mitigate the issue for another 24 days.***- The recalling firm is working on a 1.4.1 release of XTR which will address this defect.******For additional information or technical assistance, please contact: Technical Support, 1-800-522-7025 and select 2 | SoftwareDesignChange
45 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194707 | Z-1469-2022 | Sight OLO, Automated Hematology Analyzer, Catalog No. OLO-U1 | Sight OLO | 2 | CDRH | 2022-07-27 |  |  | SIGHT DIAGNOSTICS LTD | The default reference ranges for MCH, MCHC, and RDW for age group "Child 2-11y" are mixed up and incorrectly displayed on OLO devices running software version 2.63, 2.63.1, and 2.63.1.1. Due to the issue, results may incorrectly appear to be within or outside the reference range. | Class 2 Device Recall Sight OLO Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Sight OLO FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Sight OLO Date Initiated by Firm June 14, 2022 Create Date July 27, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-1469-2022 Recall Event ID 90533 510(K)Number K190898 Product Classification Counter, differential cell - Product Code GKZ Product Sight OLO, Automated Hematology Analyzer, Catalog No. OLO-U1 Code Information UDI-DI: B787OLOU10***Software versions 2.63, 2.63.1, and 2.63.1.1******Serial Number:***Q245***Q516***Q560***Q648***Q684***Q688***Q792***Q794***Q840***Q600***Q632***Q883***Q890***Q209***Q789***Q489***Q496***Q287***Q221***Q256***Q523***Q575***Q783***Q801***Q802***Q812***Q813***Q817***Q831***Q835***Q850***Q870***Q871***Q892***Q775***Q873***Q322***Q836***Q846***Q308***Q354***Q263***Q650***Q535***Q473***Q691***Q698***Q768***Q448***Q428***Q545***Q780***Q379***Q404***Q412***Q377***Q284***Q234***Q549***Q744***Q755***Q762***Q617***Q218***Q385***Q848***Q852***Q877***Q677***Q872***Q880***Q770***Q236***Q924***Q413***Q415***Q313***Q324***Q175***Q572***Q484***Q709***Q568***Q776***Q285***Q243***Q246***Q262***Q267***Q282***Q289***Q296***Q298***Q159***Q191***Q216***Q261***Q272***Q293***Q305***Q326***Q339***Q376***Q382***Q386***Q395***Q406***Q410***Q421***Q435***Q441***Q445***Q455***Q457***Q458***Q461***Q466***Q475***Q480***Q481***Q483***Q490***Q493***Q504***Q508***Q514***Q528***Q529***Q530***Q531***Q533***Q539***Q540***Q544***Q546***Q554***Q558***Q563***Q566***Q567***Q576***Q578***Q582***Q584***Q588***Q591***Q610***Q612***Q615***Q618***Q621***Q631***Q635***Q637***Q642***Q643***Q645***Q646***Q649***Q651***Q652***Q653***Q654***Q655***Q656***Q660***Q662***Q665***Q666***Q672***Q675***Q683***Q685***Q689***Q690***Q693***Q694***Q696***Q699***Q703***Q707***Q710***Q711***Q714***Q718***Q719***Q721***Q723***Q724***Q725***Q727***Q728***Q733***Q735***Q736***Q737***Q738***Q739***Q740***Q742***Q745***Q747***Q748***Q749***Q752***Q753***Q760***Q764***Q766***Q767***Q769***Q771***Q777***Q784***Q785***Q787***Q790***Q793***Q795***Q797***Q799***Q803***Q804***Q806***Q810***Q815***Q402***Q407***Q454***Q670***Q702***Q713***Q715***Q717***Q720***Q726***Q741***Q743***Q754***Q772***Q773***Q774***Q781***Q782***Q786***Q796***Q849***Q888***Q897***Q900***Q904***Q905***Q907***Q908***Q912***Q917***Q926***Q929***Q930***Q931***Q932***Q933***Q934***Q938***Q940***Q941***Q943***Q944***Q945***Q949***Q950***Q951***Q952***Q954***Q955***Q958***Q176***Q625***Q443***Q623***Q805***Q494***Q894***Q127***Q394***Q679***Q708***Q746***Q759***Q100***Q167***Q264***Q273***Q403***Q418***Q470***Q570***Q577***Q581***Q765***Q417***Q704***Q667***Q664***Q527***Q452***Q138***Q757***Q761***Q808***Q820***Q823***Q824***Q825***Q826***Q829***Q462***Q400***Q387***Q303***Q659***Q830***Q620***Q859***Q398***Q463***Q673***Q918***Q638***Q552***Q732***Q145***Q436***Q807***Q543***Q882***Q534***Q634***Q821***Q541***Q833***Q837***Q319***Q886***Q898***Q614***Q451***Q705***Q669***Q384***Q365***Q798***Q270***Q121***Q788***Q816***Q465***Q467***Q499***Q551***Q565***Q352***Q283***Q661***Q515***Q668***Q212***Q520***Q633***Q391***Q791 Recalling Firm/ Manufacturer SIGHT DIAGNOSTICS LTD 23, Derekh Menakhem Begin Tel Aviv-Yafo Israel Manufacturer Reason for Recall The default reference ranges for MCH, MCHC, and RDW for age group "Child 2-11y" are mixed up and incorrectly displayed on OLO devices running software version 2.63, 2.63.1, and 2.63.1.1. Due to the issue, results may incorrectly appear to be within or outside the reference range. FDA Determined Cause 2 Software Design Change Action On June 14, 2022, the firm notified customers via email of the recall.******Customers were informed that a software correction for this issue will be included in any software version ABOVE 2.63.1.1.******Pending the software correction through a version upgrade, the firm provided instructions to manually configure the correct default reference ranges on the device, which are also included in Section 2.5.4 of the Sight OLO Operator's Manual.******Alternatively, you can contact support@sightdx.com and coordinate a time for Sight's technical support team to remotely update your device software. Quantity in Commerce 72 US; 292 OUS Distribution Worldwide distribution - US Nationwide distribution in the states of Alabama, Arizona, ***California, Florida, Idaho, Illinois, Kansas, Kentucky, Massachusetts, Michigan, Minnesota, Mississippi, New Jersey, Ohio, Pennsylvania, South Carolina, Tennessee,***Texas, Virginia, Washington, Wisconsin and the countries of Burkina Faso, Belize, Canada, Cote D'Ivoire, Cameroon, Curacao, Germany, Spain, Finland, United Kingdom, Greece, Hong Kong, Indonesia, Ireland, Israel, India, Italy, Mexico, Portugal, Qatar, Sweden, Slovenia, Chad, Thailand, South Africa. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = GKZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On June 14, 2022, the firm notified customers via email of the recall.******Customers were informed that a software correction for this issue will be included in any software version ABOVE 2.63.1.1.******Pending the software correction through a version upgrade, the firm provided instructions to manually configure the correct default reference ranges on the device, which are also included in Section 2.5.4 of the Sight OLO Operator's Manual.******Alternatively, you can contact support@sightdx.com and coordinate a time for Sight's technical support team to remotely update your device software. | SoftwareDesignChange
46 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194130 | Z-1306-2022 | NaviCare Nurse Call/Voalte Nurse Call | NaviCare Nurse Call/ Voalte Nurse Call System | 2 | CDRH | 2022-06-30 09:03 |  |  | Baxter Healthcare Corporation | An issue has been identified with Phillips (Emergin) and Longleaf non relay (Connexall, Vocera, Cerner) wireless integrations used with NaviCare/Voalte Nurse Call, software versions 3.9.100 through 3.9.300. Calls placed from a push button call device, such as a push button switch, call cord, or from the push buttons on a room audio station, will be canceled on the nurse call system when the call is answered at the wireless phone, regardless of the call priority. | Class 2 Device Recall NaviCare Nurse Call/ Voalte Nurse Call System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall NaviCare Nurse Call/ Voalte Nurse Call System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall NaviCare Nurse Call/ Voalte Nurse Call System Date Initiated by Firm May 31, 2022 Create Date June 30, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-1306-2022 Recall Event ID 90322 Product Classification System, communication, powered - Product Code ILQ Product NaviCare Nurse Call/Voalte Nurse Call Code Information software versions: 3.9.100, 3.9.101,3.9.102, 3.9.200, 3.9.201, and 3.9.300 Recalling Firm/ Manufacturer Baxter Healthcare Corporation 1 Baxter Pkwy Deerfield IL 60015-4625 For Additional Information Contact Center for One Baxter 800-422-9837 Manufacturer Reason for Recall An issue has been identified with Phillips (Emergin) and Longleaf non relay (Connexall, Vocera, Cerner) wireless integrations used with NaviCare/Voalte Nurse Call, software versions 3.9.100 through 3.9.300. Calls placed from a push button call device, such as a push button switch, call cord, or from the push buttons on a room audio station, will be canceled on the nurse call system when the call is answered at the wireless phone, regardless of the call priority. FDA Determined Cause 2 Software Design Change Action The firm disseminated an Urgent Medical Device Correction Notice beginning on 05/31/2022 by letter. The letter described the problem and requested the following actions be taken by the customer: "Identify if your Nurse Call System utilizes one of the impacted wireless integrations identified. If so, please share the information with potential users in your organization. Hillrom recommends reviewing the wireless system settings with your wireless provider***to determine if there are configurable options for call acceptance only.***Complete the enclosed response form and return to***hillromCARY021@sedgwick.com within two weeks of receipt of this notice, to confirm***you have received this communication and understand the potential issue." ***Persons with questions were directed as follows: "If you have any questions regarding this Urgent Medical Device Correction, please contact Hillrom Technical Support, available twenty-four hours a day, seven days a week, at 1-800-445-3720, option 3, or hrccwssupport@hillrom.com" Quantity in Commerce 283 installations Distribution nationwide and Canada Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm disseminated an Urgent Medical Device Correction Notice beginning on 05/31/2022 by letter. The letter described the problem and requested the following actions be taken by the customer: "Identify if your Nurse Call System utilizes one of the impacted wireless integrations identified. If so, please share the information with potential users in your organization. Hillrom recommends reviewing the wireless system settings with your wireless provider***to determine if there are configurable options for call acceptance only.***Complete the enclosed response form and return to***hillromCARY021@sedgwick.com within two weeks of receipt of this notice, to confirm***you have received this communication and understand the potential issue." ***Persons with questions were directed as follows: "If you have any questions regarding this Urgent Medical Device Correction, please contact Hillrom Technical Support, available twenty-four hours a day, seven days a week, at 1-800-445-3720, option 3, or hrccwssupport@hillrom.com" | SoftwareDesignChange
47 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193661 | Z-1243-2022 | Sensis, Programmable Diagnostic Computer, Model Nos.***6623974 6634633***6634658 6648153***6648161 10140973***10764561 10765502***10910620*** | Sensis recording system | 2 | CDRH | 2022-06-06 13:58 | 2022-06-06 21:07 |  | Siemens Medical Solutions USA, Inc | Siemens has become aware of three potential software issues with AXIOM Sensis or Sensis/ Sensis Lite systems. This may lead to a hazardous situation for patients if treatment cannot be continued on the system and treatment needs to be continued on an alternate system. | Class 2 Device Recall Sensis recording system Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Sensis recording system FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Sensis recording system Date Initiated by Firm April 13, 2022 Date Posted June 06, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-1243-2022 Recall Event ID 90175 510(K)Number K150493 Product Classification Computer, diagnostic, programmable - Product Code DQK Product Sensis, Programmable Diagnostic Computer, Model Nos.***6623974 6634633******6634658 6648153******6648161 10140973******10764561 10765502******10910620 Code Information UDI-DI: 04056869010137******Material***Serial***6623974***40108***6623974***41202***6623974***41017***6623974***42308***6623974***6101***6623974***41021***6623974***6110***6623974***6622***6623974***6727***6634633***3021***6634633***2272***6634633***2563***6634633***2589***6634633***20380***6634633***3112***6634633***2291***6634633***3405***6634633***3406***6634633***3440***6634633***3407***6634633***3408***6634633***3437***6634633***22267***6634633***21334***6634633***3417***6634633***3564***6634633***2060***6634633***2154***6634633***2155***6634633***2201***6634633***3037***6634633***3038***6634633***3041***6634633***3166***6634633***3224***6634633***3226***6634633***3352***6634633***3534***6634633***3536***6634633***20102***6634633***20134***6634633***20137***6634633***20149***6634633***20314***6634633***22688***6634633***2520***6634633***2521***6634633***2506***6634633***2709***6634633***2739***6634633***3072***6634633***21212***6634633***2552***6634633***2744***6634633***3060***6634633***20062***6634633***20574***6634633***2442***6634633***2444***6634633***2500***6634633***2501***6634633***2025***6634633***2276***6634633***2408***6634633***3622***6634633***3650***6634633***3653***6634633***3670***6634633***3671***6634633***3676***6634633***2578***6634633***22427***6634633***20448***6634633***21034***6634633***21035***6634633***21043***6634633***22336***6634633***21037***6634633***21051***6634633***21052***6634633***21095***6634633***21100***6634633***21102***6634633***21438***6634633***22022***6634633***2515***6634633***3390***6634633***2282***6634633***2331***6634633***2332***6634633***20321***6634633***20045***6634633***20047***6634633***3333***6634633***20598***6634633***2203***6634633***3091***6634633***3097***6634633***3059***6634633***3472***6634633***22671***6634633***2434***6634633***20525***6634633***20556***6634633***20557***6634633***2022***6634633***2032***6634633***2629***6634633***20392***6634633***22406***6634633***20594***6634633***20637***6634633***22611***6634633***20162***6634633***22612***6634633***3319***6634633***21399***6634633***22037***6634633***21384***6634633***21395***6634633***21452***6634633***3475***6634633***21332***6634633***22345***6634633***22357***6634633***22366***6634633***2403***6634633***20407***6634633***20033***6634633***22509***6634633***3075***6634633***2225***6634633***20430***6634633***22282***6634633***2113***6634633***2186***6634633***2187***6634633***3049***6634633***3050***6634633***22461***6634633***22462***6634633***22415***6634633***22312***6634633***22663***6634633***22667***6634633***3069***6634633***3128***6634633***3129***6634633***20026***6634633***3316***6634633***22512***6634633***22438***6634633***2608***6634633***3529***6634633***3641***6634633***20538***6634633***21109***6634633***21383***6634633***22341***6634633***2350***6634633***2352***6634633***21297***6634633***22000***6634633***20133***6634633***2205***6634633***2224***6634633***2227***6634633***2229***6634633***21056***6634633***2293***6634633***22628***6634633***22651***6634633***22652***6634633***22670***6634633***20585***6634633***20023***6634633***20399***6634633***20385***6634633***2428***6634633***20561***6634633***22480***6634633***20319***6634633***20618***6634633***22523***6634633***22526***6634633***22528***6634633***22619***6634633***22623***6634633***22627***6634633***22646***6634633***22669***6634633***22681***6634633***22682***6634633***22720***6634633***20139***6634633***22051***6634633***3660***6634633***3101***6634658***8067***6634658***8068***6648153***16001***6648153***16000***6648153***10661***6648153***10883***6648153***16005***6648153***16004***6648153***10700***6648153***10666***6648153***10886***6648153***16100***6648153***10887***6648153***10704***6648153***10723***6648161***11053***6648161***18103***6648161***18104***6648161***11657***6648161***11882***6648161***11251***6648161***18101***6648161***11661***6648161***18105***6648161***18000***6648161***18001***6648161***18002***6648161***11855***6648161***11856***10140973***13240***10140973***13419***10140973***13450***10140973***13215***10140973***13218***10140973***13423***10140973***13272***10140973***13483***10140973***13479***10140973***13243***10140973***13461***10140973***13466***10140973***13467***10140973***13402***10140973***13212***10764561***61297***10764561***60130***10764561***61084***10764561***61108***10764561***60410***10764561***60138***10764561***60412***10764561***60110***10764561***61072***10764561***60147***10764561***60370***10764561***60372***10764561***60125***10764561***60128***10764561***60060***10764561***60116***10764561***60120***10764561***60127***10764561***60129***10764561***60134***10764561***60143***10764561***61083***10764561***61031***10764561***61330***10764561***60079***10764561***61061***10764561***61230***10764561***61232***10764561***61086***10764561***61207***10764561***60369***10764561***60308***10764561***60312***10764561***60040***10764561***61313***10764561***61128***10764561***61133***10764561***61138***10764561***60023***10764561***60024***10764561***60025***10764561***61225***10764561***61310***10764561***61143***10764561***60052***10764561***60468***10764561***60322***10764561***61201***10764561***61224***10764561***60474***10764561***60022***10764561***60089***10764561***61144***10764561***61246***10764561***61088***10764561***60166***10764561***60087***10764561***61241***10764561***60325***10764561***61299***10764561***61302***10764561***61305***10764561***61315***10764561***61309***10764561***61317***10764561***61318***10764561***61239***10764561***61320***10764561***61186***10764561***60062***10764561***61325***10764561***60436***10764561***60444***10764561***60453***10764561***60483***10764561***61064***10764561***61183***10764561***61009***10764561***61169***10764561***60058***10764561***60095***10764561***60311***10764561***61129***10764561***60395***10764561***61281***10764561***61014***10764561***60323***10764561***61008***10764561***61292***10764561***60020***10764561***61136***10764561***61141***10764561***61142***10764561***60359***10764561***61103***10764561***60171***10764561***61121***10764561***61219***10764561***61345***10764561***61198***10764561***60427***10764561***61124***10764561***61403***10764561***61048***10764561***61251***10764561***61253***10764561***61256***10764561***61266***10764561***60473***10764561***61002***10764561***61116***10764561***61394***10764561***61344***10764561***60111***10764561***60123***10764561***61081***10764561***61234***10764561***60347***10764561***60348***10764561***61151***10764561***60355***10764561***60167***10764561***61030***10764561***61089***10764561***60065***10764561***60480***10764561***60154***10764561***61071***10764561***60146***10765502***15134***10765502***15010***10765502***15105***10765502***15111***10765502***15122***10765502***15011***10765502***15121***10765502***15130***10765502***15129***10765502***15143***10765502***15144***10765502***15146***10765502***15001***10765502***15821***10765502***15116***10765502***15108***10765502***15117***10765502***15120***10765502***15112***10765502***15115***10765502***15124***10765502***15125***10765502***15132***10765502***15133***10765502***15013***10765502***15106***10765502***15150***10910620***14326***10910620***14049***10910620***14050***10910620***14058***10910620***14046***10910620***14112***10910620***14113***10910620***14118***10910620***14109***10910620***14014***10910620***14015***10910620***14117***10910620***14119***10910620***14108***10910620***14041***10910620***14308***10910620***14000***10910620***14012***10910620***14127***10910620***14103***10910620***14059***10910620***14304***10910620***14311***10910620***14106***10910620***14325***10910620***14017***10910620***14104***10910620***14004***10910620***14002***10910620***14313***10910620***14121***10910620***14301***10910620***14309***10910620***14305***10910620***14322***10910620***14105***10910620***14315***10910620***14007***10910620***14055***10910620***14110***10910620***14111***10910620***14320***10910620***14302***10910620***14009***10910620***14057*********Update 8/31/22 - the following units have been removed from the scope of this correction *********6634633***20314***6634633***20102***6634633***20149***6634633***2506***6634633***2709***6634633***2739***6634633***3072***6634633***21212***6634633***3622***6634633***3670***6634633***3676***6648161***18103***6648161***18104***10764561***61230***10764561***61232***6634633***2434***10764561***61088***10765502***15130***6634633***20594***6634633***20637***6634633***22611***6634633***3319***10765502***15146***10765502***15001***6634633***20033***6634633***22509***10140973***13272***6648161***18105***10764561***61234***6648153***16005***6634633***2203***10764561***60359***6634633***20133***6648153***16001***6648161***11053***10764561***61136***6648161***18101***6634633***21395***10765502***15013***10765502***15106***10910620***14117***10910620***14119***6634633***2442***6634633***2444***6634633***2500***6634633***2501***10764561***61031***6634633***3069*********Updated 11/9/22 - The following systems have been added to this correction*********6634633***3136***10140973***13253*********Updated 1/13/23 - Added the following system*********Material Serial***6623974***42320 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact Meredith Adams 610-219-4834 Manufacturer Reason for Recall Siemens has become aware of three potential software issues with AXIOM Sensis or Sensis/ Sensis Lite systems. This may lead to a hazardous situation for patients if treatment cannot be continued on the system and treatment needs to be continued on an alternate system. FDA Determined Cause 2 Software Design Change Action On April 13, 2022, the firm notified customers via an Urgent Medical Device Correction letter. Customers were informed of the software issues. the software in the affected systems will be updated to correct the issue. Siemens' service organization will contact customers to arrange a date to perform the corrective action. ******Please feel free to contact Siemens' service organization for an earlier appointment at 1-800-888-7436. Quantity in Commerce 443 Distribution Domestic distribution nationwide. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = DQK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On April 13, 2022, the firm notified customers via an Urgent Medical Device Correction letter. Customers were informed of the software issues. the software in the affected systems will be updated to correct the issue. Siemens' service organization will contact customers to arrange a date to perform the corrective action. ******Please feel free to contact Siemens' service organization for an earlier appointment at 1-800-888-7436. | SoftwareDesignChange
48 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193616 | Z-1223-2022 | NIM Vital Console 4.0 (P/N: NIM4CM01) NIM with Software version 1.3.2 | NIM Vital | 2 | CDRH | 2022-06-02 | 2022-06-02 21:07 |  | Medtronic Xomed, Inc. | Software anomaly was identified. | Class 2 Device Recall NIM Vital Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall NIM Vital FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall NIM Vital Date Initiated by Firm May 03, 2022 Date Posted June 02, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-1223-2022 Recall Event ID 90163 510(K)Number K200759 Product Classification Stimulator, electrical, evoked response - Product Code GWF Product NIM Vital Console 4.0 (P/N: NIM4CM01) NIM with Software version 1.3.2 Code Information Model Number: NIM4CM01; UDI/DI: 00763000395896, lot numbers: C2026207, C2026233, C2026320, C2026335, C2026354, C2026376, C2110046, C2110089, C2110131, C2110196, C2110198, C2110208, C2110217, C2110223 and 00763000002978, lot numbers: C1915048, C1915078, C1915099, C1915188, C2026207, C2026233, C2026320, C2026335, C2026354, C2026376, C2110046, C2110089, C2110131, C2110196, C2110198, C2110208, C2110217, C2110223. Recalling Firm/ Manufacturer Medtronic Xomed, Inc. 6743 Southpoint Dr N Jacksonville FL 32216-6218 For Additional Information Contact Marie Gentile 904-296-9600 Manufacturer Reason for Recall Software anomaly was identified. FDA Determined Cause 2 Software Design Change Action The firm issued a recall notification to consignees on 05/03/2022 via mail. The letter instructs the consignee to follow the IFU regarding the setup and "Lead Off Detected" error message, forward the notification to those who need to be aware within the organization and to any location which the devices have been distributed, and to complete the response form. ******An additional correction notification was sent on about 09/13/2023 to notify consignees of the update availability (NIM Vital" SW version v1.4.3) to fix this issue. Quantity in Commerce 18 units Distribution Worldwide distribution - US Nationwide distribution in the states of TX, CO, NJ and the country of France. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = GWF - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm issued a recall notification to consignees on 05/03/2022 via mail. The letter instructs the consignee to follow the IFU regarding the setup and "Lead Off Detected" error message, forward the notification to those who need to be aware within the organization and to any location which the devices have been distributed, and to complete the response form. ******An additional correction notification was sent on about 09/13/2023 to notify consignees of the update availability (NIM Vital" SW version v1.4.3) to fix this issue. | SoftwareDesignChange
49 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193652 | Z-1224-2022 | NIM Vital Patient Interface 4.0 (P/N: NIM4CPB1) NIM with Software version 1.3.2 | NIM Vital 4.0 Patient Interface | 2 | CDRH | 2022-06-02 | 2022-06-02 21:07 |  | Medtronic Xomed, Inc. | Software anomaly was identified. | Class 2 Device Recall NIM Vital 4.0 Patient Interface Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall NIM Vital 4.0 Patient Interface FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall NIM Vital 4.0 Patient Interface Date Initiated by Firm May 03, 2022 Date Posted June 02, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-1224-2022 Recall Event ID 90163 510(K)Number K200759 Product Classification Stimulator, electrical, evoked response - Product Code GWF Product NIM Vital Patient Interface 4.0 (P/N: NIM4CPB1) NIM with Software version 1.3.2 Code Information Model Number: NIM4CPB1; UDI/DI: 00763000002985, lot numbers: P1910256, P1910260, P1910402, P1910404, P2008073, P2008130, P2008131 and 00763000395902, lot numbers: P2026275, P2026290, P2026329, P2026330, P2026357, P2026388, P2026389, P2026680, P2026681, P2026718, P2026719, P2026793, P2026956, P2026957, P2026958, P2026959, P2026981, P2026994, P2026995, P2127005, P2127006, P2127013 Recalling Firm/ Manufacturer Medtronic Xomed, Inc. 6743 Southpoint Dr N Jacksonville FL 32216-6218 For Additional Information Contact Marie Gentile 904-296-9600 Manufacturer Reason for Recall Software anomaly was identified. FDA Determined Cause 2 Software Design Change Action The firm issued a recall notification to consignees on 05/03/2022 via mail. The letter instructs the consignee to follow the IFU regarding the setup and "Lead Off Detected" error message, forward the notification to those who need to be aware within the organization and to any location which the devices have been distributed, and to complete the response form. ******An additional correction notification was sent on about 09/13/2023 to notify consignees of the update availability (NIM Vital" SW version v1.4.3) to fix this issue. Quantity in Commerce 29 units Distribution Worldwide distribution - US Nationwide distribution in the states of TX, CO, NJ and the country of France. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = GWF - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm issued a recall notification to consignees on 05/03/2022 via mail. The letter instructs the consignee to follow the IFU regarding the setup and "Lead Off Detected" error message, forward the notification to those who need to be aware within the organization and to any location which the devices have been distributed, and to complete the response form. ******An additional correction notification was sent on about 09/13/2023 to notify consignees of the update availability (NIM Vital" SW version v1.4.3) to fix this issue. | SoftwareDesignChange
50 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193084 | Z-1074-2022 | 4 Series CX phototherapy units equipped with Daavlin's ClearLink Control System software versions v3.03h and 3.04v. M4 Ser CX 311-20 120-240V 50/60Hz, item # 807QS0020CX6. | 4 Series Phototherapy Unit | 2 | CDRH | 2022-05-12 |  |  | Daavlin Distributing Company | Software issue, resuming an interrupted treatment will result in swap of utilized calibration treatment distance 0" inches to "9. Resulting in over or under dosing. | Class 2 Device Recall 4 Series Phototherapy Unit Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall 4 Series Phototherapy Unit FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall 4 Series Phototherapy Unit Date Initiated by Firm April 26, 2022 Create Date May 12, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-1074-2022 Recall Event ID 89990 510(K)Number K111172 Product Classification Light, ultraviolet, dermatological - Product Code FTC Product 4 Series CX phototherapy units equipped with Daavlin's ClearLink Control System software versions v3.03h and 3.04v. M4 Ser CX 311-20 120-240V 50/60Hz, item # 807QS0020CX6. Code Information UDI-DI: +B2809807QS0020CX60L. Serial Number: QS11571PLNBCX0020 Recalling Firm/ Manufacturer Daavlin Distributing Company 205 W Bement St Bryan OH 43506-1264 For Additional Information Contact 800-322-8546 Manufacturer Reason for Recall Software issue, resuming an interrupted treatment will result in swap of utilized calibration treatment distance 0" inches to "9. Resulting in over or under dosing. FDA Determined Cause 2 Software Design Change Action Beginning April 26, 2022 Daavlin Distributing Co. is sent a recall notification letter via certified mail to customers. Customers are to acknowledge receipt of the notification and a company representative will call customers to arrange for return of the devices to update software. If you have any questions or concerns, contact the company at (800) 322-8546. Monday through Friday, 8:00 AM to 4:30 PM, EST. Quantity in Commerce 1 Distribution Worldwide distribution - US Nationwide distribution in the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, SD, TN, TX, VA, WA, WI, WV and the countries of Australia, Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = FTC - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Beginning April 26, 2022 Daavlin Distributing Co. is sent a recall notification letter via certified mail to customers. Customers are to acknowledge receipt of the notification and a company representative will call customers to arrange for return of the devices to update software. If you have any questions or concerns, contact the company at (800) 322-8546. Monday through Friday, 8:00 AM to 4:30 PM, EST. | SoftwareDesignChange
51 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193083 | Z-1073-2022 | 1 Series CX phototherapy units equipped with Daavlin's ClearLink Control System software versions v3.03h and 3.04v. Models: [1 Ser CX 305-4 120-240V 50/60Hz, item # 06OS0004CX6], ***[1 Ser CX 311-4 120-240V 50/60Hz CE, item # 807OS0004CX5], [1 Ser CX 311-4 120-240V 50/60Hz, item # 807OS0004CX6], [1 Ser CX 350-4 120-240V 50/60Hz, item # 807OS0400CX6], [1 Ser CX UVA1-4 120-240V 50/60Hz, item # 807OS0400CX6A1]************ | 1 Series Phototherapy Unit | 2 | CDRH | 2022-05-12 |  |  | Daavlin Distributing Company | Software issue, resuming an interrupted treatment will result in swap of utilized calibration treatment distance 0" inches to "9. Resulting in over or under dosing. | Class 2 Device Recall 1 Series Phototherapy Unit Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall 1 Series Phototherapy Unit FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall 1 Series Phototherapy Unit Date Initiated by Firm April 26, 2022 Create Date May 12, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-1073-2022 Recall Event ID 89990 510(K)Number K100378 Product Classification Light, ultraviolet, dermatological - Product Code FTC Product 1 Series CX phototherapy units equipped with Daavlin's ClearLink Control System software versions v3.03h and 3.04v. Models: [1 Ser CX 305-4 120-240V 50/60Hz, item # 06OS0004CX6], ******[1 Ser CX 311-4 120-240V 50/60Hz CE, item # 807OS0004CX5], [1 Ser CX 311-4 120-240V 50/60Hz, item # 807OS0004CX6], [1 Ser CX 350-4 120-240V 50/60Hz, item # 807OS0400CX6], [1 Ser CX UVA1-4 120-240V 50/60Hz, item # 807OS0400CX6A1] Code Information UDI-DI: +B2809806OS0004CX60K, +B2809807OS0004CX50K, +B2809807OS0004CX60L, +B2809807OS0400CX60L, and +B2809807OS0400CX6A10W. Serial Numbers: OS22183PLNBCX0004, OS22320PLNBCX0004, OS23725PLNBCX0004, OS24083PLNBCX0004, OS24089PLNBCX0004, OS24090PLNBCX0004, OS23532PLA1CX0400, OS22097PLNBCX0004, OS22200PLNBCX0004, OS22265PLNBCX0004, OS22270PLNBCX0004, OS22304PLNBCX0004, OS22739PLNBCX0004, OS22788PLNBCX0004, OS22795PLNBCX0004, OS23239PLNBCX0004, OS23609PLNBCX0004, OS23894PLNBCX0004, OS24017PLNBCX0004, OS24254PLNBCX0004, OS24810PLNBCX0004, OS21997PLNBCX0004, OS22060PLNBCX0004, OS22063PLNBCX0004, OS22273PLNBCX0004, OS22309PLNBCX0004, OS22637PLNBCX0004, OS22680PLNBCX0004, OS22736PLNBCX0004, OS22738PLNBCX0004, OS22770PLNBCX0004, OS22789PLNBCX0004, OS22796PLNBCX0004, OS22919PLNBCX0004, OS22977PLNBCX0004, OS22980PLNBCX0004, OS22983PLNBCX0004, OS23244PLNBCX0004, OS23252PLNBCX0004, OS23440PLNBCX0004, OS23681PLNBCX0004, OS23724PLNBCX0004, OS23762PLNBCX0004, OS23855PLNBCX0004, OS23856PLNBCX0004, OS24008PLNBCX0004, OS24077PLNBCX0004, OS24104PLNBCX0004, OS24107PLNBCX0004, OS24109PLNBCX0004, OS24113PLACX0400, OS24127PLNBCX0004, OS24308PLNBCX0004, OS24310PLNBCX0004, OS24376PLNBCX0004, OS24467PLNBCX0004, OS24478PLNBCX0004, OS24229PLACX0400, OS22663PLNBCX0004, OS23940PLNBCX0004, OS23461PLNBCX0004, OS24829PLNBCX0004, OS22154PLNBCX0004, OS22112PLNBCX0004, OS22537PLNBCX0004, OS22271PLNBCX0004, OS24078PLNBCX0004, OS22227PLNBCX0004, OS23274PLNBCX0004, OS23481PLNBCX0004, OS23525PLNBCX0004, OS23676PLNBCX0004, OS23711PLNBCX0004, OS23910PLNBCX0004, OS24296PLNBCX0004, OS24325PLNBCX0004, OS24345PLNBCX0004, OS24358PLNBCX0004, OS24828PLNBCX0004, OS22794PLNBCX0004, OS24479PLNBCX0004, OS24682PLNBCX0004, OS24760PLNBCX0004, OS24228PLNBCX0004, OS24857PLNBCX0004, OS22033PLNBCX0004, OS22293PLNBCX0004, OS23526PLNBCX0004, OS23866PLNBCX0004, OS23896PLNBCX0004, OS24500PLNBCX0004, OS22439PLNBCX0004, OS22641PLNBCX0004, OS22766PLNBCX0004, OS23502PLNBCX0004, OS23941PLNBCX0004, OS24110PLNBCX0004, OS22015PLNBCX0004, OS22452PLNBCX0004, OS22453PLNBCX0004, OS24020PLNBCX0004, OS24680PLNBCX0004, OS22090PLNBCX0004, OS22266PLNBCX0004, OS22442PLNBCX0004, OS22633PLNBCX0004, OS22640PLNBCX0004, OS22764PLNBCX0004, OS23233PLNBCX0004, OS23255PLNBCX0004, OS23496PLNBCX0004, OS23527PLNBCX0004, OS23528PLNBCX0004, OS23873PLNBCX0004, OS23911PLNBCX0004, OS23961PLNBCX0004, OS24082PLNBCX0004, OS24105PLNBCX0004, OS24132PLNBCX0004, OS24211PLNBCX0004, OS22716PLACX0400, OS22717PLBBCX0004, OS23353PLACX0400, OS23354PLACX0400, OS23352PLACX0400, OS23355PLBBCX0004, OS22103PLACX0400, OS22062PLNBCX0004, OS22258PLNBCX0004, OS22229PLNBCX0004, OS22818PLNBCX0004, OS22869PLA1CX0400, OS23253PLNBCX0004, OS24051PLNBCX0004, OS24281PLNBCX0004, OS24422PLNBCX0004, OS24423PLNBCX0004, OS23183PLNBCX0004, OS22013PLNBCX0004, OS22049PLNBCX0004, OS22075PLNBCX0004, OS22088PLNBCX0004, OS22092PLNBCX0004, OS22106PLNBCX0004, OS22263PLNBCX0004, OS22269PLNBCX0004, OS22308PLNBCX0004, OS22397PLNBCX0004, OS22440PLNBCX0004, OS22441PLNBCX0004, OS22502PLA1CX0400, OS22635PLNBCX0004, OS22643PLNBCX0004, OS22681PLNBCX0004, OS22741PLNBCX0004, OS22762PLNBCX0004, OS22763PLNBCX0004, OS22765PLNBCX0004, OS22769PLNBCX0004, OS22791PLNBCX0004, OS22812PLNBCX0004, OS22822PLNBCX0004, OS22917PLNBCX0004, OS22918PLNBCX0004, OS23049PLNBCX0004, OS23053PLNBCX0004, OS23061PLNBCX0004, OS23101PLNBCX0004, OS23237PLNBCX0004, OS23240PLNBCX0004, OS23241PLNBCX0004, OS23243PLNBCX0004, OS23246PLNBCX0004, OS23254PLNBCX0004, OS23272PLNBCX0004, OS23418PLNBCX0004, OS23463PLNBCX0004, OS23466PLNBCX0004, OS23470PLNBCX0004, OS23474PLA1CX0400, OS23480PLNBCX0004, OS23529PLNBCX0004, OS23530PLNBCX0004, OS23569PLNBCX0004, OS23620PLNBCX0004, OS23675PLNBCX0004, OS23756PLNBCX0004, OS23857PLNBCX0004, OS23867PLNBCX0004, OS23890PLNBCX0004, OS23898PLNBCX0004, OS23909PLNBCX0004, OS24106PLNBCX0004, OS24166PLNBCX0004, OS24168PLA1CX0400, OS24227PLNBCX0004, OS24312PLNBCX0004, OS24539PLNBCX0004, OS24630PLNBCX0004, OS24743PLNBCX0004, OS24789PLNBCX0004, OS24791PLNBCX0004, OS24811PLNBCX0004, OS23940PLNBCX0004, OS22777PLNBCX0004, OS22778PLNBCX0004, OS20337PLNBCX0004, OS22201PLNBCX0004, OS22303PLNBCX0004, OS22334PLNBCX0004, OS22612PLNBCX0004, OS22871PLNBCX0004, OS22873PLNBCX0004, OS23086PLNBCX0004, OS23189PLNBCX0004, OS24421PLNBCX0004. Recalling Firm/ Manufacturer Daavlin Distributing Company 205 W Bement St Bryan OH 43506-1264 For Additional Information Contact 800-322-8546 Manufacturer Reason for Recall Software issue, resuming an interrupted treatment will result in swap of utilized calibration treatment distance 0" inches to "9. Resulting in over or under dosing. FDA Determined Cause 2 Software Design Change Action Beginning April 26, 2022 Daavlin Distributing Co. is sent a recall notification letter via certified mail to customers. Customers are to acknowledge receipt of the notification and a company representative will call customers to arrange for return of the devices to update software. If you have any questions or concerns, contact the company at (800) 322-8546. Monday through Friday, 8:00 AM to 4:30 PM, EST. Quantity in Commerce 216 Distribution Worldwide distribution - US Nationwide distribution in the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, SD, TN, TX, VA, WA, WI, WV and the countries of Australia, Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = FTC - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Beginning April 26, 2022 Daavlin Distributing Co. is sent a recall notification letter via certified mail to customers. Customers are to acknowledge receipt of the notification and a company representative will call customers to arrange for return of the devices to update software. If you have any questions or concerns, contact the company at (800) 322-8546. Monday through Friday, 8:00 AM to 4:30 PM, EST. | SoftwareDesignChange
52 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192467 | Z-0998-2022 | MEDTRONIC INTERSTIM THERAPY SYSTEM FOR URINARY CONTROL, which includes the following: a) Handset with Communicator, Model Numbers: 1) TH90Q01, 2) TH90QFA, 3) TH90PFA, 4) TH90GFA, 5) TG90G01; b) InterStim X Clinician Therapy Application Software, Model Number A51300; and c) InterStim x "My Therapy" Patient Therapy Application Software, Model Number A52300 | MEDTRONIC INTERSTIM THERAPY SYSTEM FOR URINARY CONTROL | 2 | CDRH | 2022-04-22 | 2022-04-22 21:06 |  | Medtronic Neuromodulation | There is a software anomaly with the InterStim X Clinician software application with results in the data entered in the Patient Information fields not e retained and a " Data Lost" notification being displayed. | Class 2 Device Recall MEDTRONIC INTERSTIM THERAPY SYSTEM FOR URINARY CONTROL Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall MEDTRONIC INTERSTIM THERAPY SYSTEM FOR URINARY CONTROL FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall MEDTRONIC INTERSTIM THERAPY SYSTEM FOR URINARY CONTROL Date Initiated by Firm February 28, 2022 Date Posted April 22, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-0998-2022 Recall Event ID 89799 PMA Number P080025 P970004 Product Classification Stimulator, electrical, implantable, for incontinence - Product Code EZW Product MEDTRONIC INTERSTIM THERAPY SYSTEM FOR URINARY CONTROL, which includes the following: a) Handset with Communicator, Model Numbers: 1) TH90Q01, 2) TH90QFA, 3) TH90PFA, 4) TH90GFA, 5) TG90G01; b) InterStim X Clinician Therapy Application Software, Model Number A51300; and c) InterStim x "My Therapy" Patient Therapy Application Software, Model Number A52300 Code Information a) Model Numbers: ***1) TH90Q01, UDI-Device Identifier (GTIN/UPN) 00763000582470, Serial Numbers: NPF0000074, NPF0000075, NPF0000096, NPF0000097, NPF0000109, NPF0000197, NPF0000238, NPF0000263, NPF0000264, NPF0000270, NPF0000271, NPF0000284, NPF0000292, NPF0000379, NPF0000414, NPF0000422, NPF0000440, NPF0000458, NPF0000463, NPF0000480, NPF0000501, NPF0000515, NPF0000516, NPF0000517, NPF0000519, NPF0000537, NPF0000555, NPF0000618, NPF0000656, NPF0000691, NPF0000715, NPF0000784, NPF0000786, NPF0000828, NPF0000889, NPF0000908, NPF0000913, NPF0000915, NPF0000930, NPF0000931, NPF0001006, NPF0001045, NPF0001057, NPF0001058, NPF0001059, NPF0001060, NPF0001130, NPF0001135, NPF0001136, NPF0001154, NPF0001199, NPF0001267, NPF0001391, NPF0001403, NPF0001458, NPF0001485, NPF0001492, NPF0001493, NPF0001498, NPF0001502, NPF0001511, NPF0001620, NPF0001674, NPF0001713, NPF0001716, NPF0001758, NPF0001797, NPF0001824, NPF0001841, NPF0001868, NPF0001879, NPF0001893, NPF0001894, NPF0001900, NPF0001928, NPF0001931, NPF0002020, NPF0002025, NPF0002077, NPF0002078, NPF0002079, NPF0002102, NPF0002105, NPF0002117, NPF0002127, NPF0000060, NPF0000061, NPF0000063, NPF0000070, NPF0000099, NPF0000100, NPF0000116, NPF0000160, NPF0000291, NPF0000330, NPF0000465, NPF0000476, NPF0000477, NPF0000613, NPF0000614, NPF0000615, NPF0000617, NPF0000619, NPF0000654, NPF0000655, NPF0000690, NPF0000709, NPF0000711, NPF0000755, NPF0000756, NPF0000779, NPF0000804, NPF0000807, NPF0000823, NPF0000871, NPF0000880, NPF0000907, NPF0000977, NPF0000983, NPF0001000, NPF0001119, NPF0001145, NPF0001152, NPF0001158, NPF0001161, NPF0001188, NPF0001192, NPF0001208, NPF0001387, NPF0001456, NPF0001516, NPF0001554, NPF0001759, NPF0001781, NPF0002048, NPF0002076, NPF0002369, NPF0003037, NPF0003054, NPF0003386, NPF0000186, NPF0000222, NPF0000269, NPF0000289, NPF0000295, NPF0000296, NPF0000310, NPF0000319, NPF0000341, NPF0000766, NPF0001194, NPF0001209, NPF0001210, NPF0001213, NPF0001268, NPF0001278, NPF0001281, NPF0001291, NPF0001295, NPF0001319, NPF0001345, NPF0001357, NPF0001361, NPF0001362, NPF0001390, NPF0001420, NPF0001425, NPF0001555, NPF0001565, NPF0001618, NPF0001776, NPF0001898, NPF0001975, NPF0001984, NPF0002064, NPF0002080, NPF0002084, NPF0002085, NPF0002090, NPF0002144, NPF0002216, NPF0002229, NPF0002608, NPF0002622, NPF0002646, NPF0001820, NPF0001823, NPF0000064, NPF0000073, NPF0000085, NPF0000086, NPF0000088, NPF0000089, NPF0000091, NPF0000092, NPF0000095, NPF0000102, NPF0000103, NPF0000105, NPF0000107, NPF0000110, NPF0000117, NPF0000129, NPF0000130, NPF0000131, NPF0000150, NPF0000156, NPF0000158, NPF0000161, NPF0000172, NPF0000173, NPF0000174, NPF0000175, NPF0000176, NPF0000178, NPF0000179, NPF0000180, NPF0000181, NPF0000187, NPF0000192, NPF0000196, NPF0000198, NPF0000199, NPF0000203, NPF0000209, NPF0000210, NPF0000211, NPF0000212, NPF0000213, NPF0000214, NPF0000218, NPF0000221, NPF0000223, NPF0000234, NPF0000237, NPF0000244, NPF0000245, NPF0000246, NPF0000247, NPF0000248, NPF0000250, NPF0000256, NPF0000260, NPF0000266, NPF0000272, NPF0000277, NPF0000294, NPF0000298, NPF0000300, NPF0000301, NPF0000302, NPF0000303, NPF0000305, NPF0000309, NPF0000311, NPF0000312, NPF0000313, NPF0000320, NPF0000336, NPF0000340, NPF0000342, NPF0000345, NPF0000350, NPF0000353, NPF0000354, NPF0000357, NPF0000358, NPF0000364, NPF0000368, NPF0000369, NPF0000389, NPF0000393, NPF0000395, NPF0000397, NPF0000403, NPF0000404, NPF0000436, NPF0000438, NPF0000442, NPF0000447, NPF0000454, NPF0000459, NPF0000460, NPF0000461, NPF0000462, NPF0000471, NPF0000488, NPF0000497, NPF0000514, NPF0000518, NPF0000520, NPF0000521, NPF0000522, NPF0000524, NPF0000525, NPF0000526, NPF0000528, NPF0000530, NPF0000532, NPF0000539, NPF0000561, NPF0000564, NPF0000567, NPF0000568, NPF0000581, NPF0000587, NPF0000588, NPF0000606, NPF0000621, NPF0000623, NPF0000626, NPF0000638, NPF0000639, NPF0000643, NPF0000645, NPF0000650, NPF0000653, NPF0000657, NPF0000659, NPF0000660, NPF0000661, NPF0000676, NPF0000680, NPF0000683, NPF0000684, NPF0000686, NPF0000687, NPF0000689, NPF0000694, NPF0000704, NPF0000707, NPF0000712, NPF0000713, NPF0000714, NPF0000717, NPF0000720, NPF0000724, NPF0000728, NPF0000732, NPF0000737, NPF0000748, NPF0000750, NPF0000754, NPF0000758, NPF0000759, NPF0000761, NPF0000762, NPF0000763, NPF0000772, NPF0000775, NPF0000781, NPF0000783, NPF0000789, NPF0000794, NPF0000798, NPF0000805, NPF0000808, NPF0000809, NPF0000811, NPF0000812, NPF0000822, NPF0000824, NPF0000825, NPF0000826, NPF0000829, NPF0000830, NPF0000831, NPF0000832, NPF0000838, NPF0000840, NPF0000844, NPF0000846, NPF0000849, NPF0000863, NPF0000864, NPF0000865, NPF0000866, NPF0000868, NPF0000869, NPF0000870, NPF0000872, NPF0000882, NPF0000883, NPF0000886, NPF0000887, NPF0000890, NPF0000891, NPF0000892, NPF0000895, NPF0000903, NPF0000904, NPF0000906, NPF0000909, NPF0000911, NPF0000916, NPF0000918, NPF0000919, NPF0000920, NPF0000921, NPF0000923, NPF0000926, NPF0000928, NPF0000935, NPF0000948, NPF0000954, NPF0000959, NPF0000960, NPF0000962, NPF0000963, NPF0000965, NPF0000979, NPF0000980, NPF0000982, NPF0000993, NPF0000997, NPF0001001, NPF0001010, NPF0001018, NPF0001019, NPF0001024, NPF0001029, NPF0001031, NPF0001039, NPF0001040, NPF0001042, NPF0001044, NPF0001047, NPF0001049, NPF0001061, NPF0001062, NPF0001064, NPF0001067, NPF0001069, NPF0001070, NPF0001071, NPF0001072, NPF0001073, NPF0001074, NPF0001075, NPF0001076, NPF0001077, NPF0001078, NPF0001079, NPF0001080, NPF0001081, NPF0001083, NPF0001089, NPF0001090, NPF0001096, NPF0001099, NPF0001100, NPF0001105, NPF0001113, NPF0001114, NPF0001127, NPF0001131, NPF0001134, NPF0001148, NPF0001149, NPF0001150, NPF0001151, NPF0001156, NPF0001175, NPF0001178, NPF0001179, NPF0001180, NPF0001181, NPF0001182, NPF0001183, NPF0001184, NPF0001185, NPF0001187, NPF0001189, NPF0001190, NPF0001195, NPF0001205, NPF0001212, NPF0001240, NPF0001260, NPF0001264, NPF0001270, NPF0001273, NPF0001275, NPF0001277, NPF0001283, NPF0001289, NPF0001293, NPF0001294, NPF0001299, NPF0001316, NPF0001321, NPF0001327, NPF0001334, NPF0001339, NPF0001341, NPF0001342, NPF0001347, NPF0001359, NPF0001360, NPF0001366, NPF0001368, NPF0001370, NPF0001374, NPF0001378, NPF0001385, NPF0001393, NPF0001397, NPF0001398, NPF0001406, NPF0001407, NPF0001412, NPF0001413, NPF0001422, NPF0001423, NPF0001424, NPF0001426, NPF0001428, NPF0001429, NPF0001430, NPF0001431, NPF0001432, NPF0001433, NPF0001434, NPF0001436, NPF0001437, NPF0001445, NPF0001451, NPF0001452, NPF0001454, NPF0001455, NPF0001460, NPF0001461, NPF0001463, NPF0001464, NPF0001466, NPF0001467, NPF0001481, NPF0001484, NPF0001496, NPF0001499, NPF0001500, NPF0001503, NPF0001504, NPF0001508, NPF0001510, NPF0001513, NPF0001514, NPF0001520, NPF0001521, NPF0001522, NPF0001523, NPF0001526, NPF0001527, NPF0001528, NPF0001530, NPF0001532, NPF0001535, NPF0001536, NPF0001539, NPF0001540, NPF0001543, NPF0001550, NPF0001551, NPF0001556, NPF0001560, NPF0001562, NPF0001563, NPF0001570, NPF0001574, NPF0001576, NPF0001577, NPF0001578, NPF0001579, NPF0001585, NPF0001588, NPF0001604, NPF0001608, NPF0001621, NPF0001625, NPF0001626, NPF0001627, NPF0001628, NPF0001629, NPF0001633, NPF0001635, NPF0001636, NPF0001638, NPF0001645, NPF0001649, NPF0001650, NPF0001652, NPF0001653, NPF0001654, NPF0001656, NPF0001657, NPF0001668, NPF0001671, NPF0001672, NPF0001673, NPF0001681, NPF0001684, NPF0001690, NPF0001691, NPF0001692, NPF0001694, NPF0001698, NPF0001700, NPF0001703, NPF0001706, NPF0001707, NPF0001710, NPF0001714, NPF0001715, NPF0001718, NPF0001735, NPF0001741, NPF0001745, NPF0001747, NPF0001749, NPF0001750, NPF0001754, NPF0001755, NPF0001762, NPF0001763, NPF0001766, NPF0001767, NPF0001769, NPF0001775, NPF0001786, NPF0001789, NPF0001800, NPF0001803, NPF0001826, NPF0001837, NPF0001842, NPF0001843, NPF0001847, NPF0001857, NPF0001859, NPF0001876, NPF0001877, NPF0001880, NPF0001884, NPF0001886, NPF0001889, NPF0001892, NPF0001895, NPF0001906, NPF0001917, NPF0001929, NPF0001932, NPF0001934, NPF0001936, NPF0001937, NPF0001939, NPF0001940, NPF0001941, NPF0001942, NPF0001943, NPF0001946, NPF0001948, NPF0001949, NPF0001950, NPF0001951, NPF0001952, NPF0001954, NPF0001956, NPF0001957, NPF0001959, NPF0001960, NPF0001962, NPF0001970, NPF0001974, NPF0001976, NPF0001978, NPF0001981, NPF0001982, NPF0001985, NPF0001987, NPF0001988, NPF0001996, NPF0001997, NPF0001998, NPF0001999, NPF0002000, NPF0002004, NPF0002006, NPF0002008, NPF0002010, NPF0002011, NPF0002014, NPF0002015, NPF0002019, NPF0002022, NPF0002026, NPF0002034, NPF0002038, NPF0002044, NPF0002046, NPF0002049, NPF0002051, NPF0002054, NPF0002062, NPF0002071, NPF0002074, NPF0002096, NPF0002097, NPF0002098, NPF0002099, NPF0002100, NPF0002101, NPF0002106, NPF0002109, NPF0002110, NPF0002111, NPF0002113, NPF0002114, NPF0002124, NPF0002125, NPF0002131, NPF0002132, NPF0002138, NPF0002140, NPF0002142, NPF0002146, NPF0002150, NPF0002152, NPF0002154, NPF0002155, NPF0002156, NPF0002158, NPF0002163, NPF0002165, NPF0002166, NPF0002169, NPF0002172, NPF0002173, NPF0002178, NPF0002182, NPF0002183, NPF0002186, NPF0002187, NPF0002194, NPF0002196, NPF0002210, NPF0002219, NPF0002224, NPF0002230, NPF0002236, NPF0002238, NPF0002250, NPF0002256, NPF0002266, NPF0002271, NPF0002286, NPF0002288, NPF0002300, NPF0002314, NPF0002317, NPF0002327, NPF0002328, NPF0002333, NPF0002346, NPF0002348, NPF0002351, NPF0002352, NPF0002355, NPF0002361, NPF0002364, NPF0002378, NPF0002380, NPF0002381, NPF0002384, NPF0002399, NPF0002405, NPF0002406, NPF0002408, NPF0002411, NPF0002414, NPF0002415, NPF0002422, NPF0002423, NPF0002425, NPF0002432, NPF0002456, NPF0002459, NPF0002462, NPF0002463, NPF0002471, NPF0002484, NPF0002485, NPF0002486, NPF0002490, NPF0002497, NPF0002502, NPF0002506, NPF0002509, NPF0002548, NPF0002572, NPF0002573, NPF0002577, NPF0002581, NPF0002582, NPF0002584, NPF0002586, NPF0002610, NPF0002612, NPF0002619, NPF0002621, NPF0002623, NPF0002624, NPF0002626, NPF0002632, NPF0002641, NPF0002642, NPF0002647, NPF0002649, NPF0002655, NPF0002658, NPF0002660, NPF0002674, NPF0002682, NPF0002685, NPF0002686, NPF0002691, NPF0002694, NPF0002721, NPF0002741, NPF0002753, NPF0002756, NPF0002759, NPF0002760, NPF0002761, NPF0002762, NPF0002764, NPF0002765, NPF0002766, NPF0002767, NPF0002768, NPF0002770, NPF0002796, NPF0002797, NPF0002800, NPF0002803, NPF0002816, NPF0002831, NPF0002834, NPF0002836, NPF0002839, NPF0002840, NPF0002843, NPF0002844, NPF0002848, NPF0002849, NPF0002851, NPF0002852, NPF0002853, NPF0002854, NPF0002856, NPF0002858, NPF0002859, NPF0002864, NPF0002866, NPF0002868, NPF0002869, NPF0002871, NPF0002874, NPF0002875, NPF0002877, NPF0002878, NPF0002879, NPF0002881, NPF0002882, NPF0002883, NPF0002884, NPF0002885, NPF0002886, NPF0002887, NPF0002889, NPF0002895, NPF0002897, NPF0002911, NPF0002912, NPF0002913, NPF0002915, NPF0002916, NPF0002917, NPF0002918, NPF0002919, NPF0002922, NPF0002924, NPF0002927, NPF0002928, NPF0002929, NPF0002932, NPF0002933, NPF0002936, NPF0002938, NPF0002944, NPF0002952, NPF0002953, NPF0002954, NPF0002955, NPF0002958, NPF0002959, NPF0002961, NPF0002962, NPF0002964, NPF0002965, NPF0002966, NPF0002968, NPF0002969, NPF0002970, NPF0002971, NPF0002972, NPF0002973, NPF0002974, NPF0002975, NPF0002977, NPF0002978, NPF0002979, NPF0002984, NPF0002988, NPF0002989, NPF0002993, NPF0002994, NPF0002995, NPF0002996, NPF0002998, NPF0003011, NPF0003012, NPF0003016, NPF0003020, NPF0003021, NPF0003022, NPF0003023, NPF0003024, NPF0003025, NPF0003026, NPF0003027, NPF0003028, NPF0003029, NPF0003031, NPF0003032, NPF0003033, NPF0003034, NPF0003035, NPF0003036, NPF0003038, NPF0003039, NPF0003040, NPF0003045, NPF0003049, NPF0003056***2) TH90QFA, UDI-Device Identifier (GTIN/UPN) 00763000582487, Serial Numbers: NPF0000658, NPF0000662, NPF0000673, NPF0000678 ***3) TH90PFA, UDI-Device Identifier (GTIN/UPN) 00763000433871, Serial Numbers: NPT0001313, NPT0051480, NPT0021804, NPT0051392, NPT0051144, NPT0001294, NPT0049798, NPT0053021, NPT0050685, NPT0001633, NPT0021799, NPT0051778, NPT0050614, NPT0001690***4) TH90GFA, UDI-Device Identifier (GTIN/UPN) 00763000187231, Serial Numbers: NPO0014927, NPO0011973, NPO0012969, NPO0041785, NPO0013871, NPO0041444, NPO0011977, NPO0012971, NPO0012000, NPO0013811, NPO0011986, NPO0012125, NPO0011971, NPO0008660, NPO0011724, NPO0011728, NPO0011734, NPO0012060, NPO0012531, NPO0013146, NPO0013205, NPO0013212, NPO0013226, NPO0014549, NPO0042926, NPO0043071, NPO0043184, NPO0043546, NPO0012112, NPO0012804, NPO0011909, NPO0013222***5) TH90G01, UDI-Device Identifier (GTIN/UPN) 00763000058005, Serial Numbers: NPO0042790, NPO0015230, NPO0053525******b) Model Number A51300, UDI-Device Identifier (GTIN/UPN) 00763000517182***c) Model Number A52300, UDI-Device Identifier (GTIN/UPN) 00763000517199 Recalling Firm/ Manufacturer Medtronic Neuromodulation 7000 Central Ave Ne Minneapolis MN 55432-3568 For Additional Information Contact Katie Genereux 763-514-0162 Manufacturer Reason for Recall There is a software anomaly with the InterStim X Clinician software application with results in the data entered in the Patient Information fields not e retained and a " Data Lost" notification being displayed. FDA Determined Cause 2 Software Design Change Action Medtronic distributed a correction notice to impacted U.S. implanting physicians, who had implanted patients with the Model 97800 InterStim X system prior to the release of the software update or purchased the TH90Q kit (in addition to notifying, Medtronic Sales Representatives and Medtronic Patient & Technical Services) on 28-Feb-2022.******Medtronic distributed a correction notice to impacted U.S. hospital accounts, who have purchased the TH90Q kit containing the affected HH90 handset prior to the software updated release, on 04-Mar-2022, notifying consignees of this issue and the software update availability. Medtronic will also notify those consignees with the TH90PFA and TH90GFA kits where the affected software was installed. Quantity in Commerce 1005 units Distribution US Nationwide distribution in the states of FL, IL, WI, TX, CA, ND, NC, KY, AR, MO, MN, OH, CO, KS, MS, SC, LA, WA, PA, IN, SD, MI, WV, AL, GA, TN, VA, IA. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = EZW - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Medtronic distributed a correction notice to impacted U.S. implanting physicians, who had implanted patients with the Model 97800 InterStim X system prior to the release of the software update or purchased the TH90Q kit (in addition to notifying, Medtronic Sales Representatives and Medtronic Patient & Technical Services) on 28-Feb-2022.******Medtronic distributed a correction notice to impacted U.S. hospital accounts, who have purchased the TH90Q kit containing the affected HH90 handset prior to the software updated release, on 04-Mar-2022, notifying consignees of this issue and the software update availability. Medtronic will also notify those consignees with the TH90PFA and TH90GFA kits where the affected software was installed. | SoftwareDesignChange
53 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192742 | Z-0987-2022 | Artis pheno- in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). Intended use-single and biplane diagnostic imaging and interventional procedures.***Model: 10849000************ | ARTIS Pheno (VE2) | 2 | CDRH | 2022-04-21 |  |  | Siemens Medical Solutions USA, Inc | Five potential software issues affecting Artis pheno and Artis icono systems in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). The potential software errors are related to the block movements function and manual rotation of the flat detector. Depending on the status of the intervention, the limited functionality may not be sufficient to continue with treatment as planned. This may result in a situation where clinical treatment may be delayed until an additional restart is performed | Class 2 Device Recall ARTIS Pheno (VE2) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ARTIS Pheno (VE2) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ARTIS Pheno (VE2) Date Initiated by Firm February 25, 2022 Create Date April 21, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-0987-2022 Recall Event ID 89870 510(K)Number K201156 Product Classification Interventional fluoroscopic x-ray system - Product Code OWB Product Artis pheno- in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). Intended use-single and biplane diagnostic imaging and interventional procedures.******Model: 10849000 Code Information Serial Number: 164719 164712 164740 164722 164763 164735 164723 164743 164741 164728 164711 164718 164739 164703 164721 164726 164724 164725 164760 UDI: 04056869046877 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact SAME 610-219-4834 Manufacturer Reason for Recall Five potential software issues affecting Artis pheno and Artis icono systems in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). The potential software errors are related to the block movement***s function and manual rotation of the flat detector. Depending on the status of the intervention, the limited functionality may not be sufficient to continue with treatment as planned. This may result in a situation where clinical treatment may be delayed until an additional restart is performed FDA Determined Cause 2 Software Design Change Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice to all affected customers via AX005/22/S on 2/28/22. Siemens will provide a software update for all affected systems via Update Instruction AX003/22/S and AX004/22/S. Letter states reason for recall, health risk and action to take: ******We do not consider it necessary to re-examine any patients in relation to the issues (issue 1 to 5) described above. If measurements have already been performed in the past for diagnostics, please verify the results and diagnostic evaluation if applicable.******Siemens will correct the software error via Update Instruction AX003/22/S and Update Instruction AX004/22/S. Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436. Following the correction, the cause is eliminated, and any recurrence of this potential fault is prevented.******Acknowledge Receipt of this Safety Advisory Notice******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein.******If you have received this notification by email via Adobe EchoSign, kindly sign with your digital signature to acknowledge that you have read and understand the content provided. If you have received this notification via Certified Mail, please provide at least two contact names and email addresses of the appropriate personnel within your facility who should receive these notifications electronically. This information will assist in providing you with these important notifications in a more expedient and compliant manner. Please include your Siemens system names, material numbers, serial numbers, and Siemens Functional Location numbers and email this information to: recallsandrefusals.team@siemens-hea Quantity in Commerce 20 units Distribution US Nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = OWB - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice to all affected customers via AX005/22/S on 2/28/22. Siemens will provide a software update for all affected systems via Update Instruction AX003/22/S and AX004/22/S. Letter states reason for recall, health risk and action to take: ******We do not consider it necessary to re-examine any patients in relation to the issues (issue 1 to 5) described above. If measurements have already been performed in the past for diagnostics, please verify the results and diagnostic evaluation if applicable.******Siemens will correct the software error via Update Instruction AX003/22/S and Update Instruction AX004/22/S. Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436. Following the correction, the cause is eliminated, and any recurrence of this potential fault is prevented.******Acknowledge Receipt of this Safety Advisory Notice******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein.******If you have received this notification by email via Adobe EchoSign, kindly sign with your digital signature to acknowledge that you have read and understand the content provided. If you have received this notification via Certified Mail, please provide at least two contact names and email addresses of the appropriate personnel within your facility who should receive these notifications electronically. This information will assist in providing you with these important notifications in a more expedient and compliant manner. Please include your Siemens system names, material numbers, serial numbers, and Siemens Functional Location numbers and email this information to: recallsandrefusals.team@siemens-hea | SoftwareDesignChange
54 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192739 | Z-0986-2022 | Artis icono floor in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). Intended use single and biplane diagnostic imaging and interventional procedures.***Model: 11327700************ | ARTIS Icono | 2 | CDRH | 2022-04-21 |  |  | Siemens Medical Solutions USA, Inc | Five potential software issues affecting Artis pheno and Artis icono systems in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). The potential software errors are related to the block movements function and manual rotation of the flat detector. Depending on the status of the intervention, the limited functionality may not be sufficient to continue with treatment as planned. This may result in a situation where clinical treatment may be delayed until an additional restart is performed | Class 2 Device Recall ARTIS Icono Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ARTIS Icono FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ARTIS Icono Date Initiated by Firm February 25, 2022 Create Date April 21, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-0986-2022 Recall Event ID 89870 510(K)Number K193326 Product Classification Interventional fluoroscopic x-ray system - Product Code OWB Product Artis icono floor in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). Intended use single and biplane diagnostic imaging and interventional procedures.******Model: 11327700 Code Information S/N: Serial 170319 170026 170343 170367 170368 170305 170360 170307 170331 170349 170309 170310 170361 170351 170383 170023 170304 170328 170334 170336 170025 170346 170051 170358 170357 170333 170365 170332 170312 170027 170303 170378 170322 170053 170364 UDI: 04056869149325 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact SAME 610-219-4834 Manufacturer Reason for Recall Five potential software issues affecting Artis pheno and Artis icono systems in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). The potential software errors are related to the block movement***s function and manual rotation of the flat detector. Depending on the status of the intervention, the limited functionality may not be sufficient to continue with treatment as planned. This may result in a situation where clinical treatment may be delayed until an additional restart is performed FDA Determined Cause 2 Software Design Change Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice to all affected customers via AX005/22/S on 2/28/22. Siemens will provide a software update for all affected systems via Update Instruction AX003/22/S and AX004/22/S. Letter states reason for recall, health risk and action to take: ******We do not consider it necessary to re-examine any patients in relation to the issues (issue 1 to 5) described above. If measurements have already been performed in the past for diagnostics, please verify the results and diagnostic evaluation if applicable.******Siemens will correct the software error via Update Instruction AX003/22/S and Update Instruction AX004/22/S. Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436. Following the correction, the cause is eliminated, and any recurrence of this potential fault is prevented.******Acknowledge Receipt of this Safety Advisory Notice******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein.******If you have received this notification by email via Adobe EchoSign, kindly sign with your digital signature to acknowledge that you have read and understand the content provided. If you have received this notification via Certified Mail, please provide at least two contact names and email addresses of the appropriate personnel within your facility who should receive these notifications electronically. This information will assist in providing you with these important notifications in a more expedient and compliant manner. Please include your Siemens system names, material numbers, serial numbers, and Siemens Functional Location numbers and email this information to: recallsandrefusals.team@siemens-hea Quantity in Commerce 35 units Distribution US Nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = OWB - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice to all affected customers via AX005/22/S on 2/28/22. Siemens will provide a software update for all affected systems via Update Instruction AX003/22/S and AX004/22/S. Letter states reason for recall, health risk and action to take: ******We do not consider it necessary to re-examine any patients in relation to the issues (issue 1 to 5) described above. If measurements have already been performed in the past for diagnostics, please verify the results and diagnostic evaluation if applicable.******Siemens will correct the software error via Update Instruction AX003/22/S and Update Instruction AX004/22/S. Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436. Following the correction, the cause is eliminated, and any recurrence of this potential fault is prevented.******Acknowledge Receipt of this Safety Advisory Notice******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein.******If you have received this notification by email via Adobe EchoSign, kindly sign with your digital signature to acknowledge that you have read and understand the content provided. If you have received this notification via Certified Mail, please provide at least two contact names and email addresses of the appropriate personnel within your facility who should receive these notifications electronically. This information will assist in providing you with these important notifications in a more expedient and compliant manner. Please include your Siemens system names, material numbers, serial numbers, and Siemens Functional Location numbers and email this information to: recallsandrefusals.team@siemens-hea | SoftwareDesignChange
55 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192721 | Z-0985-2022 | Artis icono biplane- in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). Intended use single and biplane diagnostic imaging and interventional procedures.***Model: 11327600************ | ARTIS Icono | 2 | CDRH | 2022-04-21 |  |  | Siemens Medical Solutions USA, Inc | Five potential software issues affecting Artis pheno and Artis icono systems in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). The potential software errors are related to the block movements function and manual rotation of the flat detector. Depending on the status of the intervention, the limited functionality may not be sufficient to continue with treatment as planned. This may result in a situation where clinical treatment may be delayed until an additional restart is performed | Class 2 Device Recall ARTIS Icono Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ARTIS Icono FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ARTIS Icono Date Initiated by Firm February 25, 2022 Create Date April 21, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-0985-2022 Recall Event ID 89870 510(K)Number K193326 Product Classification Interventional fluoroscopic x-ray system - Product Code OWB Product Artis icono biplane- in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). Intended use single and biplane diagnostic imaging and interventional procedures.******Model: 11327600 Code Information S/N: 180062 180054 180019 180065 180056 180057 180041 180421 180317 180034 180072 180349 180367 180382 180060 180337 180070 180048 180083 180369 180370 180371 180372 180345 180315 180343 180080 180301 180305 180067 180330 180040 180379 180386 180360 180333 180029 180352 180336 180407 180412 180400 180397 180316 180068 180363 180084 180418 180334 180018 180350 180376 180354 180046 180047 180335 180085 180035 180339 180357 180358 180361 180362 180081 180423 180331 180411 180414 180381 180021 180342 180321 180309 180356 180378 180393 180368 180413 180341 180323 180329 180043 180049 180377 180077 180086 180053 180055 180015 180014 180066 180082 180037 180373 UDI: 04056869063317 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact SAME 610-219-4834 Manufacturer Reason for Recall Five potential software issues affecting Artis pheno and Artis icono systems in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). The potential software errors are related to the block movement***s function and manual rotation of the flat detector. Depending on the status of the intervention, the limited functionality may not be sufficient to continue with treatment as planned. This may result in a situation where clinical treatment may be delayed until an additional restart is performed FDA Determined Cause 2 Software Design Change Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice to all affected customers via AX005/22/S on 2/28/22. Siemens will provide a software update for all affected systems via Update Instruction AX003/22/S and AX004/22/S. Letter states reason for recall, health risk and action to take: ******We do not consider it necessary to re-examine any patients in relation to the issues (issue 1 to 5) described above. If measurements have already been performed in the past for diagnostics, please verify the results and diagnostic evaluation if applicable.******Siemens will correct the software error via Update Instruction AX003/22/S and Update Instruction AX004/22/S. Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436. Following the correction, the cause is eliminated, and any recurrence of this potential fault is prevented.******Acknowledge Receipt of this Safety Advisory Notice******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein.******If you have received this notification by email via Adobe EchoSign, kindly sign with your digital signature to acknowledge that you have read and understand the content provided. If you have received this notification via Certified Mail, please provide at least two contact names and email addresses of the appropriate personnel within your facility who should receive these notifications electronically. This information will assist in providing you with these important notifications in a more expedient and compliant manner. Please include your Siemens system names, material numbers, serial numbers, and Siemens Functional Location numbers and email this information to: recallsandrefusals.team@siemens-hea Quantity in Commerce 94 U.S. Distribution US Nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = OWB - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice to all affected customers via AX005/22/S on 2/28/22. Siemens will provide a software update for all affected systems via Update Instruction AX003/22/S and AX004/22/S. Letter states reason for recall, health risk and action to take: ******We do not consider it necessary to re-examine any patients in relation to the issues (issue 1 to 5) described above. If measurements have already been performed in the past for diagnostics, please verify the results and diagnostic evaluation if applicable.******Siemens will correct the software error via Update Instruction AX003/22/S and Update Instruction AX004/22/S. Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436. Following the correction, the cause is eliminated, and any recurrence of this potential fault is prevented.******Acknowledge Receipt of this Safety Advisory Notice******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein.******If you have received this notification by email via Adobe EchoSign, kindly sign with your digital signature to acknowledge that you have read and understand the content provided. If you have received this notification via Certified Mail, please provide at least two contact names and email addresses of the appropriate personnel within your facility who should receive these notifications electronically. This information will assist in providing you with these important notifications in a more expedient and compliant manner. Please include your Siemens system names, material numbers, serial numbers, and Siemens Functional Location numbers and email this information to: recallsandrefusals.team@siemens-hea | SoftwareDesignChange
56 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192927 | Z-0942-2022 | ARIA Radiation Therapy Management (RTM) VERSIONS 13.6, 15.1, 15.5, 15.6, 16.0, 16.1 | ARIA Radiation Therapy Management | 2 | CDRH | 2022-04-14 |  |  | Varian Medical Systems Imaging Laboratory GmbH | Software issue for treatment plan and image management application may result in mismatch values which could result in treatment in the wrong location | Class 2 Device Recall ARIA Radiation Therapy Management Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ARIA Radiation Therapy Management FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ARIA Radiation Therapy Management Date Initiated by Firm March 21, 2022 Create Date April 14, 2022 Recall Status 1 Completed Recall Number Z-0942-2022 Recall Event ID 89922 510(K)Number K133572 K173838 Product Classification Accelerator, linear, medical - Product Code IYE Product ARIA Radiation Therapy Management (RTM) VERSIONS 13.6, 15.1, 15.5, 15.6, 16.0, 16.1 Code Information DI Numbers: 00856100006066, 00856100006080, 00856100006097, 00856100006110, 00856100006127,00856100006165 Recalling Firm/ Manufacturer Varian Medical Systems Imaging Laboratory GmbH Tafernstrasse 7 DATTWIL Baden Switzerland For Additional Information Contact vARIAN gLOBAL hELP dESK 441 293 601267 Manufacturer Reason for Recall Software issue for treatment plan and image management application may result in mismatch values which could result in treatment in the wrong location FDA Determined Cause 2 Software Design Change Action On 3/21/22, Varian Medical Systems issued a "Urgent Medical Device Correction" notification to affected consignees. In addition to informing consignees about the recall, Varian asked consignees to take the following actions:*********Please retain a copy of this document along with your most current product labeling.******Advise the appropriate personnel working in your radiotherapy department of the content of this letter.******This document will be posted to the Varian Medical Systems customer support website: http://www.MyVarian.com.*********Please complete the attached Return Response form provided. Kindly return the completed form to returnresponse@varian.com. Quantity in Commerce 14 Distribution Worldwide distribution - US Nationwide distribution in the states of MI, KS, OH, VA, DC, FL, PA, OK, MN, AZ and the countries of South Korea, Netherlands, Poland, China. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 3/21/22, Varian Medical Systems issued a "Urgent Medical Device Correction" notification to affected consignees. In addition to informing consignees about the recall, Varian asked consignees to take the following actions:*********Please retain a copy of this document along with your most current product labeling.******Advise the appropriate personnel working in your radiotherapy department of the content of this letter.******This document will be posted to the Varian Medical Systems customer support website: http://www.MyVarian.com.*********Please complete the attached Return Response form provided. Kindly return the completed form to returnresponse@varian.com. | SoftwareDesignChange
57 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190782 | Z-0475-2022 | SOMATOM Definition Edge -Computed tomography system***Model 10590000*** | SOMATOM Definition Edge | 2 | CDRH | 2022-01-10 |  |  | Siemens Medical Solutions USA, Inc | Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis. | Class 2 Device Recall SOMATOM Definition Edge Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SOMATOM Definition Edge FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SOMATOM Definition Edge Date Initiated by Firm November 19, 2021 Create Date January 10, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-0475-2022 Recall Event ID 89208 510(K)Number K190578 Product Classification System, x-ray, tomography, computed - Product Code JAK Product SOMATOM Definition Edge -Computed tomography system******Model 10590000 Code Information Serial Numbers:***155403***84012***83047***83095***83915***83801***155351***155658***155353***83826***83754***83277***83249***83904***155521***83632***83743***155578***83931***83979***155393***155396***155413***155489***83238***83336***83984***83401***84006***83834***155671***155649***83562***155341***83564***155613***83582***155606***83596***155596***155485***155437***83934***83746***83234***83971***155573***155335***83321***155422***83332***155459***155382***83994***83026***83330***155501***155639***155535***83281***83558***83767***155600***83880***83683***155360***155361***155564***83339***155580***83244***155615***83585***83899***155504***155315***155592***83613***83852***83226***83271***83607***84011***155405***83403***83599***155373***83695***83937***83246***83388***83668***155642***155343***155554***155399***83584***83837***83591***155616***83878***83714***155463***155447***155630***83574***155318***83567***83832***83886***83662***83763***83308***155312***155557***83871***83375***83988***155629***155637***83227***83359***83031***155364***83520***83262***155502***155627***84019***83903***83901***155524***84024***155513***83798***83292***83243***83319***83294***83544***83637***83640***83705***83083***83096***83707***155354***155507***83581***83717***155634***155632***83728***83056***83950***155559***83382***83819***83041***83363***83307***155388***84020***83732***83752***155581***155584***155569***155427***83961***155329***83845***83236***83348***155374***155365***83846***83399***155601***83749***83847***155322***83786***155300***155636***83975***84016***155389***83914***83842***83775***83084***83004***155358***155623***83669***155443***155304***83014***83573***155655***155515***83796***83274***94801***83768***155575***83684***83552***83210***84001***83537***83233***83352***83569***83909***155532***83968***155348***155585***83711***83697***155572***83991***155668***83395***155356***155357***155660***155442***155582***155342***155530***83621***83859***83335***83278***83742***83018***83322***83392***155558***155350***83279***83276***83365***83970***83928***83015***83093***83327***83091***83946***83207***83938***155316***83230***83254***83680***83708***83945***155594***155621***155605***83291***155549***155328***83203***83214***83913***83896***155583***155614***83790***83875***155553***83232***83362***83824***83306***83542***84010***155384***155448***155506***155412***155493***83942***83941***84017***83981***155566***83257***155576***155495***83057***83539***83810***155579***83758***155497***83239***83398***83565***83313***83653***83957***83347***83349***84027***83323***83739***155369***83266***155510***83522***83898***83951***83387***155400***83733***155407***155371***155638***83823***155434***155309***155653***155386***155570***83611***83633***83985***155380***83967***83589******UDI: UDI 04056869006949 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact SAME 610-219-4834 Manufacturer Reason for Recall Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis. FDA Determined Cause 2 Software Design Change Action Siemens issued Urgent Medical Device Correction letter CSAN CT067/21/S on November 19, 2021, to notify affected customers of the head image degradation issue and to provide them with a workaround solution. Letter states reason for recall, health risk and action to take:***To avoid potential misdiagnosis, it is mandatory to change to the new cupping correction setting CC = 4 immediately, restart the application, and (if applicable) reconstruct the raw data again.***Cupping correction setting CC = 4 is a further development of CC = 1: The Image impression is very similar to setting CC = 1 but an important improvement reduces possible artifacts. For customers who have routinely been using CC = 0 and CC = 2 instead of CC = 1, the image impression with CC = 4 will be different from what they are***used to.***Please note that the use of cupping correction CC = 4 may cause slightly altered CT numbers of non-water equivalent body materials or tissues in reconstructions with H-kernels and enabled iBHC (i.e., setting Bone or Iodine ) compared to results obtained with settings CC = 0, 1, or 2. For any quantitative interpretation of CT numbers, e.g., the mapping of HU to electron density, a recalibration is necessary Siemens is developing a technical solution, VB20_SP6 with high priority ***.Siemens Healthineers is working with high priority on a solution for the described degradation. The solution will be provided with software version VB20_SP6. Your local service organization will contact you in due time and schedule an appropriate date for the installation of the solution. This solution will be provided to you free of charge. Please note that the Regional Unit Contact cannot schedule your appointment.***If you have any unresolved questions or require technical support, please contact your local applicable***specialist or our service organization at 1-800-888-7436.***Acknowledge Receipt of this Safety Advisory Notice***Please review, understand, and confirm acknowledgement of the informati Distribution US Nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: |  | SoftwareDesignChange
58 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190779 | Z-0472-2022 | SOMATOM Confidence -Computed tomography systems***Model 10590100 | SOMATOM Confidence | 2 | CDRH | 2022-01-10 |  |  | Siemens Medical Solutions USA, Inc | Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis. | Class 2 Device Recall SOMATOM Confidence Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SOMATOM Confidence FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SOMATOM Confidence Date Initiated by Firm November 19, 2021 Create Date January 10, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-0472-2022 Recall Event ID 89208 510(K)Number K190578 Product Classification System, x-ray, tomography, computed - Product Code JAK Product SOMATOM Confidence -Computed tomography systems******Model 10590100 Code Information Serial Numbers:***100172***100091***100112***100362***100278***100158***100143***100129***100237***100238***100191***100356***100425***100216***100307***100325***100328***100329***100326***100239***100461***100213***100070***100062***100072***100022***100107***100241***100157***100055***100340***100347***100137***100267***100190***100252***100438***100046***100197***100092***100202***100313***100311***100043***100136***100189***100283***100156***100079***100049***100056***100147***100262***100018***100240***100117***100097***100274***100242***100099***100383***100059***100289***100312***100174***100073***100016***100182***100146***100058***100170***100175***100292***100090***100271***100463***100257***100273***100067***100160***100218***100367***100087***100309***100284***100207***100405***100261***100025***100338***100272***100053***100019***100421***100280***100357***100220***100214***100393***100394***100486***100140***100085***100279***100254***100353***100298***100379******UDI: UDI 04056869011011 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact SAME 610-219-4834 Manufacturer Reason for Recall Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis. FDA Determined Cause 2 Software Design Change Action Siemens issued Urgent Medical Device Correction letter CSAN CT067/21/S on November 19, 2021, to notify affected customers of the head image degradation issue and to provide them with a workaround solution. Letter states reason for recall, health risk and action to take:***To avoid potential misdiagnosis, it is mandatory to change to the new cupping correction setting CC = 4 immediately, restart the application, and (if applicable) reconstruct the raw data again.***Cupping correction setting CC = 4 is a further development of CC = 1: The Image impression is very similar to setting CC = 1 but an important improvement reduces possible artifacts. For customers who have routinely been using CC = 0 and CC = 2 instead of CC = 1, the image impression with CC = 4 will be different from what they are***used to.***Please note that the use of cupping correction CC = 4 may cause slightly altered CT numbers of non-water equivalent body materials or tissues in reconstructions with H-kernels and enabled iBHC (i.e., setting Bone or Iodine ) compared to results obtained with settings CC = 0, 1, or 2. For any quantitative interpretation of CT numbers, e.g., the mapping of HU to electron density, a recalibration is necessary Siemens is developing a technical solution, VB20_SP6 with high priority ***.Siemens Healthineers is working with high priority on a solution for the described degradation. The solution will be provided with software version VB20_SP6. Your local service organization will contact you in due time and schedule an appropriate date for the installation of the solution. This solution will be provided to you free of charge. Please note that the Regional Unit Contact cannot schedule your appointment.***If you have any unresolved questions or require technical support, please contact your local applicable***specialist or our service organization at 1-800-888-7436.***Acknowledge Receipt of this Safety Advisory Notice***Please review, understand, and confirm acknowledgement of the informati Quantity in Commerce 108 units Distribution US Nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens issued Urgent Medical Device Correction letter CSAN CT067/21/S on November 19, 2021, to notify affected customers of the head image degradation issue and to provide them with a workaround solution. Letter states reason for recall, health risk and action to take:***To avoid potential misdiagnosis, it is mandatory to change to the new cupping correction setting CC = 4 immediately, restart the application, and (if applicable) reconstruct the raw data again.***Cupping correction setting CC = 4 is a further development of CC = 1: The Image impression is very similar to setting CC = 1 but an important improvement reduces possible artifacts. For customers who have routinely been using CC = 0 and CC = 2 instead of CC = 1, the image impression with CC = 4 will be different from what they are***used to.***Please note that the use of cupping correction CC = 4 may cause slightly altered CT numbers of non-water equivalent body materials or tissues in reconstructions with H-kernels and enabled iBHC (i.e., setting Bone or Iodine ) compared to results obtained with settings CC = 0, 1, or 2. For any quantitative interpretation of CT numbers, e.g., the mapping of HU to electron density, a recalibration is necessary Siemens is developing a technical solution, VB20_SP6 with high priority ***.Siemens Healthineers is working with high priority on a solution for the described degradation. The solution will be provided with software version VB20_SP6. Your local service organization will contact you in due time and schedule an appropriate date for the installation of the solution. This solution will be provided to you free of charge. Please note that the Regional Unit Contact cannot schedule your appointment.***If you have any unresolved questions or require technical support, please contact your local applicable***specialist or our service organization at 1-800-888-7436.***Acknowledge Receipt of this Safety Advisory Notice***Please review, understand, and confirm acknowledgement of the informati | SoftwareDesignChange
59 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190781 | Z-0474-2022 | SOMATOM Edge Plus-Computed tomography system***Model 1026700****** | SOMATOM Edge Plus | 2 | CDRH | 2022-01-10 |  |  | Siemens Medical Solutions USA, Inc | Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis. | Class 2 Device Recall SOMATOM Edge Plus Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SOMATOM Edge Plus FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SOMATOM Edge Plus Date Initiated by Firm November 19, 2021 Create Date January 10, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-0474-2022 Recall Event ID 89208 510(K)Number K190578 Product Classification System, x-ray, tomography, computed - Product Code JAK Product SOMATOM Edge Plus-Computed tomography system******Model 1026700 Code Information Serial Numbers:***122022***122185***122098***122062***122099***122121***122146***122054***122082***122080***122020***122164***122116***122008***122036***122029***122044***122172***122074***122075***122049***122194***122141***122167***122156***122183***122081***122126***122129***122090***122130***122179***122089***122085***122037***122009***122163***122177***122066***122064***122038***122161***122132***122171***122153***122168***122158***122144***122063***122086***122035***122041***122112***122114***122107***122101***122108***122109***122162***122187***122174***122125***122178***122011***122182***122052***122034***122057***122145***122147***122050***122106***122189***122180***122061***122051***122166***122170***122092***122004******UDI: UDI 04056869006949 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact SAME 610-219-4834 Manufacturer Reason for Recall Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis. FDA Determined Cause 2 Software Design Change Action Siemens issued Urgent Medical Device Correction letter CSAN CT067/21/S on November 19, 2021, to notify affected customers of the head image degradation issue and to provide them with a workaround solution. Letter states reason for recall, health risk and action to take:***To avoid potential misdiagnosis, it is mandatory to change to the new cupping correction setting CC = 4 immediately, restart the application, and (if applicable) reconstruct the raw data again.***Cupping correction setting CC = 4 is a further development of CC = 1: The Image impression is very similar to setting CC = 1 but an important improvement reduces possible artifacts. For customers who have routinely been using CC = 0 and CC = 2 instead of CC = 1, the image impression with CC = 4 will be different from what they are***used to.***Please note that the use of cupping correction CC = 4 may cause slightly altered CT numbers of non-water equivalent body materials or tissues in reconstructions with H-kernels and enabled iBHC (i.e., setting Bone or Iodine ) compared to results obtained with settings CC = 0, 1, or 2. For any quantitative interpretation of CT numbers, e.g., the mapping of HU to electron density, a recalibration is necessary Siemens is developing a technical solution, VB20_SP6 with high priority ***.Siemens Healthineers is working with high priority on a solution for the described degradation. The solution will be provided with software version VB20_SP6. Your local service organization will contact you in due time and schedule an appropriate date for the installation of the solution. This solution will be provided to you free of charge. Please note that the Regional Unit Contact cannot schedule your appointment.***If you have any unresolved questions or require technical support, please contact your local applicable***specialist or our service organization at 1-800-888-7436.***Acknowledge Receipt of this Safety Advisory Notice***Please review, understand, and confirm acknowledgement of the informati Quantity in Commerce 80 units Distribution US Nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens issued Urgent Medical Device Correction letter CSAN CT067/21/S on November 19, 2021, to notify affected customers of the head image degradation issue and to provide them with a workaround solution. Letter states reason for recall, health risk and action to take:***To avoid potential misdiagnosis, it is mandatory to change to the new cupping correction setting CC = 4 immediately, restart the application, and (if applicable) reconstruct the raw data again.***Cupping correction setting CC = 4 is a further development of CC = 1: The Image impression is very similar to setting CC = 1 but an important improvement reduces possible artifacts. For customers who have routinely been using CC = 0 and CC = 2 instead of CC = 1, the image impression with CC = 4 will be different from what they are***used to.***Please note that the use of cupping correction CC = 4 may cause slightly altered CT numbers of non-water equivalent body materials or tissues in reconstructions with H-kernels and enabled iBHC (i.e., setting Bone or Iodine ) compared to results obtained with settings CC = 0, 1, or 2. For any quantitative interpretation of CT numbers, e.g., the mapping of HU to electron density, a recalibration is necessary Siemens is developing a technical solution, VB20_SP6 with high priority ***.Siemens Healthineers is working with high priority on a solution for the described degradation. The solution will be provided with software version VB20_SP6. Your local service organization will contact you in due time and schedule an appropriate date for the installation of the solution. This solution will be provided to you free of charge. Please note that the Regional Unit Contact cannot schedule your appointment.***If you have any unresolved questions or require technical support, please contact your local applicable***specialist or our service organization at 1-800-888-7436.***Acknowledge Receipt of this Safety Advisory Notice***Please review, understand, and confirm acknowledgement of the informati | SoftwareDesignChange
60 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190780 | Z-0473-2022 | SOMATOM Definition AS-Computed tomography system***Model 8098027*** | SOMATOM Definition AS | 2 | CDRH | 2022-01-10 |  |  | Siemens Medical Solutions USA, Inc | Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis. | Class 2 Device Recall SOMATOM Definition AS Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SOMATOM Definition AS FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SOMATOM Definition AS Date Initiated by Firm November 19, 2021 Create Date January 10, 2022 Recall Status 1 Open 3 , Classified Recall Number Z-0473-2022 Recall Event ID 89208 510(K)Number K190578 Product Classification System, x-ray, tomography, computed - Product Code JAK Product SOMATOM Definition AS-Computed tomography system******Model 8098027 Code Information Serial Numbers:***96278***96887***66027***64492***64541***64519***64427***67297***95303***96267***64864***96211***59802***67241***64798***98423***65597***95389***96689***64419***64347***64860***64382***64869***65238***66734***92078***95804***96471***98402***95985***96710***95889***95778***92353***95847***95936***67236***64464***64471***65852***95798***96454***96466***95363***64927***91966***98354***95934***96279***95807***95602***95639***65288***64773***95229***65297***65159***64036***95628***66143***64317***95610***98474***95859***95939***95601***95070***95882***95596***95791***66219***65752***98356***95677***64104***66506***66312***95844***95670***66858***65382***98424***66588***92280***67399***64413***64559***64235***66611***92056***92261***92342***95684***66319***95545***96547***95702***95766***67218***67128***64859***95788***64487***96491***65370***95234***95466***95809***67098***95616***59819***66032***95372***95461***95043***66087***95582***95143***67363***96950***65768***92186***96941***95951***95100***98422***98413***66600***64030***66844***66649***66635***95299***66225***66842***66916***95584***95151***66188***66355***64022***92375***95578***96777***66467***98417***64498***64763***64372***95362***65788***65015***66150***96647***66933***96472***98366***65299***66461***66604***65115***66441***96948***95196***92184***92131***65515***66498***96380***96973***66197***66688***66164***95777***65769***95609***96424***96020***66768***65304***67233***95742***92298***64825***96413***96015***96824***65770***59825***67310***64066***66201***66406***66574***96529***95856***96544***65713***66700***65926***66503***98404***65443***96305***95881***95845***98364***96845***65804***95280***96773***96455***66126***65478***95923***95903***96747***67116***96726***95095***96370***98445***95940***64516***95784***98403***65407***65251***64720***92335***92300***66969***96750***96530***64685***66001***96038***96863***96992***95175***95457***95459***66928***65403***96341***66542***95465***92358***96984***67175***65740***66690***95451***67326***95274***95875***95983***67380***65555***64709***65291***66965***66151***67181***95891***64314***95181***66557***96633***96543***96669***66148***96335***95892***67288***67085***64446***92180***66705***96823***65418***95152***96421***67133***67048***95411***64677***67319***95960***64740***59820***98441***65755***66019***66062***92023***96014***96865***65501***66234***64221***98449***98357***95893***65927***65186***95659***59803***95193***98363***96709***96351***95413***64211***96385***95848***98359***96785***156012***66478***66850***96443***96749***92053***95437***96346***95272***64218***95207***64480***64947***95393***96893***96438***66999***95846***64463***65283***96369***95485***64835***64913***64834***64942***96316***96028***92243***66011***95647***67391***96812***66297***65818***96418***95344***92321***66618***96288***66396***96238***95080***96676***92101***96352***96444***92313***92177***96271***66473***66423***66304***66349***67010***67108***95438***95439***95751***95710***96336***96324***67229***65611***67126***65936***66958***65709***65556***92033***92296***96312***95906***96896***96021***92254***65735***95768***95715***91964***64489***64732***92169***98401***66183***65793***96937***95928***95968***98447***64051***95014***65728***96832***64282***95174***64031***66587***66789***64056***95425***67291***95415***64729***65235***64999***98411***66232***98350***66290***65229***96226***92132***65925***65148***95818***64060***95661***65455***64960***96340***66173***65145***64110***96929***66918***66808***96697***66843***92347***92307***65844***95339***95022***95748***95199***65779***96553***96303***65272***95646***95034***66021***66095***99401***96022***65151***96828***67312***96338***95302***64568***59828***65610***64248***96023***67169***67193***96545***64602***64707***66237***64096***98353***95952***65079***65831***92087***95868***96991***96942***92336***96899***64501***66308***95378***96266***65003***66792***95737***67274***96922***95912***65808***92207***92125***95361***65800***95961***95701***95376***96328***65557***65371***92315***66356***66381***96713***66821***95580***66812***64572***64967***65022***64052***65795***66221***67036***95427***65173***64407***65848***66930***66077***95782***95938***67251***95972***66116***95192***96851***65586***96250***96776***96347***96310***92058***64998***66457***66502***65154***66490***92088***95785***92099***92288***95522***64238***66110***59823***59829***64094***65659***66248***98473***66579***64796***64500***96425***95493***96061***96861***95593***64983***66931***66288***59824***98442***59822***96736***98414***95678***64247***66529***96541***92113***64322***95342***96983***64148***65760***65865***66265***66335***65854***95786***96524***65017***66813***96925***96909***98352***59812***96589***96440***98409***95592***66081***64924***64934***65701***66082***65909***64813***59814***96677***95221***95703***92108***65347***65392***92091***66925***95087***96717***92176***95517***98362***67232***96821***66076***66536***66041***64019***64767***66689***95441***66407***96857***95947***95484***65493***95521***95059***95359***65365***65803***65956***96837***64160***64605***64064***95021***95013***66023***95499***96967***64675***92189***96538***96668***66510***95924***95498***95147***98452***95808***66403***92409***91968***92000***66470***65680***65805***95575***96049***96040***95591***92323***66223***66540***96268***59815***67192***66630***95719***66280***95079***64908***65247***92057***95632***92208***92155***92165***96542***96740***66859***64297***95559***65995***67321***95797***95758***95145***96864***96281***66664***65667***96058***65157***96463***95440***95549***66434***66017***95577***96302***95523***66445***95674***65773***67083***156018***95642***95717***91996***92164***96053***66677***95520***96956***96011***95937***91974***66462***96327***96326***95402***96645***95025***96286***96700***96029***66882***91990***96604***95024***64196***65906***95927***96364***96397***67392***95886***65829***64712***64132***92045***65866***65743***65830***65694***65303***95195***92069***95640***98421***67167***95510***66771***95635***96695***65777***66860***66694***96062***65368***92098***95914***95657***65758***95513***66425***96275***96003***95529***98418***98419***64785***95716***65998***66447***65652***66959******UDI: UDI 04056869003665 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact SAME 610-219-4834 Manufacturer Reason for Recall Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis. FDA Determined Cause 2 Software Design Change Action Siemens issued Urgent Medical Device Correction letter CSAN CT067/21/S on November 19, 2021, to notify affected customers of the head image degradation issue and to provide them with a workaround solution. Letter states reason for recall, health risk and action to take:***To avoid potential misdiagnosis, it is mandatory to change to the new cupping correction setting CC = 4 immediately, restart the application, and (if applicable) reconstruct the raw data again.***Cupping correction setting CC = 4 is a further development of CC = 1: The Image impression is very similar to setting CC = 1 but an important improvement reduces possible artifacts. For customers who have routinely been using CC = 0 and CC = 2 instead of CC = 1, the image impression with CC = 4 will be different from what they are***used to.***Please note that the use of cupping correction CC = 4 may cause slightly altered CT numbers of non-water equivalent body materials or tissues in reconstructions with H-kernels and enabled iBHC (i.e., setting Bone or Iodine ) compared to results obtained with settings CC = 0, 1, or 2. For any quantitative interpretation of CT numbers, e.g., the mapping of HU to electron density, a recalibration is necessary Siemens is developing a technical solution, VB20_SP6 with high priority ***.Siemens Healthineers is working with high priority on a solution for the described degradation. The solution will be provided with software version VB20_SP6. Your local service organization will contact you in due time and schedule an appropriate date for the installation of the solution. This solution will be provided to you free of charge. Please note that the Regional Unit Contact cannot schedule your appointment.***If you have any unresolved questions or require technical support, please contact your local applicable***specialist or our service organization at 1-800-888-7436.***Acknowledge Receipt of this Safety Advisory Notice***Please review, understand, and confirm acknowledgement of the informati Quantity in Commerce 794 units Distribution US Nationwide distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens issued Urgent Medical Device Correction letter CSAN CT067/21/S on November 19, 2021, to notify affected customers of the head image degradation issue and to provide them with a workaround solution. Letter states reason for recall, health risk and action to take:***To avoid potential misdiagnosis, it is mandatory to change to the new cupping correction setting CC = 4 immediately, restart the application, and (if applicable) reconstruct the raw data again.***Cupping correction setting CC = 4 is a further development of CC = 1: The Image impression is very similar to setting CC = 1 but an important improvement reduces possible artifacts. For customers who have routinely been using CC = 0 and CC = 2 instead of CC = 1, the image impression with CC = 4 will be different from what they are***used to.***Please note that the use of cupping correction CC = 4 may cause slightly altered CT numbers of non-water equivalent body materials or tissues in reconstructions with H-kernels and enabled iBHC (i.e., setting Bone or Iodine ) compared to results obtained with settings CC = 0, 1, or 2. For any quantitative interpretation of CT numbers, e.g., the mapping of HU to electron density, a recalibration is necessary Siemens is developing a technical solution, VB20_SP6 with high priority ***.Siemens Healthineers is working with high priority on a solution for the described degradation. The solution will be provided with software version VB20_SP6. Your local service organization will contact you in due time and schedule an appropriate date for the installation of the solution. This solution will be provided to you free of charge. Please note that the Regional Unit Contact cannot schedule your appointment.***If you have any unresolved questions or require technical support, please contact your local applicable***specialist or our service organization at 1-800-888-7436.***Acknowledge Receipt of this Safety Advisory Notice***Please review, understand, and confirm acknowledgement of the informati | SoftwareDesignChange
61 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189437 | Z-0175-2022 | Cliquid MD version 3.4 software used in conjunction with Analyst MD software on SCIEX Citrine & 4500MD series liquid chromatography-tandem mass spectrometry (LC-MS/MS) systems. Model No. 5088288 | Cliquid MD | 2 | CDRH | 2021-10-26 |  | 2024-01-23 12:16 | AB SCIEX | The values of the Internal Standard (IS) concentrations are incorrectly derived when the user builds a customized test, then set the IS column to hidden and left the column empty. This may lead to the user making incorrect conclusions using incorrect results. | Class 2 Device Recall Cliquid MD Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Cliquid MD FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Cliquid MD Date Initiated by Firm October 04, 2021 Create Date October 26, 2021 Recall Status 1 Terminated 3 on January 23, 2024 Recall Number Z-0175-2022 Recall Event ID 88748 Product Classification Mass spectrometer, clinical use - Product Code DOP Product Cliquid MD version 3.4 software used in conjunction with Analyst MD software on SCIEX Citrine & 4500MD series liquid chromatography-tandem mass spectrometry (LC-MS/MS) systems. Model No. 5088288 Code Information Software version 3.4 Recalling Firm/ Manufacturer AB SCIEX 71 Four Valley Dr Concord Canada Manufacturer Reason for Recall The values of the Internal Standard (IS) concentrations are incorrectly derived when the user builds a customized test, then set the IS column to hidden and left the column empty. This may lead to the user making incorrect conclusions using incorrect results. FDA Determined Cause 2 Software Design Change Action On October 4, 2021, the firm distributed Urgent Medical Device Correction letters to affected consignees. Customers were provided with temporary actions that could be implemented for safe handling of the recalled software to prevent the reporting of incorrect results to healthcare providers resulting in incorrect conclusions:******1. Avoid setting the IS (Internal Standard) column as hidden AND leaving it empty***2. Review values in the IS Concentration column***3. Follow lab SOPs to ensure quality control******SCIEX is working on the hotfix for Cliquid MD Software version 3.4 and will contact affected customers as soon as it is ready***The hotfixes may be downloaded free of charge directly from SCIEX s website, at the following location:***https://sciex.com/software-support/software-downloads***Software release notes will provide detailed installation instructions for installing the Hotfix.******If you have any questions, please contact customer service at 1-877-740-2129, option 1, and then option 4, or your local SCIEX support number. You may also submit questions through their web portal at https://sciex.com/request-support. Quantity in Commerce 30 Distribution Worldwide distribution - US Nationwide distribution in the states of District of Columbia,***Florida, Illinois, Louisiana, Maryland, Michigan, Missouri, New Jersey, New York, North Carolina, Ohio, Pennsylvania, Tennessee, Utah, Wisconsin and the countries of Austria, Finland, Germany, Hong Kong, Italy, Japan, Norway, Spain, Switzerland, United Kingdom, Italy, France, and Japan. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On October 4, 2021, the firm distributed Urgent Medical Device Correction letters to affected consignees. Customers were provided with temporary actions that could be implemented for safe handling of the recalled software to prevent the reporting of incorrect results to healthcare providers resulting in incorrect conclusions:******1. Avoid setting the IS (Internal Standard) column as hidden AND leaving it empty***2. Review values in the IS Concentration column***3. Follow lab SOPs to ensure quality control******SCIEX is working on the hotfix for Cliquid MD Software version 3.4 and will contact affected customers as soon as it is ready***The hotfixes may be downloaded free of charge directly from SCIEX s website, at the following location:***https://sciex.com/software-support/software-downloads***Software release notes will provide detailed installation instructions for installing the Hotfix.******If you have any questions, please contact customer service at 1-877-740-2129, option 1, and then option 4, or your local SCIEX support number. You may also submit questions through their web portal at https://sciex.com/request-support. | SoftwareDesignChange
62 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188897 | Z-2394-2021 | TRUMATCH CT PIN GUIDE KIT L- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting***Part Number: 420578 | TRUMATCH CT PIN GUIDE KIT L | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. | Class 2 Device Recall TRUMATCH CT PIN GUIDE KIT L Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TRUMATCH CT PIN GUIDE KIT L FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TRUMATCH CT PIN GUIDE KIT L Date Initiated by Firm August 10, 2021 Create Date September 01, 2021 Recall Status 1 Terminated 3 on May 17, 2023 Recall Number Z-2394-2021 Recall Event ID 88522 PMA Number P830055 Product Classification Prosthesis, knee, patellofemorotibial, semi-constrained, metal/polymer, mobile bearing - Product Code NJL Product TRUMATCH CT PIN GUIDE KIT L- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting******Part Number: 420578 Code Information Lot Numbers:***27932 , 27806, 27515, 27466, ***GTIN: 10603295430797 Recalling Firm/ Manufacturer DePuy Orthopaedics, Inc. 700 Orthopaedic Dr Warsaw IN 46582-3994 For Additional Information Contact 574-267-8143 Manufacturer Reason for Recall Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. FDA Determined Cause 2 Software Design Change Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. Quantity in Commerce 4 cases (1/case) Distribution Distribution to US states of CA, CO, FL, IL, IN, NY, MI, MN, TX, and WA, Poland and UK. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = NJL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. | SoftwareDesignChange
63 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188901 | Z-2397-2021 | TRUMATCH CT FEM CUT TIB PIN R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting***Part Number: 420907 | TRUMATCH CT FEM CUT TIB PIN R | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. | Class 2 Device Recall TRUMATCH CT FEM CUT TIB PIN R Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TRUMATCH CT FEM CUT TIB PIN R FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TRUMATCH CT FEM CUT TIB PIN R Date Initiated by Firm August 10, 2021 Create Date September 01, 2021 Recall Status 1 Terminated 3 on May 17, 2023 Recall Number Z-2397-2021 Recall Event ID 88522 PMA Number P830055 Product Classification Prosthesis, knee, patellofemorotibial, semi-constrained, metal/polymer, mobile bearing - Product Code NJL Product TRUMATCH CT FEM CUT TIB PIN R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting******Part Number: 420907 Code Information Lot number/ GTIN:***27975, 27470 ******GTIN: 10603295384823 Recalling Firm/ Manufacturer DePuy Orthopaedics, Inc. 700 Orthopaedic Dr Warsaw IN 46582-3994 For Additional Information Contact 574-267-8143 Manufacturer Reason for Recall Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. FDA Determined Cause 2 Software Design Change Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. Quantity in Commerce 2 cases (1/cs) Distribution Distribution to US states of CA, CO, FL, IL, IN, NY, MI, MN, TX, and WA, Poland and UK. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = NJL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. | SoftwareDesignChange
64 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188899 | Z-2396-2021 | TRUMATCH CT PIN GUIDE FEM R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting***Part Number: 420581 | TRUMATCH CT PIN GUIDE FEM R | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. | Class 2 Device Recall TRUMATCH CT PIN GUIDE FEM R Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TRUMATCH CT PIN GUIDE FEM R FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TRUMATCH CT PIN GUIDE FEM R Date Initiated by Firm August 10, 2021 Create Date September 01, 2021 Recall Status 1 Terminated 3 on May 17, 2023 Recall Number Z-2396-2021 Recall Event ID 88522 PMA Number P830055 Product Classification Prosthesis, knee, patellofemorotibial, semi-constrained, metal/polymer, mobile bearing - Product Code NJL Product TRUMATCH CT PIN GUIDE FEM R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting******Part Number: 420581 Code Information Lot numbers:***28206, 27983 ******GTIN: 10603295430773 Recalling Firm/ Manufacturer DePuy Orthopaedics, Inc. 700 Orthopaedic Dr Warsaw IN 46582-3994 For Additional Information Contact 574-267-8143 Manufacturer Reason for Recall Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. FDA Determined Cause 2 Software Design Change Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. Quantity in Commerce 2 cases (1 case) Distribution Distribution to US states of CA, CO, FL, IL, IN, NY, MI, MN, TX, and WA, Poland and UK. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = NJL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. | SoftwareDesignChange
65 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188903 | Z-2399-2021 | TRUMATCH CT CUT GUIDE KIT R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting***Part Number: 420916*** | TRUMATCH CT CUT GUIDE KIT R | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. | Class 2 Device Recall TRUMATCH CT CUT GUIDE KIT R Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TRUMATCH CT CUT GUIDE KIT R FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TRUMATCH CT CUT GUIDE KIT R Date Initiated by Firm August 10, 2021 Create Date September 01, 2021 Recall Status 1 Terminated 3 on May 17, 2023 Recall Number Z-2399-2021 Recall Event ID 88522 PMA Number P830055 Product Classification Prosthesis, knee, patellofemorotibial, semi-constrained, metal/polymer, mobile bearing - Product Code NJL Product TRUMATCH CT CUT GUIDE KIT R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting******Part Number: 420916 Code Information Lot number/ GTIN:***28226, 28135, 27612, 27544, 27431, 27407, 26970******OUS: 27476, 27779 ***GTIN: 10603295384861; Recalling Firm/ Manufacturer DePuy Orthopaedics, Inc. 700 Orthopaedic Dr Warsaw IN 46582-3994 For Additional Information Contact 574-267-8143 Manufacturer Reason for Recall Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. FDA Determined Cause 2 Software Design Change Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. Quantity in Commerce 9 cases (1/cs) Distribution Distribution to US states of CA, CO, FL, IL, IN, NY, MI, MN, TX, and WA, Poland and UK. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = NJL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. | SoftwareDesignChange
66 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188898 | Z-2395-2021 | TRUMATCH CT PIN GUIDE KIT R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting***Part Number: 420579*** | TRUMATCH CT PIN GUIDE KIT R | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. | Class 2 Device Recall TRUMATCH CT PIN GUIDE KIT R Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TRUMATCH CT PIN GUIDE KIT R FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TRUMATCH CT PIN GUIDE KIT R Date Initiated by Firm August 10, 2021 Create Date September 01, 2021 Recall Status 1 Terminated 3 on May 17, 2023 Recall Number Z-2395-2021 Recall Event ID 88522 PMA Number P830055 Product Classification Prosthesis, knee, patellofemorotibial, semi-constrained, metal/polymer, mobile bearing - Product Code NJL Product TRUMATCH CT PIN GUIDE KIT R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting******Part Number: 420579 Code Information Lot number/ GTIN:***27929, 27855, 27667, 27591, 27521, 27332 *********GTIN 10603295430803 Recalling Firm/ Manufacturer DePuy Orthopaedics, Inc. 700 Orthopaedic Dr Warsaw IN 46582-3994 For Additional Information Contact 574-267-8143 Manufacturer Reason for Recall Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. FDA Determined Cause 2 Software Design Change Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. Quantity in Commerce 6 cases (1/Cs) Distribution Distribution to US states of CA, CO, FL, IL, IN, NY, MI, MN, TX, and WA, Poland and UK. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = NJL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. | SoftwareDesignChange
67 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188904 | Z-2400-2021 | TRUMATCH CT CUT GUIDE FEM L- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting***Part Number: 420917*** | TRUMATCH CT CUT GUIDE FEM L | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. | Class 2 Device Recall TRUMATCH CT CUT GUIDE FEM L Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TRUMATCH CT CUT GUIDE FEM L FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TRUMATCH CT CUT GUIDE FEM L Date Initiated by Firm August 10, 2021 Create Date September 01, 2021 Recall Status 1 Terminated 3 on May 17, 2023 Recall Number Z-2400-2021 Recall Event ID 88522 PMA Number P830055 Product Classification Prosthesis, knee, patellofemorotibial, semi-constrained, metal/polymer, mobile bearing - Product Code NJL Product TRUMATCH CT CUT GUIDE FEM L- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting******Part Number: 420917 Code Information Lot number/ GTIN:***27958 10603295384878 Recalling Firm/ Manufacturer DePuy Orthopaedics, Inc. 700 Orthopaedic Dr Warsaw IN 46582-3994 For Additional Information Contact 574-267-8143 Manufacturer Reason for Recall Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. FDA Determined Cause 2 Software Design Change Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. Quantity in Commerce 1 cs (1/cs) Distribution Distribution to US states of CA, CO, FL, IL, IN, NY, MI, MN, TX, and WA, Poland and UK. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = NJL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. | SoftwareDesignChange
68 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188902 | Z-2398-2021 | TRUMATCH CT CUT GUIDE KIT L- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting***Part Number: 420915*** | TRUMATCH CT CUT GUIDE KIT L | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. | Class 2 Device Recall TRUMATCH CT CUT GUIDE KIT L Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TRUMATCH CT CUT GUIDE KIT L FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TRUMATCH CT CUT GUIDE KIT L Date Initiated by Firm August 10, 2021 Create Date September 01, 2021 Recall Status 1 Terminated 3 on May 17, 2023 Recall Number Z-2398-2021 Recall Event ID 88522 PMA Number P830055 Product Classification Prosthesis, knee, patellofemorotibial, semi-constrained, metal/polymer, mobile bearing - Product Code NJL Product TRUMATCH CT CUT GUIDE KIT L- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting******Part Number: 420915 Code Information Lot number/ GTIN:***28289, 28247, 27876, 27432, 27613 (OUS)******GTIN: 10603295384854 Recalling Firm/ Manufacturer DePuy Orthopaedics, Inc. 700 Orthopaedic Dr Warsaw IN 46582-3994 For Additional Information Contact 574-267-8143 Manufacturer Reason for Recall Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment. FDA Determined Cause 2 Software Design Change Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. Quantity in Commerce 5 cases ( 1/cs) Distribution Distribution to US states of CA, CO, FL, IL, IN, NY, MI, MN, TX, and WA, Poland and UK. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = NJL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Depuy issued Urgent Medical Device Recall (Removal) Letter on 8/10/21. Letter states reson for recall, health risk and action take:***1. Examine your inventory immediately to determine if you have the subject lot and quarantine the product. DO NOT USE THE PRODUCT. The surgical plan associated with the cases in Table 1 should not be used with conventional instruments.***2. Contact your DePuy Synthes Sales Consultant to coordinate the return/credits of any affected devices or contact TruMatch Personalized Solutions Customer Support at 1-800-689-0746 / trumatchsupport@its.jnj.com between 7am and 7pm EST.***3. Review, complete, sign, and return the attached business response form (page 4 of this letter) to OneMD-Field-Actions@its.jnj.com within 3 business days of receipt of this notification. Please include in the email subject: FA 2000415: 2021 - TruMatch Personalized Solutions***4. Forward this notice to anyone in your facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the devices subject to this action).***5. If any of the subject lot has been forwarded to another facility, contact that facility and provide them with this notice.***6. Post a copy of this notice in a visible area for awareness and keep a copy for your records.***This medical device product recall has been reported to health authorities. We apologize for the inconvenience that this recall may cause and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes Sales Consultant. For Medical Information request, please visit our website: https://www.jnjmedicaldevices.com/mir. | SoftwareDesignChange
69 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187857 | Z-2093-2021 | SoftGenomics version 4.1.15.6******SoftGenomics is a laboratory workflow and information management system to be used in a clinical molecular or cytogenetics laboratory or medical research laboratory setting by knowledgeable and trained users. SoftGenomics is used to record and track activities performed in ordering, specimen processing (pre-analytical steps), testing, documentation, and reporting. | SoftGenomics | 2 | CDRH | 2021-07-19 |  | 2022-04-15 13:14 | Soft Computer Consultants, Inc. | Software showing wrong results reported, PDF not match HIS. | Class 2 Device Recall SoftGenomics Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall SoftGenomics FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall SoftGenomics Date Initiated by Firm March 31, 2021 Create Date July 19, 2021 Recall Status 1 Terminated 3 on April 15, 2022 Recall Number Z-2093-2021 Recall Event ID 88095 Product Classification Calculator/data processing module, for clinical use - Product Code JQP Product SoftGenomics version 4.1.15.6************SoftGenomics is a laboratory workflow and information management system to be used in a clinical molecular or cytogenetics laboratory or medical research laboratory setting by knowledgeable and trained users. SoftGenomics is used to record and track activities performed in ordering, specimen processing (pre-analytical steps), testing, documentation, and reporting. Code Information Version 4.1.15.6 Recalling Firm/ Manufacturer Soft Computer Consultants, Inc. 5400 Tech Data Dr Clearwater FL 33760-3116 For Additional Information Contact Technical Support 24/7 800-763-8522 Manufacturer Reason for Recall Software showing wrong results reported, PDF not match HIS. FDA Determined Cause 2 Software Design Change Action The firm sent an " URGENT: MEDICAL DEVICE RECALL" letter dated 6/2/2021 to its sole consignee of the correction on 06/02/2021 by mail. The letter described the product, problem and actions to be taken. The consignee was instructed to do the following: ******Acknowledge receipt of this letter in the TMS task******Give permission to load Hot Fix on Test environment******Perform validation activities to completion******Acknowledge Hot Fix has been moved to Live environment and is working as expected************Type of action by Company: All clients on affected lines of software have been notified of the recall. SCC will provide Hot Fixes for all affected lines that will be delivered to all affected clients. Until the software correction is delivered, clients have been informed of an alternative workflow to avoid the problem. ************The firm provided information on a software update performed on 03/31/2021 and stated that, "Hot Fix has already been delivered."************If you have any further questions, Please call 1-800-763-8522 (Technical Support 24/7). Quantity in Commerce 1 Distribution US Distribution to: CA only Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm sent an " URGENT: MEDICAL DEVICE RECALL" letter dated 6/2/2021 to its sole consignee of the correction on 06/02/2021 by mail. The letter described the product, problem and actions to be taken. The consignee was instructed to do the following: ******Acknowledge receipt of this letter in the TMS task******Give permission to load Hot Fix on Test environment******Perform validation activities to completion******Acknowledge Hot Fix has been moved to Live environment and is working as expected************Type of action by Company: All clients on affected lines of software have been notified of the recall. SCC will provide Hot Fixes for all affected lines that will be delivered to all affected clients. Until the software correction is delivered, clients have been informed of an alternative workflow to avoid the problem. ************The firm provided information on a software update performed on 03/31/2021 and stated that, "Hot Fix has already been delivered."************If you have any further questions, Please call 1-800-763-8522 (Technical Support 24/7). | SoftwareDesignChange
70 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187544 | Z-1906-2021 | Access SARS-CoV-2 lgG II Calibrator, REF C69058, IVD, CE UDI: (01)15099590742751 | Access SARS CoV-2 lgG II Calibrator | 2 | CDRH | 2021-06-21 |  | 2023-03-27 08:30 | Beckman Coulter Inc. | SARS-CoV-2 lgG II numerical results may potentially be multiplied by a factor of 1000 on systems running with assay protocol file (APF) and access assay file (AAF) versions resulting in falsely elevated numerical values. | Class 2 Device Recall Access SARS CoV2 lgG II Calibrator Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Access SARS CoV2 lgG II Calibrator FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Access SARS CoV2 lgG II Calibrator Date Initiated by Firm May 07, 2021 Create Date June 21, 2021 Recall Status 1 Terminated 3 on March 27, 2023 Recall Number Z-1906-2021 Recall Event ID 87960 Product Classification Reagent, coronavirus serological - Product Code QKO Product Access SARS-CoV-2 lgG II Calibrator, REF C69058, IVD, CE UDI: (01)15099590742751 Code Information All Lots Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact Annette Hellie 714-831-7388 Manufacturer Reason for Recall SARS-CoV-2 lgG II numerical results may potentially be multiplied by a factor of 1000 on systems running with assay protocol file (APF) and access assay file (AAF) versions resulting in falsely elevated numerical values. FDA Determined Cause 2 Software Design Change Action On May 7, 2021, Beckman Coulter issued an "URGENT MEDICAL DEVICE RECALL" to affected consignees. The notification was emailed and sent via First Class Mail. In addition to informing consignees about the recall, the firm ask consignees to take the following actions: ************1. Do not install any of the APF/AAF versions listed in the table.************2. Determine your current APF/AAF version.******" From the Main Menu: select Configure F8 System Setup F1 System Revisions F1.******" Verify the APF/AAF version displayed on the screen.************3 If the APF/AAF version is listed in the table:******" Stop using the Access SARS-CoV-2 IgG II assay. Contact the Beckman Coulter Customer Technical Support Center at 1-800-854-3633 for further instruction.******" Review all Access SARS-CoV-2 IgG II patient test results from April 13, 2021 onward.******" Do not install any additional APF/AAF versions prior to contacting Beckman Coulter.************4. If the APF/AAF version is lower than those listed in the table, you are not impacted. Continue operating your system under existing instructions.************5. Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter.************6. So that we are assured you have received this important communication, please respond within 10 days in one of the following ways:******" Electronically, if you received this communication via email.******" Manually, complete and return the enclosed Response Form.************7. If you have any questions regarding this notice, please contact our Customer Support Center:******" From our website: http://www.beckmancoulter.com******" By phone: call 1-800-854-3633 in the United States. Quantity in Commerce 3,843 units Distribution US: Alabama,******Alaska,******Arizona,******Arkansas,******California,******Florida,******Georgia,******Illinois,******Iowa,******Kentucky,******Louisiana,******Maine,******Maryland,******Massachusetts,******Michigan,******Mississippi,******Missouri,******New Hampshire,******New Jersey,******New York,******North Dakota,******Ohio,******Oklahoma,******Pennsylvania,******Puerto Rico,******Tennessee,******Texas,******Virginia,******West Virginia,******Wisconsin,******Wyoming,************OUS: Andorra, Austria, Croatia, Czech Republic, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, Israel, Italy, Japan, Jordan, Poland, Portugal, San Marino, Slovakia, Slovenia, Spain, United Arab Emirates Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On May 7, 2021, Beckman Coulter issued an "URGENT MEDICAL DEVICE RECALL" to affected consignees. The notification was emailed and sent via First Class Mail. In addition to informing consignees about the recall, the firm ask consignees to take the following actions: ************1. Do not install any of the APF/AAF versions listed in the table.************2. Determine your current APF/AAF version.******" From the Main Menu: select Configure F8 System Setup F1 System Revisions F1.******" Verify the APF/AAF version displayed on the screen.************3 If the APF/AAF version is listed in the table:******" Stop using the Access SARS-CoV-2 IgG II assay. Contact the Beckman Coulter Customer Technical Support Center at 1-800-854-3633 for further instruction.******" Review all Access SARS-CoV-2 IgG II patient test results from April 13, 2021 onward.******" Do not install any additional APF/AAF versions prior to contacting Beckman Coulter.************4. If the APF/AAF version is lower than those listed in the table, you are not impacted. Continue operating your system under existing instructions.************5. Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter.************6. So that we are assured you have received this important communication, please respond within 10 days in one of the following ways:******" Electronically, if you received this communication via email.******" Manually, complete and return the enclosed Response Form.************7. If you have any questions regarding this notice, please contact our Customer Support Center:******" From our website: http://www.beckmancoulter.com******" By phone: call 1-800-854-3633 in the United States. | SoftwareDesignChange
71 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187543 | Z-1905-2021 | Access SARS CoV-2 lgG II Reagent, REF: C69057, IVD CE UDI: (01)15099590742744 | Access SARS Cov-2 lgG II Reagent | 2 | CDRH | 2021-06-21 |  | 2023-03-27 08:30 | Beckman Coulter Inc. | SARS-CoV-2 lgG II numerical results may potentially be multiplied by a factor of 1000 on systems running with assay protocol file (APF) and access assay file (AAF) versions resulting in falsely elevated numerical values. | Class 2 Device Recall Access SARS Cov2 lgG II Reagent Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Access SARS Cov2 lgG II Reagent FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Access SARS Cov2 lgG II Reagent Date Initiated by Firm May 07, 2021 Create Date June 21, 2021 Recall Status 1 Terminated 3 on March 27, 2023 Recall Number Z-1905-2021 Recall Event ID 87960 Product Classification Reagent, coronavirus serological - Product Code QKO Product Access SARS CoV-2 lgG II Reagent, REF: C69057, IVD CE UDI: (01)15099590742744 Code Information All Lots Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact Annette Hellie 714-831-7388 Manufacturer Reason for Recall SARS-CoV-2 lgG II numerical results may potentially be multiplied by a factor of 1000 on systems running with assay protocol file (APF) and access assay file (AAF) versions resulting in falsely elevated numerical values. FDA Determined Cause 2 Software Design Change Action On May 7, 2021, Beckman Coulter issued an "URGENT MEDICAL DEVICE RECALL" to affected consignees. The notification was emailed and sent via First Class Mail. In addition to informing consignees about the recall, the firm ask consignees to take the following actions: ************1. Do not install any of the APF/AAF versions listed in the table.************2. Determine your current APF/AAF version.******" From the Main Menu: select Configure F8 System Setup F1 System Revisions F1.******" Verify the APF/AAF version displayed on the screen.************3 If the APF/AAF version is listed in the table:******" Stop using the Access SARS-CoV-2 IgG II assay. Contact the Beckman Coulter Customer Technical Support Center at 1-800-854-3633 for further instruction.******" Review all Access SARS-CoV-2 IgG II patient test results from April 13, 2021 onward.******" Do not install any additional APF/AAF versions prior to contacting Beckman Coulter.************4. If the APF/AAF version is lower than those listed in the table, you are not impacted. Continue operating your system under existing instructions.************5. Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter.************6. So that we are assured you have received this important communication, please respond within 10 days in one of the following ways:******" Electronically, if you received this communication via email.******" Manually, complete and return the enclosed Response Form.************7. If you have any questions regarding this notice, please contact our Customer Support Center:******" From our website: http://www.beckmancoulter.com******" By phone: call 1-800-854-3633 in the United States. Quantity in Commerce 13,337 units Distribution US: Alabama,******Alaska,******Arizona,******Arkansas,******California,******Florida,******Georgia,******Illinois,******Iowa,******Kentucky,******Louisiana,******Maine,******Maryland,******Massachusetts,******Michigan,******Mississippi,******Missouri,******New Hampshire,******New Jersey,******New York,******North Dakota,******Ohio,******Oklahoma,******Pennsylvania,******Puerto Rico,******Tennessee,******Texas,******Virginia,******West Virginia,******Wisconsin,******Wyoming,************OUS: Andorra, Austria, Croatia, Czech Republic, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, Israel, Italy, Japan, Jordan, Poland, Portugal, San Marino, Slovakia, Slovenia, Spain, United Arab Emirates Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On May 7, 2021, Beckman Coulter issued an "URGENT MEDICAL DEVICE RECALL" to affected consignees. The notification was emailed and sent via First Class Mail. In addition to informing consignees about the recall, the firm ask consignees to take the following actions: ************1. Do not install any of the APF/AAF versions listed in the table.************2. Determine your current APF/AAF version.******" From the Main Menu: select Configure F8 System Setup F1 System Revisions F1.******" Verify the APF/AAF version displayed on the screen.************3 If the APF/AAF version is listed in the table:******" Stop using the Access SARS-CoV-2 IgG II assay. Contact the Beckman Coulter Customer Technical Support Center at 1-800-854-3633 for further instruction.******" Review all Access SARS-CoV-2 IgG II patient test results from April 13, 2021 onward.******" Do not install any additional APF/AAF versions prior to contacting Beckman Coulter.************4. If the APF/AAF version is lower than those listed in the table, you are not impacted. Continue operating your system under existing instructions.************5. Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter.************6. So that we are assured you have received this important communication, please respond within 10 days in one of the following ways:******" Electronically, if you received this communication via email.******" Manually, complete and return the enclosed Response Form.************7. If you have any questions regarding this notice, please contact our Customer Support Center:******" From our website: http://www.beckmancoulter.com******" By phone: call 1-800-854-3633 in the United States. | SoftwareDesignChange
72 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=186007 | Z-1333-2021 | BIOCARE Medical , IntelliPath FLX, Automated Staining Instrument, Ref: "***Model Number:IPS0001US and IPS0001INTL containing software version 3.5.3.1 | IntelliPath FLX, Automated Staining Instrument | 2 | CDRH | 2021-04-01 |  | 2023-01-31 09:21 | Biocare Medical, LLC | Their is a potential that the automated staining instrument with software version 3.5.3.1 may move with random speeds in random direction across its range of motion. This could result in the instrument hitting end stops or running into other components on the working deck. This could also result in wash buffer being dispensed on random locations. ****** | Class 2 Device Recall IntelliPath FLX, Automated Staining Instrument Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall IntelliPath FLX, Automated Staining Instrument FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall IntelliPath FLX, Automated Staining Instrument Date Initiated by Firm February 25, 2021 Create Date April 01, 2021 Recall Status 1 Terminated 3 on January 31, 2023 Recall Number Z-1333-2021 Recall Event ID 87426 Product Classification Slide stainer, automated - Product Code KPA Product BIOCARE Medical , IntelliPath FLX, Automated Staining Instrument, Ref: "***Model Number:IPS0001US and IPS0001INTL containing software version 3.5.3.1 Code Information Affected US Serial Numbers: BCIP01-0021, BCIP01-0033, BCIP01-0090, BCIP01-0106, BCIP01-0112, BCIP01-0148, BCIP01-0151, BCIP01-0187, BCIP01-0189, BCIP01-0198, BCIP01-0201, BCIP01-0216, BCIP01-0221, BCIP01-0249, BCIP01-0262, BCIP01-0265, BCIP01-0266, BCIP01-0268, BCIP01-0279, BCIP01-0285, BCIP01-0307, BCIP01-0314, BCIP01-0319, BCIP01-0328, BCIP01-0341, BCIP01-0346, BCIP01-0354, BCIP01-0386, BCIP01-0387, BCIP01-0388, BCIP01-0390, BCIP01-0395, BCIP01-0396, BCIP01-0413, BCIP01-0426, BCIP01-0430, BCIP01-0431, BCIP01-0432, BCIP01-0434, BCIP01-0443, BCIP01-0450, BCIP01-0463, BCIP01-0464, BCIP01-0467, BCIP01-0482, BCIP01-0484, BCIP01-0487, BCIP01-0491, BCIP01-0492, BCIP01-0495, BCIP01-0500, BCIP01-0502, BCIP01-0508, BCIP01-0517, BCIP01-0526, BCIP01-0534, BCIP01-0543, BCIP01-0545, BCIP01-0546, BCIP01-0550, BCIP01-0553, BCIP01-0555, BCIP01-0560, BCIP01-0562, BCIP01-0570, BCIP01-0571, BCIP01-0572, BCIP01-0573 Affected OUS serial numbers: BCIP01-0454, BCIP01-0515, BCIP01-0538, BCIP01-0556 Recalling Firm/ Manufacturer Biocare Medical, LLC 60 Berry Dr Pacheco CA 94553-5601 For Additional Information Contact Theodore Morris 925-603-8002 Manufacturer Reason for Recall Their is a potential that the automated staining instrument with software version 3.5.3.1 may move with random speeds in random direction across its range of motion. This could result in the instrument hitting end stops or running into other components on the working deck. This could also result in wash buffer being dispensed on random locations. FDA Determined Cause 2 Software Design Change Action On 2/25/2021, BIOCARE Medical emailed and "Medical Device Recall" Notification to all affected consignees. In addition to informing consignees about the recall, the firm requested consignees take the following actions: ************1. Do not use the product in the Multi Batching mode.******2. Do not interrupt a run during a Single Batch mode run or access the working deck while the instrument is in operation.******3. Always run positive controls per CAP and CLIA guidance to avoid the potential for false negative result from contaminated reagents.******4. If the malfunction does occur:******" Do not open the cover and access the working deck.******" Power cycle (turn off, then on) the device to reset it.******" Look for spilled Wash Buffer solution on the lab bench and the floor.******" Carefully examine the working deck and to determine if wash buffer has been spilled. If there is evidence of spilled wash buffer near, on or in reagent vials those vials should be disposed of to prevent staining performance issues.************5. PLEASE INDICATE YOUR UNDERSTANDING AND ACKNOWLEGEMENT OF THIS NOTICE BY COMPLETING THE STEPS LISTED BELOW******1. Locate all intelliPATH FLX Automated Staining Instruments in your possession. Verify the software version to confirm your device is impacted by this notice.******2. Ensure all personnel who may utilize the intelliPATH FLX are aware of this recall.******3. Complete and return the attached Response Form to Biocare Medical to acknowledge your receipt and understanding of this Recall Notice within 10 days of receipt.******No return of product is necessary. This notification is being provided for awareness only. Biocare Medical continues to investigate this matter and will follow up with affected customers.******Adverse events or quality problems experienced with the use of this product must be reported to Biocare Medical via techsupport@biocare.net. Quantity in Commerce 72 units Distribution US: AL, AZ, CA, CT, FL, IL, MA, MD, MN, NC, OH, OR, PA, WA, WY************OUS: Germany, Netherlands Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 2/25/2021, BIOCARE Medical emailed and "Medical Device Recall" Notification to all affected consignees. In addition to informing consignees about the recall, the firm requested consignees take the following actions: ************1. Do not use the product in the Multi Batching mode.******2. Do not interrupt a run during a Single Batch mode run or access the working deck while the instrument is in operation.******3. Always run positive controls per CAP and CLIA guidance to avoid the potential for false negative result from contaminated reagents.******4. If the malfunction does occur:******" Do not open the cover and access the working deck.******" Power cycle (turn off, then on) the device to reset it.******" Look for spilled Wash Buffer solution on the lab bench and the floor.******" Carefully examine the working deck and to determine if wash buffer has been spilled. If there is evidence of spilled wash buffer near, on or in reagent vials those vials should be disposed of to prevent staining performance issues.************5. PLEASE INDICATE YOUR UNDERSTANDING AND ACKNOWLEGEMENT OF THIS NOTICE BY COMPLETING THE STEPS LISTED BELOW******1. Locate all intelliPATH FLX Automated Staining Instruments in your possession. Verify the software version to confirm your device is impacted by this notice.******2. Ensure all personnel who may utilize the intelliPATH FLX are aware of this recall.******3. Complete and return the attached Response Form to Biocare Medical to acknowledge your receipt and understanding of this Recall Notice within 10 days of receipt.******No return of product is necessary. This notification is being provided for awareness only. Biocare Medical continues to investigate this matter and will follow up with affected customers.******Adverse events or quality problems experienced with the use of this product must be reported to Biocare Medical via techsupport@biocare.net. | SoftwareDesignChange
73 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=185877 | Z-1330-2021 | Product Label - ACUSON Sequoia *** Power Input: 100-240V~, 13.0-5.4A, 50-60Hz *** (240)11148775 Model******User Manual states - ACUSON Sequoia Diagnostic Ultrasound Systems | ACUSON Sequoia ultrasound systems | 2 | CDRH | 2021-04-01 |  |  | Siemens Medical Solutions USA, Inc. | The ultrasound system averages the Mean Sac Diameter (MSD) and Gestational Sac Diameter (GSD) into the Estimated Date of Delivery (EDD) calculation. This may result in an incorrect EDD, which may influence patient management decisions regarding induction of labor and elective caesarean delivery, which may result in premature births. | Class 2 Device Recall ACUSON Sequoia ultrasound systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ACUSON Sequoia ultrasound systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ACUSON Sequoia ultrasound systems Date Initiated by Firm February 04, 2021 Create Date April 01, 2021 Recall Status 1 Open 3 , Classified Recall Number Z-1330-2021 Recall Event ID 87362 510(K)Number K193257 Product Classification System, imaging, pulsed doppler, ultrasonic - Product Code IYN Product Product Label - ACUSON Sequoia *** Power Input: 100-240V~, 13.0-5.4A, 50-60Hz *** (240)11148775 Model************User Manual states - ACUSON Sequoia Diagnostic Ultrasound Systems Code Information Model: 11148775 Equipment Serial Numbers (U.S.) 800015 800101 800103 800104 800105 800106 800107 800108 800109 800111 800112 800113 800116 800118 800119 800121 800134 800138 800139 800141 800142 800143 800149 800150 800151 800152 800153 800155 800157 800158 800161 800165 800168 800172 800174 800175 800176 800177 800178 800182 800183 800184 800185 800186 800189 800190 800191 800192 800193 800194 800195 800196 800198 800199 800201 800203 800204 800205 800208 800209 800210 800211 800215 800217 800220 800221 800223 800225 800226 800227 800228 800229 800232 800233 800234 800235 800237 800239 800241 800242 800243 800244 800247 800249 800251 800254 800256 800257 800258 800259 800260 800263 800265 800267 800268 800270 800273 800274 800278 800280 800281 800285 800286 800291 800292 800293 800295 800300 800301 800302 800303 800304 800305 800308 800313 800314 800317 800319 800320 800321 800324 800328 800333 800335 800336 800339 800341 800342 800348 800349 800355 800360 800361 800363 800366 800369 800370 800375 800376 800383 800390 800393 800397 800405 800408 800411 800418 800424 800427 800428 800429 800433 800435 800436 800440 800451 800462 800463 800464 800466 800467 800469 800473 800474 800475 800480 800481 800486 800489 800490 800495 800498 800505 800506 800509 800511 800515 800516 800518 800521 800523 800525 800528 800532 800535 800540 800546 800549 800550 800552 800555 800556 800559 800560 800564 800566 800567 800579 800581 800583 800590 800592 800594 800597 800598 800607 800609 800611 800618 800623 800627 800628 800630 800631 800634 800638 800639 800640 800645 800648 800650 800651 800652 800653 800657 800661 800662 800667 800668 800671 800672 800673 800674 800675 800677 800680 800683 800686 800688 800693 800694 800696 800701 800703 800707 800708 800735 800738 800739 800741 800742 800743 800745 800747 800749 800752 800754 800756 800758 800760 800761 800762 800763 800771 800775 800786 800787 800788 800789 800791 800800 800805 800806 800807 800809 800813 800821 800834 800835 800854 800859 800861 800862 800863 800865 800866 800868 800869 800870 800871 800873 800874 800876 800878 800879 800880 800881 800882 800883 800884 800885 800887 800889 800890 800896 800901 800903 800904 800906 800907 800908 800912 800913 800914 800915 800917 800918 800920 800931 800933 800936 800940 800944 800945 800947 800952 800955 800959 800961 800963 800964 800979 800980 800993 800994 800998 801002 801003 801004 801005 801006 801007 801008 801009 801010 801011 801013 801014 801015 801016 801017 801018 801019 801021 801022 801023 801025 801026 801033 801034 801047 801080 801084 801087 801088 801089 801116 801117 801121 801123 801124 801125 801126 801127 801128 801129 801130 801131 801133 801134 801135 801136 801137 801138 801139 801140 801141 801144 801145 801147 801150 801152 801153 801156 801165 801166 801174 801177 801178 801191 801192 801193 801194 801196 801200 801202 801209 801211 801213 801223 801225 801230 801232 801234 801244 801245 801247 801248 801249 801259 801261 801271 801278 801279 801286 801300 801302 801306 801310 801311 801313 801314 801315 801316 801321 801324 801325 801326 801328 801330 801334 801336 801339 801345 801346 801347 801348 801352 801353 801356 801357 801358 801360 801363 801365 801366 801371 801372 801378 801381 801386 801390 801396 801400 801402 801405 801406 801408 801413 801414 801421 801423 801424 801426 801427 801430 801432 801435 801436 801437 801439 801441 801442 801443 801444 801447 801449 801450 801451 801452 801454 801458 801462 801463 801464 801465 801470 801471 801473 801474 801475 801476 801481 801483 801484 801488 801491 801492 801494 801495 801497 801499 801501 801502 801503 801506 801509 801510 801514 801515 801516 801517 801519 801521 801524 801527 801528 801529 801530 801531 801532 801534 801535 801536 801537 801538 801539 801540 801541 801542 801543 801544 801545 801548 801549 801550 801552 801553 801554 801555 801556 801558 801559 801564 801565 801567 801572 801574 801576 801579 801594 801596 801598 801605 801609 801610 801611 801629 801631 801632 801636 801639 801641 801642 801643 801645 801647 801648 801649 801650 801655 801660 801663 801675 801678 801680 801685 801708 801714 801759 801760 801768 801771 801812 801903 801911 802004 802008 802017 802019 802022 802028 802029 802032 802033 802093 802109 802125 802128 802146 802164 802217 802245 802263 Equipment Serial Numbers (O.U.S.): 800344 800438 800932 801787 801851 801020 801927 801982 801988 802069 802073 802078 802079 802081 802083 802084 802086 802123 802131 802135 802140 802147 802180 802181 802183 802185 802186 802187 802188 802189 800114 800115 800117 800120 800122 800123 800126 800127 800129 800130 800133 800135 800136 800140 800147 800148 800159 800162 800163 800166 800167 800169 800170 800171 800181 800187 800197 800206 800207 800212 800214 800218 800219 800224 800230 800231 800236 800245 800246 800248 800250 800252 800264 800266 800271 800277 800279 800282 800284 800288 800298 800309 800311 800312 800315 800323 800329 800334 800337 800338 800340 800343 800345 800346 800350 800356 800358 800362 800364 800365 800367 800368 800377 800379 800380 800382 800388 800391 800392 800398 800399 800400 800401 800404 800406 800409 800410 800412 800413 800416 800417 800421 800423 800425 800426 800430 800431 800432 800434 800439 800447 800449 800450 800456 800457 800458 800472 800477 800478 800483 800484 800488 800492 800494 800500 800503 800508 800517 800520 800527 800529 800530 800531 800536 800542 800543 800544 800548 800557 800561 800563 800572 800575 800584 800585 800586 800591 800599 800600 800601 800602 800605 800606 800610 800612 800613 800614 800617 800620 800621 800622 800624 800625 800632 800636 800637 800642 800644 800646 800647 800649 800654 800658 800659 800663 800664 800669 800678 800682 800695 800698 800699 800700 800702 800705 800706 800730 800731 800734 800737 800750 800753 800764 800766 800767 800769 800776 800779 800780 800783 800785 800793 800794 800802 800803 800804 800817 800818 800820 800822 800825 800828 800830 800832 800836 800838 800840 800841 800845 800846 800847 800850 800852 800855 800856 800857 800860 800875 800892 800894 800897 800899 800900 800909 800911 800916 800919 800921 800922 800923 800924 800926 800927 800928 800930 800934 800937 800941 800942 800946 800948 800949 800950 800965 800968 800974 800976 800977 800978 800981 800982 800984 800988 800992 800996 800997 800999 801012 801024 801032 801036 801037 801041 801044 801061 801065 801070 801074 801077 801081 801082 801083 801086 801149 801151 801162 801170 801175 801205 801212 801218 801238 801243 801252 801253 801254 801255 801257 801269 801303 801305 801308 801309 801317 801318 801331 801332 801337 801342 801343 801344 801354 801355 801359 801364 801367 801368 801370 801373 801375 801376 801380 801383 801384 801391 801392 801393 801399 801407 801410 801415 801417 801420 801422 801428 801446 801453 801456 801466 801467 801469 801472 801480 801487 801493 801496 801498 801505 801507 801508 801511 801513 801520 801522 801526 801533 801546 801547 801551 801560 801561 801562 801563 801569 801575 801577 801578 801584 801586 801589 801591 801592 801595 801597 801602 801603 801606 801607 801624 801626 801638 801644 801646 801651 801652 801653 801654 801656 801661 801664 801669 801670 801673 801736 801761 801762 801763 801764 801765 801766 801769 801772 801773 801774 801777 801778 801779 801780 801781 801782 801783 801786 801788 801791 801792 801793 801794 801795 801796 801797 801798 801799 801800 801801 801802 801803 801805 801806 801809 801810 801811 801813 801815 801816 801817 801818 801819 801820 801821 801823 801824 801828 801829 801830 801831 801835 801836 801837 801838 801839 801842 801844 801846 801847 801848 801849 801853 801859 801862 801875 801876 801889 801892 801898 801901 801902 801907 801908 801915 801933 801934 801935 801936 801937 801965 801971 801972 801973 801976 801978 801983 801984 801991 801992 801994 801995 802003 802007 802009 802010 802012 802015 802018 802020 802021 802023 802024 802070 802072 802080 802087 802089 802096 802100 802103 802104 802107 802108 802111 802112 802113 802114 802115 802117 802118 802119 802122 802126 802127 802129 802130 802133 802134 802136 802137 802143 802193 802213 802219 802220 802229 802230 802241 802251 802252 802258 802269 802273 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc. 22010 SE 51st St Issaquah WA 98029-7298 For Additional Information Contact Mr. Khalil Thomas 425-392-9180 Manufacturer Reason for Recall The ultrasound system averages the Mean Sac Diameter (MSD) and Gestational Sac Diameter (GSD) into the Estimated Date of Delivery (EDD) calculation. This may result in an incorrect EDD, which may influence patient management decisions regarding induction of labor and elective caesarean delivery, which may result in premature births. FDA Determined Cause 2 Software Design Change Action On 02/04/2021, the firm sent an "CUSTOMER SAFETY ADVISORY NOTIFICATION" email to customer informing them that the ultrasound system averages the Mean Sac Diameter (MSD) and Gestational Sac Diameter (GSD) into the Estimated Date of Delivery (EDD) calculation. This may result in an incorrect EDD. An incorrect EDD may influence patient management decisions regarding induction of labor and elective caesarean delivery, which may result in premature births.************-Customers are instructed to avoid this issue, delete any MSD or GSD values from the patient report if other biometric parameters are measured.******-Recalling Firm recommends a review of any early obstetric or gynecology ultrasound results for current pregnancies which produced a gestational age using MSD or GSD as an input.******-The issue will be corrected by the Recalling Firm with an updated to the software and a Customer Service Engineer will contact customers to schedule a facility visit to update the software or updates will be issued via the Remote Update Handling process.******-Customers are instructed to inform their personnel accordingly.************Additional questions or assistance - contact Quality & Regulatory at 732-331-6672 or email khalil.thomas@siemens-healthineers.com. Quantity in Commerce 1,172 systems Distribution U.S.: AK, AL, AR, AZ, CA, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, Puerto Rico, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY************O.U.S.: U.A.E., Algeria, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Indonesia, Iran, Israel, Italy, Japan, Kazakhstan, Kuwait, Mexico, Morocco, Netherlands, New Zealand, Oman, Philippines, Poland, Portugal, Qatar, Republic Korea, Romania, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom, and Vietnam Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = IYN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 02/04/2021, the firm sent an "CUSTOMER SAFETY ADVISORY NOTIFICATION" email to customer informing them that the ultrasound system averages the Mean Sac Diameter (MSD) and Gestational Sac Diameter (GSD) into the Estimated Date of Delivery (EDD) calculation. This may result in an incorrect EDD. An incorrect EDD may influence patient management decisions regarding induction of labor and elective caesarean delivery, which may result in premature births.************-Customers are instructed to avoid this issue, delete any MSD or GSD values from the patient report if other biometric parameters are measured.******-Recalling Firm recommends a review of any early obstetric or gynecology ultrasound results for current pregnancies which produced a gestational age using MSD or GSD as an input.******-The issue will be corrected by the Recalling Firm with an updated to the software and a Customer Service Engineer will contact customers to schedule a facility visit to update the software or updates will be issued via the Remote Update Handling process.******-Customers are instructed to inform their personnel accordingly.************Additional questions or assistance - contact Quality & Regulatory at 732-331-6672 or email khalil.thomas@siemens-healthineers.com. | SoftwareDesignChange
74 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184992 | Z-0785-2021 | TomoTherapy Treatment Delivery System with iDMS - Product Usage: used as an integrated system for the planning and precise delivery of radiation therapy, stereotactic radiotherapy, or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. | TomoTherapy Treatment Delivery System with iDMS | 2 | CDRH | 2021-01-07 |  | 2022-09-26 05:49 | Accuray Incorporated | "MLC tickle error" may result in the delivered dose to effectively rotate from the planned dose. | Class 2 Device Recall TomoTherapy Treatment Delivery System with iDMS Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TomoTherapy Treatment Delivery System with iDMS FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TomoTherapy Treatment Delivery System with iDMS Date Initiated by Firm December 10, 2020 Create Date January 07, 2021 Recall Status 1 Terminated 3 on September 26, 2022 Recall Number Z-0785-2021 Recall Event ID 86993 510(K)Number K171837 Product Classification Accelerator, linear, medical - Product Code IYE Product TomoTherapy Treatment Delivery System with iDMS - Product Usage: used as an integrated system for the planning and precise delivery of radiation therapy, stereotactic radiotherapy, or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. Code Information Product Codes: 104430, 1018286, H-0000-0003, 1018283. System Version: 1.2 (also referred to as 5.6.2). Serial Numbers: 110105, 110219, 110221, 110267, 110284, 110290, 110308, 110368, Recalling Firm/ Manufacturer Accuray Incorporated 1240 Deming Way Madison WI 53717-1954 For Additional Information Contact Daniel Biank 608-824-2971 Manufacturer Reason for Recall "MLC tickle error" may result in the delivered dose to effectively rotate from the planned dose. FDA Determined Cause 2 Software Design Change Action Urgent Medical Device Recall "MLC tickle error" issue during TomoHelical deliver was sent to customers beginning 12/10/2020. Accuray is planning software releases that will address this issue. All TomoTherapy systems at affected software versions will be upgraded to these versions once released. Quantity in Commerce 8 units Distribution Worlwide distribution - US Nationwide distribution including in the states of CA, FL, IL, ND, NJ, OK, UT and the countries of Australia, China, Italy, Switzerland, United Arab Emirates. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Urgent Medical Device Recall "MLC tickle error" issue during TomoHelical deliver was sent to customers beginning 12/10/2020. Accuray is planning software releases that will address this issue. All TomoTherapy systems at affected software versions will be upgraded to these versions once released. | SoftwareDesignChange
75 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184985 | Z-0784-2021 | TomoTherapy Treatment System - Product Usage: used as an integrated system for the planning and precise delivery of radiation therapy, stereotactic radiotherapy, or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. | TomoTherapy Treatment System | 2 | CDRH | 2021-01-07 |  | 2022-09-26 05:49 | Accuray Incorporated | "MLC tickle error" may result in the delivered dose to effectively rotate from the planned dose. | Class 2 Device Recall TomoTherapy Treatment System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall TomoTherapy Treatment System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall TomoTherapy Treatment System Date Initiated by Firm December 10, 2020 Create Date January 07, 2021 Recall Status 1 Terminated 3 on September 26, 2022 Recall Number Z-0784-2021 Recall Event ID 86993 510(K)Number K121934 Product Classification Accelerator, linear, medical - Product Code IYE Product TomoTherapy Treatment System - Product Usage: used as an integrated system for the planning and precise delivery of radiation therapy, stereotactic radiotherapy, or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. Code Information Product Codes: H-0000-0003, 104430, 1018284. System Version 5.1.5. Serial Numbers: 110210, 110319, 110340, 110366, 110383, 110418, 110468 Recalling Firm/ Manufacturer Accuray Incorporated 1240 Deming Way Madison WI 53717-1954 For Additional Information Contact Daniel Biank 608-824-2971 Manufacturer Reason for Recall "MLC tickle error" may result in the delivered dose to effectively rotate from the planned dose. FDA Determined Cause 2 Software Design Change Action Urgent Medical Device Recall "MLC tickle error" issue during TomoHelical deliver was sent to customers beginning 12/10/2020. Accuray is planning software releases that will address this issue. All TomoTherapy systems at affected software versions will be upgraded to these versions once released. Quantity in Commerce 7 units Distribution Worlwide distribution - US Nationwide distribution including in the states of CA, FL, IL, ND, NJ, OK, UT and the countries of Australia, China, Italy, Switzerland, United Arab Emirates. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Urgent Medical Device Recall "MLC tickle error" issue during TomoHelical deliver was sent to customers beginning 12/10/2020. Accuray is planning software releases that will address this issue. All TomoTherapy systems at affected software versions will be upgraded to these versions once released. | SoftwareDesignChange
76 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180659 | Z-2956-2020 | NovaPACS Diagnostic Viewer. NovaPACS is a picture archiving and communication system software that retrieves, distributes, and display images and data from different modalities including CT, MR, US, CR/DX, NM, PT, and XA. | NovaPACS | 2 | CDRH | 2020-09-11 | 2020-09-12 18:07 | 2021-06-30 13:27 | Novarad Corporation | The firm received a report of an atypical dataset being generated. When using the cross-localization feature and images from a modality that generates asymmetrical images, the cross-localization reference line may be inaccurately placed on any of the corresponding images that are open. | Class 2 Device Recall NovaPACS Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall NovaPACS FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall NovaPACS Date Initiated by Firm September 10, 2019 Date Posted September 12, 2020 Recall Status 1 Terminated 3 on June 30, 2021 Recall Number Z-2956-2020 Recall Event ID 85347 510(K)Number K171754 Product Classification System, image processing, radiological - Product Code LLZ Product NovaPACS Diagnostic Viewer. NovaPACS is a picture archiving and communication system software that retrieves, distributes, and display images and data from different modalities including CT, MR, US, CR/DX, NM, PT, and XA. Code Information NovaPACS versions 8.6.8, 8.6.9, 8.6.11, 8.6.12, 8.6.14, 8.7.4, 8.7.6, and 8.7.8 Recalling Firm/ Manufacturer Novarad Corporation 752 E 1180 S Ste 200 American Fork UT 84003-3561 For Additional Information Contact 801-642-1001 Manufacturer Reason for Recall The firm received a report of an atypical dataset being generated. When using the cross-localization feature and images from a modality that generates asymmetrical images, the cross-localization reference line may be inaccurately placed on any of the corresponding images that are open. FDA Determined Cause 2 Software Design Change Action The firm issued their ***Urgent Medical Device Correction Notice*** on 09/24/2019 and letters were sent via certified mail. ******The issue affected versions 8.6.8, 8.6.9, 8.6.11, 8.6.12, 8.6.14, 8.7.4, 8.7.6, and 8.7.8. Customers were instructed to view in the NovaPACS Study Browser***'s menu what version they are running by selecting "Help" and then selecting "About".******The firm informs that the issue has been fixed in versions 8.6.14.62 and 8.7.8.12 and will be in all future releases of the software. It is further explained that you can eliminate the possibility of the malfunction occurring by contacting Technical Support to arrange for an upgrade to a corrected version of the software. This correction is available to all affected customers at no cost and will be included in all future upgrades. ******Customers will need to FAX, mail, or email the Acknowledgement and Receipt Form back to the firm. For questions or concerns regarding the notification or would like to request an upgrade, you can contact Technical Support at (801) 221-5895. Quantity in Commerce 197 units Distribution Worldwide distribution - U.S. Nationwide distribution including in the states of AK, AZ, CA, CO, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. The countries of Canada, Colombia, Dominican Republic, Indonesia, Jordan, Mexico, Peru, Phillipines, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm issued their ***Urgent Medical Device Correction Notice*** on 09/24/2019 and letters were sent via certified mail. ******The issue affected versions 8.6.8, 8.6.9, 8.6.11, 8.6.12, 8.6.14, 8.7.4, 8.7.6, and 8.7.8. Customers were instructed to view in the NovaPACS Study Browser***'s menu what version they are running by selecting "Help" and then selecting "About".******The firm informs that the issue has been fixed in versions 8.6.14.62 and 8.7.8.12 and will be in all future releases of the software. It is further explained that you can eliminate the possibility of the malfunction occurring by contacting Technical Support to arrange for an upgrade to a corrected version of the software. This correction is available to all affected customers at no cost and will be included in all future upgrades. ******Customers will need to FAX, mail, or email the Acknowledgement and Receipt Form back to the firm. For questions or concerns regarding the notification or would like to request an upgrade, you can contact Technical Support at (801) 221-5895. | SoftwareDesignChange
77 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=182679 | Z-2886-2020 | Match Point System******The Match Point System is intended to by uses as a surgical instrument to assist in the intraoperative positioning of glenoid components used with total and reverse shoulder arthroplasty by referencing anatomic landmarks of the shoulder that are identifiable on preoperative CT-image scans. | Match Point System Patient specific instrumentation for shoulder arthroplasty | 2 | CDRH | 2020-08-26 |  | 2020-12-23 11:52 | Materialise N.V. | The procedure side indicated in the top header of the Shoulder Case Planning Report incorrectly indicated the procedure side as "Left" for surgical plans for a "right" procedure side. | Class 2 Device Recall Match Point System Patient specific instrumentation for shoulder arthroplasty Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Match Point System Patient specific instrumentation for shoulder arthroplasty FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Match Point System Patient specific instrumentation for shoulder arthroplasty Date Initiated by Firm July 17, 2020 Create Date August 26, 2020 Recall Status 1 Terminated 3 on December 23, 2020 Recall Number Z-2886-2020 Recall Event ID 86113 510(K)Number K153602 K170893 Product Classification Prosthesis, shoulder, semi-constrained, metal/polymer cemented - Product Code KWS Product Match Point System************The Match Point System is intended to by uses as a surgical instrument to assist in the intraoperative positioning of glenoid components used with total and reverse shoulder arthroplasty by referencing anatomic landmarks of the shoulder that are identifiable on preoperative CT-image scans. Code Information ID #s: 804-03-057, 804-25-148, 804-03-058, 804,07,022 UDI #s: E314804030570, E314804251480, E314804070220 Recalling Firm/ Manufacturer Materialise N.V. Technologielaan 15 Heverlee Belgium For Additional Information Contact Customer Service Dept/ 734-259-7014 Manufacturer Reason for Recall The procedure side indicated in the top header of the Shoulder Case Planning Report incorrectly indicated the procedure side as "Left" for surgical plans for a "right" procedure side. FDA Determined Cause 2 Software Design Change Action The firm, materialise, sent Recall letters dated 17 July 2020 informing customers of the issue. This letter informs the customer that Materialise has corrected the Shoulder Case Planning Report in the SurgiCase online system.************If you have any questions, contact Global Quality and Regulatory Manager Medical, at 734-259-7017 or email: Jenny.Jones@materialise.com. Quantity in Commerce 128 Distribution Nationwide distribution to: US including PR and International to UK and Australia. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = KWS - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, materialise, sent Recall letters dated 17 July 2020 informing customers of the issue. This letter informs the customer that Materialise has corrected the Shoulder Case Planning Report in the SurgiCase online system.************If you have any questions, contact Global Quality and Regulatory Manager Medical, at 734-259-7017 or email: Jenny.Jones@materialise.com. | SoftwareDesignChange
78 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=182369 | Z-2750-2020 | CARESCAPE ONE, Physiological Patient Monitor - Product Usage: is both a multi-parameter physiological patient monitor and an accessory to a multi-parameter patient monitor intended for use in multiple areas and intra-hospital transport within a professional healthcare facility. | CARESCAPE ONE | 2 | CDRH | 2020-08-04 | 2020-08-04 21:47 | 2024-08-19 16:43 | GE Healthcare, LLC | CARESCAPE ONE may not provide visual and audible alarms for Ventricular Fibrillation (V Fib), if V Fib occurs at the time CARESCAPE ONE is docked to a CARESCAPE B450/B650/B850 host monitor. | Class 2 Device Recall CARESCAPE ONE Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall CARESCAPE ONE FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall CARESCAPE ONE Date Initiated by Firm June 22, 2020 Date Posted August 04, 2020 Recall Status 1 Terminated 3 on August 19, 2024 Recall Number Z-2750-2020 Recall Event ID 85979 510(K)Number K200494 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product CARESCAPE ONE, Physiological Patient Monitor - Product Usage: is both a multi-parameter physiological patient monitor and an accessory to a multi-parameter patient monitor intended for use in multiple areas and intra-hospital transport within a professional healthcare facility. Code Information Model: 2087075-001 Serial Numbers: SNA20210164SA, SNA20210031SA, SNA20210045SA, SNA20210064SA, SNA20210319SA, SNA20210330SA, SNA20210348SA, SNA20210361SA, SNA20210362SA, SNA20210375SA, SNA20210376SA, SNA20210377SA, SNA20220003SA, SNA20220005SA, SNA20220007SA, SNA20220018SA, SNA20220022SA, SNA20220026SA, SNA20220045SA, SNA20220047SA, SNA18140024SA, SNA18340072SA, SNA18370025SA, SNA18380001SA, SNA18390009SA, SNA18430002SA, SNA18430059SA, SNA18430088SA, SNA18430116SA, SNA19390023SA, SNA19400056SA, SNA19400143SA, SNA19440048SA, SNA19440069SA, SNA19440071SA, SNA19440122SA, SNA19450008SA, SNA19450025SA, SNA19450128SA, SNA19450136SA, SNA19500072SA, SNA19500073SA, SNA19500080SA, SNA19500084SA, SNA19500086SA, SNA19510001SA, SNA19510005SA, SNA19510007SA, SNA19510012SA, SNA19510027SA, SNA19510053SA, SNA19510055SA, SNA19510103SA, SNA20020119SA, SNA20030070SA, SNA20030071SA, SNA20170022SA, SNA20170080SA, SNA20170082SA, SNA20170112SA, SNA20170113SA, SNA20170116SA, SNA20170117SA, SNA20170118SA, SNA20170120SA, SNA20170131SA, SNA20170134SA, SNA20170136SA, SNA20170139SA, SNA20170145SA, SNA20170160SA, SNA20170161SA, SNA20170164SA, SNA20170166SA, SNA20170168SA, SNA20170169SA, SNA20170171SA, SNA20170173SA, SNA20170176SA, SNA20170181SA, SNA20170182SA, SNA20170194SA, SNA20170197SA, SNA20170198SA, SNA20170232SA, SNA20170237SA, SNA20170264SA, SNA20180002SA, SNA20180008SA, SNA20180009SA, SNA20180011SA, SNA20180012SA, SNA20180017SA, SNA20180018SA, SNA20180023SA, SNA20180025SA, SNA20180026SA, SNA20180027SA, SNA20180028SA, SNA20180029SA, SNA20180030SA, SNA20180031SA, SNA20180032SA, SNA20180040SA, SNA20180042SA, SNA20180043SA, SNA20180044SA, SNA20180071SA, SNA20180073SA, SNA20180074SA, SNA20180075SA, SNA20180076SA, SNA20180080SA, SNA20180096SA, SNA20180101SA, SNA20180102SA, SNA20180109SA, SNA20180117SA, SNA20180144SA, SNA20180161SA, SNA20180162SA, SNA20190321SA, SNA20190331SA, SNA18140004SA, SNA18140005SA, SNA18140007SA, SNA18240025SA, SNA18240026SA, SNA18240030SA, SNA18240032SA, SNA18240044SA, SNA18240051SA, SNA18240054SA, SNA18240079SA, SNA18250010SA, SNA18250012SA, SNA18250017SA, SNA18250023SA, SNA18250026SA, SNA18250033SA, SNA18250034SA, SNA18250036SA, SNA18250044SA, SNA18250061SA, SNA18250072SA, SNA18250075SA, SNA18250078SA, SNA18250080SA, SNA18250081SA, SNA18250082SA, SNA18250086SA, SNA18250092SA, SNA18250093SA, SNA18440051SA, SNA18450094SA, SNA18450098SA, SNA18460013SA, SNA18460040SA, SNA18460072SA, SNA18460075SA, SNA18460083SA, SNA18460086SA, SNA18460089SA, SNA18460101SA, SNA19040008SA, SNA19050016SA, SNA19050025SA, SNA19050028SA, SNA19050030SA, SNA19050031SA, SNA19050033SA, SNA19050038SA, SNA19050043SA, SNA19050044SA, SNA19050045SA, SNA19050046SA, SNA19050049SA, SNA19050050SA, SNA19070071SA, SNA19070072SA, SNA19070073SA, SNA19070074SA, SNA19490074SA, SNA19500037SA, SNA19500043SA, SNA19500054SA, SNA19500059SA, SNA19510003SA, SNA20140029SA, SNA20140031SA, SNA20190324SA, SNA19500131SA, SNA20030120SA, SNA20040052SA, SNA20170179SA, SNA20170195SA, SNA20170203SA, SNA20180054SA, SNA20180057SA, SNA20180059SA, SNA20180077SA, SNA20180100SA, SNA20180103SA, SNA20180104SA, SNA20180113SA, SNA20180115SA, SNA20180142SA, SNA20180145SA, SNA20180146SA, SNA20180147SA, SNA20180154SA, SNA20180158SA, SNA20180163SA, SNA20180168SA, SNA20180183SA, SNA20180202SA, SNA20190024SA, SNA20190026SA, SNA20190030SA, SNA20190034SA, SNA20190038SA, SNA20190056SA, SNA20190057SA, SNA20190059SA, SNA20190060SA, SNA20190063SA, SNA20190106SA, SNA20190107SA, SNA20190109SA, SNA20190112SA, SNA20190114SA, SNA20190115SA, SNA20190116SA, SNA20190117SA, SNA20190119SA, SNA20190120SA, SNA20190121SA, SNA20190122SA, SNA20190124SA, SNA20190126SA, SNA20190127SA, SNA20190130SA, SNA20190131SA, SNA20190132SA, SNA20190133SA, SNA20190134SA, SNA20190137SA, SNA20190141SA, SNA20190142SA, SNA20190143SA, SNA20190144SA, SNA20190145SA, SNA20190146SA, SNA20190147SA, SNA20190148SA, SNA20190149SA, SNA20190152SA, SNA20190153SA, SNA20190155SA, SNA20190156SA, SNA20190159SA, SNA20190161SA, SNA20190174SA, SNA20190201SA, SNA20190204SA, SNA20190206SA, SNA20190209SA, SNA20190210SA, SNA20190214SA, SNA20190216SA, SNA20190217SA, SNA20190218SA, SNA20190219SA, SNA20190220SA, SNA20190221SA, SNA20190222SA, SNA20190223SA, SNA20190229SA, SNA20190230SA, SNA20190231SA, SNA20190232SA, SNA20190233SA, SNA20190235SA, SNA20190236SA, SNA20190237SA, SNA20190238SA, SNA20190239SA, SNA20190240SA, SNA20190247SA, SNA20190279SA, SNA20190289SA, SNA20190328SA, SNA20190339SA, SNA20200031SA, SNA20200081SA, SNA20200403SA. Recalling Firm/ Manufacturer GE Healthcare, LLC 3000 N Grandview Blvd Waukesha WI 53188-1615 For Additional Information Contact 800-437-1171 Manufacturer Reason for Recall CARESCAPE ONE may not provide visual and audible alarms for Ventricular Fibrillation (V Fib), if V Fib occurs at the time CARESCAPE ONE is docked to a CARESCAPE B450/B650/B850 host monitor. FDA Determined Cause 2 Software Design Change Action You can continue to use your CARESCAPE ONE to monitor patients. Follow the instructions below each time the CARESCAPE ONE is docked to a CARESCAPE B450/B650/B850 host monitor.******1. As per the User Manual, closely observe and monitor the patient***s clinical condition and ECG waveforms on the CARESCAPE B450/B650/B850 host monitor for V Fib during the Learning phase and for several minutes after the Learning phase to allow the algorithm to reach optimal detection performance.******2. As per the User Manual, if the Learning message on the host monitor does not clear after 30 seconds, select the Relearn QRS option from the ECG menu on the host monitor following these steps.****** Select the HR parameter window.****** Select Advanced > Page 1.****** Select Relearn QRS.******GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction.******After the CARESCAPE ONE Monitor has been updated, discontinue usage of any previous versions of CARESCAPE ONE software. Destroy any software media containing previous versions of CARESCAPE ONE software, including reimage, upgrade kits, and FRUs. If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. Quantity in Commerce 8,414 units Distribution Worldwide distribution - US Nationwide distribution including in the states of MT, NC, TX and WI. The countries of Australia, Austria, Bahrain,Belgium, Bolivia, Brazil ,Canada, Chile, Check Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, French Guiana, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kenya, Korea, Kuwait, Latvia, Lebanon, Lithuania, Macedonia, Malaysia, Malta, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Republic of Serbia, Russia, Saudi Arabia, Singapore, Saint Martin, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, United Arab Emirates, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action You can continue to use your CARESCAPE ONE to monitor patients. Follow the instructions below each time the CARESCAPE ONE is docked to a CARESCAPE B450/B650/B850 host monitor.******1. As per the User Manual, closely observe and monitor the patient***s clinical condition and ECG waveforms on the CARESCAPE B450/B650/B850 host monitor for V Fib during the Learning phase and for several minutes after the Learning phase to allow the algorithm to reach optimal detection performance.******2. As per the User Manual, if the Learning message on the host monitor does not clear after 30 seconds, select the Relearn QRS option from the ECG menu on the host monitor following these steps.****** Select the HR parameter window.****** Select Advanced > Page 1.****** Select Relearn QRS.******GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction.******After the CARESCAPE ONE Monitor has been updated, discontinue usage of any previous versions of CARESCAPE ONE software. Destroy any software media containing previous versions of CARESCAPE ONE software, including reimage, upgrade kits, and FRUs. If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. | SoftwareDesignChange
79 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=181805 | Z-2537-2020 | VistaSolution, physiological signal transmitter. A software graphical user interface for use by healthcare professionals to display physiological data collected by the wireless remote monitoring system in healthcare settings. This is a secondary, adjunct patient monitor and is not intended to replace existing standard-of-care patient monitoring. | VistaSolution Version 2.1 | 2 | CDRH | 2020-07-08 | 2020-07-10 10:21 | 2021-05-13 09:59 | Vitalconnect Inc. | A healthcare provider (HCP) noticed a patient had a lowered historic SpO2 reading that the software should have triggered a notification to the care team, but no notification was found. A software defect was encountered causing the lack of notification. The defect could result in exposing a patient to potential hypoxia for a short window of time. The firm is correcting the software to bring the performance back into compliance with specifications. | Class 2 Device Recall VistaSolution Version 2.1 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VistaSolution Version 2.1 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VistaSolution Version 2.1 Date Initiated by Firm May 21, 2020 Date Posted July 10, 2020 Recall Status 1 Terminated 3 on May 13, 2021 Recall Number Z-2537-2020 Recall Event ID 85802 510(K)Number K193343 Product Classification Transmitters and receivers, physiological signal, radiofrequency - Product Code DRG Product VistaSolution, physiological signal transmitter. A software graphical user interface for use by healthcare professionals to display physiological data collected by the wireless remote monitoring system in healthcare settings. This is a secondary, adjunct patient monitor and is not intended to replace existing standard-of-care patient monitoring. Code Information Catalog number: 6200, Vista Solution version 2.1, UDI: 00850490007139 Recalling Firm/ Manufacturer Vitalconnect Inc. 224 Airport Pkwy Ste 300 San Jose CA 95110-1022 For Additional Information Contact Tony Wilder 408-963-4600 Manufacturer Reason for Recall A healthcare provider (HCP) noticed a patient had a lowered historic SpO2 reading that the software should have triggered a notification to the care team, but no notification was found. A software defect was encountered causing the lack of notification. The defect could result in exposing a patient to potential hypoxia for a short window of time. The firm is correcting the software to bring the performance back into compliance with specifications. FDA Determined Cause 2 Software Design Change Action On May 21, 2020, the firm sent a "Urgent: Medical Device Recall" notification to all affected customers via E-mail and mail. The letter stated, that the third party Sp02 continuous device notifications may not indicate an alert as intended. Therefore, the health care provider may not be notified to check the patient. ******Potential Hazards: A patient may have hypoxia and the designated health care provider may not be notified. The firm will be turning off the ability to set custom notifications containing continuous SP02 until the issue is resolved. The firm tells customers to monitor the patients for SPO2, as if the patients are not connected to the recalled device.************The customers are informed to acknowledge the receipt of this notification by E-mail or via mail to the firm, email to: cmerrell@vitalconnect.com and if any questions, then contact: Tony Wilder at phone: (408) 963-4600 or support@vitalconnect.com Quantity in Commerce 7 Distribution U.S. Nationwide distribution including in the states of NJ, NY, VA, OUS: None Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = DRG - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On May 21, 2020, the firm sent a "Urgent: Medical Device Recall" notification to all affected customers via E-mail and mail. The letter stated, that the third party Sp02 continuous device notifications may not indicate an alert as intended. Therefore, the health care provider may not be notified to check the patient. ******Potential Hazards: A patient may have hypoxia and the designated health care provider may not be notified. The firm will be turning off the ability to set custom notifications containing continuous SP02 until the issue is resolved. The firm tells customers to monitor the patients for SPO2, as if the patients are not connected to the recalled device.************The customers are informed to acknowledge the receipt of this notification by E-mail or via mail to the firm, email to: cmerrell@vitalconnect.com and if any questions, then contact: Tony Wilder at phone: (408) 963-4600 or support@vitalconnect.com | SoftwareDesignChange
80 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=181744 | Z-2424-2020 | CARESCAPE Patient Data Module (PDM) used with the following monitors: CARESCAPE B450/B650/B850, SOLAR 8000M/I and Transport Pro, Model Numbers:***2042084-001***2094504-001***2042084-001-046050***2042084-001-208621***2042084-001-045915***2042084-001-215164***2042084-001-078474***2042084-001-053034***2042084-001-215165***2042084-001-079462***2042084-001-047389***2042084-001-112198***2042084-001-041673***2042084-001-227238***2042084-001-071564***2042084-001-503824***2042084-001-119839***2042084-001-569269***2042084-001-041674***2042084-001-047549***2042084-001-514651***2042084-001-163572***2042084-001-583721***2042084-001-427592***2042084-001-527358***2042084-001-583723***2042084-001-317166***2042084-001-583724***2042084-001-583715***2042084-001-583718***2042084-001-055156***2042084-001-583722***2042084-001-041835***2042084-001-074124***2042084-001-583706***2042084-001-583707***2042084-001-519926***2042084-001-583717***2042084-001-583713***2042084-001-207863***2042084-001-460604***2042084-001-079192***2042084-001-197663***2042084-001-045872***2042084-001-053195***2042084-001-092010***2042084-001-560824***2042084-001-052372***2042084-001-071088***2042084-001-046667***2042084-001-213006***2042084-001-086202***2042084-001-301821***2042084-001-045720***2042084-001-060454***2042084-001-418294***2042084-001-088907***2042084-001-065769***2042084-001-589530***2042084-001-623792***2042084-001-187563***2042084-001-320333***2042084-001-357143***2042084-001-148302***To be provided***2042084-001-082825***2042084-001-057016***2042084-001-153601***2042084-001-161552***2042084-001-140229***2042084-001-611704***2042084-001-626688***2042084-001-389208***2042084-001-657836***2042084-001-225214***2042084-001-503162***2042084-001-477881***2042084-001-316025***2042084-001-046826***2042084-001-245255***2042084-001-196543***2042084-001-557725***2042084-001-505890***2042084-001-218596***2042084-001-336980***2042084-001-524525***2042084-001-600032***2042084-001-529732***2042084-001-195742***2042084-001-195782***2042084-001-573889***2042084-001-573890***2042084-001-573886***2042084-001-573887***2042084-001-195822***2042084-001-195862***2042084-001-048420***2042084-001-065538***2042084-001-055814***2042084-001-341260***2042084-001-149220***2042084-001-079193***2042084-001-229282***2042084-001-629486***2042084-001-058695***2042084-001-061170***2042084-001-629548***2042084-001-080803***2042084-001-596662***2042084-001-250146***2042084-001-202515***2042084-001-515709***2042084-001-579491***2042084-001-505608***2042084-001-065010***2042084-001-619171*** | CARESCAPE Patient Data Module | 2 | CDRH | 2020-06-23 | 2020-06-24 21:00 |  | GE Healthcare, LLC | The CARESCAPE Patient Data Module (PDM) is used with the following monitors: CARESCAPE B450/B650/B850, SOLAR 8000M/I and Transport Pro. If Pace detection is turned ON, on the monitor, and if an Automated External Defibrillator (AED) is used to perform defibrillation, the low frequency signals deployed by the AED, cause the PDM to issue a false flat line ECG waveform and alarm Asystole to be displayed on the monitor. | Class 2 Device Recall CARESCAPE Patient Data Module Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall CARESCAPE Patient Data Module FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall CARESCAPE Patient Data Module Date Initiated by Firm May 18, 2020 Date Posted June 24, 2020 Recall Status 1 Open 3 , Classified Recall Number Z-2424-2020 Recall Event ID 85765 510(K)Number K071073 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product CARESCAPE Patient Data Module (PDM) used with the following monitors: CARESCAPE B450/B650/B850, SOLAR 8000M/I and Transport Pro, Model Numbers:******2042084-001******2094504-001******2042084-001-046050******2042084-001-208621******2042084-001-045915******2042084-001-215164******2042084-001-078474******2042084-001-053034******2042084-001-215165******2042084-001-079462******2042084-001-047389******2042084-001-112198******2042084-001-041673******2042084-001-227238******2042084-001-071564******2042084-001-503824******2042084-001-119839******2042084-001-569269******2042084-001-041674******2042084-001-047549******2042084-001-514651******2042084-001-163572******2042084-001-583721******2042084-001-427592******2042084-001-527358******2042084-001-583723******2042084-001-317166******2042084-001-583724******2042084-001-583715******2042084-001-583718******2042084-001-055156******2042084-001-583722******2042084-001-041835******2042084-001-074124******2042084-001-583706******2042084-001-583707******2042084-001-519926******2042084-001-583717******2042084-001-583713******2042084-001-207863******2042084-001-460604******2042084-001-079192******2042084-001-197663******2042084-001-045872******2042084-001-053195******2042084-001-092010******2042084-001-560824******2042084-001-052372******2042084-001-071088******2042084-001-046667******2042084-001-213006******2042084-001-086202******2042084-001-301821******2042084-001-045720******2042084-001-060454******2042084-001-418294******2042084-001-088907******2042084-001-065769******2042084-001-589530******2042084-001-623792******2042084-001-187563******2042084-001-320333******2042084-001-357143******2042084-001-148302******To be provided******2042084-001-082825******2042084-001-057016******2042084-001-153601******2042084-001-161552******2042084-001-140229******2042084-001-611704******2042084-001-626688******2042084-001-389208******2042084-001-657836******2042084-001-225214******2042084-001-503162******2042084-001-477881******2042084-001-316025******2042084-001-046826******2042084-001-245255******2042084-001-196543******2042084-001-557725******2042084-001-505890******2042084-001-218596******2042084-001-336980******2042084-001-524525******2042084-001-600032******2042084-001-529732******2042084-001-195742******2042084-001-195782******2042084-001-573889******2042084-001-573890******2042084-001-573886******2042084-001-573887******2042084-001-195822******2042084-001-195862******2042084-001-048420******2042084-001-065538******2042084-001-055814******2042084-001-341260******2042084-001-149220******2042084-001-079193******2042084-001-229282******2042084-001-629486******2042084-001-058695******2042084-001-061170******2042084-001-629548******2042084-001-080803******2042084-001-596662******2042084-001-250146******2042084-001-202515******2042084-001-515709******2042084-001-579491******2042084-001-505608******2042084-001-065010******2042084-001-619171 Code Information All Model Numbers that contain software versions 2.6 or 2.7 Recalling Firm/ Manufacturer GE Healthcare, LLC 3000 N Grandview Blvd Waukesha WI 53188-1615 For Additional Information Contact GE Healthcare Service 800-437-1171 Manufacturer Reason for Recall The CARESCAPE Patient Data Module (PDM) is used with the following monitors: CARESCAPE B450/B650/B850, SOLAR 8000M/I and Transport Pro. If Pace detection is turned ON, on the monitor, and if an Automated External Defibrillator (AED) is used to perform defibrillation, the low frequency signals deployed by the AED, cause the PDM to issue a false flat line ECG waveform and alarm Asystole to be displayed on the monitor. FDA Determined Cause 2 Software Design Change Action The firm notified its consignees by letter on 05/18/2020. The letter explained the issue, how it occurs, cautions and a workaround. The notices were disseminated to the following entities: Director of Biomedical/Clinical Engineering, Chief of Nursing, Health Care Administrator/Risk Manager. Quantity in Commerce 89,109 units Distribution Worldwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm notified its consignees by letter on 05/18/2020. The letter explained the issue, how it occurs, cautions and a workaround. The notices were disseminated to the following entities: Director of Biomedical/Clinical Engineering, Chief of Nursing, Health Care Administrator/Risk Manager. | SoftwareDesignChange
81 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=181430 | Z-2403-2020 | Merge Healthcare Merge Application Server Software Release and Merge Healthcare Merge Cardio Workstation Software Release - Product Usage: VERICIS is a system intended to be used to acquire, store, print, transfer, and archive clinical information from Camtronics and other vendors systems including images, Hemodynamic studies and reports, measurements (via import from DICOM Structured Reporting, text files or optical character recognition of measurements captured on images) and cardiology signal (waveform) data. | VERICIS CARDIOVASCULAR IMAGE AND INFORMATION SYSTEM, MODEL 4.0 | 2 | CDRH | 2020-06-19 |  | 2023-06-23 13:27 | Merge Healthcare, Inc. | An error in the calculation of the Michigan Reference Ranges can generate an inaccurate Z-Score calculation. The error may reduce the sensitivity of the Z-Score in detecting an abnormality resulting in a false negative. | Class 2 Device Recall VERICIS CARDIOVASCULAR IMAGE AND INFORMATION SYSTEM, MODEL 4.0 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VERICIS CARDIOVASCULAR IMAGE AND INFORMATION SYSTEM, MODEL 4.0 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VERICIS CARDIOVASCULAR IMAGE AND INFORMATION SYSTEM, MODEL 4.0 Date Initiated by Firm May 08, 2020 Create Date June 19, 2020 Recall Status 1 Terminated 3 on June 23, 2023 Recall Number Z-2403-2020 Recall Event ID 85633 510(K)Number K051649 Product Classification System, image processing, radiological - Product Code LLZ Product Merge Healthcare Merge Application Server Software Release and Merge Healthcare Merge Cardio Workstation Software Release - Product Usage: VERICIS is a system intended to be used to acquire, store, print, transfer, and archive clinical information from Camtronics and other vendors systems including images, Hemodynamic studies and reports, measurements (via import from DICOM Structured Reporting, text files or optical character recognition of measurements captured on images) and cardiology signal (waveform) data. Code Information Merge Cardio versions 10.X and 11.X with Knowledge Base versions 3.3, 3.4, 3.5 or 5.0 installed on or before February 21, 2020. Version 10.0 Pre-UDI Application server PN 88-00528-00; Workstation PN:88-00534-00 Version 10.1.2 Pre-UDI Application server PN 88-01075-00; Workstation PN 88-01078-00 Version 10.2 P2 Pre-UDI Application Server PN 88-01203-00 Version 10.2 P1&P2 Pre-UDI Software Release PN 87-00247-00 Version 10.3 UDI: (01) 00842000100041(10)10.3.0.1137(11)170104 Application server PN: 88-01233-00; Workstation PN 88-01239-00 Version 10.3 Patch 1 UDI: (01)00842000100041(10)10.3.0.2009(11)170309 Software Release PN 88-01319-00; Workstation PN: 88-01320-00 Version 10.3.1 UDI: (01) 00842000100041 (10) 10312135 (11) 171024 Application Server PN: 88-01426-00; Workstation PN 88-01430-00 Version 11.0.2 UDI: (01)00842000100416(10)11.0.2.1523(11)180809 Application Server Installation PN: 88-01570-00; Workstation Installation PN 88-01571-00 Version 11.0.3 UDI: (01)00842000100416(10)11.0.3.1601(11)181019 Workstation Installation PN: 88-01589-00 Version 11.1.1 UDI: (01)00842000100751(10)11.1.1(11)190724 Application Server 30-01650-00; Workstation PN:30-01654-00 Recalling Firm/ Manufacturer Merge Healthcare, Inc. 900 Walnut Ridge Dr Hartland WI 53029-8347 For Additional Information Contact 262-367-0700 Manufacturer Reason for Recall An error in the calculation of the Michigan Reference Ranges can generate an inaccurate Z-Score calculation. The error may reduce the sensitivity of the Z-Score in detecting an abnormality resulting in a false negative. FDA Determined Cause 2 Software Design Change Action IBM Watson Health Imaging, also known as Merge Healthcare, sent an "Urgent Field Safety Notice" on May 6, 2020. The letter states that it is safe to continue to use the device as is, with the understanding that the Z-Score ranges are incorrect and should not be used to make clinical determinations. Z-Scores can be removed upon request by contacting Customer Support at the contact information below.************The will provide a software correction to address the issue. A representative will be in contact with distributors to ensure all affected units have been updated.************All potential users of Merge Cardio versions 10.X and 11.X with Pediatric Echo Knowledge Base with Michigan or Hybrid Z-Scores versions 3.3, 3.4, 3.5 or 5.0 should receive the notice.************Customers with questions regarding the notice, please contact Customer Support at WHISupport@us.ibm.com or 877-741-5369. Quantity in Commerce 62 units Distribution US Nationwide distribution including in the states of IL, TX, OR, AZ, OH, TN, CA, UT, MO, GA, NM, OK, VA, IN, CT, and VT. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action IBM Watson Health Imaging, also known as Merge Healthcare, sent an "Urgent Field Safety Notice" on May 6, 2020. The letter states that it is safe to continue to use the device as is, with the understanding that the Z-Score ranges are incorrect and should not be used to make clinical determinations. Z-Scores can be removed upon request by contacting Customer Support at the contact information below.************The will provide a software correction to address the issue. A representative will be in contact with distributors to ensure all affected units have been updated.************All potential users of Merge Cardio versions 10.X and 11.X with Pediatric Echo Knowledge Base with Michigan or Hybrid Z-Scores versions 3.3, 3.4, 3.5 or 5.0 should receive the notice.************Customers with questions regarding the notice, please contact Customer Support at WHISupport@us.ibm.com or 877-741-5369. | SoftwareDesignChange
82 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180938 | Z-2251-2020 | Accelerator APS Aliquoter Module, List number: 07L40 - Product Usage: The ACCELERATOR APS Workcell is a modular system designed to automate sample handling and processing in the clinical laboratory. The system allows consolidations of multiple clinical chemistry and immunoassay analytical instruments into a unified workstation. | Accelerator APS Aliquoter Module | 2 | CDRH | 2020-06-03 | 2020-06-03 21:00 | 2021-04-28 18:12 | Abbott Laboratories | In the case of an AQM clot detection error a portion of the distilled water from the Aliquoter Module hydraulic circuit may be dispensed into the Primary Tube leading to sample dilution, which could lead to incorrect results. | Class 2 Device Recall Accelerator APS Aliquoter Module Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Accelerator APS Aliquoter Module FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Accelerator APS Aliquoter Module Date Initiated by Firm April 02, 2020 Date Posted June 03, 2020 Recall Status 1 Terminated 3 on April 28, 2021 Recall Number Z-2251-2020 Recall Event ID 85449 510(K)Number K093318 Product Classification Electrode, ion specific, potassium - Product Code CEM Product Accelerator APS Aliquoter Module, List number: 07L40 - Product Usage: The ACCELERATOR APS Workcell is a modular system designed to automate sample handling and processing in the clinical laboratory. The system allows consolidations of multiple clinical chemistry and immunoassay analytical instruments into a unified workstation. Code Information All serial numbers in the field Recalling Firm/ Manufacturer Abbott Laboratories 1915 Hurd Dr Irving TX 75038 Manufacturer Reason for Recall In the case of an AQM clot detection error a portion of the distilled water from the Aliquoter Module hydraulic circuit may be dispensed into the Primary Tube leading to sample dilution, which could lead to incorrect results. FDA Determined Cause 2 Software Design Change Action On 04/02/2020, Product Correction notices were mailed to customers via overnight service. ************Customers were asked to do the following to avoid the risk of contamination, take the following precaution:******- Discard the Primary Tube flagged with Clot Detection error or manage it according to your laboratory guidelines considering that it may be diluted.******- Call Customer Support in case the frequency of the Clot Detection Error increases (more than 5 consecutive Clot Detection Errors).******- Please complete the included Abbott Customer Reply Form.******- If you have forwarded the product listed above to other laboratories, please inform them of this Product Correction and provide to them a copy of this letter.******- Retain this letter for your laboratory records.************This information will be added to the Operations Manual, which is expected in Q2 2020.******If you or any of the health care providers you serve have any questions regarding this information, please contact Customer Service at 1-877-4ABBOTT (available 24 hours a day, 7 days a week). Quantity in Commerce 11 Distribution Worldwide distribution - U.S. Nationwide distribution including in the states of NC, ND, SC, TN.************The countries of United Kingdom, France, New Zealand, Pakistan, Turkey. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = CEM - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 04/02/2020, Product Correction notices were mailed to customers via overnight service. ************Customers were asked to do the following to avoid the risk of contamination, take the following precaution:******- Discard the Primary Tube flagged with Clot Detection error or manage it according to your laboratory guidelines considering that it may be diluted.******- Call Customer Support in case the frequency of the Clot Detection Error increases (more than 5 consecutive Clot Detection Errors).******- Please complete the included Abbott Customer Reply Form.******- If you have forwarded the product listed above to other laboratories, please inform them of this Product Correction and provide to them a copy of this letter.******- Retain this letter for your laboratory records.************This information will be added to the Operations Manual, which is expected in Q2 2020.******If you or any of the health care providers you serve have any questions regarding this information, please contact Customer Service at 1-877-4ABBOTT (available 24 hours a day, 7 days a week). | SoftwareDesignChange
83 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180837 | Z-1947-2020 | Quickie/Zippie powered wheelchair. Q700M wheelchair. A battery operated powered wheelchair, which is controlled by Power Module RNET 120A Expandable (Item No.: 250ER10, Part No.: 123966). | Quickie/Zippie Powered Wheelchair | 2 | CDRH | 2020-05-11 | 2020-05-11 21:00 | 2021-04-28 08:49 | Sunrise Medical (US) LLC | Due to programming errors in the Controller, an increase in the set motor parameters can be made by the end user beyond tested safe limits. | Class 2 Device Recall Quickie/Zippie Powered Wheelchair Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Quickie/Zippie Powered Wheelchair FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Quickie/Zippie Powered Wheelchair Date Initiated by Firm February 13, 2020 Date Posted May 11, 2020 Recall Status 1 Terminated 3 on April 28, 2021 Recall Number Z-1947-2020 Recall Event ID 85431 510(K)Number K142457 Product Classification Wheelchair, powered - Product Code ITI Product Quickie/Zippie powered wheelchair. Q700M wheelchair. A battery operated powered wheelchair, which is controlled by Power Module RNET 120A Expandable (Item No.: 250ER10, Part No.: 123966). Code Information Q700M wheelchair S/N Q7MP-083941. UDI:(01)05022408059799. Power Module RNET 120A Expandable (Item No.: 250ER10, Part No.: 123966). Controller label: D50903.15 CR19100531. Recalling Firm/ Manufacturer Sunrise Medical (US) LLC 2842 N Business Park Ave Fresno CA 93727-1328 For Additional Information Contact Mr. Devin I. McElroy 559-294-2374 Manufacturer Reason for Recall Due to programming errors in the Controller, an increase in the set motor parameters can be made by the end user beyond tested safe limits. FDA Determined Cause 2 Software Design Change Action An URGENT: MEDICAL DEVICE RECALL letter was sent the firm's dealer. The dealer was instructed on February 4, 2020, to replace the controller unit for the device. Quantity in Commerce 1 Distribution Distributed to US Consignee in MT. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = ITI - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action An URGENT: MEDICAL DEVICE RECALL letter was sent the firm's dealer. The dealer was instructed on February 4, 2020, to replace the controller unit for the device. | SoftwareDesignChange
84 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180537 | Z-1927-2020 | The Cogent Hemodynamic Monitoring System. REF58400-000, GTIN 00840619099459; REF 58400-000R, GTIN 0087709094314; REF 58403, GTIN 00840619079505; REF 58401, GTIN 00840619079499 - Product Usage: intended for patients for whom the monitoring of continuous cardiac output and calculated hemodynamic parameters is indicated for diagnostic and prognostic evaluation by a clinician. The target population includes Critical Care Patients, Trauma Patients, and Cardiac Surgery Patients. | Cogent Hemodynamic Monitoring System | 2 | CDRH | 2020-05-07 |  | 2023-05-03 11:33 | ICU Medical, Inc. | Firm identified software issues which leads to the patient ID and patient information not being stored in the system. | Class 2 Device Recall Cogent Hemodynamic Monitoring System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Cogent Hemodynamic Monitoring System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Cogent Hemodynamic Monitoring System Date Initiated by Firm February 14, 2020 Create Date May 07, 2020 Recall Status 1 Terminated 3 on May 03, 2023 Recall Number Z-1927-2020 Recall Event ID 85329 510(K)Number K152006 Product Classification Computer, diagnostic, pre-programmed, single-function - Product Code DXG Product The Cogent Hemodynamic Monitoring System. REF58400-000, GTIN 00840619099459; REF 58400-000R, GTIN 0087709094314; REF 58403, GTIN 00840619079505; REF 58401, GTIN 00840619079499 - Product Usage: intended for patients for whom the monitoring of continuous cardiac output and calculated hemodynamic parameters is indicated for diagnostic and prognostic evaluation by a clinician. The target population includes Critical Care Patients, Trauma Patients, and Cardiac Surgery Patients. Code Information All serial numbers of Cogent HMS, List Number(s): 58400-000; 58400-000R; 58403-000; 58401-000 with Software Versions 1.3.3.42 and prior are impacted. Recalling Firm/ Manufacturer ICU Medical, Inc. 951 Calle Amanecer San Clemente CA 92673-6212 For Additional Information Contact ICU Medical 844-654-7780 Manufacturer Reason for Recall Firm identified software issues which leads to the patient ID and patient information not being stored in the system. FDA Determined Cause 2 Software Design Change Action On 14 February 2020, the firm sent notification of the recall to direct consignees of record. Delivery will be confirmed via traceable mailing. All consignees who further distributed the product were asked to notify their customers Quantity in Commerce 123 units Distribution US Nationwide distribution including in the states of AR, CA, FL, IL, MD, NC, NE, NY, TX, VT, WA and WI. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = DXG - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 14 February 2020, the firm sent notification of the recall to direct consignees of record. Delivery will be confirmed via traceable mailing. All consignees who further distributed the product were asked to notify their customers | SoftwareDesignChange
85 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180396 | Z-1899-2020 | Varian Multileaf Collimator [MLC]; Millenium Multileaf Collimator [52, 80, 120 leaf]; High Definition 120 Multileaf Collimator Model Number(s): product code: H54 and H56; MLC software version 8.5 - Product Usage: is intended to assist the clinician in the delivery of external beam radiation to defined target volumes during radiosurgery and radiotherapy while sparing surrounding normal tissue and critical organs from excess radiation.****** | Varian Multileaf Collimator with C-Series Linac | 2 | CDRH | 2020-05-05 |  | 2021-01-21 11:00 | Varian Medical Systems, Inc. | After a recent upgrade to the collimator software version 8.5, the firm became aware that the multi-leaf collimator leaves did not move during an arc treatment. The issue is related to a transient carriage primary- secondary interlock due to carriage fault on the MLC that immediately preceded the initiation of the arc treatment. The issue will occur only with MLC software version 8.5 and only affects conformal arc treatments, including VMAT and RapidArc. There have been no reports of adverse health consequences due to this issue.*************** | Class 2 Device Recall Varian Multileaf Collimator with CSeries Linac Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Varian Multileaf Collimator with CSeries Linac FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Varian Multileaf Collimator with CSeries Linac Date Initiated by Firm August 12, 2019 Create Date May 05, 2020 Recall Status 1 Terminated 3 on January 21, 2021 Recall Number Z-1899-2020 Recall Event ID 85269 510(K)Number K181404 Product Classification Accelerator, linear, medical - Product Code IYE Product Varian Multileaf Collimator [MLC]; Millenium Multileaf Collimator [52, 80, 120 leaf]; High Definition 120 Multileaf Collimator Model Number(s): product code: H54 and H56; MLC software version 8.5 - Product Usage: is intended to assist the clinician in the delivery of external beam radiation to defined target volumes during radiosurgery and radiotherapy while sparing surrounding normal tissue and critical organs from excess radiation. Code Information Model Number(s): product code: H54 and H56; MLC software version 8.5. UDIs: HD 120 MLC Head Assembly. UDI: 00899475002318 80 Leaf Assy, MLC Head, UDI: 00899475002325 120 Leaf Assy, MLC Head, UDI: 00899475002332 Recalling Firm/ Manufacturer Varian Medical Systems, Inc. 911 Hansen Way Palo Alto CA 94304-1028 For Additional Information Contact K.Jeffrey Semone 800-544-4636 Manufacturer Reason for Recall After a recent upgrade to the collimator software version 8.5, the firm became aware that the multi-leaf collimator leaves did not move during an arc treatment. The issue is related to a transient carriage primary- secondary interlock due to carriage fault on the MLC that immediately preceded the initiation of the arc treatment. The issue will occur only with MLC software version 8.5 and only affects conformal arc treatments, including VMAT and RapidArc. There have been no reports of adverse health consequences due to this issue. FDA Determined Cause 2 Software Design Change Action Urgent Medical Device Correction Field Safety Notice letters were sent to customer beginning August 12, 2019. Letters were sent to U.S. and Canadian customers via Fed-Ex, and via local third-party courier or email for the rest of the world. Users were instructed to: 1. monitor MLC leaf movement on the 4DITC console screen, 2. observe the start of every arc treatment and verify that the leaves are moving as intended, and 3. stop treatment immediately if MLC leaves are not moving and contact Varian to assess the MLC status. Users may contact Varian Support for help to verify whether carriage faults have been recorded for the MLC. If carriage fault is confirmed in the TerminalLog file, users were instructed to contact Varian Service Representative to assess the MLC status.******Customers were requested to complete and return the response form provided to returnresponse@varian.com, and to contact their local Varian Medical Systems service manager with any questions. The notification letter includes Varian Support Information with the following contact information: phone number for North America Customers: 1 888.827.4265; Global email address support@varian.com; and website contact address for customers outside of North America: https://www.myvarian.com/s/contactus?lang=en. Quantity in Commerce 151 units Distribution Worldwide distribution - US Nationwide distribution including in the states of TN, FL, IN, PA, NV, WA, DC, CA, AL, IL, MD, VA, TX, MN, AR, GA, OR, MS, MO, NY, SC, DE, AZ, NH, MI, NC, OH, PA and OK. The country of Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Urgent Medical Device Correction Field Safety Notice letters were sent to customer beginning August 12, 2019. Letters were sent to U.S. and Canadian customers via Fed-Ex, and via local third-party courier or email for the rest of the world. Users were instructed to: 1. monitor MLC leaf movement on the 4DITC console screen, 2. observe the start of every arc treatment and verify that the leaves are moving as intended, and 3. stop treatment immediately if MLC leaves are not moving and contact Varian to assess the MLC status. Users may contact Varian Support for help to verify whether carriage faults have been recorded for the MLC. If carriage fault is confirmed in the TerminalLog file, users were instructed to contact Varian Service Representative to assess the MLC status.******Customers were requested to complete and return the response form provided to returnresponse@varian.com, and to contact their local Varian Medical Systems service manager with any questions. The notification letter includes Varian Support Information with the following contact information: phone number for North America Customers: 1 888.827.4265; Global email address support@varian.com; and website contact address for customers outside of North America: https://www.myvarian.com/s/contactus?lang=en. | SoftwareDesignChange
86 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179188 | Z-1563-2020 | VNS Therapy Programmer, Model 3000 used with Models 103, 104, 105, 106, or 8103 VNS Therapy generators. | VNS Therapy Programmer | 2 | CDRH | 2020-03-24 |  |  | LivaNova USA Inc | False positive warning may occur after: 1) VNS Generator interrogated at 0mA normal output current, 2) Generator programmed to non-0mA output current, 3) In-session re-interrogation performed. Users instructed to lower output current and widen pulse width. Only system diagnostic testing evaluates output current. Users may conclude device malfunction, could lead to medical/surgical intervention. | Class 2 Device Recall VNS Therapy Programmer Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VNS Therapy Programmer FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VNS Therapy Programmer Date Initiated by Firm February 04, 2020 Create Date March 24, 2020 Recall Status 1 Open 3 , Classified Recall Number Z-1563-2020 Recall Event ID 84723 PMA Number P970003 Product Classification Stimulator, autonomic nerve, implanted for epilepsy - Product Code LYJ Product VNS Therapy Programmer, Model 3000 used with Models 103, 104, 105, 106, or 8103 VNS Therapy generators. Code Information Model 3000 version 1.6 Recalling Firm/ Manufacturer LivaNova USA Inc 100 Cyberonics Blvd Houston TX 77058-2069 For Additional Information Contact 281-228-7200 Manufacturer Reason for Recall False positive warning may occur after: 1) VNS Generator interrogated at 0mA normal output current, 2) Generator programmed to non-0mA output current, 3) In-session re-interrogation performed. Users instructed to lower output current and widen pulse width. Only system diagnostic testing evaluates output current. Users may conclude device malfunction, could lead to medical/surgical intervention. FDA Determined Cause 2 Software Design Change Action On 02/04/20, Urgent Medical Device Correction notices were mailed to Doctors, who were informed the following:************Which Patients are Potentially Impacted?******Because the false positive warning message only appears if the generator is initially interrogated at 0mA, there are only certain scenarios where the issue may appear:******- After initial implant, when the device is programmed on for the first time;******- Following an MRI, after the device was programmed to 0mA per labeling instructions; or******- Resuming therapy following previous device disablement due to an adverse event or device issue.************Actions to be taken by the Physician:******1. If you obtain the low output current warning message, follow the Physician's Manual instructions to perform a System Diagnostics test after initial interrogation and programming before utilizing in-session re-interrogation.******2. During programming visits, always perform System Diagnostics testing immediately after any programming operations are performed to ensure the device is delivering programmed stimulation as intended.******2. Please complete and return the attached Customer Response Form.************Please ensure that this notice is communicated to all personnel within your organization who need to be aware of it.************To report a complaint or for technical support, contact the Customer Quality team at cservices@livanova.com or (866) 882-8804.************Actions to be taken by Company:******An update to the software is in development to eliminate this false positive low output current message. Information on availability of the software update will be shared in future communications.************New users will be trained on proper use of programming system, including the appropriate sequence of use of diagnostic testing. Quantity in Commerce 2683 Distribution U.S.: LA, ID, AZ, UT, KY, VA, TN, NY, SC, IL, IN, NC, WA, CA, OH, FL, PA, NJ, CO, TX, RI, MD, HI, OR, MA, NE, MO, OK, MS, MN, SD, VT, AR, WV, WI, MI, AL, KS, DE, NM, GA, ND, AK, MT, ME, NV, DC, MO, WY, IA, CT.************Canada, and Israel. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. PMA Database PMAs with Product Code = LYJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 02/04/20, Urgent Medical Device Correction notices were mailed to Doctors, who were informed the following:************Which Patients are Potentially Impacted?******Because the false positive warning message only appears if the generator is initially interrogated at 0mA, there are only certain scenarios where the issue may appear:******- After initial implant, when the device is programmed on for the first time;******- Following an MRI, after the device was programmed to 0mA per labeling instructions; or******- Resuming therapy following previous device disablement due to an adverse event or device issue.************Actions to be taken by the Physician:******1. If you obtain the low output current warning message, follow the Physician's Manual instructions to perform a System Diagnostics test after initial interrogation and programming before utilizing in-session re-interrogation.******2. During programming visits, always perform System Diagnostics testing immediately after any programming operations are performed to ensure the device is delivering programmed stimulation as intended.******2. Please complete and return the attached Customer Response Form.************Please ensure that this notice is communicated to all personnel within your organization who need to be aware of it.************To report a complaint or for technical support, contact the Customer Quality team at cservices@livanova.com or (866) 882-8804.************Actions to be taken by Company:******An update to the software is in development to eliminate this false positive low output current message. Information on availability of the software update will be shared in future communications.************New users will be trained on proper use of programming system, including the appropriate sequence of use of diagnostic testing. | SoftwareDesignChange
87 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179419 | Z-1346-2020 | COHERENCE Oncologist, Model Nos. 07333680 & 07351898 | COHERENCE Oncologist | 2 | CDRH | 2020-02-21 |  | 2021-01-04 14:33 | Siemens Medical Solutions USA, Inc | A potential safety issue arises when an offset calculation is followed by a filter operation; correcting the image alignment after this sequence will result in incorrect offset values which could lead to incorrect repositioning of the patient and dose to wrong location. | Class 2 Device Recall COHERENCE Oncologist Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall COHERENCE Oncologist FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall COHERENCE Oncologist Date Initiated by Firm January 16, 2020 Create Date February 21, 2020 Recall Status 1 Terminated 3 on January 04, 2021 Recall Number Z-1346-2020 Recall Event ID 84795 510(K)Number K060633 Product Classification Accelerator, linear, medical - Product Code IYE Product COHERENCE Oncologist, Model Nos. 07333680 & 07351898 Code Information Software versions 2.0.49, 2.0.50, 2.0.51 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact Rebecca Tudor 610-219-4834 Manufacturer Reason for Recall A potential safety issue arises when an offset calculation is followed by a filter operation; correcting the image alignment after this sequence will result in incorrect offset values which could lead to incorrect repositioning of the patient and dose to wrong location. FDA Determined Cause 2 Software Design Change Action The firm started notifying customers on January 16, 2020, with letters titled Urgent Medical Device Correction. Customers were advised of the product safety issue. To avoid the possibility of an incorrect alignment of the patient, the user must ensure that filtered images arrive on the COHERENCE Oncologist, or apply any additional filters to the image BEFORE using the "interactive shift" or "landmark registration" tools.************The firm will be offering a software solution for the issue free of charge. The firm's letter stated that based on their investigation, customers may continue to use their devices. Quantity in Commerce 64 Worldwide/26 Domestic Distribution Domestic distribution to CA, FL, GA, IA, MA, MI, NE, NJ, NV, NY, OH, OK, PA, SD, TN, WV. Worldwide foreign distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm started notifying customers on January 16, 2020, with letters titled Urgent Medical Device Correction. Customers were advised of the product safety issue. To avoid the possibility of an incorrect alignment of the patient, the user must ensure that filtered images arrive on the COHERENCE Oncologist, or apply any additional filters to the image BEFORE using the "interactive shift" or "landmark registration" tools.************The firm will be offering a software solution for the issue free of charge. The firm's letter stated that based on their investigation, customers may continue to use their devices. | SoftwareDesignChange
88 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179418 | Z-1330-2020 | DP-30 Digital Ultrasonic Diagnostic Imaging System, Part No. 1155E-PA00004 - Product Usage: The Digital Ultrasonic Diagnostic Imaging System is applicable for adults, pregnant women, pediatric patients and neonates. It is intended for use in fetal, abdominal, pediatric, small organ (breast, thyroid, testes), neonatal cephalic, adult cephalic, trans-rectal, trans-vaginal, muscular-skeletal (conventional, superficial), cardiac (adult, pediatric), peripheral vascular and urology exams. | DP-30 Digital Ultrasonic Diagnostic Imaging System | 2 | CDRH | 2020-02-20 | 2020-02-05 20:03 | 2020-08-21 09:09 | Mindray DS USA, Inc. dba Mindray North America | The DP-30 displays an incorrect needle-guide bracket when used with the 65EC10EA model***transducer. | Class 2 Device Recall DP30 Digital Ultrasonic Diagnostic Imaging System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall DP30 Digital Ultrasonic Diagnostic Imaging System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall DP30 Digital Ultrasonic Diagnostic Imaging System Date Initiated by Firm January 13, 2020 Date Posted February 05, 2020 Recall Status 1 Terminated 3 on August 21, 2020 Recall Number Z-1330-2020 Recall Event ID 84794 510(K)Number K183660 Product Classification System, imaging, pulsed doppler, ultrasonic - Product Code IYN Product DP-30 Digital Ultrasonic Diagnostic Imaging System, Part No. 1155E-PA00004 - Product Usage: The Digital Ultrasonic Diagnostic Imaging System is applicable for adults, pregnant women, pediatric patients and neonates. It is intended for use in fetal, abdominal, pediatric, small organ (breast, thyroid, testes), neonatal cephalic, adult cephalic, trans-rectal, trans-vaginal, muscular-skeletal (conventional, superficial), cardiac (adult, pediatric), peripheral vascular and urology exams. Code Information DP-30 Ultrasound with software version 03.00.00 and 03.01.00 Recalling Firm/ Manufacturer Mindray DS USA, Inc. dba Mindray North America 800 Macarthur Blvd Mahwah NJ 07430-2001 For Additional Information Contact Diane Arpino 201-995-8407 Manufacturer Reason for Recall The DP-30 displays an incorrect needle-guide bracket when used with the 65EC10EA model******transducer. FDA Determined Cause 2 Software Design Change Action On January 13, 2020, the firm distributed Urgent Medical Device Correction letters to affected distributors. Distributors were informed that the system displays the incorrect needle-guide bracket when used with the 65EC10EA model transducer.************Customers were advised to avoid use of NGB-004 needle-guide with transducer model 65EC10EA until a software update has been applied to the system. The system may be used normally for all other functions. The firm recommended that distributors immediately notify any clients to which they have distributed affected systems (DP-30 systems with software version 03.00.00 and 03.01.00 distributed between 06/25/2018******and 06/28/2019), or provided a software update (USB with V03.01.00, part number 110-002261-05).************Within the next several weeks, Mindray will provide a software update at no cost. Should you******require technical assistance while performing the update, please contact Mindray***s Technical******Support team at 800-288-2121, option 2, Monday through Friday, 8:30am - 5:30pm PT.******The replacement USB drive contains instructions on how to perform the software update, under******the ***Manuals*** folder in the ***DP-10 & DP-20 & DP-30 Series System Recovery Guide*** file.************If you have questions regarding this recall, you may contact Diane Arpino, Director, Quality Operations and Regulatory Affairs, at 201-995-8407, or via email at d.arpino@mindray.com. Quantity in Commerce 53 DP-30 systems and 18 USB's containing the affected software Distribution US Nationwide distribution in the states of CA, GA, IL, MD, NY, SC, and WA. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 13, 2020, the firm distributed Urgent Medical Device Correction letters to affected distributors. Distributors were informed that the system displays the incorrect needle-guide bracket when used with the 65EC10EA model transducer.************Customers were advised to avoid use of NGB-004 needle-guide with transducer model 65EC10EA until a software update has been applied to the system. The system may be used normally for all other functions. The firm recommended that distributors immediately notify any clients to which they have distributed affected systems (DP-30 systems with software version 03.00.00 and 03.01.00 distributed between 06/25/2018******and 06/28/2019), or provided a software update (USB with V03.01.00, part number 110-002261-05).************Within the next several weeks, Mindray will provide a software update at no cost. Should you******require technical assistance while performing the update, please contact Mindray***s Technical******Support team at 800-288-2121, option 2, Monday through Friday, 8:30am - 5:30pm PT.******The replacement USB drive contains instructions on how to perform the software update, under******the ***Manuals*** folder in the ***DP-10 & DP-20 & DP-30 Series System Recovery Guide*** file.************If you have questions regarding this recall, you may contact Diane Arpino, Director, Quality Operations and Regulatory Affairs, at 201-995-8407, or via email at d.arpino@mindray.com. | SoftwareDesignChange
89 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178080 | Z-0670-2020 | Monaco Radiation Treatment Planning System (RTP) System | Radiation Treatment Planning System | 2 | CDRH | 2019-12-12 |  |  | Elekta Inc | Monaco is using the incorrect energy when optimizing and calculating dose. | Class 2 Device Recall Radiation Treatment Planning System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Radiation Treatment Planning System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Radiation Treatment Planning System Date Initiated by Firm November 22, 2019 Create Date December 12, 2019 Recall Status 1 Open 3 , Classified Recall Number Z-0670-2020 Recall Event ID 84330 510(K)Number K183037 K190178 Product Classification System, planning, radiation therapy treatment - Product Code MUJ Product Monaco Radiation Treatment Planning System (RTP) System Code Information Software Builds: 5.40.00, 5.50.00 and 5.51.00. UDI GTIN: (01)00858164002190(10) 5.40.00, (01)00858164002244(10) 5.50.00, and (01)00858164002268(10) 5.51.00 Recalling Firm/ Manufacturer Elekta Inc 1450 Beale St Ste 205 Saint Charles MO 63303 For Additional Information Contact 800-878-4267 Manufacturer Reason for Recall Monaco is using the incorrect energy when optimizing and calculating dose. FDA Determined Cause 2 Software Design Change Action An Important Field Safety Notification letter is being emailed to affected customers beginning 11/22/2019. The letter identifies affected product, provides a description of the problem, and recommends user not use multiple energies within the same plan until a software fix has been applied to the device. A response form was asked to be returned. Affected customers are advised to contact local Elekta service teams for questions. Quantity in Commerce 76 (7 US, 69 OUS) Distribution IA, IN, NY, OH, TX, WI.******Internationally distributed to: Australia, Canada, China, Denmark, Germany, Greece, Hong Kong, Italy, Netherlands, Sweden, Switzerland, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = MUJ 510(K)s with Product Code = MUJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action An Important Field Safety Notification letter is being emailed to affected customers beginning 11/22/2019. The letter identifies affected product, provides a description of the problem, and recommends user not use multiple energies within the same plan until a software fix has been applied to the device. A response form was asked to be returned. Affected customers are advised to contact local Elekta service teams for questions. | SoftwareDesignChange
90 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=174163 | Z-2230-2019 | iGUIDE System with software build iGuide 2.2.0, 2.2.1, and 2.2.2 | iGUIDE System | 2 | CDRH | 2019-08-09 |  | 2024-05-06 07:53 | Elekta, Inc. | Potential for iGUIDE to incorrectly monitor the 3D position. | Class 2 Device Recall iGUIDE System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall iGUIDE System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall iGUIDE System Date Initiated by Firm June 19, 2019 Create Date August 09, 2019 Recall Status 1 Terminated 3 on May 06, 2024 Recall Number Z-2230-2019 Recall Event ID 83180 510(K)Number K072079 Product Classification Accelerator, linear, medical - Product Code IYE Product iGUIDE System with software build iGuide 2.2.0, 2.2.1, and 2.2.2 Code Information Version 2.2.0, Product Code/GTIN Number: 04056719001704; Version 2.2.1, Product Code/GTIN Number: 04056719001742; Version 2.2.2, Product Code/GTIN Number: 04056719002039 Recalling Firm/ Manufacturer Elekta, Inc. 400 Perimeter Center Ter NE Ste 50 Atlanta GA 30346-1227 For Additional Information Contact Alejandra Carrillo 770-670-2359 Manufacturer Reason for Recall Potential for iGUIDE to incorrectly monitor the 3D position. FDA Determined Cause 2 Software Design Change Action Elekta, Inc. notified customers on about 06/19/2019 via URGENT IMPORTANT FIELD SAFETY NOTIFICATION" letter. Customers were instructed to follow the below steps: ******1. Log on to iGUIDE and move the 3D position. If the 3D position is equal to the PARK position, the system is not affected. ******2. If the 3D position is not equal to the PARK position, log on to iGUIDE with Operator and Administrator rights. (If no user is available, go to the iGUIDE User Manager and add an appropriate user.). ******3. Move the HexaPOD to the 3D Position (If 3D position is equal to the DRIVE position, both checkmarks will appear). ******4. Open the ***HexaPOD properties*** (Left click the HexaPOD symbol on the system overview), which will display the HexaPOD properties. ******5. Enter a rotation of 1.00 degree for any of the axis (RX, RY, RZ) and click ***Go to position.*** ******6. Execute the pending movement. If the checkmark for the 3D position disappears and the EXTERNAL INHIBIT is set, your system is not affected by the comma / dot issue. No further actions are necessary. If the 3D position monitoring does not consider the 1.00 degree and the checkmark for the 3D position is still displayed and the EXTERNAL INHIBIT is not set, your system is affected! If the customer wants to still use a 3D workflow without imaging, do not rely on the 3D position indicator in iGUIDE. The HexaPOD should be moved to the PARK position first, then the 3D workflow can be started moving the HexaPOD back to the 3d position. ************Customers were also instructed to place the recall notice in a place accessible to all users (e.g. Instructions for Use, until this action is closed) and to advise the appropriate personnel, working with this product, on the content of this letter. ************A technical solution (iGUIDE patch) will be provided to correct the behavior.******Until the iGUIDE patch is available, an interim solution can be implemented (change comma to dot) if your system is affected. Please consult your local Elekta represent Quantity in Commerce 374 devices Distribution Nationwide distribution to AK, AL, CA, CO, CT, FL, GA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, ND, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, ******WA, WI. ************Worldwide distribution to Algeria, Argentina, Australia, Austria, Belgium, Botswana, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Denmark, Egypt, France, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Poland, Portugal, Romania, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Elekta, Inc. notified customers on about 06/19/2019 via URGENT IMPORTANT FIELD SAFETY NOTIFICATION" letter. Customers were instructed to follow the below steps: ******1. Log on to iGUIDE and move the 3D position. If the 3D position is equal to the PARK position, the system is not affected. ******2. If the 3D position is not equal to the PARK position, log on to iGUIDE with Operator and Administrator rights. (If no user is available, go to the iGUIDE User Manager and add an appropriate user.). ******3. Move the HexaPOD to the 3D Position (If 3D position is equal to the DRIVE position, both checkmarks will appear). ******4. Open the ***HexaPOD properties*** (Left click the HexaPOD symbol on the system overview), which will display the HexaPOD properties. ******5. Enter a rotation of 1.00 degree for any of the axis (RX, RY, RZ) and click ***Go to position.*** ******6. Execute the pending movement. If the checkmark for the 3D position disappears and the EXTERNAL INHIBIT is set, your system is not affected by the comma / dot issue. No further actions are necessary. If the 3D position monitoring does not consider the 1.00 degree and the checkmark for the 3D position is still displayed and the EXTERNAL INHIBIT is not set, your system is affected! If the customer wants to still use a 3D workflow without imaging, do not rely on the 3D position indicator in iGUIDE. The HexaPOD should be moved to the PARK position first, then the 3D workflow can be started moving the HexaPOD back to the 3d position. ************Customers were also instructed to place the recall notice in a place accessible to all users (e.g. Instructions for Use, until this action is closed) and to advise the appropriate personnel, working with this product, on the content of this letter. ************A technical solution (iGUIDE patch) will be provided to correct the behavior.******Until the iGUIDE patch is available, an interim solution can be implemented (change comma to dot) if your system is affected. Please consult your local Elekta represent | SoftwareDesignChange
91 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=174647 | Z-2063-2019 | Merge OrthoPACS versions 7.3, 7.3.1, 8.0 and 8.0.1 | Merge OrthoPACS | 2 | CDRH | 2019-07-23 |  | 2023-03-17 14:53 | Merge Healthcare, Inc. | Potential for prior thumbnails to not display in reverse chronological order and images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. ******Merge PACS provides image manipulation tools to enable users to view and compare images such as: linking, MPR, MIP, 3D image fusion/registration of CT, MR and PET; as well as CVR (Color Volume Rendering), measurements (linear distances, angles, areas, SUV, etc.), and annotations (for example, outline and label regions of interest, label spinal vertebrae). | Class 2 Device Recall Merge OrthoPACS Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Merge OrthoPACS FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Merge OrthoPACS Date Initiated by Firm July 11, 2019 Create Date July 23, 2019 Recall Status 1 Terminated 3 on March 17, 2023 Recall Number Z-2063-2019 Recall Event ID 83300 510(K)Number K173475 Product Classification System, image processing, radiological - Product Code LLZ Product Merge OrthoPACS versions 7.3, 7.3.1, 8.0 and 8.0.1 Code Information Version 7.3 (UDI Number (01)00842000100706(10)7.3.0.158618(11)180409); 7.3.1 (UDI Number (01)00842000100706(10)7.3.1.161404(11)180830); 8.0 (UDI Number (01)00842000100775(10)8.0.0.162051(11)181026); 8.0.1 (UDI Number (01)00842000100775(10)8.0.1.165704(11)190528) Recalling Firm/ Manufacturer Merge Healthcare, Inc. 900 Walnut Ridge Dr Hartland WI 53029-8347 For Additional Information Contact 262-367-0700 Manufacturer Reason for Recall Potential for prior thumbnails to not display in reverse chronological order and images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. ************Merge PACS provides image manipulation tools to enable users to view and compare images such as: linking, MPR, MIP, 3D image fusion/registration of CT, MR and PET; as well as CVR (Color Volume Rendering), measurements (linear distances, angles, areas, SUV, etc.), and annotations (for example, outline and label regions of interest, label spinal vertebrae). FDA Determined Cause 2 Software Design Change Action Merge Healthcare notified customers on about 07/11/2019 via "Urgent Field Safety Notice" letter. The notice informed customers of two possible issues that may occur on Merge PACS and Merge OrthoPACS: 1. Prior thumbnails may not display in reverse chronological order. 2. Images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. ************Use of an out of order image may result in the implementation of an incorrect or unnecessary treatment plan such as unnecessary medication, change in dosage of existing medication, or unnecessary medical procedures. It may also lead to progression of an existing patient condition, which may result in the progression of a disease state. ************Merge Healthcare will provide a software correction for the issue. A Merge Healthcare representative will contact each customer schedule the correction. ************In the interim, Merge Healthcare instructed customers that they can continue to use their Merge PACS/OrthoPACS devices by following the below instructions:******1. When pulling a series from thumbnail bar, confirm that the date of each series is visible in the viewport. The date will appear when overlays are configured to display the date. An administrator can assist in configuring Merge PACS/OrthoPACS if needed. ******2. When utilizing the mammography scroll stacking tool, confirm that the date of each series is visible in the viewport. The date will appear when overlays are configured to display the date. An administrator can assist in configuring Merge PACS/OrthoPACS if needed.************Customers were instructed to use the date, rather than the order the series are presented in, to determine chronological order.************Additional instructions included to ensure all potential users of Merge PACS and Merge OrthoPACS have received this notice. Questions regarding the recall notice can be directed to Merge Customer Support at 877-741-5369. Quantity in Commerce 73 devices Distribution Nationwide distribution to Arizona, Arkansas, California, Colorado, Florida, Georgia, Hawaii, Illinois, Iowa, Kentucky, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New York, North Carolina, Ohio, Pennsylvania, Rhode Island, South Carolina, Texas, Utah, Virginia, Wisconsin. ************International distribution to Belgium and Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Merge Healthcare notified customers on about 07/11/2019 via "Urgent Field Safety Notice" letter. The notice informed customers of two possible issues that may occur on Merge PACS and Merge OrthoPACS: 1. Prior thumbnails may not display in reverse chronological order. 2. Images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. ************Use of an out of order image may result in the implementation of an incorrect or unnecessary treatment plan such as unnecessary medication, change in dosage of existing medication, or unnecessary medical procedures. It may also lead to progression of an existing patient condition, which may result in the progression of a disease state. ************Merge Healthcare will provide a software correction for the issue. A Merge Healthcare representative will contact each customer schedule the correction. ************In the interim, Merge Healthcare instructed customers that they can continue to use their Merge PACS/OrthoPACS devices by following the below instructions:******1. When pulling a series from thumbnail bar, confirm that the date of each series is visible in the viewport. The date will appear when overlays are configured to display the date. An administrator can assist in configuring Merge PACS/OrthoPACS if needed. ******2. When utilizing the mammography scroll stacking tool, confirm that the date of each series is visible in the viewport. The date will appear when overlays are configured to display the date. An administrator can assist in configuring Merge PACS/OrthoPACS if needed.************Customers were instructed to use the date, rather than the order the series are presented in, to determine chronological order.************Additional instructions included to ensure all potential users of Merge PACS and Merge OrthoPACS have received this notice. Questions regarding the recall notice can be directed to Merge Customer Support at 877-741-5369. | SoftwareDesignChange
92 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=174646 | Z-2062-2019 | Merge PACS versions 7.3, 7.3.1, 8.0 and 8.0.1 | Merge PACS | 2 | CDRH | 2019-07-23 |  | 2023-03-17 14:53 | Merge Healthcare, Inc. | Potential for prior thumbnails to not display in reverse chronological order and images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. ******Merge PACS provides image manipulation tools to enable users to view and compare images such as: linking, MPR, MIP, 3D image fusion/registration of CT, MR and PET; as well as CVR (Color Volume Rendering), measurements (linear distances, angles, areas, SUV, etc.), and annotations (for example, outline and label regions of interest, label spinal vertebrae). | Class 2 Device Recall Merge PACS Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Merge PACS FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Merge PACS Date Initiated by Firm July 11, 2019 Create Date July 23, 2019 Recall Status 1 Terminated 3 on March 17, 2023 Recall Number Z-2062-2019 Recall Event ID 83300 510(K)Number K173475 Product Classification System, image processing, radiological - Product Code LLZ Product Merge PACS versions 7.3, 7.3.1, 8.0 and 8.0.1 Code Information Version 7.3 (UDI Number (01)00842000100690(10)7.3.0.158618(11)180409); 7.3.1 (UDI Number (01)00842000100690(10)7.3.1.161404(11)180830); 8.0 (UDI Number (01) 00842000100737 (10) 8.0.0.162051(11)181026); 8.0.1 (UDI Number (01) 00842000100737(10)8.0.1.165704(11)190528) Recalling Firm/ Manufacturer Merge Healthcare, Inc. 900 Walnut Ridge Dr Hartland WI 53029-8347 For Additional Information Contact 262-367-0700 Manufacturer Reason for Recall Potential for prior thumbnails to not display in reverse chronological order and images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. ************Merge PACS provides image manipulation tools to enable users to view and compare images such as: linking, MPR, MIP, 3D image fusion/registration of CT, MR and PET; as well as CVR (Color Volume Rendering), measurements (linear distances, angles, areas, SUV, etc.), and annotations (for example, outline and label regions of interest, label spinal vertebrae). FDA Determined Cause 2 Software Design Change Action Merge Healthcare notified customers on about 07/11/2019 via "Urgent Field Safety Notice" letter. The notice informed customers of two possible issues that may occur on Merge PACS and Merge OrthoPACS: 1. Prior thumbnails may not display in reverse chronological order. 2. Images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. ************Use of an out of order image may result in the implementation of an incorrect or unnecessary treatment plan such as unnecessary medication, change in dosage of existing medication, or unnecessary medical procedures. It may also lead to progression of an existing patient condition, which may result in the progression of a disease state. ************Merge Healthcare will provide a software correction for the issue. A Merge Healthcare representative will contact each customer schedule the correction. ************In the interim, Merge Healthcare instructed customers that they can continue to use their Merge PACS/OrthoPACS devices by following the below instructions:******1. When pulling a series from thumbnail bar, confirm that the date of each series is visible in the viewport. The date will appear when overlays are configured to display the date. An administrator can assist in configuring Merge PACS/OrthoPACS if needed. ******2. When utilizing the mammography scroll stacking tool, confirm that the date of each series is visible in the viewport. The date will appear when overlays are configured to display the date. An administrator can assist in configuring Merge PACS/OrthoPACS if needed.************Customers were instructed to use the date, rather than the order the series are presented in, to determine chronological order.************Additional instructions included to ensure all potential users of Merge PACS and Merge OrthoPACS have received this notice. Questions regarding the recall notice can be directed to Merge Customer Support at 877-741-5369. Quantity in Commerce 179 devices Distribution Nationwide distribution to Arizona, Arkansas, California, Colorado, Florida, Georgia, Hawaii, Illinois, Iowa, Kentucky, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New York, North Carolina, Ohio, Pennsylvania, Rhode Island, South Carolina, Texas, Utah, Virginia, Wisconsin. ************International distribution to Belgium and Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Merge Healthcare notified customers on about 07/11/2019 via "Urgent Field Safety Notice" letter. The notice informed customers of two possible issues that may occur on Merge PACS and Merge OrthoPACS: 1. Prior thumbnails may not display in reverse chronological order. 2. Images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. ************Use of an out of order image may result in the implementation of an incorrect or unnecessary treatment plan such as unnecessary medication, change in dosage of existing medication, or unnecessary medical procedures. It may also lead to progression of an existing patient condition, which may result in the progression of a disease state. ************Merge Healthcare will provide a software correction for the issue. A Merge Healthcare representative will contact each customer schedule the correction. ************In the interim, Merge Healthcare instructed customers that they can continue to use their Merge PACS/OrthoPACS devices by following the below instructions:******1. When pulling a series from thumbnail bar, confirm that the date of each series is visible in the viewport. The date will appear when overlays are configured to display the date. An administrator can assist in configuring Merge PACS/OrthoPACS if needed. ******2. When utilizing the mammography scroll stacking tool, confirm that the date of each series is visible in the viewport. The date will appear when overlays are configured to display the date. An administrator can assist in configuring Merge PACS/OrthoPACS if needed.************Customers were instructed to use the date, rather than the order the series are presented in, to determine chronological order.************Additional instructions included to ensure all potential users of Merge PACS and Merge OrthoPACS have received this notice. Questions regarding the recall notice can be directed to Merge Customer Support at 877-741-5369. | SoftwareDesignChange
93 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172008 | Z-1489-2019 | Horizon Cardiology Hemo | Horizon Cardiology Hemo | 2 | CDRH | 2019-05-24 |  | 2025-03-04 11:07 | Change Healthcare Israel Ltd. | users are not notified of procedure medication discrepancies between the Vitals and Meds, the Procedure Notes and Patient Common Data screens in Horizon/McKesson Cardiology Hemo" (Hemo). The discrepancy may also affect administered medication data in reports generated from Hemo or the Horizon/McKesson Cardiology Physician s Report as well as customers, who have implemented an outbound interface of procedure results. | Class 2 Device Recall Horizon Cardiology Hemo Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Horizon Cardiology Hemo FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Horizon Cardiology Hemo Date Initiated by Firm March 25, 2019 Create Date May 24, 2019 Recall Status 1 Terminated 3 on March 04, 2025 Recall Number Z-1489-2019 Recall Event ID 82600 510(K)Number K131497 Product Classification Computer, diagnostic, programmable - Product Code DQK Product Horizon Cardiology Hemo Code Information Software version 11.x and 12.x Recalling Firm/ Manufacturer Change Healthcare Israel Ltd. 26, Ha-Rokmim Holon Israel Manufacturer Reason for Recall users are not notified of procedure medication discrepancies between the Vitals and Meds, the Procedure Notes and Patient Common Data screens in Horizon/McKesson Cardiology Hemo" (Hemo). The discrepancy may also affect administered medication data in reports generated from Hemo or the Horizon/McKesson Cardiology Physician s Report as well as customers, who have implemented an outbound interface of procedure results. FDA Determined Cause 2 Software Design Change Action Change Healthcare notified consignees on about 03/25/2019 via "Urgent Field Safety Notice" letter sent through email. The notification informed consignees that the issue can occur under the following 3 circumstances: 1. When a user intentionally updates the Physician Report, overriding procedure medication******derived from the Hemo report, or 2. When a user intentionally updates the medication information charted for Pre- or Intra- procedure stage, from the Intra- or Post-procedure stage (respectively), or 3. Due to an unexpected system error. ************Instructions included that prior to closing a Hemo case, ensure the list of administered medication information is correct on the Vitals and Meds, Procedure Notes, and Patient Common Data screens, and when opening, finalizing or closing the Physician Report, review the procedure medications to ensure that all medication information is correct. Customers were requested to immediately contact Change Healthcare to acknowledge they have read and understand the Urgent Field Safety Notice and to arrange for installation of the software update/upgrade. They were also requested to ensure distribution of the Urgent Field Safety Notice to all personnel within your organization who should be aware of the recall or to other affiliated parties that may be affected by the Urgent Field Safety Notice.************Customer Support contacts are as follows: ******1. US/Canada: 1-877-654-4366; ******2. UK / Ireland: +44 (0) 208 9527399; ******3. Australia: 1800 625 435; ******4. Israel: +972 52 433 3366 Quantity in Commerce 125 devices Distribution Nationwide distribution to AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, ******OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, WY. ************International distribution to Australia, Canada, Germany, Ireland, Israel, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = DQK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Change Healthcare notified consignees on about 03/25/2019 via "Urgent Field Safety Notice" letter sent through email. The notification informed consignees that the issue can occur under the following 3 circumstances: 1. When a user intentionally updates the Physician Report, overriding procedure medication******derived from the Hemo report, or 2. When a user intentionally updates the medication information charted for Pre- or Intra- procedure stage, from the Intra- or Post-procedure stage (respectively), or 3. Due to an unexpected system error. ************Instructions included that prior to closing a Hemo case, ensure the list of administered medication information is correct on the Vitals and Meds, Procedure Notes, and Patient Common Data screens, and when opening, finalizing or closing the Physician Report, review the procedure medications to ensure that all medication information is correct. Customers were requested to immediately contact Change Healthcare to acknowledge they have read and understand the Urgent Field Safety Notice and to arrange for installation of the software update/upgrade. They were also requested to ensure distribution of the Urgent Field Safety Notice to all personnel within your organization who should be aware of the recall or to other affiliated parties that may be affected by the Urgent Field Safety Notice.************Customer Support contacts are as follows: ******1. US/Canada: 1-877-654-4366; ******2. UK / Ireland: +44 (0) 208 9527399; ******3. Australia: 1800 625 435; ******4. Israel: +972 52 433 3366 | SoftwareDesignChange
94 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172011 | Z-1490-2019 | McKesson Cardiology Hemo | McKesson Cardiology Hemo | 2 | CDRH | 2019-05-24 |  | 2025-03-04 11:07 | Change Healthcare Israel Ltd. | users are not notified of procedure medication discrepancies between the Vitals and Meds, the Procedure Notes and Patient Common Data screens in Horizon/McKesson Cardiology Hemo" (Hemo). The discrepancy may also affect administered medication data in reports generated from Hemo or the Horizon/McKesson Cardiology Physician s Report as well as customers, who have implemented an outbound interface of procedure results. | Class 2 Device Recall McKesson Cardiology Hemo Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall McKesson Cardiology Hemo FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall McKesson Cardiology Hemo Date Initiated by Firm March 25, 2019 Create Date May 24, 2019 Recall Status 1 Terminated 3 on March 04, 2025 Recall Number Z-1490-2019 Recall Event ID 82600 510(K)Number K131497 Product Classification Computer, diagnostic, programmable - Product Code DQK Product McKesson Cardiology Hemo Code Information Software version 13.x Recalling Firm/ Manufacturer Change Healthcare Israel Ltd. 26, Ha-Rokmim Holon Israel Manufacturer Reason for Recall users are not notified of procedure medication discrepancies between the Vitals and Meds, the Procedure Notes and Patient Common Data screens in Horizon/McKesson Cardiology Hemo" (Hemo). The discrepancy may also affect administered medication data in reports generated from Hemo or the Horizon/McKesson Cardiology Physician s Report as well as customers, who have implemented an outbound interface of procedure results. FDA Determined Cause 2 Software Design Change Action Change Healthcare notified consignees on about 03/25/2019 via "Urgent Field Safety Notice" letter sent through email. The notification informed consignees that the issue can occur under the following 3 circumstances: 1. When a user intentionally updates the Physician Report, overriding procedure medication******derived from the Hemo report, or 2. When a user intentionally updates the medication information charted for Pre- or Intra- procedure stage, from the Intra- or Post-procedure stage (respectively), or 3. Due to an unexpected system error. ************Instructions included that prior to closing a Hemo case, ensure the list of administered medication information is correct on the Vitals and Meds, Procedure Notes, and Patient Common Data screens, and when opening, finalizing or closing the Physician Report, review the procedure medications to ensure that all medication information is correct. Customers were requested to immediately contact Change Healthcare to acknowledge they have read and understand the Urgent Field Safety Notice and to arrange for installation of the software update/upgrade. They were also requested to ensure distribution of the Urgent Field Safety Notice to all personnel within your organization who should be aware of the recall or to other affiliated parties that may be affected by the Urgent Field Safety Notice.************Customer Support contacts are as follows: ******1. US/Canada: 1-877-654-4366; ******2. UK / Ireland: +44 (0) 208 9527399; ******3. Australia: 1800 625 435; ******4. Israel: +972 52 433 3366 Quantity in Commerce 1,383 devices Distribution Nationwide distribution to AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, ******OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, WY. ************International distribution to Australia, Canada, Germany, Ireland, Israel, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = DQK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Change Healthcare notified consignees on about 03/25/2019 via "Urgent Field Safety Notice" letter sent through email. The notification informed consignees that the issue can occur under the following 3 circumstances: 1. When a user intentionally updates the Physician Report, overriding procedure medication******derived from the Hemo report, or 2. When a user intentionally updates the medication information charted for Pre- or Intra- procedure stage, from the Intra- or Post-procedure stage (respectively), or 3. Due to an unexpected system error. ************Instructions included that prior to closing a Hemo case, ensure the list of administered medication information is correct on the Vitals and Meds, Procedure Notes, and Patient Common Data screens, and when opening, finalizing or closing the Physician Report, review the procedure medications to ensure that all medication information is correct. Customers were requested to immediately contact Change Healthcare to acknowledge they have read and understand the Urgent Field Safety Notice and to arrange for installation of the software update/upgrade. They were also requested to ensure distribution of the Urgent Field Safety Notice to all personnel within your organization who should be aware of the recall or to other affiliated parties that may be affected by the Urgent Field Safety Notice.************Customer Support contacts are as follows: ******1. US/Canada: 1-877-654-4366; ******2. UK / Ireland: +44 (0) 208 9527399; ******3. Australia: 1800 625 435; ******4. Israel: +972 52 433 3366 | SoftwareDesignChange
95 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172923 | Z-1384-2019 | UniCel DxH 900 Coulter Cellular Analysis System, REF number C11478******Product Usage:***Note: The DxH 900 hematology analyzer is CLIA categorized under document CR180294 and is currently marketed as a member of the instrument family of the DxH 800 cleared under K140911, per FDA s guidance, Guidance for Industry and FDA Staff, Replacement Reagent and Instrument Family Policy . DxH 800 intended use states: The UniCel DxH 800 Analyzer is a quantitative multi-parameter, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel¿ DxH 800 Analyzer identifies and enumerates the parameters indicated below on the following sample types: Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV, NE%, NE#, LY%, LY#, MO%, MO#, EO%, EO#, BA%, BA#, NRBC%, NRBC#, RET%, RET#, MRV, IRF Pre-Diluted Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV Body Fluids (cerebrospinal, serous and synovial) o TNC and RBC*** | UniCel DxH 900 Coulter Cellular Analysis System | 1 | CDRH | 2019-05-23 | 2019-05-23 21:00 | 2022-03-31 14:47 | Beckman Coulter Inc. | Complaints received of sporadic erroneously elevated platelet results without flags or system messages. Thrombocytopenia may go unrecognized. | Class 1 Device Recall UniCel DxH 900 Coulter Cellular Analysis System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall UniCel DxH 900 Coulter Cellular Analysis System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall UniCel DxH 900 Coulter Cellular Analysis System Date Initiated by Firm July 30, 2018 Date Posted May 23, 2019 Recall Status 1 Terminated 3 on March 31, 2022 Recall Number Z-1384-2019 Recall Event ID 82329 510(K)Number K140911 Product Classification Counter, differential cell - Product Code GKZ Product UniCel DxH 900 Coulter Cellular Analysis System, REF number C11478************Product Usage:******Note: The DxH 900 hematology analyzer is CLIA categorized under document CR180294 and is currently marketed as a member of the instrument family of the DxH 800 cleared under K140911, per FDA s guidance, Guidance for Industry and FDA Staff, Replacement Reagent and Instrument Family Policy . DxH 800 intended use states: The UniCel DxH 800 Analyzer is a quantitative multi-parameter, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer identifies and enumerates the parameters indicated below on the following sample types: Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV, NE%, NE#, LY%, LY#, MO%, MO#, EO%, EO#, BA%, BA#, NRBC%, NRBC#, RET%, RET#, MRV, IRF Pre-Diluted Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV Body Fluids (cerebrospinal, serous and synovial) o TNC and RBC Code Information All current software versions are impacted Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact Mr. Christopher Riley 714-993-5321 Manufacturer Reason for Recall Complaints received of sporadic erroneously elevated platelet results without flags or system messages. Thrombocytopenia may go unrecognized. FDA Determined Cause 2 Software Design Change Action Urgent Medical Device Recall letters, dated July 30, 2018, were sent via email and postal mail on August 1, 2018 (for those customers not on their email notification system). On 05/22/19, second Urgent Medical Device (UMDC) Recall letters, and Important Safety Information notices, dated May 20, 2019 (replacing BEC UMDC letter dated July 30, 2018), were sent via mail and/or phone and/or e-mail informing customers perform the following actions:****** ******- Run samples on an instrument not subject to this recall to confirm the platelet results. ******- If an alternative instrument is not available, use the following quality control measures to aid in identification of discrepant platelet results:************a) Perform manual scanning/estimate of platelets on a peripheral smear and compare with instrument results. Note that this method will identify samples with marked to moderate thrombocytopenia but may not identify smaller discrepancies.******b) Repeat testing of samples in a workflow configuration may facilitate the identification of discrepancies. If an erroneous result is detected, review results from adjacent samples, i.e., those tested on the instrument both before and after the erroneous result.******c) Additional instrument and/or LIS features including reference ranges, XM (exponentially-weighted moving average) and delta checks may be informative.******d) Follow your laboratory's standard operating procedure to confirm unexpected results.************- Ongoing investigation indicates that the probable root cause is the sweep flow disruption that may occur following the "Clear RBC Apertures" procedure. This potential root cause is currently under further investigation. Customers should discontinue using this procedure. If you suspect that your instrument has a clogged aperture that will not clear, discontinue use of the analyzer, contact Customer Support Center and request service.************- Communicate to the ordering physicians the need to avoid patient treatment based solely on any single test resu Quantity in Commerce 341 Distribution Worldwide Distribution - US Nationwide including Puerto Rico. There was government and military distribution. ************Foreign distribution was made to Andorra, Australia, Austria, Belgium, Canada, China, Croatia, Egypt, France, Germany, Greece, Hungary, India, Israel, Italy, Japan, Korea, Kuwait, Latvia, Lithuania, Mexico, Morocco, Myanmar, Netherlands Antilles, New Zealand, Portugal, Romania, Russia, Saudi Arabia, Serbia, Slovenia, Spain, Switzerland, Trinidad and Tobago, U.A.E, United Kingdom, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = GKZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Urgent Medical Device Recall letters, dated July 30, 2018, were sent via email and postal mail on August 1, 2018 (for those customers not on their email notification system). On 05/22/19, second Urgent Medical Device (UMDC) Recall letters, and Important Safety Information notices, dated May 20, 2019 (replacing BEC UMDC letter dated July 30, 2018), were sent via mail and/or phone and/or e-mail informing customers perform the following actions:****** ******- Run samples on an instrument not subject to this recall to confirm the platelet results. ******- If an alternative instrument is not available, use the following quality control measures to aid in identification of discrepant platelet results:************a) Perform manual scanning/estimate of platelets on a peripheral smear and compare with instrument results. Note that this method will identify samples with marked to moderate thrombocytopenia but may not identify smaller discrepancies.******b) Repeat testing of samples in a workflow configuration may facilitate the identification of discrepancies. If an erroneous result is detected, review results from adjacent samples, i.e., those tested on the instrument both before and after the erroneous result.******c) Additional instrument and/or LIS features including reference ranges, XM (exponentially-weighted moving average) and delta checks may be informative.******d) Follow your laboratory's standard operating procedure to confirm unexpected results.************- Ongoing investigation indicates that the probable root cause is the sweep flow disruption that may occur following the "Clear RBC Apertures" procedure. This potential root cause is currently under further investigation. Customers should discontinue using this procedure. If you suspect that your instrument has a clogged aperture that will not clear, discontinue use of the analyzer, contact Customer Support Center and request service.************- Communicate to the ordering physicians the need to avoid patient treatment based solely on any single test resu | SoftwareDesignChange
96 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171362 | Z-1383-2019 | UniCel DxH 600 Coulter Cellular Analysis System, REF number B23858.******Product Usage:***Note: The DxH 600 hematology analyzer is the same instrument as the DxH 800 (K140911), except that it is a benchtop unit with no instrument cart, computer stand, and no connectivity capabilities. DxH 800 intended use states: The UniCel DxH 800 Analyzer is a quantitative multi-parameter, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel¿ DxH 800 Analyzer identifies and enumerates the parameters indicated below on the following sample types: Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV, NE%, NE#, LY%, LY#, MO%, MO#, EO%, EO#, BA%, BA#, NRBC%, NRBC#, RET%, RET#, MRV, IRF Pre-Diluted Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV Body Fluids (cerebrospinal, serous and synovial) o TNC and RBC*** | UniCel DxH 600 Coulter Cellular Analysis System | 1 | CDRH | 2019-05-23 | 2019-05-23 21:00 | 2022-03-31 14:47 | Beckman Coulter Inc. | Complaints received of sporadic erroneously elevated platelet results without flags or system messages. Thrombocytopenia may go unrecognized. | Class 1 Device Recall UniCel DxH 600 Coulter Cellular Analysis System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall UniCel DxH 600 Coulter Cellular Analysis System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall UniCel DxH 600 Coulter Cellular Analysis System Date Initiated by Firm July 30, 2018 Date Posted May 23, 2019 Recall Status 1 Terminated 3 on March 31, 2022 Recall Number Z-1383-2019 Recall Event ID 82329 510(K)Number K140911 Product Classification Counter, differential cell - Product Code GKZ Product UniCel DxH 600 Coulter Cellular Analysis System, REF number B23858.************Product Usage:******Note: The DxH 600 hematology analyzer is the same instrument as the DxH 800 (K140911), except that it is a benchtop unit with no instrument cart, computer stand, and no connectivity capabilities. DxH 800 intended use states: The UniCel DxH 800 Analyzer is a quantitative multi-parameter, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer identifies and enumerates the parameters indicated below on the following sample types: Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV, NE%, NE#, LY%, LY#, MO%, MO#, EO%, EO#, BA%, BA#, NRBC%, NRBC#, RET%, RET#, MRV, IRF Pre-Diluted Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV Body Fluids (cerebrospinal, serous and synovial) o TNC and RBC Code Information All current software versions are impacted Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact Mr. Christopher Riley 714-993-5321 Manufacturer Reason for Recall Complaints received of sporadic erroneously elevated platelet results without flags or system messages. Thrombocytopenia may go unrecognized. FDA Determined Cause 2 Software Design Change Action Urgent Medical Device Recall letters, dated July 30, 2018, were sent via email and postal mail on August 1, 2018 (for those customers not on their email notification system). On 05/22/19, second Urgent Medical Device (UMDC) Recall letters, and Important Safety Information notices, dated May 20, 2019 (replacing BEC UMDC letter dated July 30, 2018), were sent via mail and/or phone and/or e-mail informing customers perform the following actions:****** ******- Run samples on an instrument not subject to this recall to confirm the platelet results. ******- If an alternative instrument is not available, use the following quality control measures to aid in identification of discrepant platelet results:************a) Perform manual scanning/estimate of platelets on a peripheral smear and compare with instrument results. Note that this method will identify samples with marked to moderate thrombocytopenia but may not identify smaller discrepancies.******b) Repeat testing of samples in a workflow configuration may facilitate the identification of discrepancies. If an erroneous result is detected, review results from adjacent samples, i.e., those tested on the instrument both before and after the erroneous result.******c) Additional instrument and/or LIS features including reference ranges, XM (exponentially-weighted moving average) and delta checks may be informative.******d) Follow your laboratory's standard operating procedure to confirm unexpected results.************- Ongoing investigation indicates that the probable root cause is the sweep flow disruption that may occur following the "Clear RBC Apertures" procedure. This potential root cause is currently under further investigation. Customers should discontinue using this procedure. If you suspect that your instrument has a clogged aperture that will not clear, discontinue use of the analyzer, contact Customer Support Center and request service.************- Communicate to the ordering physicians the need to avoid patient treatment based solely on any single test resu Quantity in Commerce 2,059 devices Distribution Worldwide Distribution - US Nationwide including Puerto Rico. There was government and military distribution. ************Foreign distribution was made to Andorra, Australia, Austria, Belgium, Canada, China, Croatia, Egypt, France, Germany, Greece, Hungary, India, Israel, Italy, Japan, Korea, Kuwait, Latvia, Lithuania, Mexico, Morocco, Myanmar, Netherlands Antilles, New Zealand, Portugal, Romania, Russia, Saudi Arabia, Serbia, Slovenia, Spain, Switzerland, Trinidad and Tobago, U.A.E, United Kingdom, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = GKZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Urgent Medical Device Recall letters, dated July 30, 2018, were sent via email and postal mail on August 1, 2018 (for those customers not on their email notification system). On 05/22/19, second Urgent Medical Device (UMDC) Recall letters, and Important Safety Information notices, dated May 20, 2019 (replacing BEC UMDC letter dated July 30, 2018), were sent via mail and/or phone and/or e-mail informing customers perform the following actions:****** ******- Run samples on an instrument not subject to this recall to confirm the platelet results. ******- If an alternative instrument is not available, use the following quality control measures to aid in identification of discrepant platelet results:************a) Perform manual scanning/estimate of platelets on a peripheral smear and compare with instrument results. Note that this method will identify samples with marked to moderate thrombocytopenia but may not identify smaller discrepancies.******b) Repeat testing of samples in a workflow configuration may facilitate the identification of discrepancies. If an erroneous result is detected, review results from adjacent samples, i.e., those tested on the instrument both before and after the erroneous result.******c) Additional instrument and/or LIS features including reference ranges, XM (exponentially-weighted moving average) and delta checks may be informative.******d) Follow your laboratory's standard operating procedure to confirm unexpected results.************- Ongoing investigation indicates that the probable root cause is the sweep flow disruption that may occur following the "Clear RBC Apertures" procedure. This potential root cause is currently under further investigation. Customers should discontinue using this procedure. If you suspect that your instrument has a clogged aperture that will not clear, discontinue use of the analyzer, contact Customer Support Center and request service.************- Communicate to the ordering physicians the need to avoid patient treatment based solely on any single test resu | SoftwareDesignChange
97 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171361 | Z-1382-2019 | UniCel DxH 800 Coulter Cellular Analysis System, REF numbers 629029, B24465, B24802, B68304, B66445, and B63322.******Product Usage:***The UniCel DxH 800 Analyzer is a quantitative multi-parameter, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel¿ DxH 800 Analyzer identifies and enumerates the parameters indicated below on the following sample types: Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV, NE%, NE#, LY%, LY#, MO%, MO#, EO%, EO#, BA%, BA#, NRBC%, NRBC#, RET%, RET#, MRV, IRF Pre-Diluted Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV Body Fluids (cerebrospinal, serous and synovial) o TNC and RBC*** | UniCel DxH 800 Coulter Cellular Analysis System | 1 | CDRH | 2019-05-23 | 2019-05-23 21:00 | 2022-03-31 14:47 | Beckman Coulter Inc. | Complaints received of sporadic erroneously elevated platelet results without flags or system messages. Thrombocytopenia may go unrecognized. | Class 1 Device Recall UniCel DxH 800 Coulter Cellular Analysis System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 1 Device Recall UniCel DxH 800 Coulter Cellular Analysis System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 1 Device Recall UniCel DxH 800 Coulter Cellular Analysis System Date Initiated by Firm July 30, 2018 Date Posted May 23, 2019 Recall Status 1 Terminated 3 on March 31, 2022 Recall Number Z-1382-2019 Recall Event ID 82329 510(K)Number K140911 Product Classification Counter, differential cell - Product Code GKZ Product UniCel DxH 800 Coulter Cellular Analysis System, REF numbers 629029, B24465, B24802, B68304, B66445, and B63322.************Product Usage:******The UniCel DxH 800 Analyzer is a quantitative multi-parameter, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel DxH 800 Analyzer identifies and enumerates the parameters indicated below on the following sample types: Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV, NE%, NE#, LY%, LY#, MO%, MO#, EO%, EO#, BA%, BA#, NRBC%, NRBC#, RET%, RET#, MRV, IRF Pre-Diluted Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV Body Fluids (cerebrospinal, serous and synovial) o TNC and RBC Code Information All current software versions are impacted Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 For Additional Information Contact Mr. Christopher Riley 714-993-5321 Manufacturer Reason for Recall Complaints received of sporadic erroneously elevated platelet results without flags or system messages. Thrombocytopenia may go unrecognized. FDA Determined Cause 2 Software Design Change Action Urgent Medical Device Recall letters, dated July 30, 2018, were sent via email and postal mail on August 1, 2018 (for those customers not on their email notification system). On 05/22/19, second Urgent Medical Device (UMDC) Recall letters, and Important Safety Information notices, dated May 20, 2019 (replacing BEC UMDC letter dated July 30, 2018), were sent via mail and/or phone and/or e-mail informing customers perform the following actions:****** ******- Run samples on an instrument not subject to this recall to confirm the platelet results. ******- If an alternative instrument is not available, use the following quality control measures to aid in identification of discrepant platelet results:************a) Perform manual scanning/estimate of platelets on a peripheral smear and compare with instrument results. Note that this method will identify samples with marked to moderate thrombocytopenia but may not identify smaller discrepancies.******b) Repeat testing of samples in a workflow configuration may facilitate the identification of discrepancies. If an erroneous result is detected, review results from adjacent samples, i.e., those tested on the instrument both before and after the erroneous result.******c) Additional instrument and/or LIS features including reference ranges, XM (exponentially-weighted moving average) and delta checks may be informative.******d) Follow your laboratory's standard operating procedure to confirm unexpected results.************- Ongoing investigation indicates that the probable root cause is the sweep flow disruption that may occur following the "Clear RBC Apertures" procedure. This potential root cause is currently under further investigation. Customers should discontinue using this procedure. If you suspect that your instrument has a clogged aperture that will not clear, discontinue use of the analyzer, contact Customer Support Center and request service.************- Communicate to the ordering physicians the need to avoid patient treatment based solely on any single test resu Quantity in Commerce 5,414 devices Distribution Worldwide Distribution - US Nationwide including Puerto Rico. There was government and military distribution. ************Foreign distribution was made to Andorra, Australia, Austria, Belgium, Canada, China, Croatia, Egypt, France, Germany, Greece, Hungary, India, Israel, Italy, Japan, Korea, Kuwait, Latvia, Lithuania, Mexico, Morocco, Myanmar, Netherlands Antilles, New Zealand, Portugal, Romania, Russia, Saudi Arabia, Serbia, Slovenia, Spain, Switzerland, Trinidad and Tobago, U.A.E, United Kingdom, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = GKZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Urgent Medical Device Recall letters, dated July 30, 2018, were sent via email and postal mail on August 1, 2018 (for those customers not on their email notification system). On 05/22/19, second Urgent Medical Device (UMDC) Recall letters, and Important Safety Information notices, dated May 20, 2019 (replacing BEC UMDC letter dated July 30, 2018), were sent via mail and/or phone and/or e-mail informing customers perform the following actions:****** ******- Run samples on an instrument not subject to this recall to confirm the platelet results. ******- If an alternative instrument is not available, use the following quality control measures to aid in identification of discrepant platelet results:************a) Perform manual scanning/estimate of platelets on a peripheral smear and compare with instrument results. Note that this method will identify samples with marked to moderate thrombocytopenia but may not identify smaller discrepancies.******b) Repeat testing of samples in a workflow configuration may facilitate the identification of discrepancies. If an erroneous result is detected, review results from adjacent samples, i.e., those tested on the instrument both before and after the erroneous result.******c) Additional instrument and/or LIS features including reference ranges, XM (exponentially-weighted moving average) and delta checks may be informative.******d) Follow your laboratory's standard operating procedure to confirm unexpected results.************- Ongoing investigation indicates that the probable root cause is the sweep flow disruption that may occur following the "Clear RBC Apertures" procedure. This potential root cause is currently under further investigation. Customers should discontinue using this procedure. If you suspect that your instrument has a clogged aperture that will not clear, discontinue use of the analyzer, contact Customer Support Center and request service.************- Communicate to the ordering physicians the need to avoid patient treatment based solely on any single test resu | SoftwareDesignChange
98 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=170821 | Z-0973-2019 | Stryker Sustainability Solutions (SSS) reprocessed Arthroscopic Shavers****** Arthroscopic shavers can be used to abrade, cut and excise tissue and bone; remove loose fragments; and, shave away debris in arthroscopic surgeries, as well as surgeries of the jaw and sinuses. Intended for use in orthopedic surgical procedures of the joints, jaw or sinuses where the cutting and removal of soft and hard tissue or bone is needed in patients requiring orthopedic surgery. | Stryker Sustainability Solutions (SSS) | 2 | CDRH | 2019-03-05 |  | 2020-09-17 13:46 | Stryker Sustainability Solutions | Affected lots of reprocessed arthroscopic shavers may display the incorrect catalog number and speed settings when connected to the generator, require manual mode for use on CORE generators, or display an error message on Crossfire 1 and 2 generators which does not allow the device to be used. | Class 2 Device Recall Stryker Sustainability Solutions (SSS) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Stryker Sustainability Solutions (SSS) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Stryker Sustainability Solutions (SSS) Date Initiated by Firm July 27, 2018 Create Date March 05, 2019 Recall Status 1 Terminated 3 on September 17, 2020 Recall Number Z-0973-2019 Recall Event ID 82063 510(K)Number K012605 Product Classification Arthroscope - Product Code HRX Product Stryker Sustainability Solutions (SSS) reprocessed Arthroscopic Shavers************ Arthroscopic shavers can be used to abrade, cut and excise tissue and bone; remove loose fragments; and, shave away debris in arthroscopic surgeries, as well as surgeries of the jaw and sinuses. Intended for use in orthopedic surgical procedures of the joints, jaw or sinuses where the cutting and removal of soft and hard tissue or bone is needed in patients requiring orthopedic surgery. Code Information Catalog Numbers: 375-450-000 375-940-000, 375-532-000 375-940-012, 375-534-000 375-941-000, 375-542-000 375-941-012, 375-544-000 375-941-500, 375-545-000 375-950-000, 375-548-000 375-950-012, 375-552-000 375-950-112, 375-554-000 375-951-000, 375-555-000 375-951-012, 375-562-000 375-951-100, 375-564-000 375-951-112, 375-565-000 375-951-200 and 375-747-000 375-952-000. Lots from affected catalog numbers: 4601435***7687551***7954336***8069955***8143951***8213505***8290940 4866597***7692818***7957638***8078776***8143979***8216233***8296920 5275122***7695149***7962443***8078965***8144105***8220563***8297887 5810401***7696653***7964118***8079076***8144770***8220748***8297970 6165680***7702505***7964966***8082376***8146583***8223629***8298035 6552450***7725942***7965089***8082638***8148337***8223910***8298128 6563177***7743624***7969196***8082746***8150492***8224388***8298625 6634381***7755555***7978089***8095770***8150617***8224621***8305181 6694100***7766506***7978931***8096123***8150664***8228215***8305467 6705350***7767267***7981959***8096255***8152993***8228340***8305524 6712065***7768175***7986332***8096613***8153141***8228408***8305595 6761681***7782229***7991417***8099709***8153248***8231445***8308196 6819315***7821595***7997091***8100287***8153419***8231599***8308261 7066139***7829681***8004687***8100406***8153658***8231736***8310523 7076865***7831750***8007843***8104913***8154010***8231864***8320330 7119659***7838095***8016985***8104966***8159807***8232000***8320617 7129915***7839027***8022071***8105004***8159895***8235618***8323092 7184490***7840722***8028142***8105058***8161407***8244348***8323194 7196683***7842014***8028208***8105299***8163785***8244353***8323457 7252265***7842476***8028220***8105378***8163860***8244492***8323899 7370055***7842578***8028556***8113641***8163873***8244519***8330243 7383670***7851733***8028637***8113734***8166754***8247857***8339095 7403808***7852107***8036566***8113809***8167287***8251682***8339146 7413359***7852182***8036815***8113845***8170405***8251719***8352883 7428476***7865156***8040157***8117025***8170763***8251844***8359800 7455721***7868429***8040323***8117115***8172102***8251994***8367003 7505081***7869132***8045052***8121520***8177364***8252345***8367185 7505505***7874618***8045121***8121667***8181204***8257718***8377347 7535206***7879255***8045270***8121786***8181289***8265008***8388727 7544440***7879286***8048728***8125248***8181408***8265217***8397197 7567055***7903097***8049139***8125474***8181431***8267426***8400191 7597084***7903772***8049230***8129960***8188037***8267517***8418674 7615431***7911516***8053354***8130103***8188419***8272680***8423234 7624503***7924275***8053383***8130220***8192179***8273138***8423268 7624733***7929803***8056987***8130853***8196392***8273172***8424154 7655856***7933324***8064615***8131232***8196545***8273275***8424316 7674248***7933542***8064847***8131395***8196681***8276053***8424407 7674643***7938522***8065027***8131601***8196702***8287528***8440175 7685438***7941629***8069917***8143807***8213465***8290290***8462684 Recalling Firm/ Manufacturer Stryker Sustainability Solutions 1810 W Drake Dr Tempe AZ 85283-4327 For Additional Information Contact Marissa Richmond 480-763-5300 Ext. 5034 Manufacturer Reason for Recall Affected lots of reprocessed arthroscopic shavers may display the incorrect catalog number and speed settings when connected to the generator, require manual mode for use on CORE generators, or display an error message on Crossfire 1 and 2 generators which does not allow the device to be used. FDA Determined Cause 2 Software Design Change Action On July 26, 2018, Stryker Sustainability Solutions sent an "URGENT MEDICAL DEVICE RECALL" Customer Notification Letter to affected consignees asking them to follow the instructions in the letter. The letter also included a risk evaluation these devices may pose to end users and their patients. According to the Urgent Medical Device Recall Customer Notification Letter there are no harms associated with the use of affected reprocessed arthroscopic shavers.******1) Noncompatible devices are rendered inactive with console and alternative surgical methods are readily available.******2) Incorrect spin speed does not result in any harm to the patient************The letter also informed customers to; 1) discontinue use of affected reprocessed arthroscopic shavers; 2) complete the recall effectiveness check form and indicate if any of the affected product remains in their inventory; 3)Return the completed form either via e-mail to SSSPFA@Stryker.com or to their Stryker Sustainability Solutions sales rep, or via fax to 480.763.5345; 4) If no product is in inventory customers are also being told to complete the recall effectiveness check form and send in as described above; 5) finally any adverse events or quality problems experienced with the use of this product should be reported to: The FDA***s MedWatch Adverse Event Reporting program either online at https://www.accessdata.fda.gov/scripts/medwatch/index.cfm?action=reporting.home, by regular mail or by fax; Stryker***s Sustainability Solutions division complaint hotline: +1 888 763 8803 or http://www.stryker.com/productexperience Quantity in Commerce 14,265 units Distribution Nationwide Distribution: AL, AR, AZ, CA, CO, DC, DE, FL, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, WY and Puerto Rico. OUS Israel and Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = HRX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On July 26, 2018, Stryker Sustainability Solutions sent an "URGENT MEDICAL DEVICE RECALL" Customer Notification Letter to affected consignees asking them to follow the instructions in the letter. The letter also included a risk evaluation these devices may pose to end users and their patients. According to the Urgent Medical Device Recall Customer Notification Letter there are no harms associated with the use of affected reprocessed arthroscopic shavers.******1) Noncompatible devices are rendered inactive with console and alternative surgical methods are readily available.******2) Incorrect spin speed does not result in any harm to the patient************The letter also informed customers to; 1) discontinue use of affected reprocessed arthroscopic shavers; 2) complete the recall effectiveness check form and indicate if any of the affected product remains in their inventory; 3)Return the completed form either via e-mail to SSSPFA@Stryker.com or to their Stryker Sustainability Solutions sales rep, or via fax to 480.763.5345; 4) If no product is in inventory customers are also being told to complete the recall effectiveness check form and send in as described above; 5) finally any adverse events or quality problems experienced with the use of this product should be reported to: The FDA***s MedWatch Adverse Event Reporting program either online at https://www.accessdata.fda.gov/scripts/medwatch/index.cfm?action=reporting.home, by regular mail or by fax; Stryker***s Sustainability Solutions division complaint hotline: +1 888 763 8803 or http://www.stryker.com/productexperience | SoftwareDesignChange
99 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=170402 | Z-0847-2019 | VITEK 2 AST-N351 Test Kit | VITEK 2 AST-N351 Test Kit | 2 | CDRH | 2019-02-16 |  | 2021-08-16 06:48 | bioMerieux, Inc. | False Positive ESBL Phenotype****** | Class 2 Device Recall VITEK 2 ASTN351 Test Kit Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VITEK 2 ASTN351 Test Kit FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VITEK 2 ASTN351 Test Kit Date Initiated by Firm December 13, 2018 Create Date February 16, 2019 Recall Status 1 Terminated 3 on August 16, 2021 Recall Number Z-0847-2019 Recall Event ID 81953 Product Classification System, test, automated, antimicrobial susceptibility, short incubation - Product Code LON Product VITEK 2 AST-N351 Test Kit Code Information Ref: 421257 ALL LOTS Recalling Firm/ Manufacturer bioMerieux, Inc. 100 Rodolphe St Durham NC 27712-9402 For Additional Information Contact 919-620-2000 Manufacturer Reason for Recall False Positive ESBL Phenotype FDA Determined Cause 2 Software Design Change Action On 12/13/2018 bioMerieux submitted a communication to their distributor in the UK stating the following:************Our Ref: 4191-1 FSCA******Dear Valued bioMrieux Customer,************Our records indicate that your laboratory uses the VITEK 2 System in conjunction with the VITEK 2 AST-N351 test kit. ************Description of Issue:******bioMrieux has received reports from VITEK 2 customers that an unexpected ESBL (Extended Spectrum Beta Lactamase) phenotype has been proposed for some Escherichia coli strains in conjunction with the VITEK 2 AST-N351 (ref. 421257) test kit.************When an ESBL phenotype is proposed by VITEK 2 Systems Software AES (Advanced Expert System), the result is stopped for review by the user. The intent of the review is that the laboratory personnel take measures to confirm the proposed ESBL phenotype makes sense based on all information available, and perform confirmatory testing if deemed appropriate. ************Internal bioMrieux investigations concluded that within the VITEK 2 Systems Software version 8.01/9.01 AES (Advanced Expert System) module, the MIC (Minimum Inhibitory Concentration) ranges for Escherichia coli and pipericillin/tazobactam were changed for the ***Acquired Penicillinase*** and ***Inhibitor Resistant Penicillinase*** phenotypes. These changes consisted of removing MICs >=128 mg/L as possible MICs for both phenotypes. ************This change revealed an issue in the AES definition of the cefuroxime MIC range for Escherichia coli and ESBLs (the MIC range is too broad; cefuroxime MIC <=8 mg/L should prevent the ESBL phenotype proposal). This, in addition to the specific antibiotic configuration of the AST-N351 card and absence of the ESBL test, is the root cause of the reported increase in ESBL phenotype proposal for the AST-N351 card.************Implementation of a bioART" rule with the following criteria will prevent inappropriate proposal of the ESBL phenotype related to this anomaly (reference Appendix A for detailed instructions):******"***If the card is AST-N351******"***And the ph Quantity in Commerce 30,292 cartons (20 cards per carton) Distribution UK and Ireland Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On 12/13/2018 bioMerieux submitted a communication to their distributor in the UK stating the following:************Our Ref: 4191-1 FSCA******Dear Valued bioMrieux Customer,************Our records indicate that your laboratory uses the VITEK 2 System in conjunction with the VITEK 2 AST-N351 test kit. ************Description of Issue:******bioMrieux has received reports from VITEK 2 customers that an unexpected ESBL (Extended Spectrum Beta Lactamase) phenotype has been proposed for some Escherichia coli strains in conjunction with the VITEK 2 AST-N351 (ref. 421257) test kit.************When an ESBL phenotype is proposed by VITEK 2 Systems Software AES (Advanced Expert System), the result is stopped for review by the user. The intent of the review is that the laboratory personnel take measures to confirm the proposed ESBL phenotype makes sense based on all information available, and perform confirmatory testing if deemed appropriate. ************Internal bioMrieux investigations concluded that within the VITEK 2 Systems Software version 8.01/9.01 AES (Advanced Expert System) module, the MIC (Minimum Inhibitory Concentration) ranges for Escherichia coli and pipericillin/tazobactam were changed for the ***Acquired Penicillinase*** and ***Inhibitor Resistant Penicillinase*** phenotypes. These changes consisted of removing MICs >=128 mg/L as possible MICs for both phenotypes. ************This change revealed an issue in the AES definition of the cefuroxime MIC range for Escherichia coli and ESBLs (the MIC range is too broad; cefuroxime MIC <=8 mg/L should prevent the ESBL phenotype proposal). This, in addition to the specific antibiotic configuration of the AST-N351 card and absence of the ESBL test, is the root cause of the reported increase in ESBL phenotype proposal for the AST-N351 card.************Implementation of a bioART" rule with the following criteria will prevent inappropriate proposal of the ESBL phenotype related to this anomaly (reference Appendix A for detailed instructions):******"***If the card is AST-N351******"***And the ph | SoftwareDesignChange
100 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=170000 | Z-0784-2019 | Anti-HAV IgM test System, Material No. 11820591160******Product Usage:***The Roche Elecsys Anti-HAV IgM immunoassay is used for the in vitro qualitative detection of IgM antibodies to hepatitis A virus (anti-HAV IgM) in human serum and plasma (potassium EDTA, lithium or sodium heparin, sodium citrate). The assay is intended for use as an aid in the laboratory diagnosis of an acute or recently acquired hepatitis A virus infection. Assay results, in conjunction with other laboratory results and clinical information, may be used to provide presumptive evidence of infection with hepatitis A virus in persons with signs and symptoms of hepatitis and in persons at risk for hepatitis A infection. The electrochemiluminescence immunoassay ECLIA is intended for use on Elecsys and cobas e immunoassay analyzers.*** | Elecsys Anti HAV IgM | 2 | CDRH | 2019-01-29 |  | 2020-04-10 13:50 | Roche Diagnostics Corporation | Potential for erroneous result messages for the Anti-HAV IgM assay when run on the cobas e 602 module. Discrepant result reporting may result. | Class 2 Device Recall Elecsys Anti HAV IgM Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Elecsys Anti HAV IgM FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Elecsys Anti HAV IgM Date Initiated by Firm December 17, 2018 Create Date January 29, 2019 Recall Status 1 Terminated 3 on April 10, 2020 Recall Number Z-0784-2019 Recall Event ID 81823 510(K)Number K093955 Product Classification Hepatitis a test (antibody and igm antibody) - Product Code LOL Product Anti-HAV IgM test System, Material No. 11820591160************Product Usage:******The Roche Elecsys Anti-HAV IgM immunoassay is used for the in vitro qualitative detection of IgM antibodies to hepatitis A virus (anti-HAV IgM) in human serum and plasma (potassium EDTA, lithium or sodium heparin, sodium citrate). The assay is intended for use as an aid in the laboratory diagnosis of an acute or recently acquired hepatitis A virus infection. Assay results, in conjunction with other laboratory results and clinical information, may be used to provide presumptive evidence of infection with hepatitis A virus in persons with signs and symptoms of hepatitis and in persons at risk for hepatitis A infection. The electrochemiluminescence immunoassay ECLIA is intended for use on Elecsys and cobas e immunoassay analyzers. Code Information APP-file version 04.10-101 on the cobas e 602 module. Recalling Firm/ Manufacturer Roche Diagnostics Corporation 9115 Hague Rd Indianapolis IN 46256-1025 For Additional Information Contact 800-428-2336 Manufacturer Reason for Recall Potential for erroneous result messages for the Anti-HAV IgM assay when run on the cobas e 602 module. Discrepant result reporting may result. FDA Determined Cause 2 Software Design Change Action Roche issued URGENT MEDICAL DEVICE CORRECTION notices and Fax Forms dated 12/17/2018 to customers via UPS courier service. **************Actions to be Taken by Customers/Users Currently Using APP-file Version 04.10 - 101**************1. If your LIS host stores the interpretation and COI information for Anti-HAV IgM, all COI results in the gray zone of 0.90 *** 1.10 should be reviewed. Implement the following action for all gray zone results:************If discrepant interpretation of results is suspected, re-testing may be advisable in concordance with the recommendations in the assay method sheet and relevant clinical information.************2. Review the APP-file version downloaded for the Anti-HAV IgM application. The most current******APP-file version is (05.30 - 402). Download this version of the APP-file for the Anti-HAV IgM assay. ************You can determine the APP-file version for Anti-HAV IgM currently installed on your cobas e 602 module by going to Utility > Application > Select Anti-HAV IgM test > Analyze tab, bottom right corner of the screen.**************Actions Required********- Follow the instructions provided in the Actions to be Taken by Customers/Users Currently Using APP-file Version 04.10 - 101 ******- Confirm your cobas e 602 module is using APP-file 05.30 - 402. If it is not, immediately update to APP-file 05.30 - 402.******- Complete all sections of the fax form (TP-00518) and fax it to 1-855-207-2758 or email it to 6189roche@stericycle.com.******- File the Urgent Medical Device Correction for future use.************Customers with questions please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 Quantity in Commerce 70,639 units Distribution US Nationwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LOL - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Roche issued URGENT MEDICAL DEVICE CORRECTION notices and Fax Forms dated 12/17/2018 to customers via UPS courier service. **************Actions to be Taken by Customers/Users Currently Using APP-file Version 04.10 - 101**************1. If your LIS host stores the interpretation and COI information for Anti-HAV IgM, all COI results in the gray zone of 0.90 *** 1.10 should be reviewed. Implement the following action for all gray zone results:************If discrepant interpretation of results is suspected, re-testing may be advisable in concordance with the recommendations in the assay method sheet and relevant clinical information.************2. Review the APP-file version downloaded for the Anti-HAV IgM application. The most current******APP-file version is (05.30 - 402). Download this version of the APP-file for the Anti-HAV IgM assay. ************You can determine the APP-file version for Anti-HAV IgM currently installed on your cobas e 602 module by going to Utility > Application > Select Anti-HAV IgM test > Analyze tab, bottom right corner of the screen.**************Actions Required********- Follow the instructions provided in the Actions to be Taken by Customers/Users Currently Using APP-file Version 04.10 - 101 ******- Confirm your cobas e 602 module is using APP-file 05.30 - 402. If it is not, immediately update to APP-file 05.30 - 402.******- Complete all sections of the fax form (TP-00518) and fax it to 1-855-207-2758 or email it to 6189roche@stericycle.com.******- File the Urgent Medical Device Correction for future use.************Customers with questions please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 | SoftwareDesignChange
101 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=168911 | Z-0614-2019 | Xper Flex Cardio Physiomonitoring System, Software versions 1.5.16.0322 and 1.6.1.0390, Model Numbers 453564241901, 453564241911, 453564483321, 453564483331, 453564621791, 453564621801, 453564669081 | Xper Flex Cardio Physiomonitoring System | 2 | CDRH | 2018-12-14 |  | 2024-09-20 14:45 | Philips Electronics North America Corporation | There are 3 issues:***1) Under certain conditions, real time waveforms may not be accurately plotted or displayed. ***2) The oxygen saturation (SpO2) numeric value displayed on the device may freeze.***3) In auto cycle mode, if the NIBP communication is lost, the pump would not cycle and the non-invasive blood pressure (NIBP) numeric value displayed on the device would not update. | Class 2 Device Recall Xper Flex Cardio Physiomonitoring System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Xper Flex Cardio Physiomonitoring System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Xper Flex Cardio Physiomonitoring System Date Initiated by Firm November 15, 2018 Create Date December 14, 2018 Recall Status 1 Terminated 3 on September 20, 2024 Recall Number Z-0614-2019 Recall Event ID 81527 510(K)Number K101571 Product Classification Computer, diagnostic, programmable - Product Code DQK Product Xper Flex Cardio Physiomonitoring System, Software versions 1.5.16.0322 and 1.6.1.0390, Model Numbers 453564241901, 453564241911, 453564483321, 453564483331, 453564621791, 453564621801, 453564669081 Code Information US1251000031 US1311000006 US1407000009 US1419000044 US1438000118 US1446000275 US1547001419 US1610001769 US1625002302 US1638002539 US1450000330 US1746001254 US1311000077 US1417000003 US1511000589 US1319000051 US1537001287 US1736000931 US1604001618 US1623002140 153 US1251000021 US1203010332 US1604001569 US1407000057 US1324000017 420 421 428 251 616 US1234000045 14 US1537001306 45 US1603001509 US1617002012 US1438000113 US1334000026 US1239000047 US1203010392 US1417000036 1213 US1319000082 US1608001676 US1611001852 US1324000025 US1731000712 US1609001745 US1334000028 US1248000022 US1638002517 US1638002591 US1644002662 US1644002668 US1644002682 US1546001353 US1203010414 US1424000022 US1523000858 US1547001364 US1547001374 US1547001394 US1547001395 US1610001753 US1626002429 US1201010146 US1202010179 US1244000031 US1244000058 US1505000520 US1534001212 US1738001102 US1203010380 US1644002701 US1505000552 US1523000883 US1604001624 US1609001735 US1611001878 US1611001892 US1638002619 US1644002625 US1644002635 US1644002684 US1319000043 US1738002731 US1738002766 US1738002769 US1738002776 US1738002782 US1738002786 US1738002798 US1738002799 US1319000061 US1738002727 US1738002739 US1738002743 US1534001213 US1609001737 US1644002629 US1644002697 US1738002736 US1738002749 US1738002756 US1738002759 US1331000014 US1738002730 US1738002732 US1623002113 US1638002549 US1638002607 US1644002655 US1644002674 US1644002678 US1644002679 US1644002680 US1644002681 US1644002691 US1644002696 US1644002699 US1738002729 US1738002733 US1738002751 US1619002085 US1324000001 US1638002570 US1644002683 US1644002698 US1108010051 US1438000122 US1505000553 US1619002099 US1623002117 US1623002189 US1624002199 US1738002728 US1738002750 US1738002763 US1738002764 US1644002634 US1644002644 US1738002734 US1108010061 US1234000017 US1424000039 US1517000675 US1518000800 US1537001309 US1604001612 US1505000534 US1638002593 US1638002622 US1644002641 US1644002693 US1511000617 US1638002569 US1644002653 US1644002666 US1644002695 US1731000671 US1441000153 US1109010105 US1203010326 US1244000014 US1311000058 US1337000034 US1337000047 US1431000020 US1450000362 US1224000035 US1224000042 US1450000355 1169 US1319000071 US1424000049 US1505000516 US1417000022 US1324000021 US1517000729 US1505000521 875 US1424000044 US1529000998 US1625002276 US1518000787 US1609001716 US1319000036 US1720000394 US1234000018 US1407000019 US1625002236 US1311000080 US1625002246 US1337000073 19 US1529001058 US1533001136 US1304000016 US1334000016 US1529001019 US1624002224 US1408000049 US1337000072 US1109010083 US1107010010 US1109010111 US1652000079 US1511000631 US1529001008 US1611001899 US1337000012 1287 US1611001880 US1202010156 US1626002463 US1407000023 US1419000055 US1443000212 931 951 3 1622 952 95 402 US1433000041 501 834 217 US1109010065 287 208 357 934 2042 25 861 438 604 573 503 512 1436 1591 46 225 US1248000079 35 1307 630 526 524 1026 1277 1347 1279 US1604001609 672 65 66 399 US1438000150 US1550001462 US1623002148 US1529001047 US1201010120 US1419000033 US1537001282 US1734000888 US1731000740 US1608001684 US1738001023 US1751001372 358 US1608001661 US1251000044 US1412000048 US1324000026 433 US1511000628 US1724000514 US1319000069 US1529001048 US1234000056 US1331000060 US1709000218 US1311000013 US1517000657 US1337000003 US1343000078 US1202010155 US1231000066 US1251000020 US1337000022 US1337000042 US1408000001 US1433000080 US1438000133 US1511000575 US1511000609 US1511000620 US1525000947 US1729000625 US1623002146 US1547001360 US1443000187 US1716000341 US1239000015 US1537001261 184 US1224000033 US1234000069 US1234000087 US1343000021 US1417000017 US1419000018 US1502000416 US1511000608 US1537001248 US1731000705 US1231000029 US123100080 US1234000005 US1304000005 US1319000008 US1319000083 US1431000012 US1502000430 US1529001011 US1609001728 US1517000692 US1343000081 US1234000052 436 US1251000054 817 1068 128 1329 1297 356 US1325000023 US1505000504 US1224000013 US1518000745 249 US1234000082 US1239000017 US1443000210 US1533001102 US1619002102 US1611001862 US1239000048 US1319000072 US1234000079 US1529000994 US1244000044 US1523000840 US1234000063 US1251000051 US1304000006 US1412000010 US1424000051 US1431000004 US1443000170 US1450000308 US1505000526 US1505000547 US1511000600 US1518000779 US1523000891 US1533001117 US1547001386 US1611001911 US1644002690 US1738002804 US1231000065 US1518000761 1406 1410 US1438000102 US1251000097 US1319000012 US1203010411 US1203010250 US1203010254 US1203010255 US1203010259 US1203010263 US1203010268 US1203010269 US1203010273 US1203010274 US1203010275 US1203010277 US1203010282 US1203010292 US1203010299 US1203010301 US1203010304 US1203010312 US1203010318 US1203010343 US1203010365 US1203010369 US1203010372 US1203010399 US1203010406 US1203010409 US1203010415 US1203010416 US1203010419 US1203010424 US1210000003 US1210000004 US1224000005 US1224000028 US1224000031 US1231000001 US1231000011 US1231000021 US1231000039 US1231000041 US1231000043 US1231000054 US1231000056 US1231000068 US1231000072 US1231000077 US1231000087 US1231000088 US1231000089 US1234000010 US1234000015 US1234000022 US1234000023 US1234000026 US1234000027 US1234000029 US1234000034 US1234000038 US1234000044 US1234000054 US1234000058 US1234000065 US1234000074 US1234000075 US1234000076 US1234000081 US1234000085 US1239000003 US1244000030 US1244000045 US1244000047 US1244000051 US1244000059 US1244000063 US1248000003 US1248000069 US1248000076 US1248000077 US1251000001 US1251000004 US1251000009 US1251000032 US1251000034 US1251000036 US1251000042 US1251000056 US1251000064 US1251000094 US1304000001 US1304000003 US1304000015 US1304000022 US1304000028 US1304000031 US1304000032 US1304000037 US1304000038 US1311000001 US1311000015 US1311000019 US1311000031 US1311000037 US1311000040 US1311000041 US1311000046 US1311000052 US1311000055 US1311000059 US1311000060 US1311000061 US1311000063 US1311000064 US1311000066 US1311000068 US1311000070 US1311000073 US1311000074 US1311000075 US1311000078 US1311000079 US1319000034 US1319000044 US1319000046 US1319000048 US1319000050 US1319000053 US1319000056 US1319000062 US1319000068 US1319000073 US1319000076 US1319000081 US1319000088 US1319000090 US1319000092 US1319000093 US1319000096 US1319000099 US1319000100 US1328000010 US1328000011 US1331000005 US1331000027 US1331000035 US1331000036 US1331000039 US1331000045 US1331000048 US1331000055 US1331000058 US1331000066 US1334000003 US1334000005 US1334000006 US1334000012 US1334000014 US1334000015 US1334000018 US1334000023 US1334000029 US1334000034 US1334000046 US1334000048 US1337000002 US1337000004 US1337000010 US1337000027 US1337000031 US1337000037 US1337000044 US1337000046 US1337000058 US1337000067 US1343000007 US1343000010 US1343000014 US1343000015 US1343000017 US1343000018 US1343000023 US1343000029 US1343000030 US1343000032 US1343000035 US1343000036 US1343000037 US1343000042 US1343000043 US1343000045 US1343000046 US1343000047 US1343000048 US1343000052 US1343000053 US1343000055 US1343000062 US1343000063 US1343000064 US1343000074 US1343000076 US1343000080 US1343000082 US1352000049 US1407000001 US1407000005 US1407000008 US1407000011 US1407000017 US1407000018 US1407000025 US1407000028 US1407000032 US1407000033 US1407000035 US1407000043 US1407000044 US1407000045 US1407000061 US1407000062 US1408000004 US1408000009 US1408000012 US1408000017 US1408000019 US1408000029 US1408000038 US1408000039 US1408000040 US1408000042 US1408000043 US1408000045 US1408000046 US1408000047 US1412000001 US1412000002 US1412000004 US1412000007 US1412000014 US1412000015 US1412000017 US1412000018 US1412000019 US1412000022 US1412000023 US1412000024 US1412000029 US1412000031 US1412000034 US1412000035 US1412000039 US1412000040 US1412000047 US1412000050 US1417000001 US1417000033 US1417000039 US1419000003 US1419000008 US1419000013 US1419000016 US1419000022 US1419000034 US1419000045 US1419000048 US1419000054 US1419000057 US1419000060 US1424000001 US1424000006 US1424000011 US1424000016 US1424000019 US1424000026 US1424000028 US1424000029 US1424000030 US1424000031 US1424000034 US1424000038 US1424000045 US1424000053 US1424000055 US1431000002 US1431000007 US1431000008 US1431000009 US1431000010 US1431000015 US1431000018 US1431000019 US1431000021 US1431000022 US1431000030 US1431000031 US1431000032 US1431000033 US1431000036 US1431000039 US1433000048 US1433000051 US1433000055 US1433000056 US1433000057 US1433000058 US1433000060 US1433000061 US1433000063 US1433000069 US1433000071 US1433000073 US1433000075 US1433000077 US1433000078 US1433000082 US1433000084 US1433000089 US1433000090 US1438000092 US1438000093 US1438000094 US1438000095 US1438000096 US1438000097 US1438000098 US1438000099 US1438000101 US1438000104 US1438000109 US1438000112 US1438000114 US1438000117 US1438000123 US1438000124 US1438000127 US1438000130 US1438000131 US1438000137 US1438000138 US1438000147 US1443000175 US1443000189 US1443000193 US1443000231 US1443000232 US1443000234 US1446000237 US1446000239 US1446000244 US1446000250 US1446000251 US1446000252 US1446000253 US1446000258 US1446000268 US1446000270 US1446000271 US1446000272 US1446000277 US1446000280 US1446000281 US1446000283 US1446000288 US1446000289 US1446000290 US1446000293 US1446000296 US1446000297 US1446000298 US1446000300 US1446000302 US1446000303 US1446000304 US1446000306 US1446000307 US1450000309 US1450000311 US1450000314 US1450000317 US1450000319 US1450000321 US1450000324 US1450000328 US1450000329 US1450000331 US1450000332 US1450000333 US1450000334 US1450000335 US1450000337 US1450000339 US1450000340 US1450000341 US1450000343 US1450000348 US1450000350 US1450000352 US1450000353 US1450000356 US1450000358 US1450000359 US1450000360 US1450000363 US1450000365 US1450000367 US1450000368 US1450000370 US1450000371 US1450000372 US1450000374 US1450000377 US1502000378 US1502000382 US1502000383 US1502000385 US1502000386 US1502000393 US1502000394 US1502000398 US1502000403 US1502000404 US1502000406 US1502000407 US1502000408 US1502000410 US1502000413 US1502000415 US1502000418 US1502000421 US1502000425 US1502000438 US1502000456 US1502000457 US1502000458 US1502000461 US1502000468 US1502000469 US1502000476 US1505000495 US1505000539 US1511000594 US1511000601 US1511000606 US1511000625 US1511000630 US1511000632 US1511000636 US1511000641 US1517000653 US1517000664 US1517000718 US1517000720 US1517000721 US1517000722 US1517000723 US1517000728 US1518000748 US1518000749 US1518000750 US1518000754 US1518000757 US1518000760 US1518000767 US1518000769 US1518000799 US1523000808 US1523000814 US1523000818 US1523000819 US1523000821 US1523000824 US1523000826 US1523000829 US1523000830 US1523000831 US1523000832 US1523000834 US1523000835 US1523000843 US1523000845 US1523000846 US1523000847 US1523000850 US1523000851 US1523000852 US1523000853 US1523000856 US1523000857 US1523000860 US1523000865 US1523000868 US1523000871 US1523000874 US1523000876 US1523000877 US1523000878 US1523000880 US1523000899 US1525000913 US1525000918 US1525000928 US1525000932 US1525000937 US1525000948 US1525000951 US1525000953 US1525000957 US1525000965 US1525000967 US1525000969 US1525000971 US1525000977 US1525000980 US1525000981 US1525000982 US1525000983 US1529000996 US1529001002 US1529001003 US1529001007 US1529001009 US1529001015 US1529001018 US1529001020 US1529001022 US1529001023 US1529001026 US1529001030 US1529001034 US1529001035 US1529001039 US1529001040 US1529001041 US1529001043 US1529001060 US1529001061 US1529001062 US1529001065 US1529001066 US1529001069 US1529001070 US1529001071 US1529001074 US1529001091 US1533001101 US1533001105 US1533001107 US1533001108 US1533001109 US1533001112 US1533001113 US1533001120 US1533001122 US1533001123 US1533001131 US1533001135 US1533001138 US1533001141 US1533001142 US1533001143 US1533001145 US1533001149 US1534001152 US1534001154 US1534001156 US1534001161 US1534001176 US1534001178 US1534001191 US1534001197 US1534001198 US1534001200 US1534001206 US1534001207 US1534001219 US1534001222 US1534001225 US1534001234 US1534001239 US1534001244 US1537001249 US1537001250 US1537001251 US1537001252 US1537001253 US1537001254 US1537001255 US1537001256 US1537001257 US1537001262 US1537001264 US1537001267 US1537001269 US1537001270 US1537001271 US1537001274 US1537001278 US1537001285 US1537001290 US1537001291 US1537001292 US1537001304 US1537001313 US1537001319 US1537001322 US1537001326 US1537001327 US1537001328 US1537001329 US1537001332 US1537001333 US1537001334 US1537001335 US1537001336 US1537001337 US1537001339 US1546001354 US1547001359 US1547001361 US1547001363 US1547001375 US1547001382 US1547001418 US1547001422 US1547001424 US1547001425 US1547001428 US1547001434 US1547001436 US1550001444 US1550001452 US1550001456 US1550001458 US1550001459 US1550001461 US1550001481 US1550001490 US1550001492 US1550001493 US1550001496 US1550001497 US1550001499 US1603001513 US1603001517 US1603001523 US1603001524 US1603001534 US1603001537 US1603001540 US1603001542 US1603001543 US1603001544 US1603001545 US1603001546 US1603001548 US1603001553 US1603001554 US1603001558 US1603001559 US1603001560 US1603001561 US1603001562 US1603001563 US1603001564 US1604001567 US1604001577 US1604001579 US1604001589 US1608001629 US1608001660 US1608001667 US1608001668 US1608001674 US1608001681 US1609001686 US1609001694 US1609001711 US1609001713 US1609001720 US1609001721 US1609001736 US1609001740 US1609001741 US1609001742 US1609001744 US1610001746 US1610001751 US1610001752 US1610001754 US1610001756 US1610001758 US1610001759 US1610001763 US1610001765 US1610001766 US1610001767 US1610001770 US1610001777 US1610001784 US1610001790 US1610001795 US1610001799 US1610001800 US1610001802 US1611001810 US1611001811 US1611001812 US1611001813 US1611001818 US1611001822 US1611001823 US1611001824 US1611001825 US1611001827 US1611001830 US1611001833 US1611001836 US1611001837 US1611001840 US1611001842 US1611001843 US1611001849 US1611001851 US1611001853 US1611001857 US1611001859 US1611001861 US1611001864 US1611001865 US1611001866 US1611001873 US1611001874 US1611001879 US1611001881 US1611001884 US1611001887 US1611001894 US1611001898 US1611001903 US1611001905 US1611001907 US1611001912 US1611001915 US1611001921 US1613001939 US1613001941 US1613001942 US1613001946 US1617002031 US1619002046 US1619002052 US1619002053 US1619002056 US1619002063 US1619002065 US1619002066 US1619002069 US1619002072 US1619002073 US1619002074 US1619002075 US1619002086 US1619002087 US1619002088 US1619002089 US1619002091 US1619002092 US1619002093 US1619002094 US1619002095 US1619002096 US1619002100 US1619002101 US1623002106 US1623002111 US1623002130 US1623002135 US1623002138 US1623002142 US1623002145 US1623002164 US1623002169 US1623002174 US1623002177 US1623002185 US1623002187 US1623002188 US1623002190 US1623002195 US1624002200 US1624002204 US1624002215 US1624002220 US1625002229 US1625002235 US1625002248 US1625002252 US1625002254 US1625002258 US1625002265 US1625002270 US1625002280 US1625002283 US1625002285 US1625002288 US1625002296 US1625002299 US1625002300 US1625002301 US1625002303 US1625002304 US1625002305 US1625002306 US1625002308 US1625002309 US1625002310 US1625002313 US1625002314 US1625002316 US1625002319 US1625002321 US1625002323 US1625002325 US1625002326 US1625002327 US1625002328 US1625002329 US1625002331 US1625002332 US1625002334 US1625002335 US1625002336 US1625002337 US1625002339 US1625002341 US1625002342 US1625002343 US1626002395 US1626002400 US1626002414 US1626002428 US1626002437 US1626002440 US1626002441 US1626002442 US1626002445 US1626002448 US1626002449 US1626002450 US1626002452 US1626002453 US1626002456 US1626002467 US1638002482 US1638002484 US1638002487 US1638002496 US1638002503 US1638002509 US1638002512 US1638002513 US1638002516 US1638002519 US1638002520 US1638002522 US1638002526 US1638002530 US1638002531 US1638002532 US1638002535 US1638002536 US1638002538 US1638002540 US1638002541 US1638002542 US1638002544 US1638002546 US1638002552 US1638002553 US1638002557 US1638002564 US1638002605 US1638002608 US1638002609 US1638002611 US1638002613 US1638002621 US1644002637 US1644002639 US1644002640 US1644002645 US1644002646 US1644002652 US1644002654 US1644002656 US1644002657 US1644002659 US1644002660 US1644002661 US1644002665 US1644002667 US1644002671 US1644002672 US1644002673 US1644002676 US1644002677 US1644002686 US1644002687 US1644002688 US1644002692 US1644002694 US1644002700 US1652000018 US1652000022 US1652000046 US1652000050 US1652000097 US1708000104 US1708000106 US1708000121 US1708000123 US1708000124 US1708000125 US1708000128 US1708000129 US1708000130 US1708000131 US1708000136 US1708000137 US1708000139 US1708000141 US1708000143 US1708000147 US1708000148 US1708000151 US1708000153 US1708000155 US1708000156 US1708000157 US1708000158 US1708000160 US1708000162 US1708000165 US1708000166 US1708000167 US1708000169 US1708000172 US1708000176 US1708000179 US1708000180 US1708000182 US1708000185 US1708000187 US1708000188 US1708000189 US1708000191 US1708000192 US1708000195 US1708000196 US1708000197 US1708000199 US1708000201 US1709000204 US1709000208 US1709000211 US1709000215 US1709000227 US1709000229 US1709000242 US1709000243 US1709000246 US1709000247 US1709000257 US1709000262 US1709000267 US1709000270 US1709000274 US1716000291 US1716000292 US1716000298 US1716000302 US1716000314 US1716000317 US1716000318 US1716000319 US1716000320 US1716000325 US1716000328 US1716000330 US1716000331 US1716000334 US1716000336 US1716000337 US1716000338 US1720000359 US1720000375 US1720000379 US1720000380 US1720000381 US1720000383 US1720000385 US1720000388 US1720000390 US1720000392 US1720000393 US1720000395 US1720000396 US1720000398 US1720000399 US1720000401 US1720000402 US1720000403 US1720000404 US1720000406 US1720000408 US1720000411 US1720000413 US1720000414 US1720000415 US1720000418 US1720000420 US1720000426 US1720000438 US1720000440 US1720000443 US1720000445 US1720000446 US1720000449 US1720000450 US1720000451 US1720000452 US1720000454 US1720000460 US1720000461 US1724000465 US1724000467 US1724000468 US1724000469 US1724000471 US1724000472 US1724000476 US1724000477 US1724000479 US1724000483 US1724000504 US1724000527 US1724000532 US1724000536 US1724000556 US1724000557 US1724000560 US1729000562 US1729000563 US1729000566 US1729000581 US1729000582 US1729000583 US1729000587 US1729000593 US1729000594 US1729000597 US1729000598 US1729000602 US1729000610 US1729000620 US1729000621 US1729000627 US1729000633 US1729000635 US1729000648 US1729000653 US1729000654 US1729000658 US1729000659 US1729000660 US1729000661 US1731000663 US1731000667 US1731000668 US1731000669 US1731000670 US1731000672 US1731000674 US1731000678 US1731000679 US1731000680 US1731000681 US1731000683 US1731000684 US1731000686 US1731000687 US1731000688 US1731000689 US1731000690 US1731000692 US1731000693 US1731000694 US1731000698 US1731000699 US1731000700 US1731000701 US1731000704 US1731000706 US1731000707 US1731000709 US1731000713 US1731000716 US1731000717 US1731000718 US1731000719 US1731000721 US1731000723 US1731000725 US1731000726 US1731000730 US1731000731 US1731000735 US1731000736 US1731000738 US1731000739 US1731000741 US1731000743 US1731000748 US1731000749 US1731000751 US1731000755 US1731000757 US1731000758 US1731000759 US1731000760 US1734000765 US1734000766 US1734000767 US1734000770 US1734000773 US1734000774 US1734000779 US1734000784 US1734000787 US1734000788 US1734000793 US1734000794 US1734000795 US1734000797 US1734000799 US1734000800 US1734000802 US1734000804 US1734000805 US1734000809 US1734000810 US1734000811 US1734000812 US1734000813 US1734000815 US1734000818 US1734000819 US1734000820 US1734000821 US1734000823 US1734000827 US1734000829 US1734000831 US1734000832 US1734000851 US1734000856 US1734000860 US1734000861 US1734000865 US1734000868 US1734000869 US1734000872 US1734000873 US1734000875 US1734000885 US1734000886 US1734000887 US1734000889 US1734000890 US1734000896 US1734000897 US1734000903 US1734000904 US1734000905 US1734000910 US1736000916 US1736000924 US1736000927 US1736000929 US1736000934 US1736000935 US1736000937 US1736000938 US1736000940 US1736000941 US1736000945 US1736000948 US1736000949 US1736000951 US1736000953 US1736000955 US1736000957 US1736000958 US1736000962 US1736000963 US1736000968 US1736000969 US1736000970 US1736000972 US1736000975 US1736000976 US1736000978 US1736000980 US1736000981 US1736000983 US1736000984 US1736000986 US1736000987 US1736000988 US1736000990 US1736000991 US1736000994 US1736000995 US1736000996 US1736000999 US1736001001 US1736001005 US1736001007 US1736001008 US1736001011 US1738001012 US1738001013 US1738001014 US1738001015 US1738001016 US1738001017 US1738001019 US1738001021 US1738001022 US1738001025 US1738001027 US1738001028 US1738001030 US1738001033 US1738001035 US1738001037 US1738001038 US1738001039 US1738001041 US1738001042 US1738001047 US1738001048 US1738001053 US1738001054 US1738001056 US1738001061 US1738001064 US1738001065 US1738001069 US1738001070 US1738001071 US1738001073 US1738001075 US1738001078 US1738001080 US1738001081 US1738001082 US1738001084 US1738001088 US1738001089 US1738001093 US1738001096 US1738001097 US1738001098 US1738001099 US1738001101 US1738001103 US1738001107 US1738001108 US1738001109 US1738001110 US1738001111 US1738001112 US1738001113 US1738001114 US1738001115 US1738001117 US1738001119 US1738001121 US1738001122 US1738001123 US1738001127 US1738001128 US1738001131 US1738001132 US1744001139 US1746001146 US1746001150 US1746001158 US1746001159 US1746001160 US1746001163 US1746001164 US1746001167 US1746001173 US1746001175 US1746001176 US1746001182 US1746001191 US1746001194 US1746001198 US1746001201 US1746001221 US1746001226 US1746001236 US1751001292 US1751001293 US1751001294 US1751001296 US1751001298 US1751001299 US1751001302 US1751001303 US1751001304 US1751001306 US1751001307 US1751001309 US1751001310 US1751001311 US1751001319 US1751001322 US1751001325 US1751001328 US1751001331 US1751001333 US1751001334 US1751001336 US1751001337 US1751001339 US1751001340 US1751001341 US1751001344 US1751001346 US1751001347 US1751001351 US1751001352 US1751001353 US1751001354 US1751001355 US1751001357 US1751001359 US1751001360 US1751001363 US1751001365 US1751001370 US1751001375 US1751001377 US1751001381 US1751001382 US1751001386 US1751001387 US1751001388 US1751001389 US1810001396 US1810001398 US1810001401 US1810001402 US1810001403 US1810001405 US1810001408 US1810001411 US1810001414 US1810001415 US1810001420 US1810001424 US1810001429 US1810001434 US1810001436 US1810001437 US1810001438 US1810001443 US1810001447 US1810001450 US1810001452 US1810001456 US1810001457 US1810001458 US1810001465 US1810001466 US1810001467 US1810001468 US1810001469 US1810001470 US1810001472 US1810001474 US1810001475 US1810001476 US1810001478 US1810001479 US1810001480 US1810001481 US1810001482 US1810001483 US1810001485 US1810001486 US1810001488 US1810001494 US1810001499 US1810001500 US1810001501 US1810001505 US1810001507 US1810001513 US1810001521 US1810001526 US1810001533 US1814000001 US1814000002 US1814000003 US1814000004 US1814000006 US1814000007 US1814000008 US1814000011 US1814000012 US1814000013 US1814000014 US1814000016 US1814000017 US1814000018 US1814000019 US1814000020 US1814000022 US1814000023 US1814000024 US1814000027 US1814000028 US1814000029 US1814000030 US1814000031 US1814000032 US1814000033 US1814000034 US1814000035 US1814000036 US1814000037 US1814000038 US1814000039 US1814000040 US1814000041 US1814000043 US1814000044 US1814000045 US1814000046 US1814000047 US1814000048 US1814000049 US1814000050 US1814000051 US1814000052 US1814000053 US1814000054 US1814000055 US1814000056 US1814000058 US1814000059 US1814000060 US1814000061 US1814000062 US1814000063 US1814000064 US1814000065 US1814000066 US1814000067 US1814000068 US1814000069 US1814000070 US1814000072 US1814000074 US1814000078 US1814000081 US1814000082 US1814000083 US1814000084 US1814000085 US1814000086 US1814000087 US1814000088 US1814000089 US1814000092 US1814000094 US1814000096 US1823001536 US1823001541 US1823001559 US1823001566 US1450000310 US1352000004 US1523000841 US1248000001 US1108010058 US1203010256 US1203010295 US1203010320 US1334000008 US1352000066 US1523000822 US1523000886 US1525000989 US1537001317 US1625002269 US1644002650 US1652000002 US1109010093 US1109010108 US1202010169 US1534001155 US1534001172 US1534001236 US1202010165 US1202010216 US1203010276 US1203010349 US1611001841 US1407000010 278 US1505000483 US1533001139 US1239000040 776 US1319000037 US1304000013 US1109010101 US1424000015 US1731000673 US1511000595 US1352000014 US1324000049 US1744001137 US1709000261 US1224000029 US1244000013 US1352000008 US1537001320 US1505000505 US1407000015 US1603001552 1038 US1446000301 US1337000053 US1537001289 US1109010064 US1505000512 US1324000009 US1518000807 US1523000833 US1331000041 US1202010229 US1625002307 US1231000048 US1511000583 US1424000012 77 US1623002184 42 328 US1319000049 943 US1608001680 394 638 US1450000312 8 US1746001240 US1611001889 US1511000621 US1652000043 US1716000307 US1534001169 US1352000019 US1720000369 US1720000457 US1319000067 US1446000246 US1517000662 US1419000006 US1324000043 US1547001383 US1547001431 US1738002773 US1203010338 US1248000083 US1352000064 US1709000260 US1331000059 US1603001510 US1729000647 US1511000576 US1729000564 US1611001890 US1244000024 US1625002289 US1626002392 US1511000619 US1419000051 US1203010403 US1523000895 US1244000033 US1529000995 US1244000011 US1638002620 US1337000070 US1511000616 US1604001591 US1203010359 US1623002139 889 US1202010183 US1107010036 785 US1319000080 US1319000065 US1604001568 US1234000053 US1523000892 US1443000185 US1234000040 US1625002295 US1337000074 US1203010402 US1529001025 US1603001556 US1202010247 US1251000074 US1319000017 US1441000155 US1248000011 US1407000052 US1224000045 US1311000072 US1419000010 US1517000688 US1604001611 US1609001689 US1610001761 US1625002291 US1734000762 US1736000992 US1419000007 US1319000058 US1248000005 US1716000305 US1311000071 US1244000038 US1611001904 536 81 626 20 362 US1547001369 US1724000520 US1746001277 US1523000844 US1517000652 US1624002218 764 US1446000236 US1523000823 US1529001038 US1626002462 US1529000993 US1626002387 US1319000089 US1731000715 US1319000022 US1203010251 US1203010279 US1203010280 US1203010308 US1203010330 US1203010334 US1203010346 US1203010352 US1203010358 US1203010360 US1203010366 US1203010371 US1203010373 US1203010375 US1203010382 US1203010398 US1203010400 US1203010407 US1203010413 US1239000027 US1736000952 US1734000898 US1734000907 US1734000909 US1224000050 US1304000043 US1304000046 US1311000007 US1311000009 US1311000017 US1550001470 US1231000061 US1239000009 US1251000063 US1412000043 US1738001126 US1610001776 US1407000013 US1611001885 US1611001923 US1613001945 US1108010053 US1109010070 US1109010075 US1311000042 US1331000013 US1343000034 US1505000507 US1529001016 US1751001338 US1109010072 US1201010130 US1202010217 US1234000060 US1337000033 US1337000038 US1343000024 US1352000030 US1334000021 US1417000027 US1234000066 US1107010006 US1107010016 US1107010018 US1107010019 US1107010022 US1107010031 US1108010062 US1109010090 US1109010104 US1109010106 US1202010153 US1202010164 US1202010168 US1202010171 US1202010173 US1202010242 US1202010243 US1203010261 US1203010267 US1203010289 US1203010306 US1203010307 US1203010340 US1203010345 US1203010377 US1203010391 US1203010421 US1224000016 US1224000048 US1231000005 US1231000013 US1231000018 US1231000022 US1231000038 US1231000070 US1231000085 US1234000013 US1234000032 US1234000049 US1234000073 US1234000077 US1234000084 US1234000086 US1239000004 US1239000013 US1239000014 US1239000018 US1239000024 US1239000030 US1239000036 US1239000037 US1239000042 US1239000045 US1239000049 US1239000051 US1239000052 US1244000002 US1244000009 US1244000016 US1244000017 US1244000023 US1244000029 US1244000032 US1244000034 US1244000037 US1244000040 US1244000043 US1244000055 US1248000012 US1248000021 US1248000023 US1248000044 US1248000048 US1248000071 US1248000080 US1248000082 US1251000005 US1251000008 US1251000013 US1251000018 US1251000024 US1251000026 US1251000033 US1251000052 US1251000065 US1251000076 US1251000077 US1251000079 US1251000082 US1251000087 US1251000088 US1304000009 US1304000012 US1304000040 US1311000045 US1311000065 US1311000069 US1319000005 US1319000009 US1319000016 US1319000030 US1319000035 US1319000059 US1319000064 US1319000077 US1319000091 US1324000007 US1324000008 US1324000014 US1324000015 US1324000016 US1324000031 US1324000037 US1324000039 US1324000041 US1325000006 US1325000022 US1325000028 US1328000002 US1328000009 US1328000015 US1331000026 US1331000043 US1334000011 US1334000013 US1334000017 US1334000020 US1334000031 US1337000001 US1337000005 US1337000013 US1337000016 US1337000017 US1337000035 US1337000039 US1337000041 US1337000045 US1337000055 US1337000056 US1337000069 US1337000084 US1343000003 US1343000027 US1343000038 US1343000040 US1343000054 US1343000059 US1343000068 US1343000072 US1343000073 US1343000079 US1352000005 US1352000016 US1352000031 US1352000034 US1352000037 US1352000043 US1352000061 US1407000007 US1407000014 US1407000024 US1407000037 US1408000003 US1408000006 US1408000014 US1408000033 US1408000034 US1412000012 US1412000030 US1412000044 US1417000005 US1417000011 US1417000013 US1417000032 US1419000015 US1419000043 US1419000046 US1424000004 US1424000007 US1431000016 US1431000034 US1433000046 US1433000052 US1438000091 US1438000135 US1441000154 US1443000203 US1443000215 US1443000216 US1446000292 US1450000347 US1502000384 US1502000405 US1502000422 US1502000443 US1502000450 US1502000463 US1502000465 US1502000467 US1505000485 US1505000530 US1505000538 US1505000555 US1511000580 US1511000602 US1511000603 US1511000607 US1511000610 US1511000611 US1511000614 US1511000626 US1511000643 US1511000646 US1517000658 US1517000668 US1517000670 US1517000684 US1517000724 US1517000739 US1518000747 US1518000752 US1518000756 US1518000764 US1518000768 US1518000777 US1518000785 US1518000796 US1523000870 US1523000890 US1525000902 US1525000923 US1525000955 US1529001028 US1529001045 US1529001053 US1529001056 US1529001075 US1529001081 US1533001100 US1533001106 US1533001124 US1534001202 US1534001205 US1534001224 US1537001273 US1537001286 US1537001296 US1537001307 US1546001352 US1547001357 US1547001385 US1547001390 US1547001404 US1547001408 US1547001430 US1550001475 US1550001477 US1603001521 US1603001547 US1603001565 US1604001578 US1604001592 US1604001608 US1608001627 US1608001633 US1608001659 US1609001690 US1609001705 US1609001723 US1610001755 US1610001762 US1610001768 US1610001778 US1611001854 US1611001855 US1613001968 US1617002011 US1623002134 US1623002136 US1623002165 US1623002178 US1623002181 US1624002201 US1624002203 US1625002238 US1625002261 US1625002290 US1625002317 US1625002340 US1626002439 US1638002571 US1644002647 US1652000034 US1652000086 US1724000500 US1724000543 US1729000607 US1729000630 US1729000631 US1729000652 US1731000696 US1731000702 US1731000728 US1734000778 US1734000790 US1734000796 US1734000801 US1734000876 US1736000933 US1736000965 US1736000982 US1736001003 US1736001009 US1738001062 US1738001095 US1738001120 US1738002746 US1738002772 US1746001153 US1746001211 US1746001213 US1746001217 US1746001237 US1746001248 US1746001252 US1751001373 US1810001417 US1810001461 US1810001492 US1603001535 US1638002555 US1652000040 US1239000002 US1511000645 US1517000655 US1734000850 US1734000899 US1734000908 US1736000919 US1736000925 US1738001020 US1352000046 US1107010005 US1324000011 US1352000006 US1352000022 US1352000044 US1352000047 US1352000048 US1408000035 US1518000789 US1738002783 US1638002590 US1638002596 US1319000038 US1239000019 US1709000217 US1525000939 US1547001421 US1547001426 US1450000345 US151 |  | SoftwareDesignChange
102 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167437 | Z-0462-2019 | Xper Flex Cardio Physiomonitoring System, Software version 1.6.0.0388 | Xper Flex Cardio Physiomonitoring System | 2 | CDRH | 2018-11-14 |  |  | Invivo Corporation | The Invasive Blood Pressure (IBP) numeric values will freeze on the display when any of the following functions are performed: change an IBP Site Label using the Site Label icon or the keyboard, activate or deactivate the mean pressure waveform display option (On/Off) using the Invasive Pressure Options Dialog or the keyboard, activate or deactivate the waveform hide option (On/Off) using the Invasive Pressure Options Dialog, activate or deactivate the pressure waveform filter option (On/Off) using the keyboard, and/or perform a Pullback using the Pullback icon or the keyboard*** | Class 2 Device Recall Xper Flex Cardio Physiomonitoring System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Xper Flex Cardio Physiomonitoring System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Xper Flex Cardio Physiomonitoring System Date Initiated by Firm September 24, 2018 Create Date November 14, 2018 Recall Status 1 Completed Recall Number Z-0462-2019 Recall Event ID 81092 510(K)Number K101571 Product Classification Computer, diagnostic, programmable - Product Code DQK Product Xper Flex Cardio Physiomonitoring System, Software version 1.6.0.0388 Code Information Model Number 453564669081; System Serial Number (UDI): US1814000001 ((01)00884838091474(21)US1814000001), US1814000002 ((01)00884838091474(21)US1814000002), US1814000003 ((01)00884838091474(21)US1814000003), US1814000004 ((01)00884838091474(21)US1814000004), US1814000006 ((01)00884838091474(21)US1814000006), US1814000007 ((01)00884838091474(21)US1814000007), US1814000008 ((01)00884838091474(21)US1814000008), US1814000011 ((01)00884838091474(21)US1814000011), US1814000012 ((01)00884838091474(21)US1814000012), US1814000013 ((01)00884838091474(21)US1814000013), US1814000014 ((01)00884838091474(21)US1814000014), US1814000016 ((01)00884838091474(21)US1814000016), US1814000017 ((01)00884838091474(21)US1814000017), US1814000018 ((01)00884838091474(21)US1814000018), US1814000019 ((01)00884838091474(21)US1814000019), US1814000020 ((01)00884838091474(21)US1814000020), US1814000022 ((01)00884838091474(21)US1814000022), US1814000023 ((01)00884838091474(21)US1814000023), US1814000024 ((01)00884838091474(21)US1814000024), US1814000027 ((01)00884838091474(21)US1814000027), US1814000028 ((01)00884838091474(21)US1814000028), US1814000029 ((01)00884838091474(21)US1814000029), US1814000030 ((01)00884838091474(21)US1814000030), US1814000031 ((01)00884838091474(21)US1814000031), US1814000032 ((01)00884838091474(21)US1814000032), US1814000033 ((01)00884838091474(21)US1814000033), US1814000034 ((01)00884838091474(21)US1814000034), US1814000035 ((01)00884838091474(21)US1814000035), US1814000036 ((01)00884838091474(21)US1814000036), US1814000037 ((01)00884838091474(21)US1814000037), US1814000038 ((01)00884838091474(21)US1814000038), US1814000039 ((01)00884838091474(21)US1814000039), US1814000040 ((01)00884838091474(21)US1814000040), US1814000041 ((01)00884838091474(21)US1814000041), US1814000043 ((01)00884838091474(21)US1814000043), US1814000044 ((01)00884838091474(21)US1814000044), US1814000045 ((01)00884838091474(21)US1814000045), US1814000046 ((01)00884838091474(21)US1814000046), US1814000047 ((01)00884838091474(21)US1814000047), US1814000048 ((01)00884838091474(21)US1814000048), US1814000049 ((01)00884838091474(21)US1814000049), US1814000050 ((01)00884838091474(21)US1814000050), US1814000051 ((01)00884838091474(21)US1814000051), US1814000052 ((01)00884838091474(21)US1814000052), US1814000053 ((01)00884838091474(21)US1814000053), US1814000054 ((01)00884838091474(21)US1814000054), US1814000055 ((01)00884838091474(21)US1814000055), US1814000056 ((01)00884838091474(21)US1814000056), US1814000058 ((01)00884838091474(21)US1814000058), US1814000059 ((01)00884838091474(21)US1814000059), US1814000060 ((01)00884838091474(21)US1814000060), US1814000061 ((01)00884838091474(21)US1814000061), US1814000062 ((01)00884838091474(21)US1814000062), US1814000063 ((01)00884838091474(21)US1814000063), US1814000064 ((01)00884838091474(21)US1814000064), US1814000065 ((01)00884838091474(21)US1814000065), US1814000066 ((01)00884838091474(21)US1814000066), US1814000067 ((01)00884838091474(21)US1814000067), US1814000068 ((01)00884838091474(21)US1814000068), US1814000069 ((01)00884838091474(21)US1814000069), US1814000070 ((01)00884838091474(21)US1814000070), US1814000072 ((01)00884838091474(21)US1814000072), US1814000074 ((01)00884838091474(21)US1814000074), US1814000078 ((01)00884838091474(21)US1814000078), US1814000081 ((01)00884838091474(21)US1814000081), US1814000082 ((01)00884838091474(21)US1814000082), US1814000083 ((01)00884838091474(21)US1814000083), US1814000084 ((01)00884838091474(21)US1814000084), US1814000085 ((01)00884838091474(21)US1814000085), US1814000086 ((01)00884838091474(21)US1814000086), US1814000087 ((01)00884838091474(21)US1814000087), US1814000088 ((01)00884838091474(21)US1814000088), US1814000089 ((01)00884838091474(21)US1814000089), US1814000092 ((01)00884838091474(21)US1814000092), US1814000094 ((01)00884838091474(21)US1814000094), US1814000096 ((01)00884838091474(21)US1814000096). Recalling Firm/ Manufacturer Invivo Corporation 12151 Research Pkwy Ste 200 Orlando FL 32826-3286 For Additional Information Contact 800-722-9377 Manufacturer Reason for Recall The Invasive Blood Pressure (IBP) numeric values will freeze on the display when any of the following functions are performed: change an IBP Site Label using the Site Label icon or the keyboard, activate or deactivate the mean pressure waveform display option (On/Off) using the Invasive Pressure Options Dialog or the keyboard, activate or deactivate the waveform hide option (On/Off) using the Invasive Pressure Options Dialog, activate or deactivate the pressure waveform filter option (On/Off) using the keyboard, and/or perform a Pullback using the Pullback icon or the keyboard FDA Determined Cause 2 Software Design Change Action Customers were notified on about 09/20/2018 via an "URGENT - Medical Device Correction" letter. An initial letter was sent to distributors with instructions to provide the Field Safety Notice to all customers the devices were further distributed to. ************The "URGENT - Medical Device Correction" letter provided guidance on how to identify the Service Number and serial number on the Flex Cardio device, which is on the bottom right corner of the back of the device. Actions to be taken by the customer/user until Philips contacts the customer to arrange for a software update to address this issue, is if the IBP numeric values become frozen and do not update in real-time, opening the Waveform Setup Dialog and then immediately pressing the OK button will restore the proper display of the IBP numeric values. ************Alternative methods of performing the functions listed in the Problem Description section that do not cause the problem to occur include the following: ******1. Change an IBP Site Label using the Pressure Setup Dialog or a macro; ******2. Activate or deactivate the mean pressure waveform display option (On/Off) by using the Waveform Setup Dialog, the Pressure Setup Dialog or a macro; ******3. Activate or deactivate the waveform hide option (On/Off) using the Waveform Setup Dialog, the Pressure Setup Dialog, the keyboard or a macro; ******4. Activate or deactivate the pressure waveform filter option (On/Off) using a macro. ************A Philips representative will contact each customer regarding their affected Flex Cardio devices. A software update will be required on both the affected Flex Cardio device and on the host workstation***s driver. These actions will be implemented free of charge by Philips. Quantity in Commerce 76 units Distribution Distributed to the Netherlands Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 510(K) Database 510(K)s with Product Code = DQK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were notified on about 09/20/2018 via an "URGENT - Medical Device Correction" letter. An initial letter was sent to distributors with instructions to provide the Field Safety Notice to all customers the devices were further distributed to. ************The "URGENT - Medical Device Correction" letter provided guidance on how to identify the Service Number and serial number on the Flex Cardio device, which is on the bottom right corner of the back of the device. Actions to be taken by the customer/user until Philips contacts the customer to arrange for a software update to address this issue, is if the IBP numeric values become frozen and do not update in real-time, opening the Waveform Setup Dialog and then immediately pressing the OK button will restore the proper display of the IBP numeric values. ************Alternative methods of performing the functions listed in the Problem Description section that do not cause the problem to occur include the following: ******1. Change an IBP Site Label using the Pressure Setup Dialog or a macro; ******2. Activate or deactivate the mean pressure waveform display option (On/Off) by using the Waveform Setup Dialog, the Pressure Setup Dialog or a macro; ******3. Activate or deactivate the waveform hide option (On/Off) using the Waveform Setup Dialog, the Pressure Setup Dialog, the keyboard or a macro; ******4. Activate or deactivate the pressure waveform filter option (On/Off) using a macro. ************A Philips representative will contact each customer regarding their affected Flex Cardio devices. A software update will be required on both the affected Flex Cardio device and on the host workstation***s driver. These actions will be implemented free of charge by Philips. | SoftwareDesignChange
103 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=166225 | Z-2968-2018 | McKesson Cardiology Hemo, Release SW version 13.0.***Used for physiological monitoring, image and data processing.****** McKesson Cardiology Hemo is intended for complete physiological/hemodynamic monitoring, clinical data acquisition, medical image and data processing, and analytical assessment. McKesson Cardiology Hemo is also intended for patient/procedural data management, such as documentation, logging, reporting, trending, storing, reviewing, carrying out clinical calculations and exporting various representations of the acquired data. | McKesson Cardiology Hemo | 2 | CDRH | 2018-08-31 |  | 2021-11-23 12:57 | McKesson Israel Ltd. | Change Healthcare has identified an issue where, under certain circumstances, the Real Time Monitor (RTM) may not display physiological signals. | Class 2 Device Recall McKesson Cardiology Hemo Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall McKesson Cardiology Hemo FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall McKesson Cardiology Hemo Date Initiated by Firm March 12, 2018 Create Date August 31, 2018 Recall Status 1 Terminated 3 on November 23, 2021 Recall Number Z-2968-2018 Recall Event ID 80723 510(K)Number K131497 Product Classification Computer, diagnostic, programmable - Product Code DQK Product McKesson Cardiology Hemo, Release SW version 13.0.******Used for physiological monitoring, image and data processing.************ McKesson Cardiology Hemo is intended for complete physiological/hemodynamic monitoring, clinical data acquisition, medical image and data processing, and analytical assessment. McKesson Cardiology Hemo is also intended for patient/procedural data management, such as documentation, logging, reporting, trending, storing, reviewing, carrying out clinical calculations and exporting various representations of the acquired data. Code Information software versions: 13.0 HF1, 13.0HF2, 13.0HF3 Recalling Firm/ Manufacturer McKesson Israel Ltd. Or Towers Bldg. B 4, Nehoshet St. Tel Aviv Israel Manufacturer Reason for Recall Change Healthcare has identified an issue where, under certain circumstances, the Real Time Monitor (RTM) may not display physiological signals. FDA Determined Cause 2 Software Design Change Action The firm, Change Healthcare Israel Ltd., sent an "Urgent Field Safety Notice" dated on 03/12/2018 to its customers. The notice explained the product, issue and actions to be taken. The customers were instructed to: distribute the notice to all those within your organization who need to be aware; alert other parties affiliated with your organization, which are affected by this field safety notice;******and maintain awareness of this notice and resulting action until the defect has been corrected, to ensure effectiveness of the corrective action.************Change Healthcare has developed software updates to address the problem.************If you have any questions, contact Change Healthcare Customer Support (US) 1-800-654-4366 or (UK) 440(0) 208-952-799 to arrange for the installation of software update/upgrade. Quantity in Commerce 31 units Distribution US Distribution to states of: CA, CO, CT, FL, GA, KS, IN, LA, NJ, NY, OK, OH, SC, and TX; and internationally to: UK. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = DQK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, Change Healthcare Israel Ltd., sent an "Urgent Field Safety Notice" dated on 03/12/2018 to its customers. The notice explained the product, issue and actions to be taken. The customers were instructed to: distribute the notice to all those within your organization who need to be aware; alert other parties affiliated with your organization, which are affected by this field safety notice;******and maintain awareness of this notice and resulting action until the defect has been corrected, to ensure effectiveness of the corrective action.************Change Healthcare has developed software updates to address the problem.************If you have any questions, contact Change Healthcare Customer Support (US) 1-800-654-4366 or (UK) 440(0) 208-952-799 to arrange for the installation of software update/upgrade. | SoftwareDesignChange
104 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164966 | Z-2392-2018 | Reliance 1227 Cart and Utensil Washer/Disinfector, FW03101, FW03102, FW03S003******Product Usage:***The Reliance 1227 Cart and Utensil Washer/Disinfector is intended for use in the cleaning and low-level disinfection of bedpans and urinals, basins, carts, beds, theatre shoes and other miscellaneous reusable items used in the care of patients.*** | Reliance 1227 Cart and Utensil Washer/Disinfector | 2 | CDRH | 2018-07-09 |  | 2019-11-05 08:18 | Steris Corporation | The firm has become aware that the Reliance 1227 Cart and Utensil Washer/Disinfector s Chemical Low Level alarm, intended to prevent the user from initiating a cycle when a low chemical level situation occurs, does not operate as intended. Currently, if a low chemical level situation occurs, the alarm will only be generated at the unit s next power-up or when accessing service mode. The failure of the Reliance 1227 to identify that the chemicals used for cleaning are low or empty could result in bedpans and urinals, basins, case carts, beds, theater shoes and other miscellaneous reusable patient care items not being properly cleaned before disinfection or reuse. | Class 2 Device Recall Reliance 1227 Cart and Utensil Washer/Disinfector Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Reliance 1227 Cart and Utensil Washer/Disinfector FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Reliance 1227 Cart and Utensil Washer/Disinfector Date Initiated by Firm May 23, 2018 Create Date July 09, 2018 Recall Status 1 Terminated 3 on November 05, 2019 Recall Number Z-2392-2018 Recall Event ID 80255 Product Classification Disinfector, medical devices - Product Code MEC Product Reliance 1227 Cart and Utensil Washer/Disinfector, FW03101, FW03102, FW03S003************Product Usage:******The Reliance 1227 Cart and Utensil Washer/Disinfector is intended for use in the cleaning and low-level disinfection of bedpans and urinals, basins, carts, beds, theatre shoes and other miscellaneous reusable items used in the care of patients. Code Information 3632511020 3634311005 3607312011 3615312001 3611012005 3627012001 3626212001 3633112020 3632612001 3601413013 3629612014 3602113013 3605113001 3615912011 3606613001A 3607413001 3612713001 3614913001 3616213001 3608613001 3623313001 3624113001 3626813001 3627313019 3628113001 3628213002 3629113001 3631613011 3635013015 3601514001 3603614001 3604114024 3604914002 3630113020 3613214010 3611814020 3616714023 3630813004 3628214010 3629414002 3630814002 3631614001 3631414016 3632914001 3632414003 3634914022 3604415042 3605415029 3606115033 3604715027 3608215021 3615215026 3613915001 3619415002 3623015006 3623715001 3627115035 3627315001 3627815020 3629215021 3629515001 3633015001 3632215001 3633615001 3634215001 3632315001 3626415043 3601116014 3602516001 3634315001 3604116001 3604616003 3604716001 3604816001 3606916011 3613716013 3618716001 3600616001 3625616008 3630116001 3632616007 3634216002 3633316017 3635416008 3625716001 3635616001 3603217002 3632116001 3635016001 3604517001 3604617001 3605417001 3606117001 3628616001 3605317001 3605117029 3607417001 3611417015 3612217002 3610917001 3611617001 3613617001 3613717001 3619817008 3619117006 3619917001 3619317001 3617717005 3630317038 3634117001 3601118001 3634517031 3606418014 3606518001 3608718001 3609318016 3610118001 3616814019 Recalling Firm/ Manufacturer Steris Corporation 5960 Heisley Rd Mentor OH 44060-1834 For Additional Information Contact Mrs. Lindsay Orlando 440-392-7601 Manufacturer Reason for Recall The firm has become aware that the Reliance 1227 Cart and Utensil Washer/Disinfector s Chemical Low Level alarm, intended to prevent the user from initiating a cycle when a low chemical level situation occurs, does not operate as intended. Currently, if a low chemical level situation occurs, the alarm will only be generated at the unit s next power-up or when accessing service mode. The failure of the Reliance 1227 to identify that the chemicals used for cleaning are low or empty could result in bedpans and urinals, basins, case carts, beds, theater shoes and other miscellaneous reusable patient care items not being properly cleaned before disinfection or reuse. FDA Determined Cause 2 Software Design Change Action On May 23, 2018, the firm notified affected customers by Urgent Device Field Correction letters distributed via FedEx. ************The letter stated that the system software had been updated to correct the issue, and that a STERIS service technician would be contacting each facility to perform the system software update.************In the interim of the software update, users were instructed to continue to verify the unit***s chemistry levels daily as part of the facility***s Quality Management System. Chemistry containers with low chemistry volume should be replaced prior to initiating the next cycle.************If you have questions regarding this matter, please contact Katilin Slaper, Product Manager, at (440) 392-7431, STERIS Customer Service at 1-800-548-4873, or your local STERIS Representative. Quantity in Commerce 118 Distribution Worldwide Distribution - US Nationwide in the states of AL, AK, CA, CO, CN, DE, FL, GA, IL, IA, KS, KY, LA, MA, MI, NJ, NY, ND, OH, OK, OR, PA, RI, SC, TN, TX, VT, VA, WI, and WV. Foreign distribution to Canada. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On May 23, 2018, the firm notified affected customers by Urgent Device Field Correction letters distributed via FedEx. ************The letter stated that the system software had been updated to correct the issue, and that a STERIS service technician would be contacting each facility to perform the system software update.************In the interim of the software update, users were instructed to continue to verify the unit***s chemistry levels daily as part of the facility***s Quality Management System. Chemistry containers with low chemistry volume should be replaced prior to initiating the next cycle.************If you have questions regarding this matter, please contact Katilin Slaper, Product Manager, at (440) 392-7431, STERIS Customer Service at 1-800-548-4873, or your local STERIS Representative. | SoftwareDesignChange
105 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164139 | Z-1597-2018 | Biograph Horizon - PET/CT, PETsyngo VJ20A Software Nuclear medicine/ xray diagnostic scanner. | Biograph Horizon PET/CT | 2 | CDRH | 2018-05-02 |  | 2018-10-12 | Siemens Medical Solutions USA, Inc. | Error introduced into PET images acquired and reconstructed with VJ20A software. Array values are indexed improperly when the norm file is created during QC. During data reconstruction, incorrect values are being applied. This can lead to a gradient in the image. The severity of the error is directly related to the positioning of the PET QC phantom relative to the center of the field of view. | Class 2 Device Recall Biograph Horizon PET/CT Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Biograph Horizon PET/CT FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Biograph Horizon PET/CT Date Initiated by Firm January 26, 2018 Create Date May 02, 2018 Recall Status 1 Terminated 3 on October 12, 2018 Recall Number Z-1597-2018 Recall Event ID 79988 510(K)Number K170904 Product Classification positron emmission and x-ray computed tomography - Product Code JAK Product Biograph Horizon - PET/CT, PETsyngo VJ20A Software Nuclear medicine/ xray diagnostic scanner. Code Information Software version VJ20A. Code No. 10532746, 10532748. Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc. 2501 Barrington Rd Hoffman Estates IL 60192-2061 For Additional Information Contact 847-304-7700 Manufacturer Reason for Recall Error introduced into PET images acquired and reconstructed with VJ20A software. Array values are indexed improperly when the norm file is created during QC. During data reconstruction, incorrect values are being applied. This can lead to a gradient in the image. The severity of the error is directly related to the positioning of the PET QC phantom relative to the center of the field of view. FDA Determined Cause 2 Software Design Change Action Customers were notified via letter on approximately 01/26/2018. Instructions included to ensure proper positioning of the PET QC Phantom when performing the PET Quality Control Procedure by following section 6.3.1 in the Biograph Horizon Operator's Manual for Examination and Acquisition, complete the software update to PETsyngo VJ20B, ensure the recall notice is placed in the Biograph Horizon Operator's Manual for Examination and Acquisition and disseminated to all operators of the Biograph Horizon. If the equipment has been sold to another customer, they are instructed to provide the recall notice to the new owner. Quantity in Commerce 42 units Distribution Worldwide distribution. US nationwide, Australia, France, India, Japan, and Netherlands. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were notified via letter on approximately 01/26/2018. Instructions included to ensure proper positioning of the PET QC Phantom when performing the PET Quality Control Procedure by following section 6.3.1 in the Biograph Horizon Operator's Manual for Examination and Acquisition, complete the software update to PETsyngo VJ20B, ensure the recall notice is placed in the Biograph Horizon Operator's Manual for Examination and Acquisition and disseminated to all operators of the Biograph Horizon. If the equipment has been sold to another customer, they are instructed to provide the recall notice to the new owner. | SoftwareDesignChange
106 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161733 | Z-0829-2018 | 2008T, Hemodialysis Delivery System******Product Indicated for acute and chronic dialysis therapy. | 2008T Hemodialysis Delivery System | 2 | CDRH | 2018-03-01 | 2018-02-20 21:01 | 2021-07-28 12:53 | Fresenius Medical Care Renal Therapies Group, LLC | While reviewing documentation for the next software release of the 2008T, an R&D technician identified that setting the UF goal to '0' introduces a discrepancy between the UF rate displayed and the actual UF pump rate. The software anomaly is also applicable to the 2008T Machines that contained the 2.63 Bug Fix per DCAF 17-088. 2008T Upgrade kits that contained the 2.63 Bug Fix are also affected. Additionally, SW version 2.64 (in design freeze) is also impacted by this anomaly. | Class 2 Device Recall 2008T Hemodialysis Delivery System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall 2008T Hemodialysis Delivery System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall 2008T Hemodialysis Delivery System Date Initiated by Firm October 05, 2017 Date Posted February 20, 2018 Recall Status 1 Terminated 3 on July 28, 2021 Recall Number Z-0829-2018 Recall Event ID 79239 510(K)Number K080964 K113427 K120017 K120505 K121341 Product Classification Dialyzer, high permeability with or without sealed dialysate system - Product Code KDI Product 2008T, Hemodialysis Delivery System************Product Indicated for acute and chronic dialysis therapy. Code Information 2008T machines: 190766 2008T Hemodialysis System w/Bibag 190895 2008T GEN 2 Bibag without CDX 190766 = 00840861100910 (+Serial Number) 190895 = 00840861100927 (+Serial Number)ux=&Su Recalling Firm/ Manufacturer Fresenius Medical Care Renal Therapies Group, LLC 920 Winter St Waltham MA 02451-1521 For Additional Information Contact Beth St.Germain*** 781-699-2243 Manufacturer Reason for Recall While reviewing documentation for the next software release of the 2008T, an R&D technician identified that setting the UF goal to '0' introduces a discrepancy between the UF rate displayed and the actual UF pump rate. The software anomaly is also applicable to the 2008T Machines that contained the 2.63 Bug Fix per DCAF 17-088. 2008T Upgrade kits that contained the 2.63 Bug Fix are also affected. Additionally, SW version 2.64 (in design freeze) is also impacted by this anomaly. FDA Determined Cause 2 Software Design Change Action Fresenius Medical Care sent an Urgent Medical Device Correction letter dated April 24, 2017 to affected customers. the letter identified the affected product, problem and actions to be taken. For questions contact Technical Services at 1-800-227-2572. Quantity in Commerce 15 machines Distribution US Nationwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = KDI 510(K)s with Product Code = KDI 510(K)s with Product Code = KDI - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Fresenius Medical Care sent an Urgent Medical Device Correction letter dated April 24, 2017 to affected customers. the letter identified the affected product, problem and actions to be taken. For questions contact Technical Services at 1-800-227-2572. | SoftwareDesignChange
107 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161520 | Z-0808-2018 | Syngo.via systems with SW VB20A | Syngo | 2 | CDRH | 2018-03-01 |  | 2018-10-01 | Siemens Medical Solutions USA, Inc | A software update was released on August 22, 2017. Performing this software installation and upgrading syngo.via system to the software version VB20A_HF02 will resolve various bugs and inconsistent diagnosis with the old system. | Class 2 Device Recall Syngo Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Syngo FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Syngo Date Initiated by Firm August 31, 2017 Create Date March 01, 2018 Recall Status 1 Terminated 3 on October 01, 2018 Recall Number Z-0808-2018 Recall Event ID 79178 510(K)Number K123375 Product Classification System, image processing, radiological - Product Code LLZ Product Syngo.via systems with SW VB20A Code Information 10496180 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact Anastasia Sokolova 610-219-4834 Manufacturer Reason for Recall A software update was released on August 22, 2017. Performing this software installation and upgrading syngo.via system to the software version VB20A_HF02 will resolve various bugs and inconsistent diagnosis with the old system. FDA Determined Cause 2 Software Design Change Action On August 25, 2017 an URGENT letter was issued to customers notify of a system update for Syngo.via systems with SW VB20A. The software was released on August 22, 2017. Several bugs and inconsistent diagnoses were detected and this software is intended to correct these. If you have any questions, call Siemens Customer Care Centre at 888-826-9702 (24/7). Quantity in Commerce 30 Distribution Nationwide Distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On August 25, 2017 an URGENT letter was issued to customers notify of a system update for Syngo.via systems with SW VB20A. The software was released on August 22, 2017. Several bugs and inconsistent diagnoses were detected and this software is intended to correct these. If you have any questions, call Siemens Customer Care Centre at 888-826-9702 (24/7). | SoftwareDesignChange
108 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=160086 | Z-0724-2018 | Integrity R1.2 | Integrity | 2 | CDRH | 2018-02-26 |  | 2024-03-15 16:39 | Elekta Limited | Potential for positional errors following automatic table movement (ATM). | Class 2 Device Recall Integrity Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Integrity FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Integrity Date Initiated by Firm November 20, 2017 Create Date February 26, 2018 Recall Status 1 Terminated 3 on March 15, 2024 Recall Number Z-0724-2018 Recall Event ID 78663 510(K)Number K112613 Product Classification Accelerator, linear, medical - Product Code IYE Product Integrity R1.2 Code Information Product code 90I--YE, no serial number Recalling Firm/ Manufacturer Elekta Limited Fleming Way Crawley United Kingdom For Additional Information Contact Alejandra Carrillo 404-2024333 Manufacturer Reason for Recall Potential for positional errors following automatic table movement (ATM). FDA Determined Cause 2 Software Design Change Action Customers were notified via letter on about [DATE]. Instructions include to post the notice in a place accessible to all users until the action is closed and to advise appropriate personnel working with the affected product of the recall action. Additionally Elekta plans to release Integrity R4.0.0, which should identify positional errors over 5 mm, in the first half of 2018. Quantity in Commerce 539 units total Distribution Distributed domestically to AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. ************Distributed internationally to ALBANIA, ALGERIA, ANTIGUA AND BARBUDA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOSNIA AND HERZEGOVINIA, BOTSWANA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, LAOS, LITHUANIA, MALAYSIA, MALTA, MEXICO, MOROCCO, MYANMAR, NAMIBIA, NEPAL, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA******SUDAN, SURINAME, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UK, UKRAINE, VENEZUELA, VIETNAM, ZIMBABWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were notified via letter on about [DATE]. Instructions include to post the notice in a place accessible to all users until the action is closed and to advise appropriate personnel working with the affected product of the recall action. Additionally Elekta plans to release Integrity R4.0.0, which should identify positional errors over 5 mm, in the first half of 2018. | SoftwareDesignChange
109 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=160085 | Z-0723-2018 | Agility with Integrity R3.2 | Agility Accelerator, linear, medical | 2 | CDRH | 2018-02-26 |  | 2024-03-15 16:39 | Elekta Limited | Potential for positional errors following automatic table movement (ATM). | Class 2 Device Recall Agility Accelerator, linear, medical Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Agility Accelerator, linear, medical FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Agility Accelerator, linear, medical Date Initiated by Firm November 20, 2017 Create Date February 26, 2018 Recall Status 1 Terminated 3 on March 15, 2024 Recall Number Z-0723-2018 Recall Event ID 78663 510(K)Number K123808 Product Classification Accelerator, linear, medical - Product Code IYE Product Agility with Integrity R3.2 Code Information Product code 90I--YE, no serial numbers Recalling Firm/ Manufacturer Elekta Limited Fleming Way Crawley United Kingdom For Additional Information Contact Alejandra Carrillo 404-2024333 Manufacturer Reason for Recall Potential for positional errors following automatic table movement (ATM). FDA Determined Cause 2 Software Design Change Action Customers were notified via letter on about [DATE]. Instructions include to post the notice in a place accessible to all users until the action is closed and to advise appropriate personnel working with the affected product of the recall action. Additionally Elekta plans to release Integrity R4.0.0, which should identify positional errors over 5 mm, in the first half of 2018. Quantity in Commerce 539 units total Distribution Distributed domestically to AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. ************Distributed internationally to ALBANIA, ALGERIA, ANTIGUA AND BARBUDA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOSNIA AND HERZEGOVINIA, BOTSWANA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, LAOS, LITHUANIA, MALAYSIA, MALTA, MEXICO, MOROCCO, MYANMAR, NAMIBIA, NEPAL, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA******SUDAN, SURINAME, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UK, UKRAINE, VENEZUELA, VIETNAM, ZIMBABWE Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were notified via letter on about [DATE]. Instructions include to post the notice in a place accessible to all users until the action is closed and to advise appropriate personnel working with the affected product of the recall action. Additionally Elekta plans to release Integrity R4.0.0, which should identify positional errors over 5 mm, in the first half of 2018. | SoftwareDesignChange
110 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=159835 | Z-0263-2018 | Xper Flex Cardio Physiomonitoring system.******Used to facilitate invasive investigations of heart and vascular disease when non-invasive indicators warrant such. | Xper Flex Cardio Physiomonitoring system | 2 | CDRH | 2017-12-19 |  | 2024-09-20 14:33 | Invivo Corporation | Intermittent communication between the host system and the FlexCardio | Class 2 Device Recall Xper Flex Cardio Physiomonitoring system Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Xper Flex Cardio Physiomonitoring system FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Xper Flex Cardio Physiomonitoring system Date Initiated by Firm August 18, 2017 Create Date December 19, 2017 Recall Status 1 Terminated 3 on September 20, 2024 Recall Number Z-0263-2018 Recall Event ID 78585 510(K)Number K101571 Product Classification Monitor,physiological,patient(without arrhythmia detection or alarms) - Product Code MWI Product Xper Flex Cardio Physiomonitoring system.************Used to facilitate invasive investigations of heart and vascular disease when non-invasive indicators warrant such. Code Information All Flex Cardio devices, Revision A, B Code Number's: 453564241901, 453564241911, 453564483321, 453564483331, 989803185601, 9898031930191, 860335 and 860338 Recalling Firm/ Manufacturer Invivo Corporation 12151 Research Pkwy Ste 200 Orlando FL 32826-3286 For Additional Information Contact Your Local Philips Representative 407-275-3220 Manufacturer Reason for Recall Intermittent communication between the host system and the FlexCardio FDA Determined Cause 2 Software Design Change Action On August 24, 2017 Philips send the following communications to distributors and clients.******Dear Distributor,************Philips is initiating an action related to the Xper Flex Cardio Patient Monitoring System.************The attached Field Safety Notice is intended to inform customers about:************"***what the problem is and under what circumstances it can occur******"***the actions that should be taken by the customer / user in order to prevent risks for patients, users or servicers******"***the actions planned by Philips to correct the problem.************It is imperative that all affected end users of the Xper Flex Cardio receive the attached Field Safety Notice. Because Philips sells these products through distributors, including your organization, we may not have the information necessary to contact all end users. Therefore, please send the attached Field Safety Notice to each customer to whom you have distributed one of the affected devices. A description of the affected device can be found in the Field Safety Notice attached.******--******Dear Customer,************A problem has been detected in the Philips Xper Flex Cardio Patient Monitoring System (***Flex Cardio***) that, if it were to recur, could pose a risk for patients or users. This Field Safety Notice is intended to inform you about:************"***what the problem is and under what circumstances it can occur******"***the actions that should be taken by the customer / user in order to prevent risks for patients or users******"***the actions planned by Philips to correct the problem.******Because of intermittent communication between the host system and the Flex Cardio, it is possible for any of the following conditions to occur, which could result in delays in diagnosis or treatment:************"***Boom Monitor (display) may not display all active waveform and/or vital sign data.******"***Delayed audible and visual alarms at the Flex Cardio and Boom Monitor respectively, due to a mismatch between the alarm limits on the host system and the Flex Cardio/Boom Monitor.******"***After patient admission and ca Quantity in Commerce 4,535 Distribution AK, AL, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, KS, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and Puerto Rico******Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Bulgaria, Cambodia, Canada, Cayman Island, Chile, China, Colombia, Congo, Cote D'Ivoire, Cuba, Cyprus, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Honduras, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Jordan, Kazakhstan, Kenya, Korea, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Maldives, Mauritius, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Caledonia, New Zealand, Norway, Oman, Pakistan, Palestine, Panama, Papua, Paraguay, Philippines, Poland, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Vietnam, and Yemen Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MWI - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On August 24, 2017 Philips send the following communications to distributors and clients.******Dear Distributor,************Philips is initiating an action related to the Xper Flex Cardio Patient Monitoring System.************The attached Field Safety Notice is intended to inform customers about:************"***what the problem is and under what circumstances it can occur******"***the actions that should be taken by the customer / user in order to prevent risks for patients, users or servicers******"***the actions planned by Philips to correct the problem.************It is imperative that all affected end users of the Xper Flex Cardio receive the attached Field Safety Notice. Because Philips sells these products through distributors, including your organization, we may not have the information necessary to contact all end users. Therefore, please send the attached Field Safety Notice to each customer to whom you have distributed one of the affected devices. A description of the affected device can be found in the Field Safety Notice attached.******--******Dear Customer,************A problem has been detected in the Philips Xper Flex Cardio Patient Monitoring System (***Flex Cardio***) that, if it were to recur, could pose a risk for patients or users. This Field Safety Notice is intended to inform you about:************"***what the problem is and under what circumstances it can occur******"***the actions that should be taken by the customer / user in order to prevent risks for patients or users******"***the actions planned by Philips to correct the problem.******Because of intermittent communication between the host system and the Flex Cardio, it is possible for any of the following conditions to occur, which could result in delays in diagnosis or treatment:************"***Boom Monitor (display) may not display all active waveform and/or vital sign data.******"***Delayed audible and visual alarms at the Flex Cardio and Boom Monitor respectively, due to a mismatch between the alarm limits on the host system and the Flex Cardio/Boom Monitor.******"***After patient admission and ca | SoftwareDesignChange
111 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=157969 | Z-0202-2018 | CARESCAPE Patient Data Module (PDM) v2.6 software used with CARESCAPE Monitors B850, B650, and B450 v2.0.7 or earlier software***The PDM is intended to provide uninterrupted acquisition of physiologic parameter data on adult, pediatric and neonatal patients during non-transport/bedside and transport patient care episodes. Physiological parameter data acquired by the PDM includes EGG, invasive blood pressure, non-invasive blood pressure, pulse oximetry, temperature, cardiac output and respiration. This device acquires, processes and stores information for all aforementioned parameters and transmits this information to a transport or bedside central processing unit for viewing and alarm surveillance purposes. | CARESCAPE | 2 | CDRH | 2017-12-06 |  | 2021-04-01 10:13 | GE Medical Systems Information Technologies, Inc. | Device does not produce a visual or audible impedance respiration APN alarm when an impedance respiration apnea event occurs. | Class 2 Device Recall CARESCAPE Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall CARESCAPE FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall CARESCAPE Date Initiated by Firm July 13, 2017 Create Date December 06, 2017 Recall Status 1 Terminated 3 on April 01, 2021 Recall Number Z-0202-2018 Recall Event ID 77818 510(K)Number K071073 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product CARESCAPE Patient Data Module (PDM) v2.6 software used with CARESCAPE Monitors B850, B650, and B450 v2.0.7 or earlier software******The PDM is intended to provide uninterrupted acquisition of physiologic parameter data on adult, pediatric and neonatal patients during non-transport/bedside and transport patient care episodes. Physiological parameter data acquired by the PDM includes EGG, invasive blood pressure, non-invasive blood pressure, pulse oximetry, temperature, cardiac output and respiration. This device acquires, processes and stores information for all aforementioned parameters and transmits this information to a transport or bedside central processing unit for viewing and alarm surveillance purposes. Code Information Model No. 2042084-001. Serial No. SA314096842GA, SA314212130GA, SA313430730GA, SA313430743GA, SA313430731GA, SA313430735GA, SA313430771GA, SA313430795GA, SA313430801GA, SA313430806GA, SA313430813GA, SA313441307GA, SA313441311GA, SA313441399GA, SA313441414GA, SA313441474GA, SA314086697GA, SA313272250GA, SA313378480GA, SA313419589GA, SA313430691GA, SA313430697GA, SA313430700GA, SA313430705GA, SA313430706GA, SA313430711GA, SA313430713GA, SA313430715GA, SA313430723GA, SA313430724GA, SA313430751GA, SA313430752GA, SA313430753GA, SA313430755GA, SA313430757GA, SA313430759GA, SA313430761GA, SA313430762GA, SA313430763GA, SA313430764GA, SA313430766GA, SA313430767GA, SA313430777GA, SA313430779GA, SA313430780GA, SA313430781GA, SA313430783GA, SA313430784GA, SA313430789GA, SA313430790GA, SA313430792GA, SA313430793GA, SA313430794GA, SA313430796GA, SA313430797GA, SA313430798GA, SA313430799GA, SA313430802GA, SA313430804GA, SA313430805GA, SA313430807GA, SA313430808GA, SA313430809GA, SA313430812GA,SA313430815GA, SA313430819GA, SA313430822GA, SA313430823GA, SA313430824GA, SA313430825GA, SA313430827GA, SA313430828GA, SA313430875GA, SA313430907GA, SA313430915GA, SA313430916GA, SA313431027GA, SA313431034GA, SA313431040GA, SA313431043GA, SA313431045GA, SA313431047GA, SA313431048GA, SA313431062GA, SA313431063GA, SA313431064GA, SA313431065GA, SA313431066GA, SA313431067GA, SA313431069GA, SA313431073GA, SA313431074GA, SA313431079GA, SA313431081GA, SA313431082GA, SA313431083GA, SA313431089GA, SA313431091GA, SA313431098GA, SA313431099GA, SA313431104GA, SA313431158GA, SA313441304GA, SA313441305GA, SA313441315GA, SA313441334GA, SA313441362GA, SA313441363GA, SA313441367GA, SA313441369GA, SA313441370GA, SA313441372GA, SA313441374GA, SA313441377GA, SA313441382GA, SA313441391GA, SA313441417GA, SA313441438GA, SA313441440GA, SA313441460GA, SA313441461GA, SA313441462GA, SA313441464GA, SA313441465GA, SA313441471GA, SA313441476GA, SA313441478GA, SA313441621GA, SA313441622GA, SA313441623GA, SA313441684GA, SA313441686GA, SA313441687GA, SA314065742GA, SA314065752GA, SA314065754GA, SA314065756GA, SA314065757GA, SA314065759GA, SA314065760GA, SA314065763GA, SA314065765GA, SA314065766GA, SA314065767GA, SA314065768GA, SA314065785GA, SA314065788GA, SA314065791GA, SA314065793GA, SA314075804GA, SA314075806GA, SA314075807GA, SA314075808GA, SA314075809GA, SA314075810GA, SA314075812GA, SA314075814GA, SA314075823GA, SA314075825GA, SA314075826GA, SA314075827GA, SA314075828GA, SA314075829GA, SA314075830GA, SA314075839GA, SA314075850GA, SA314075853GA, SA314075855GA, SA314075857GA, SA314075858GA, SA314075859GA, SA314075860GA, SA314075861GA, SA314075862GA, SA314075863GA, SA314075864GA, SA314086544GA, SA314086555GA, SA314086559GA, SA314086566GA, SA314086628GA, SA314086651GA, SA314086661GA, SA314086702GA, SA314086716GA, SA314096753GA, SA314096756GA, SA314096757GA, SA314096760GA, SA314096762GA, SA314096764GA, SA314096765GA, SA314096775GA, SA314096782GA, SA314096784GA, SA314096785GA, SA314096790GA, SA314096794GA, SA314096804GA, SA314096808GA, SA314096809GA, SA314096810GA, SA314096814GA, SA314096817GA, SA314096818GA, SA314096819GA, SA314096821GA, SA314096822GA, SA314096823GA, SA314096825GA, SA314096828GA, SA314096829GA, SA314096837GA, SA314096838GA, SA314096839GA, SA314096840GA, SA314096843GA, SA314096852GA, SA314096853GA, SA314096854GA, SA314096855GA, SA314096868GA, SA314096870GA, SA314096933GA, SA314096937GA, SA314096938GA, SA314096939GA, SA314096942GA, SA314096944GA, SA314096946GA, SA314096947GA, SA314096948GA, SA314096949GA, SA314096951GA, SA314096952GA, SA314096957GA, SA314096958GA, SA314096960GA, SA314096962GA, SA314096999GA, SA314097025GA, SA314097071GA, SA314097080GA, SA314097086GA, SA314097152GA, SA314097159GA, SA314097161GA, SA314107694GA, SA314107701GA, SA314107703GA, SA314108152GA, SA314149726GA, SA314159806GA, SA314160061GA, SA314160065GA, SA314160074GA, SA314160079GA, SA314160082GA, SA314160088GA, SA314160098GA, SA314160099GA, SA314160105GA, SA314170129GA, SA314170133GA.SA314170135GA, SA314170136GA, SA314170137GA, SA314170139GA, SA314170140GA, SA314170141GA, SA314170142GA, SA314170143GA, SA314170145GA, SA314170147GA, SA314170148GA, SA314170170GA, SA314170176GA, SA314170179GA, SA314170277GA, SA314170312GA, SA314170313GA, SA314170314GA, SA314170316GA, SA314170317GA, SA314170318GA, SA314170319GA, SA314170320GA, SA314170321GA, SA314170322GA, SA314170323GA, SA314170324GA, SA314170325GA, SA314170326GA, SA314170328GA, SA314170329GA, SA314170330GA, SA314170331GA, SA314170332GA, SA314170333GA, SA314170334GA, SA314170335GA, SA314170336GA, SA314170337GA, SA314170338GA, SA314170339GA, SA314170340GA, SA314170341GA, SA314170367GA, SA314170368GA, SA314170370GA, SA314170371GA, SA314170374GA, SA314170375GA, SA314170381GA, SA314170382GA, SA314170383GA, SA314170385GA, SA314170391GA, SA314170392GA, SA314170393GA, SA314170394GA, SA314170395GA, SA314190788GA, SA314190793GA, SA314190794GA, SA314190796GA, SA314190799GA, SA314190800GA, SA314190813GA, SA314190814GA, SA314190815GA, SA314190817GA, SA314190818GA, SA314190819GA, SA314190820GA, SA314190821GA, SA314190822GA, SA314190823GA, SA314190824GA, SA314190825GA, SA314190826GA, SA314190827GA, SA314190828GA, SA314190829GA, SA314190830GA, SA314190831GA, SA314190832GA, SA314190833GA, SA314190834GA, SA314190835GA, SA314190836GA, SA314190837GA, SA314190839GA, SA314190840GA, SA314190841GA, SA314190842GA, SA314190843GA, SA314190844GA, SA314190845GA, SA314190846GA, SA314190847GA, SA314190848GA, SA314190849GA, SA314190851GA, SA314190852GA, SA314190853GA, SA314190854GA, SA314190855GA, SA314190856GA, SA314190858GA, SA314190860GA, SA314190861GA, SA314190864GA, SA314190865GA, SA314190866GA, SA314190867GA, SA314190868GA, SA314190869GA, SA314190871GA, SA314190874GA, SA314190875GA, SA314190876GA, SA314190877GA, SA314190878GA, SA314190880GA, SA314190882GA, SA314190883GA, SA314190884GA, SA314190887GA, SA314190888GA, SA314190889GA, SA314190891GA, SA314190893GA, SA314190898GA, SA314190900GA, SA314190903GA, SA314190912GA, SA314190913GA, SA314190914GA, SA314190916GA, SA314190918GA, SA314190919GA, SA314190921GA, SA314190922GA, SA314190924GA, SA314190927GA, SA314190928GA, SA314190929GA, SA314190931GA, SA314190937GA, SA314190938GA, SA314190954GA, SA314190956GA, SA314190957GA, SA314190958GA, SA314190967GA, SA314190976GA, SA314190979GA, SA314191039GA, SA314191061GA, SA314191072GA, SA314191081GA, SA314191091GA, SA314191107GA, SA314191165GA, SA314191172GA, SA314191174GA, SA314201207GA, SA314201208GA, SA314201211GA, SA314201212GA, SA314201217GA, SA314201219GA, SA314201221GA, SA314201222GA, SA314201227GA, SA314201231GA, SA314201235GA, SA314201236GA, SA314201238GA, SA314201241GA, SA314201242GA, SA314201243GA, SA314201247GA, SA314201248GA, SA314201249GA, SA314201251GA, SA314201253GA, SA314201255GA, SA314201256GA, SA314201257GA, SA314201261GA, SA314201262GA, SA314201266GA, SA314201267GA, SA314201270GA, SA314201273GA, SA314201277GA, SA314201278GA, SA314201280GA, SA314201282GA, SA314201284GA, SA314201285GA, SA314201287GA, SA314201291GA, SA314201294GA, SA314201307GA, SA314201308GA, SA314201312GA, SA314201314GA, SA314201317GA, SA314201325GA, SA314201327GA, SA314201330GA, SA314201331GA, SA314201332GA, SA314201338GA, SA314201344GA, SA314201348GA, SA314201352GA, SA314201354GA, SA314201360GA, SA314201361GA, SA314201362GA, SA314201367GA, SA314201368GA, SA314201369GA, SA314201370GA, SA314201371GA, SA314201373GA, SA314201376GA, SA314201378GA, SA314201458GA, SA314201461GA, SA314201534GA, SA314201552GA, SA314201572GA, SA314201578GA, SA314201646GA, SA314201648GA, SA314201672GA, SA314201691GA, SA314201798GA, SA314211920GA, SA314211952GA, SA314211954GA, SA314211961GA, SA314211977GA, SA314211983GA, SA314211986GA, SA314211989GA, SA314211992GA, SA314211995GA, SA314211996GA, SA314212002GA, SA314212004GA, SA314212124GA, SA314212126GA, SA314212128GA, SA314212134GA, SA314212135GA, SA314212137GA, SA314212138GA, SA314212140GA, SA314212144GA, SA314212147GA, SA314212148GA, SA314212152GA, SA314212154GA, SA314212158GA, SA314212165GA, SA314212166GA, SA314212167GA, SA314212169GA, SA314212171GA, SA314212173GA, SA314212178GA, SA314212179GA, SA314212181GA, SA314212182GA, SA314212183GA, SA314222351GA, SA314222352GA, SA314222353GA, SA314222355GA, SA314222359GA, SA314222361GA, SA314222372GA, SA314222373GA, SA314222386GA, SA314222395GA, SA314222396GA, SA314222401GA, SA314222404GA, SA314222411GA, SA314222424GA, SA314222438GA, SA314222456GA, SA314222485GA, SA314222509GA, SA314222572GA, SA314222601GA, SA314222612GA, SA314222615GA, SA314222764GA, SA314222776GA, SA314222780GA, SA314222803GA, SA314222853GA, SA314222864GA, SA314222867GA, SA314222868GA, SA314222878GA, SA314222880GA, SA314222887GA, SA314222892GA, SA314222900GA, SA314222907GA, SA314222908GA, SA314222911GA, SA314222913GA, SA314222914GA, SA314222918GA, SA314222921GA, SA314222922GA, SA314222924GA, SA314222928GA, SA314222931GA, SA314222933GA, SA314222936GA, SA314222937GA, SA314222938GA, SA314222943GA, SA314222945GA, SA314222946GA, SA314222949GA, SA314222950GA, SA314222958GA, SA314222969GA, SA314222970GA, SA314243216GA, SA314243226GA, SA314243227GA, SA314243229GA, SA314243231GA, SA314243239GA, SA314243240GA, SA314243244GA, SA314243245GA, SA314243250GA, SA314243256GA, SA314243257GA, SA314243272GA, SA314243274GA, SA314243276GA, SA314243277GA, SA314243280GA, SA314243282GA, SA314243283GA, SA314243284GA, SA314243285GA, SA314243287GA, SA314243289GA, SA314243291GA, SA314243294GA, SA314243301GA, SA314243305GA, SA314243327GA, SA314243335GA, SA314243337GA, SA314243348GA, SA314243403GA, SA314273878GA, SA314273897GA, SA314273898GA, SA314273996GA, SA314273998GA, SA314273999GA, SA314274000GA, SA314274001GA, SA314274002GA, SA314274003GA, SA314274006GA, SA314274009GA, SA314274011GA, SA314274012GA, SA314274013GA, SA314274014GA, SA314274015GA, SA314274016GA, SA314274017GA, SA314274026GA, SA314274027GA, SA314274029GA, SA314274030GA, SA314274031GA, SA314274033GA, SA314274034GA, SA314274035GA, SA314274036GA, SA314274038GA, SA314274039GA, SA314274040GA, SA314274041GA, SA314274042GA, SA314274049GA, SA314274051GA, SA314274055GA, SA314274070GA, SA314274071GA, SA314086548GA, SA314065751GA, SA314065753GA, SA314096820GA, SA314086646GA, SA314096955GA, SA313012912GA, SA313023225GA, SA313023334GA, SA313043939GA, SA313043982GA, SA313043986GA, SA313043992GA, SA313043993GA, SA313043995GA, SA313044003GA, SA313044124GA, SA313044126GA, SA313044127GA, SA313044128GA, SA313044130GA, SA313064165GA, SA313064204GA, SA313064213GA, SA313064215GA, SA313064219GA, SA313064221GA, SA313064224GA, SA313064227GA, SA313064228GA, SA313064264GA, SA313064286GA, SA313064289GA, SA313064291GA, SA313064304GA, SA313064308GA, SA313064312GA, SA313064313GA, SA313064317GA, SA313064328GA, SA313064329GA, SA313064332GA, SA313064335GA, SA313064384GA, SA313064388GA, SA313064389GA, SA313064390GA, SA313064395GA, SA313064396GA, SA313064397GA, SA313064398GA, SA313064399GA, SA313064400GA, SA313064404GA, SA313064405GA, SA313064406GA, SA313064407GA, SA313064408GA, SA313064409GA, SA313064411GA, SA313064412GA, SA313064413GA, SA313064414GA, SA313064416GA, SA313064417GA, SA313064418GA, SA313064419GA, SA313064420GA, SA313064421GA, SA313064422GA, SA313064423GA, SA313064424GA, SA313116790GA, SA307504196GA, SA307504206GA, SA308206740GA, SA308398683GA, SA308409193GA, SA308511505GA, SA309062299GA, SA309324364GA, SA309324390GA, SA309324403GA, SA309324407GA, SA310139080GA, SA310139087GA, SA310139268GA, SA311111666GA, SA311121772GA, SA311121791GA, SA311142804GA, SA311142898GA, SA312070143GA, SA307382574GA, SA307514340GA, SA309385607GA, SA314014581GA, SA314108142GA, SA314118914GA, SA307483786GA, SA307504190GA, SA307504209GA, SA307514325GA, SA307514328GA, SA307504198GA, SA307504203GA, SA307514331GA, SA307382531GA, SA307382601GA, SA307382612GA, SA307443225GA, SA307443239GA, SA307443250GA, SA307443253GA, SA307453391GA, SA307453411GA, SA307453414GA, SA307453418GA, SA307453551GA, SA307453558GA, SA307453604GA, SA307453606GA, SA307483768GA, SA307493938GA, SA307493950GA, SA307504117GA, SA307504128GA, SA307504133GA, SA307504172GA, SA307443370GA, SA307504110GA, SA308084774GA, SA308084758GA, SA308084759GA, SA307362382GA, SA307443365GA, SA307453598GA, SA307483774GA, SA307483811GA, SA307493939GA, SA307504165GA, SA307504197GA, SA308491249GA, SA308501408GA, SA308511502GA, SA308511517GA, SA308511519GA, SA308511524GA, SA308511533GA, SA309062267GA, SA309062271GA, SA309062273GA, SA309062278GA, SA309062279GA, SA309062281GA, SA309062282GA, SA309062283GA, SA309062285GA, SA309062298GA, SA309062302GA, SA309062303GA, SA309122752GA, SA309122765GA, SA309122771GA, SA309122784GA, SA309122791GA, SA309132802GA, SA309304218GA, SA309324355GA, SA309324358GA, SA309324359GA, SA309324362GA, SA309324363GA, SA309324367GA, SA309324370GA, SA309324372GA, SA309324374GA, SA309324378GA, SA309324379GA, SA309324381GA, SA309324388GA, SA309324397GA, SA309324398GA, SA309324405GA, SA309324406GA, SA309324419GA, SA309324420GA, SA309324421GA, SA309324422GA, SA309324423GA, SA309324435GA, SA309324436GA, SA309375492GA, SA309375507GA, SA309385546GA, SA309385547GA, SA309385552GA, SA309385553GA, SA309385556GA, SA309385567GA, SA309385579GA, SA309385605GA, SA309385608GA, SA309385640GA, SA309385538GA, SA309385584GA, SA309385627GA, SA310169523GA, SA310139074GA, SA310139082GA, SA310139083GA, SA310139265GA, SA310139267GA, SA310139270GA, SA310139284GA, SA310260041GA, SA310260051GA, SA310260054GA, SA310260055GA, SA310260062GA, SA310260081GA, SA310139068GA, SA310372025GA, SA310372144GA, SA310372152GA, SA310372158GA, SA310372275GA, SA310372280GA, SA310372304GA, SA310372310GA, SA311111692GA, SA311121788GA, SA311121790GA, SA311121796GA, SA311121801GA, SA311121814GA, SA311121815GA, SA311121819GA, SA311121829GA, SA311121832GA, SA311121838GA, SA311121839GA, SA311121840GA, SA311121844GA, SA311121845GA, SA311132372GA, SA311132389GA, SA311132529GA, SA311132616GA, SA311142713GA, SA311142716GA, SA311142745GA, SA311142798GA, SA311142806GA, SA311153334GA, SA311132563GA, SA311132565GA, SA311132573GA, SA311132586GA, SA311142882GA, SA311142904GA, SA311111665GA, SA311111668GA, SA311111670GA, SA311111678GA, SA311111682GA, SA311111683GA, SA311111684GA, SA311111690GA, SA311111693GA, SA311121752GA, SA311121759GA, SA311121763GA, SA311121769GA, SA311121785GA, SA311121793GA, SA311121794GA, SA311121795GA, SA311121824GA, SA311121825GA, SA311132483GA, SA311132610GA, SA311132620GA, SA312070057GA, SA312070068GA, SA312070139GA, SA312070164GA, SA312070198GA, SA312091226GA, SA312502087GA, SA312480512GA, SA313084589GA, SA313084590GA, SA313106217GA, SA313208952GA, SA313208961GA, SA313208970GA, SA313095467GA, SA313095487GA, SA313116845GA, SA313116855GA, SA313127579GA, SA313198738GA, SA313219011GA, SA313219017GA, SA313219035GA, SA313219037GA, SA313219047GA, SA313219050GA, SA313219059GA, SA313219060GA, SA313219064GA, SA313219090GA, SA313272023GA, SA313272065GA, SA313272066GA, SA313272070GA, SA313272314GA, SA313272319GA, SA313272322GA, SA313282756GA, SA313282785GA, SA313292916GA, SA313345833GA, SA313356688GA, SA313356691GA, SA313388585GA, SA313335282GA, SA313335293GA, SA313335329GA, SA313345532GA, SA313345533GA, SA313345548GA, SA313345550GA, SA313345592GA, SA313345600GA, SA313345602GA, SA313345603GA, SA313345612GA, SA313345617GA, SA313345619GA, SA313345627GA, SA313345651GA, SA313345786GA, SA314159763GA, SA314222816GA, SA314191108GA, SA314201240GA, SA314201244GA, SA314201250GA, SA314201254GA, SA314201260GA, SA314201548GA, SA314222884GA, SA314222906GA, SA307493932GA, SA310068858GA, SA310068880GA, SA313346104GA, SA313084910GA, SA313095431GA, SA313105821GA, SA313127701GA, SA313148001GA, SA313345542GA, SA313345545GA, SA313345551GA, SA313345564GA, SA313345606GA, SA313324988GA, SA313324997GA, SA313335009GA, SA313335067GA, SA313482611GA, SA313378461GA, SA313431206GA, SA313282833GA, SA313292929GA, SA307504207GA, SA307504208GA, SA307514337GA, SA308388614GA, SA308084783GA, SA308095015GA, SA308095074GA, SA308094934GA, SA308094942GA, SA308094949GA, SA308094965GA, SA308094978GA, SA308094944GA, SA308084729GA, SA308084756GA, SA307504218GA.SA307504220GA, SA307514330GA, SA308115265GA, SA308115276GA, SA308125325GA, SA308125326GA, SA308166092GA, SA308206718GA, SA308206736GA, SA308206747GA, SA308347781GA, SA308347786GA, SA308347788GA, SA308347794GA, SA308398795GA, SA308398802GA, SA308398803GA, SA308398810GA, SA308398812GA, SA308398825GA, SA308398826GA, SA308398828GA, SA308398830GA, SA308398855GA, SA308398857GA, SA308398903GA, SA308368167GA, SA308398813GA, SA308409190GA, SA308409192GA, SA308429324GA, SA308429332GA, SA308429334GA, SA308429335GA, SA308429336GA, SA308429343GA, SA308429344GA, SA308429348GA, SA308429358GA, SA308429318GA, SA308449516GA, SA308449572GA, SA308459659GA, SA308491161GA, SA308491177GA, SA308095000GA, SA308095013GA, SA307504201GA, SA309385544GA, SA309122778GA, SA308398858GA, SA308511525GA, SA313377985GA, SA314222792GA, SA313345653GA, SA307504223GA, SA307504200GA, SA310455121GA, SA310518483GA, SA310518509GA, SA310528775GA, SA310528925GA, SA310539311GA, SA311142719GA, SA310528915GA, SA310260402GA, SA310455119GA, SA310528795GA, SA311090954GA, SA310260320GA, SA310260383GA, SA310300791GA, SA310300795GA, SA310300798GA, SA310300806GA, SA310341381GA, SA310361785GA, SA310382838GA, SA310392886GA, SA310393202GA, SA310434122GA, SA310507218GA, SA310518198GA, SA310518203GA, SA310518208GA, SA310518262GA, SA310518396GA, SA310518399GA, SA310518452GA, SA310518453GA, SA310518463GA, SA310518479GA, SA310518492GA, SA310518493GA, SA310518495GA, SA310518501GA, SA310518521GA, SA310518533GA, SA310528733GA, SA310528778GA, SA310528813GA, SA310528842GA, SA310528914GA, SA310528918GA, SA310528928GA, SA310528952GA, SA310528955GA, SA310528959GA, SA310528960GA, SA310529061GA, SA310529102GA, SA310539239GA, SA314409312GA, SA316420518SA, SA316420545SA, SA317030027SA, SA317040589SA, SA313127634GA, SA309466839GA, SA313430874GA, SA315482785SA, SA316025061SA, SA316025074SA, SA316025147SA, SA316025157SA, SA316035166SA, SA316035168SA, SA316035177SA, SA316035188SA, SA3xxxxxxx1GA, SA3xxxxxxx2GA, SA3xxxxxxx3GA. Recalling Firm/ Manufacturer GE Medical Systems Information Technologies, Inc. 8200 W Tower Ave Milwaukee WI 53223-3219 For Additional Information Contact Technical Support 800-558-7044 Manufacturer Reason for Recall Device does not produce a visual or audible impedance respiration APN alarm when an impedance respiration apnea event occurs. FDA Determined Cause 2 Software Design Change Action GE Healthcare sent an Urgent Medical Device Correction letter dated July 12, 2017, to all affected consignees via certified mail. Consignees were advised to enable "Resp Rate (Impedance)" alarms and set the low limits to the factory defaults (NICU: 20/minute; Adult ICU: 5/minute). Consignees were also advised to enable audible alarms under Alarm Priorities for RR (Impedance) high/low. GE will provide a correction free of charge once it is available. GE will contact consignees to arrange device corrections. Consignees were instructed to call 1-800-558-7044 or a local representative with any questions pertaining to this recall. Quantity in Commerce 1,196 (1,176 US; 20 OUS) Distribution Worldwide Distribution - US including NY, WI; Foreign: Australia, France, Germany. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action GE Healthcare sent an Urgent Medical Device Correction letter dated July 12, 2017, to all affected consignees via certified mail. Consignees were advised to enable "Resp Rate (Impedance)" alarms and set the low limits to the factory defaults (NICU: 20/minute; Adult ICU: 5/minute). Consignees were also advised to enable audible alarms under Alarm Priorities for RR (Impedance) high/low. GE will provide a correction free of charge once it is available. GE will contact consignees to arrange device corrections. Consignees were instructed to call 1-800-558-7044 or a local representative with any questions pertaining to this recall. | SoftwareDesignChange
112 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=157814 | Z-0161-2018 | ROSA Surgical Device 2.5.8 ****** It is intended to be used in the operating room for the spatial positioning and orientation of an instrument holder or tool guide. Guidance is based on a pre-operative plan developed with three-dimensional imaging software, and uses fiducial markers or optical registration. The system is intended for use by neurosurgeons to guide standard neurosurgical instruments. It is indicated for any neurosurgical condition in which the use of stereotactic surgery may be appropriate. | Computer-Assisted Surgical Device | 2 | CDRH | 2017-11-21 |  | 2020-06-16 10:13 | Zimmer Biomet, Inc. | Software corrections reactivating the cooperative endoscopy mode. | Class 2 Device Recall ComputerAssisted Surgical Device Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ComputerAssisted Surgical Device FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ComputerAssisted Surgical Device Date Initiated by Firm June 04, 2014 Create Date November 21, 2017 Recall Status 1 Terminated 3 on June 16, 2020 Recall Number Z-0161-2018 Recall Event ID 77750 510(K)Number K092239 K101791 Product Classification Neurological stereotaxic Instrument - Product Code HAW Product ROSA Surgical Device 2.5.8 ************ It is intended to be used in the operating room for the spatial positioning and orientation of an instrument holder or tool guide. Guidance is based on a pre-operative plan developed with three-dimensional imaging software, and uses fiducial markers or optical registration. The system is intended for use by neurosurgeons to guide standard neurosurgical instruments. It is indicated for any neurosurgical condition in which the use of stereotactic surgery may be appropriate. Code Information Serial No. RO10011, RO14033, RO14031, RO14035, RO13027, RO13023, RO10014, RO10009, RO10010, RO14030, RO11016, RO11017, RO12022, RO14029, RO14032, RO08003, RO09004, RO10007, RO10013, RO12018, RO12019, RO13025, RO14034, RO11015, RO13024, RO09005, RO13026, RO12021, and RO12020. Recalling Firm/ Manufacturer Zimmer Biomet, Inc. 1800 W Center St Warsaw IN 46580-2304 For Additional Information Contact Kevin Escaule 574-267-6131 Manufacturer Reason for Recall Software corrections reactivating the cooperative endoscopy mode. FDA Determined Cause 2 Software Design Change Action The firm, Zimmer Biomet Inc., sent Field Service Technicians to consignee locations to perform system upgrade. All US units have been corrected.************If you any question, please contact Post Market Surveillance and Regulatory Compliance Director at 574-372-4487 or email to: kevin.escapule@zimmberbiomet.com. Quantity in Commerce 29 devices(8 US and 21 OUS) Distribution Worldwide distribution: US (nationwide) in states of: AR, GA, MI, OH, and TX; and countries of:Canada, China, France, Germany, Israel, Italy, Russia, Saudi Arabia, and Spain. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = HAW 510(K)s with Product Code = HAW - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, Zimmer Biomet Inc., sent Field Service Technicians to consignee locations to perform system upgrade. All US units have been corrected.************If you any question, please contact Post Market Surveillance and Regulatory Compliance Director at 574-372-4487 or email to: kevin.escapule@zimmberbiomet.com. | SoftwareDesignChange
113 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=159623 | Z-0082-2018 | Xper Flex Cardio PhysioMonitoring system ***The Xper Flex Cardio physiomonitoring system is used to facilitate invasive investigation of heart and vascular disease when non-invasive indicators warrant such. Its users, responsible to interpret the data made available, will be professional health care providers, such as physicians, registered nurses, cardiovascular technicians, radiology technicians, and other clinical health care providers. The Xper Flex Cardio physiomonitoring system may be used to display and analyze surface ECG (electrocardiogram), respiration, invasive pressure, Sp02 (Pulse Oximetry), End Tidal C02 (carbon dioxide) and non-invasive pressure waveforms; surface body temperature and thermal cardiac output curves. The system is capable of processing/analyzing information such as multi-channel ECG signals and displaying a graphical ST segment map. The Xper Flex Cardio physiomonitoring system makes measurements that assist physicians and other clinicians to evaluate a patient's overall cardiopulmonary performance, conductive system and general vascular status. | Flex Cardio Patient Monitoring System | 2 | CDRH | 2017-11-09 |  | 2024-09-20 14:38 | Invivo Corporation | Intermittent communication between the host system and the Flex Cardio | Class 2 Device Recall Flex Cardio Patient Monitoring System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Flex Cardio Patient Monitoring System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Flex Cardio Patient Monitoring System Date Initiated by Firm August 11, 2017 Create Date November 09, 2017 Recall Status 1 Terminated 3 on September 20, 2024 Recall Number Z-0082-2018 Recall Event ID 78465 510(K)Number K101571 Product Classification Monitor,physiological,patient(without arrhythmia detection or alarms) - Product Code MWI Product Xper Flex Cardio PhysioMonitoring system ******The Xper Flex Cardio physiomonitoring system is used to facilitate invasive investigation of heart and vascular disease when non-invasive indicators warrant such. Its users, responsible to interpret the data made available, will be professional health care providers, such as physicians, registered nurses, cardiovascular technicians, radiology technicians, and other clinical health care providers. The Xper Flex Cardio physiomonitoring system may be used to display and analyze surface ECG (electrocardiogram), respiration, invasive pressure, Sp02 (Pulse Oximetry), End Tidal C02 (carbon dioxide) and non-invasive pressure waveforms; surface body temperature and thermal cardiac output curves. The system is capable of processing/analyzing information such as multi-channel ECG signals and displaying a graphical ST segment map. The Xper Flex Cardio physiomonitoring system makes measurements that assist physicians and other clinicians to evaluate a patient's overall cardiopulmonary performance, conductive system and general vascular status. Code Information Model 453564621791 and 453564621801 - ALL Flex Cardio Devices, Revision C Recalling Firm/ Manufacturer Invivo Corporation 12151 Research Pkwy Ste 200 Orlando FL 32826-3286 For Additional Information Contact Contact your Philip representative 407-275-3220 Manufacturer Reason for Recall Intermittent communication between the host system and the Flex Cardio FDA Determined Cause 2 Software Design Change Action On August 2017 the firm sent letters to their consignees indicating the following:******This document contains important information for the continued safe and proper use of your equipment************Please review the following information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication.************Please retain a copy with the equipment Instruction for Use.******-------------------------------------------******A problem has been detected in the Philips Xper Flex Cardio Patient Monitoring System (***Flex Cardio***) that, if it were to recur, could pose a risk for patients or users. This Field Safety Notice is intended to inform you about:************"***what the problem is and under what circumstances it can occur******"***the actions that should be taken by the customer / user in order to prevent risks for patients or users******"***the actions planned by Philips to correct the problem.******Because of intermittent communication between the host system and the Flex Cardio, it is possible for any of the following conditions to occur:******"***Inability of the Flex Cardio to connect to the host system.******"***Boom Monitor (display) may not display all active waveform and/or vital sign data.******"***Delayed audible and visual alarms at the Flex Cardio and Boom Monitor respectively, due to a mismatch between the alarm limits on the host system and the Flex Cardio/Boom Monitor.******"***After patient admission and case initiation at the host system, the case may not be fully transmitted to the Flex Cardio. In this case, the Boom Monitor in the exam/procedure room will display vital sign data, but the Flex Cardio will not provide audible alarms.******"***After patient discharge and case termination at the host system, case termination may not be transmitted to the Flex Cardio. In this case, the Boom Monitor in the exam/procedure room will continue to display vital sign data, and the Flex Cardio will continue to provide audible alarms.************Our re Quantity in Commerce 529 Distribution Worldwide Distribution - US including AK, AL, AZ, CA, CO, DE, FL, GA, IA, ID, IL, KY, LA, MA, MI, MO, NC, NE, NM, NY, OH, OR, OK, PA, SD, TN, TX, VA, WA, WV and Internationally to Germany, Hungary, India, Netherlands, Russia, Saudi Arabia, South Africa and Tunisia. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MWI - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On August 2017 the firm sent letters to their consignees indicating the following:******This document contains important information for the continued safe and proper use of your equipment************Please review the following information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication.************Please retain a copy with the equipment Instruction for Use.******-------------------------------------------******A problem has been detected in the Philips Xper Flex Cardio Patient Monitoring System (***Flex Cardio***) that, if it were to recur, could pose a risk for patients or users. This Field Safety Notice is intended to inform you about:************"***what the problem is and under what circumstances it can occur******"***the actions that should be taken by the customer / user in order to prevent risks for patients or users******"***the actions planned by Philips to correct the problem.******Because of intermittent communication between the host system and the Flex Cardio, it is possible for any of the following conditions to occur:******"***Inability of the Flex Cardio to connect to the host system.******"***Boom Monitor (display) may not display all active waveform and/or vital sign data.******"***Delayed audible and visual alarms at the Flex Cardio and Boom Monitor respectively, due to a mismatch between the alarm limits on the host system and the Flex Cardio/Boom Monitor.******"***After patient admission and case initiation at the host system, the case may not be fully transmitted to the Flex Cardio. In this case, the Boom Monitor in the exam/procedure room will display vital sign data, but the Flex Cardio will not provide audible alarms.******"***After patient discharge and case termination at the host system, case termination may not be transmitted to the Flex Cardio. In this case, the Boom Monitor in the exam/procedure room will continue to display vital sign data, and the Flex Cardio will continue to provide audible alarms.************Our re | SoftwareDesignChange
114 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153847 | Z-1711-2017 | UniCel DxC 600i SYNCHRON Access Clinical System, Part No. A25637, A25638******Product Usage:***The Access 2 and UniCel DxI Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids.*** | Unicel DxC 600i SYNCHRON Access Clinical Systems | 2 | CDRH | 2017-04-04 | 2017-03-15 21:02 | 2018-03-01 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access 2 Family of Immunoassay Systems because a circuit board that controls motor functions can affect device positioning on the pipettor and incubator belt while the instrument is running. The position errors cause the instrument to stop running and cancel tests that are in process. | Class 2 Device Recall Unicel DxC 600i SYNCHRON Access Clinical Systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Unicel DxC 600i SYNCHRON Access Clinical Systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Unicel DxC 600i SYNCHRON Access Clinical Systems Date Initiated by Firm March 07, 2017 Date Posted March 15, 2017 Recall Status 1 Terminated 3 on March 01, 2018 Recall Number Z-1711-2017 Recall Event ID 76688 510(K)Number K922823 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product UniCel DxC 600i SYNCHRON Access Clinical System, Part No. A25637, A25638************Product Usage:******The Access 2 and UniCel DxI Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. Code Information Serial No. (Part No. A25637) 800999,801019 Serial No. (Part No. A25638)800995,800996,800997,800998,801000,801001,801002,801003,801004, 801005,801006,801007,801008,801009,801010,801011,801012,801013, 801014,801015,801016,801017,801018,801020,801021,801022,801023, 801024,801025,801026,801027 Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 Manufacturer Reason for Recall Beckman Coulter is recalling the Access 2 Family of Immunoassay Systems because a circuit board that controls motor functions can affect device positioning on the pipettor and incubator belt while the instrument is running. The position errors cause the instrument to stop running and cancel tests that are in process. FDA Determined Cause 2 Software Design Change Action An Urgent Medical Device Recall letter was sent on 3/7/17 to inform customers that Beckman Coulter has determined through internal investigation that a circuit board that controls motor functions in the instrument systems listed above can affect device positioning on the pipettor and incubator belt while the instrument is running. Customers are informed of the impact, action to be taken, and the resolution. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions are instructed to contact Customer Technical Support******Center:****** From our website: http://www.beckmancoulter.com****** By phone: call 1-800-854-3633 in the United States and Canada.****** Outside the United States and Canada, contact your local Beckman Coulter******representative. Quantity in Commerce 851 units total for all products (291 units in US) Distribution Affected geographies: Algeria, Austria, Australia, Bangladesh, Bulgaria, Brazil, Canada, China, Costa Rica, Cyprus, Colombia, Czech Republic, El Salvador, France, Germany, Ghana, Greece, Guatemala, India, Indonesia, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Lithuania, Malawi, Mexico, Morocco, Myanmar, Nepal, Oman, Panama, Philippines, Pakistan, Paraguay, Poland, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Serbia, South Africa, South Korea, Spain, Switzerland, Tunisia, Turkey, United Kingdom, United States, Vietnam, Zimbabwe Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action An Urgent Medical Device Recall letter was sent on 3/7/17 to inform customers that Beckman Coulter has determined through internal investigation that a circuit board that controls motor functions in the instrument systems listed above can affect device positioning on the pipettor and incubator belt while the instrument is running. Customers are informed of the impact, action to be taken, and the resolution. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions are instructed to contact Customer Technical Support******Center:****** From our website: http://www.beckmancoulter.com****** By phone: call 1-800-854-3633 in the United States and Canada.****** Outside the United States and Canada, contact your local Beckman Coulter******representative. | SoftwareDesignChange
115 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153845 | Z-1709-2017 | Access 2 Immunoassay Systems, Part No. 81600N******Product Usage:***The Access 2 and UniCel DxI Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. | Access 2 Immunoassay System | 2 | CDRH | 2017-04-04 | 2017-03-15 21:02 | 2018-03-01 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access 2 Family of Immunoassay Systems because a circuit board that controls motor functions can affect device positioning on the pipettor and incubator belt while the instrument is running. The position errors cause the instrument to stop running and cancel tests that are in process. | Class 2 Device Recall Access 2 Immunoassay System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Access 2 Immunoassay System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Access 2 Immunoassay System Date Initiated by Firm March 07, 2017 Date Posted March 15, 2017 Recall Status 1 Terminated 3 on March 01, 2018 Recall Number Z-1709-2017 Recall Event ID 76688 510(K)Number K922823 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product Access 2 Immunoassay Systems, Part No. 81600N************Product Usage:******The Access 2 and UniCel DxI Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. Code Information Serial No. 511527 - 511678; 511681 - 512212 Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 Manufacturer Reason for Recall Beckman Coulter is recalling the Access 2 Family of Immunoassay Systems because a circuit board that controls motor functions can affect device positioning on the pipettor and incubator belt while the instrument is running. The position errors cause the instrument to stop running and cancel tests that are in process. FDA Determined Cause 2 Software Design Change Action An Urgent Medical Device Recall letter was sent on 3/7/17 to inform customers that Beckman Coulter has determined through internal investigation that a circuit board that controls motor functions in the instrument systems listed above can affect device positioning on the pipettor and incubator belt while the instrument is running. Customers are informed of the impact, action to be taken, and the resolution. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions are instructed to contact Customer Technical Support******Center:****** From our website: http://www.beckmancoulter.com****** By phone: call 1-800-854-3633 in the United States and Canada.****** Outside the United States and Canada, contact your local Beckman Coulter******representative. Quantity in Commerce 851 units total for all products (291 units in US) Distribution Affected geographies: Algeria, Austria, Australia, Bangladesh, Bulgaria, Brazil, Canada, China, Costa Rica, Cyprus, Colombia, Czech Republic, El Salvador, France, Germany, Ghana, Greece, Guatemala, India, Indonesia, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Lithuania, Malawi, Mexico, Morocco, Myanmar, Nepal, Oman, Panama, Philippines, Pakistan, Paraguay, Poland, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Serbia, South Africa, South Korea, Spain, Switzerland, Tunisia, Turkey, United Kingdom, United States, Vietnam, Zimbabwe Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action An Urgent Medical Device Recall letter was sent on 3/7/17 to inform customers that Beckman Coulter has determined through internal investigation that a circuit board that controls motor functions in the instrument systems listed above can affect device positioning on the pipettor and incubator belt while the instrument is running. Customers are informed of the impact, action to be taken, and the resolution. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions are instructed to contact Customer Technical Support******Center:****** From our website: http://www.beckmancoulter.com****** By phone: call 1-800-854-3633 in the United States and Canada.****** Outside the United States and Canada, contact your local Beckman Coulter******representative. | SoftwareDesignChange
116 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153846 | Z-1710-2017 | Access 2 assay Immunoassay System, Remanufactured, Part No. 386220******Product Usage:***The Access 2 and UniCel DxI Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. | Access 2 Immunoassay System. | 2 | CDRH | 2017-04-04 | 2017-03-15 21:02 | 2018-03-01 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access 2 Family of Immunoassay Systems because a circuit board that controls motor functions can affect device positioning on the pipettor and incubator belt while the instrument is running. The position errors cause the instrument to stop running and cancel tests that are in process. | Class 2 Device Recall Access 2 Immunoassay System. Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Access 2 Immunoassay System. FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Access 2 Immunoassay System. Date Initiated by Firm March 07, 2017 Date Posted March 15, 2017 Recall Status 1 Terminated 3 on March 01, 2018 Recall Number Z-1710-2017 Recall Event ID 76688 510(K)Number K922823 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product Access 2 assay Immunoassay System, Remanufactured, Part No. 386220************Product Usage:******The Access 2 and UniCel DxI Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. Code Information Serial No. 501012,501110,501190,501224,501294,501371,501394,501403,501440, 501443,501490,501501,501673,501681,501693,501861,501920,501926, 501936,501970,502165,502196,502319,502331,502332,502368,502458, 502475,502506,502543,502549,502556,502574,502640,502647,502651, 502791,502843,502848,503153,503279,503313,503354,503384,503392, 503424,503543,503627,503678,504046,504071,504072,504092,504112, 504195,504322,504365,504370,504756,504767,504902,505042,505195, 505251,505254,505644,505795,505840,505947,506002,506026,506108, 506119,506153,506187,506371,506406,506560,506586,506664,506708, 506766,506786,506837,506846,507039,507120,507161,507162,507214, 507229,507248,507282,507353,507397,507480,507511,507604,507619, 507842,507929,508019,508096,508179,508227,508313,508334,508412, 508478,508509,508516,508767,509051,509087,509113,509164,509253, 509297,509632,509837,510049,511042,511047,511180,511383,700024, 800321,800341,800443,800608,800611,800650,800849 Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 Manufacturer Reason for Recall Beckman Coulter is recalling the Access 2 Family of Immunoassay Systems because a circuit board that controls motor functions can affect device positioning on the pipettor and incubator belt while the instrument is running. The position errors cause the instrument to stop running and cancel tests that are in process. FDA Determined Cause 2 Software Design Change Action An Urgent Medical Device Recall letter was sent on 3/7/17 to inform customers that Beckman Coulter has determined through internal investigation that a circuit board that controls motor functions in the instrument systems listed above can affect device positioning on the pipettor and incubator belt while the instrument is running. Customers are informed of the impact, action to be taken, and the resolution. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions are instructed to contact Customer Technical Support******Center:****** From our website: http://www.beckmancoulter.com****** By phone: call 1-800-854-3633 in the United States and Canada.****** Outside the United States and Canada, contact your local Beckman Coulter******representative. Quantity in Commerce 851 units total for all products (291 units in US) Distribution Affected geographies: Algeria, Austria, Australia, Bangladesh, Bulgaria, Brazil, Canada, China, Costa Rica, Cyprus, Colombia, Czech Republic, El Salvador, France, Germany, Ghana, Greece, Guatemala, India, Indonesia, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Lithuania, Malawi, Mexico, Morocco, Myanmar, Nepal, Oman, Panama, Philippines, Pakistan, Paraguay, Poland, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Serbia, South Africa, South Korea, Spain, Switzerland, Tunisia, Turkey, United Kingdom, United States, Vietnam, Zimbabwe Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action An Urgent Medical Device Recall letter was sent on 3/7/17 to inform customers that Beckman Coulter has determined through internal investigation that a circuit board that controls motor functions in the instrument systems listed above can affect device positioning on the pipettor and incubator belt while the instrument is running. Customers are informed of the impact, action to be taken, and the resolution. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions are instructed to contact Customer Technical Support******Center:****** From our website: http://www.beckmancoulter.com****** By phone: call 1-800-854-3633 in the United States and Canada.****** Outside the United States and Canada, contact your local Beckman Coulter******representative. | SoftwareDesignChange
117 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152742 | Z-1681-2017 | Merge Eye Care Systems, consisting of Merge Eye Station, Merge Eye Care PACS, and Merge Eye Station Import Utility (ESIU). | Merge Eye Care Systems | 2 | CDRH | 2017-03-24 | 2017-03-13 21:00 | 2020-06-22 17:51 | Merge Healthcare, Inc. | The system will completely stop allowing the downloading or acquisition of new images due to incompatible software application versions. | Class 2 Device Recall Merge Eye Care Systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Merge Eye Care Systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Merge Eye Care Systems Date Initiated by Firm April 04, 2016 Date Posted March 13, 2017 Recall Status 1 Terminated 3 on June 22, 2020 Recall Number Z-1681-2017 Recall Event ID 76309 510(K)Number K110006 K913929 Product Classification System, image management, ophthalmic - Product Code NFJ Product Merge Eye Care Systems, consisting of Merge Eye Station, Merge Eye Care PACS, and Merge Eye Station Import Utility (ESIU). Code Information All versions of Eye Care PACS, Eye Station, and ESIU. Recalling Firm/ Manufacturer Merge Healthcare, Inc. 900 Walnut Ridge Dr Hartland WI 53029-8347 For Additional Information Contact Merge Service and Support 877-741-5369 Manufacturer Reason for Recall The system will completely stop allowing the downloading or acquisition of new images due to incompatible software application versions. FDA Determined Cause 2 Software Design Change Action The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. Quantity in Commerce 1,627 sites potentially have the affected software Distribution Distribution was made to medical facilities nationwide. Government distribution was made. Foreign distribution was made to Canada, as well as other countries. There was no military distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = NFJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. | SoftwareDesignChange
118 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152040 | Z-1055-2017 | Digital RID Plate Reader and Software******Product Code: AD400 | Digital RID Plate Reader | 2 | CDRH | 2017-01-19 | 2017-01-19 21:00 | 2017-04-04 | The Binding Site Group, Ltd. | If a control ring is marked after reading, the software will not flag results that are out of the specified QC range. | Class 2 Device Recall Digital RID Plate Reader Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Digital RID Plate Reader FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Digital RID Plate Reader Date Initiated by Firm October 09, 2012 Date Posted January 19, 2017 Recall Status 1 Terminated 3 on April 04, 2017 Recall Number Z-1055-2017 Recall Event ID 76043 Product Classification Plates and equipment, radial immunodiffusion - Product Code JZQ Product Digital RID Plate Reader and Software************Product Code: AD400 Code Information UDI 05051700000376 Recalling Firm/ Manufacturer The Binding Site Group, Ltd. 8 Calthorpe Road Birmingham United Kingdom For Additional Information Contact Germaine Schwider 858-4539177 Manufacturer Reason for Recall If a control ring is marked after reading, the software will not flag results that are out of the specified QC range. FDA Determined Cause 2 Software Design Change Action On 10/2012 notification letter was sent to customers requesting they take the following actions:************1. If have not already installed software version Vl.0.3 then please continue to work with version V1.0.2 until the above issues are resolved.******2. If have already installed software version V1.0.3, when processing results always mark the controlprior to reading the ring.******3. If are uncertain about the validity of any assays you have performed using software version Vl.0.3 please review the controlresults obtained and compare them to the specified control range for the RID assay and;******"***If the control results are acceptable then your assay is valid and results will be unaffected.******"***If the control results are not acceptable your assay should be considered invalid and your patient results should be reviewed and repeated as appropriate.************On 02/2013 a second notification was sent in which customers were informed the new software was released.************Customers were asked to take the following actions;******1. Upon receipt install the new software and uninstall the previou version******2. Install following the onscreen instructions. The installation process has not changed from previous versions. The instruction guide received with the RID Plate reader is still the current version.******3. After receipt of the new software disc complete and return the Eback form to The Binding Site.************For further assistance contact Technical Services at The Binding Site via the email address: technicalservices@bindingsite.co.uk. Distribution U.S. distribution to the following; MA, GA, NY, IA, CA, IL, UT, NJ, NC, MD, MN.************Foreign distribution to the following; UK, Germany, Spain, France. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: |  | SoftwareDesignChange
119 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151856 | Z-0991-2017 | ConFORMIS ITOTAL PS-IPOLYIMPLANT KIT-RIGHT***Catalog Number:TPS-111-1111***The iTotal¿ PS Knee Replacement System is intended for use as a total knee replacement in patients with knee joint pain and disability whose conditions cannot be solely addressed by the use of a prosthetic device that treats only one or two of the three knee compartments, such as a unicondylar, patellofemoral or bicompartmental prosthesis | The iTotal¿ PS Knee Replacement System | 2 | CDRH | 2017-01-10 | 2017-01-10 21:00 | 2017-01-11 | ConforMIS, Inc. | iView report displayed femoral cut measurements in the incorrect unit of measure (inches vs. mm). | Class 2 Device Recall The iTotal PS Knee Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall The iTotal PS Knee Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall The iTotal PS Knee Replacement System Date Initiated by Firm October 31, 2016 Date Posted January 10, 2017 Recall Status 1 Terminated 3 on January 11, 2017 Recall Number Z-0991-2017 Recall Event ID 75969 510(K)Number K161668 Product Classification Prosthesis, knee, patellofemorotibial, semi-constrained, cemented, polymer/metal/polymer - Product Code JWH Product ConFORMIS ITOTAL PS-IPOLYIMPLANT KIT-RIGHT******Catalog Number:TPS-111-1111******The iTotal PS Knee Replacement System is intended for use as a total knee replacement in patients with knee joint pain and disability whose conditions cannot be solely addressed by the use of a prosthetic device that treats only one or two of the three knee compartments, such as a unicondylar, patellofemoral or bicompartmental prosthesis Code Information Serial Number: 0384180 Recalling Firm/ Manufacturer ConforMIS, Inc. 28 Crosby Dr Bedford MA 01730-1402 For Additional Information Contact SAME 781-345-9001 Manufacturer Reason for Recall iView report displayed femoral cut measurements in the incorrect unit of measure (inches vs. mm). FDA Determined Cause 2 Software Design Change Action ConForMIS intiated the recall by telephone on October 31, 2016, to the Sales Rep to request return of the kit.******For questions regarding this recall call 781-345-9001. Quantity in Commerce 1 unit Distribution Distribution to TX only Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JWH - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action ConForMIS intiated the recall by telephone on October 31, 2016, to the Sales Rep to request return of the kit.******For questions regarding this recall call 781-345-9001. | SoftwareDesignChange
120 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151797 | Z-0962-2017 | Amplia MRI" CRT -D SureScan", ***Amplia MRI" Quad CRT -D SureScan", ***Claria MRI" CRT -D SureScan", ***Claria MRI" Quad CRT -D SureScan. implantable medical devices that automatically detect and treat episodes of ventricular fibrillation, ventricular tachycardia, fast ventricular tachycardia, and bradyarrhythmia. | AdaptivCRT (aCRT) | 2 | CDRH | 2017-01-07 |  | 2019-07-16 16:26 | Medtronic Inc., Cardiac Rhythm and Heart Failure | During internal testing conducted as part of next generation product development, it was discovered that the AdaptivCRT (aCRT) feature of Amplia MRI and Claria MRI CRT -D devices may appear to be enabled when it is not active. | Class 2 Device Recall AdaptivCRT (aCRT) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall AdaptivCRT (aCRT) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall AdaptivCRT (aCRT) Date Initiated by Firm December 09, 2016 Create Date January 07, 2017 Recall Status 1 Terminated 3 on July 16, 2019 Recall Number Z-0962-2017 Recall Event ID 75943 PMA Number P010031S550 P010031S513 Product Classification Defibrillator, automatic implantable cardioverter, with cardiac resynchronization (CRT-D) - Product Code NIK Product Amplia MRI" CRT -D SureScan", ******Amplia MRI" Quad CRT -D SureScan", ******Claria MRI" CRT -D SureScan", ******Claria MRI" Quad CRT -D SureScan. implantable medical devices that automatically detect and treat episodes of ventricular fibrillation, ventricular tachycardia, fast ventricular tachycardia, and bradyarrhythmia. Code Information Serial Numbers of AMPLIA & CLARIA Released from Manufacturing as of 22-Nov-2016*** UPN***Serial 00643169543362***RPA200003H 00643169543362***RPA200004H 00643169543362***RPA200005H 00643169543362***RPA200006H 00643169543362***RPA200008H 00643169543362***RPA200009H 00643169543362***RPA200010H 00643169543362***RPA200011H 00643169543362***RPA200012H 00643169543362***RPA200013H 00643169543362***RPA200014H 00643169543362***RPA200015H 00643169543362***RPA200016H 00643169543362***RPA200017H 00643169543362***RPA200018H 00643169543362***RPA200019H 00643169543362***RPA200021H 00643169543362***RPA200022H 00643169543362***RPA200023H 00643169543362***RPA200024H 00643169543362***RPA200025H 00643169543362***RPA200027H 00643169543362***RPA200028H 00643169543362***RPA200029H 00643169543362***RPA200030H 00643169543362***RPA200031H 00643169543362***RPA200032H 00643169543362***RPA200033H 00643169543362***RPA200034H 00643169543362***RPA200035H 00643169543362***RPA200036H 00643169543362***RPA200037H 00643169543362***RPA200039H 00643169543362***RPA200040H 00643169543362***RPA200041H 00643169543362***RPA200042H 00643169543362***RPA200043H 00643169543362***RPA200044H 00643169543362***RPA200045H 00643169543362***RPA200046H 00643169543362***RPA200047H 00643169543362***RPA200048H 00643169543362***RPA200049H 00643169543362***RPA200050H 00643169543362***RPA200051H 00643169543362***RPA200052H 00643169543362***RPA200053H 00643169543362***RPA200054H 00643169543362***RPA200055H 00643169543362***RPA200056H 00643169543362***RPA200057H 00643169543362***RPA200058H 00643169543362***RPA200059H 00643169543362***RPA200060H 00643169543362***RPA200061H 00643169543362***RPA200062H 00643169543362***RPA200063H 00643169543362***RPA200064H 00643169543362***RPA200065H 00643169543362***RPA200066H 00643169543362***RPA200067H 00643169543362***RPA200068H 00643169543362***RPA200069H 00643169543362***RPA200071H 00643169543362***RPA200072H 00643169543362***RPA200073H 00643169543362***RPA200075H 00643169543362***RPA200076H 00643169543362***RPA200077H 00643169543362***RPA200078H 00643169543362***RPA200079H 00643169543362***RPA200080H 00643169543362***RPA200081H 00643169543362***RPA200082H 00643169543362***RPA200083H 00643169543362***RPA200084H 00643169543362***RPA200085H 00643169543362***RPA200086H 00643169543362***RPA200087H 00643169543362***RPA200088H 00643169543362***RPA200089H 00643169543362***RPA200090H 00643169543362***RPA200091H 00643169543362***RPA200092H 00643169543362***RPA200093H 00643169543362***RPA200094H 00643169543362***RPA200095H 00643169543362***RPA200096H 00643169543362***RPA200097H 00643169543362***RPA200098H 00643169543362***RPA200099H 00643169543362***RPA200100H 00643169543362***RPA200101H 00643169543362***RPA200102H 00643169543362***RPA200103H 00643169543362***RPA200104H 00643169543362***RPA200105H 00643169543362***RPA200106H 00643169543362***RPA200107H 00643169543362***RPA200108H 00643169543362***RPA200109H 00643169543362***RPA200110H 00643169543362***RPA200111H 00643169543362***RPA200112H 00643169543362***RPA200113H 00643169543362***RPA200114H 00643169543362***RPA200115H 00643169543362***RPA200116H 00643169543362***RPA200117H 00643169543362***RPA200118H 00643169543362***RPA200119H 00643169543362***RPA200120H 00643169543362***RPA200121H 00643169543362***RPA200122H 00643169543362***RPA200123H 00643169543362***RPA200124H 00643169543362***RPA200125H 00643169543362***RPA200126H 00643169543362***RPA200127H 00643169543362***RPA200128H 00643169543362***RPA200129H 00643169543362***RPA200130H 00643169543362***RPA200131H 00643169543362***RPA200132H 00643169543362***RPA200133H 00643169543362***RPA200134H 00643169543362***RPA200135H 00643169543362***RPA200136H 00643169543362***RPA200137H 00643169543362***RPA200138H 00643169543362***RPA200139H 00643169543362***RPA200140H 00643169543362***RPA200141H 00643169543362***RPA200142H 00643169543362***RPA200143H 00643169543362***RPA200144H 00643169543362***RPA200145H 00643169543362***RPA200146H 00643169543362***RPA200147H 00643169543362***RPA200148H 00643169543362***RPA200149H 00643169543362***RPA200150H 00643169543362***RPA200151H 00643169543362***RPA200152H 00643169543362***RPA200153H 00643169543362***RPA200154H 00643169543362***RPA200155H 00643169543362***RPA200156H 00643169543362***RPA200157H 00643169543362***RPA200158H 00643169543362***RPA200159H 00643169543362***RPA200160H 00643169543362***RPA200161H 00643169543362***RPA200162H 00643169543362***RPA200164H 00643169543362***RPA200165H 00643169543362***RPA200166H 00643169543362***RPA200167H 00643169543362***RPA200168H 00643169543362***RPA200169H 00643169543362***RPA200170H 00643169543362***RPA200172H 00643169543362***RPA200173H 00643169543362***RPA200174H 00643169543362***RPA200175H 00643169543362***RPA200176H 00643169543362***RPA200177H 00643169543362***RPA200178H 00643169543362***RPA200179H 00643169543362***RPA200180H 00643169543362***RPA200181H 00643169543362***RPA200182H 00643169543362***RPA200183H 00643169543362***RPA200184H 00643169543362***RPA200185H 00643169543362***RPA200186H 00643169543362***RPA200187H 00643169543362***RPA200188H 00643169543362***RPA200189H 00643169543362***RPA200190H 00643169543362***RPA200191H 00643169543362***RPA200192H 00643169543362***RPA200193H 00643169543362***RPA200194H 00643169543362***RPA200195H 00643169543362***RPA200196H 00643169543362***RPA200197H 00643169543362***RPA200198H 00643169543362***RPA200199H 00643169543362***RPA200200H 00643169543362***RPA200201H 00643169543362***RPA200202H 00643169543362***RPA200203H 00643169543362***RPA200204H 00643169543362***RPA200205H 00643169543362***RPA200206H 00643169543362***RPA200207H 00643169543362***RPA200208H 00643169543362***RPA200209H 00643169543362***RPA200210H 00643169543362***RPA200211H 00643169543362***RPA200212H 00643169543362***RPA200213H 00643169543362***RPA200214H 00643169543362***RPA200215H 00643169543362***RPA200216H 00643169543362***RPA200217H 00643169543362***RPA200218H 00643169543362***RPA200219H 00643169543362***RPA200220H 00643169543362***RPA200221H 00643169543362***RPA200222H 00643169543362***RPA200223H 00643169543362***RPA200224H 00643169543362***RPA200225H 00643169543362***RPA200226H 00643169543362***RPA200227H 00643169543362***RPA200228H 00643169543362***RPA200230H 00643169543362***RPA200231H 00643169543362***RPA200232H 00643169543362***RPA200233H 00643169543362***RPA200234H 00643169543362***RPA200235H 00643169543362***RPA200236H 00643169543362***RPA200237H 00643169543362***RPA200238H 00643169543362***RPA200239H 00643169543362***RPA200240H 00643169543362***RPA200241H 00643169543362***RPA200242H 00643169543362***RPA200243H 00643169543362***RPA200244H 00643169543362***RPA200245H 00643169543362***RPA200246H 00643169543362***RPA200247H 00643169543362***RPA200248H 00643169543362***RPA200249H 00643169543362***RPA200250H 00643169543362***RPA200251H 00643169543362***RPA200252H 00643169543362***RPA200253H 00643169543362***RPA200254H 00643169543362***RPA200258H 00643169543362***RPA200259H 00643169543362***RPA200260H 00643169543362***RPA200261H 00643169543362***RPA200262H 00643169543362***RPA200263H 00643169543362***RPA200264H 00643169543362***RPA200265H 00643169543362***RPA200266H 00643169543362***RPA200268H 00643169543362***RPA200269H 00643169543362***RPA200270H 00643169543362***RPA200271H 00643169543362***RPA200272H 00643169543362***RPA200273H 00643169543362***RPA200275H 00643169543362***RPA200276H 00643169543362***RPA200277H 00643169543362***RPA200278H 00643169543362***RPA200279H 00643169543362***RPA200280H 00643169543362***RPA200281H 00643169543362***RPA200282H 00643169543362***RPA200284H 00643169543362***RPA200285H 00643169543362***RPA200287H 00643169543362***RPA200291H 00643169543362***RPA200297H 00643169543362***RPA200298H 00643169543362***RPA200299H 00643169543362***RPA200300H 00643169543362***RPA200301H 00643169543362***RPA200302H 00643169543362***RPA200303H 00643169543362***RPA200304H 00643169543362***RPA200305H 00643169543362***RPA200306H 00643169543362***RPA200307H 00643169543362***RPA200309H 00643169543362***RPA200310H 00643169543362***RPA200311H 00643169543362***RPA200312H 00643169543362***RPA200313H 00643169543362***RPA200314H 00643169543362***RPA200315H 00643169543362***RPA200316H 00643169543362***RPA200317H 00643169543362***RPA200318H 00643169543362***RPA200319H 00643169543362***RPA200320H 00643169543362***RPA200321H 00643169543362***RPA200322H 00643169543362***RPA200324H 00643169543362***RPA200325H 00643169543362***RPA200326H 00643169543362***RPA200327H 00643169543362***RPA200328H 00643169543362***RPA200329H 00643169543362***RPA200330H 00643169543362***RPA200331H 00643169543362***RPA200332H 00643169543362***RPA200333H 00643169543362***RPA200334H 00643169543362***RPA200335H 00643169543362***RPA200336H 00643169543362***RPA200337H 00643169543362***RPA200338H 00643169543362***RPA200339H 00643169543362***RPA200341H 00643169543362***RPA200343H 00643169543362***RPA200344H 00643169543362***RPA200345H 00643169543362***RPA200346H 00643169543362***RPA200348H 00643169543362***RPA200349H 00643169543362***RPA200350H 00643169543362***RPA200351H 00643169543362***RPA200352H 00643169543362***RPA200353H 00643169543362***RPA200354H 00643169543362***RPA200355H 00643169543362***RPA200356H 00643169543362***RPA200357H 00643169543362***RPA200358H 00643169543362***RPA200359H 00643169543362***RPA200360H 00643169543362***RPA200361H 00643169543362***RPA200362H 00643169543362***RPA200363H 00643169543362***RPA200364H 00643169543362***RPA200365H 00643169543362***RPA200366H 00643169543362***RPA200367H 00643169543362***RPA200368H 00643169543362***RPA200369H 00643169543362***RPA200372H 00643169543362***RPA200373H 00643169543362***RPA200374H 00643169543362***RPA200375H 00643169543362***RPA200376H 00643169543362***RPA200377H 00643169543362***RPA200378H 00643169543362***RPA200379H 00643169543362***RPA200380H 00643169543362***RPA200381H 00643169543362***RPA200382H 00643169543362***RPA200383H 00643169543362***RPA200384H 00643169543362***RPA200385H 00643169543362***RPA200386H 00643169543362***RPA200387H 00643169543362***RPA200388H 00643169543362***RPA200389H 00643169543362***RPA200390H 00643169543362***RPA200391H 00643169543362***RPA200392H 00643169543362***RPA200393H 00643169543362***RPA200394H 00643169543362***RPA200395H 00643169543362***RPA200396H 00643169543362***RPA200397H 00643169543362***RPA200398H 00643169543362***RPA200399H 00643169543362***RPA200400H 00643169543362***RPA200401H 00643169543362***RPA200402H 00643169543362***RPA200403H 00643169543362***RPA200404H 00643169543362***RPA200405H 00643169543362***RPA200406H 00643169543362***RPA200407H 00643169543362***RPA200408H 00643169543362***RPA200409H 00643169543362***RPA200410H 00643169543362***RPA200411H 00643169543362***RPA200412H 00643169543362***RPA200413H 00643169543362***RPA200414H 00643169543362***RPA200415H 00643169543362***RPA200416H 00643169543362***RPA200417H 00643169543362***RPA200418H 00643169543362***RPA200419H 00643169543362***RPA200420H 00643169543362***RPA200421H 00643169543362***RPA200422H 00643169543362***RPA200423H 00643169543362***RPA200425H 00643169543362***RPA200426H 00643169543362***RPA200427H 00643169543362***RPA200428H 00643169543362***RPA200429H 00643169543362***RPA200430H 00643169543362***RPA200431H 00643169543362***RPA200432H 00643169543362***RPA200433H 00643169543362***RPA200434H 00643169543362***RPA200435H 00643169543362***RPA200436H 00643169543362***RPA200438H 00643169543362***RPA200439H 00643169543362***RPA200440H 00643169543362***RPA200441H 00643169543362***RPA200442H 00643169543362***RPA200443H 00643169543362***RPA200445H 00643169543362***RPA200446H 00643169543362***RPA200447H 00643169543362***RPA200448H 00643169543362***RPA200450H 00643169543362***RPA200452H 00643169543362***RPA200453H 00643169543362***RPA200454H 00643169543362***RPA200455H 00643169543362***RPA200456H 00643169543362***RPA200457H 00643169543362***RPA200458H 00643169543362***RPA200459H 00643169543362***RPA200460H 00643169543362***RPA200461H 00643169543362***RPA200462H 00643169543362***RPA200463H 00643169543362***RPA200464H 00643169543362***RPA200465H 00643169543362***RPA200466H 00643169543362***RPA200467H 00643169543362***RPA200469H 00643169543362***RPA200470H 00643169543362***RPA200471H 00643169543362***RPA200472H 00643169543362***RPA200473H 00643169543362***RPA200474H 00643169543362***RPA200475H 00643169543362***RPA200476H 00643169543362***RPA200477H 00643169543362***RPA200478H 00643169543362***RPA200479H 00643169543362***RPA200480H******00643169543362***RPA200481H 00643169543362***RPA200482H 00643169543362***RPA200484H 00643169543362***RPA200485H 00643169543362***RPA200486H 00643169543362***RPA200487H 00643169543362***RPA200488H 00643169543362***RPA200489H 00643169543362***RPA200490H 00643169543362***RPA200491H 00643169543362***RPA200492H 00643169543362***RPA200493H 00643169543362***RPA200494H 00643169543362***RPA200496H 00643169543362***RPA200497H 00643169543362***RPA200498H 00643169543362***RPA200499H 00643169543362***RPA200500H 00643169543362***RPA200501H 00643169543362***RPA200503H 00643169543362***RPA200504H 00643169543362***RPA200507H 00643169543362***RPA200509H 00643169543362***RPA200510H 00643169543362***RPA200511H 00643169543362***RPA200512H 00643169543362***RPA200513H 00643169543362***RPA200514H 00643169543362***RPA200515H 00643169543362***RPA200516H 00643169543362***RPA200517H 00643169543362***RPA200518H 00643169543362***RPA200519H 00643169543362***RPA200520H 00643169543362***RPA200522H 00643169543362***RPA200523H 00643169543362***RPA200524H 00643169543362***RPA200525H 00643169543362***RPA200526H 00643169543362***RPA200528H 00643169543362***RPA200529H 00643169543362***RPA200530H 00643169543362***RPA200531H 00643169543362***RPA200533H 00643169543362***RPA200534H 00643169543362***RPA200536H 00643169543362***RPA200537H 00643169543362***RPA200538H 00643169543362***RPA200539H 00643169543362***RPA200540H 00643169543362***RPA200541H 00643169543362***RPA200542H 00643169543362***RPA200543H 00643169543362***RPA200544H 00643169543362***RPA200545H 00643169543362***RPA200546H 00643169543362***RPA200547H 00643169543362***RPA200548H 00643169543362***RPA200550H 00643169543362***RPA200551H 00643169543362***RPA200552H 00643169543362***RPA200553H 00643169543362***RPA200554H 00643169543362***RPA200555H 00643169543362***RPA200556H 00643169543362***RPA200557H 00643169543362***RPA200558H 00643169543362***RPA200559H 00643169543362***RPA200560H 00643169543362***RPA200561H 00643169543362***RPA200562H 00643169543362***RPA200563H 00643169543362***RPA200564H 00643169543362***RPA200565H 00643169543362***RPA200566H 00643169543362***RPA200567H 00643169543362***RPA200568H 00643169543362***RPA200569H 00643169543362***RPA200570H 00643169543362***RPA200571H 00643169543362***RPA200572H 00643169543362***RPA200573H 00643169543362***RPA200574H 00643169543362***RPA200575H 00643169543362***RPA200576H 00643169543362***RPA200577H 00643169543362***RPA200579H 00643169543362***RPA200580H 00643169543362***RPA200581H 00643169543362***RPA200582H 00643169543362***RPA200583H 00643169543362***RPA200584H 00643169543362***RPA200585H 00643169543362***RPA200587H 00643169543362***RPA200588H 00643169543362***RPA200589H 00643169543362***RPA200590H 00643169543362***RPA200591H 00643169543362***RPA200592H 00643169543362***RPA200594H 00643169543362***RPA200595H 00643169543362***RPA200596H 00643169543362***RPA200597H 00643169543362***RPA200598H 00643169543362***RPA200599H 00643169543362***RPA200600H 00643169543362***RPA200601H 00643169543362***RPA200602H 00643169543362***RPA200603H 00643169543362***RPA200604H 00643169543362***RPA200605H 00643169543362***RPA200606H 00643169543362***RPA200607H 00643169543362***RPA200608H 00643169543362***RPA200609H 00643169543362***RPA200610H 00643169543362***RPA200611H 00643169543362***RPA200612H 00643169543362***RPA200613H 00643169543362***RPA200614H 00643169543362***RPA200615H 00643169543362***RPA200616H 00643169543362***RPA200617H 00643169543362***RPA200618H 00643169543362***RPA200619H 00643169543362***RPA200620H 00643169543362***RPA200621H 00643169543362***RPA200622H 00643169543362***RPA200623H 00643169543362***RPA200624H 00643169543362***RPA200625H 00643169543362***RPA200626H 00643169543362***RPA200627H 00643169543362***RPA200628H 00643169543362***RPA200629H 00643169543362***RPA200630H 00643169543362***RPA200631H 00643169543362***RPA200632H 00643169543362***RPA200633H 00643169543362***RPA200634H 00643169543362***RPA200635H 00643169543362***RPA200636H 00643169543362***RPA200637H 00643169543362***RPA200638H 00643169543362***RPA200639H 00643169543362***RPA200640H 00643169543362***RPA200641H 00643169543362***RPA200642H 00643169543362***RPA200643H 00643169543362***RPA200644H 00643169543362***RPA200645H 00643169543362***RPA200646H 00643169543362***RPA200647H 00643169543362***RPA200648H 00643169543362***RPA200649H 00643169543362***RPA200651H 00643169543362***RPA200652H 00643169543362***RPA200653H 00643169543362***RPA200654H 00643169543362***RPA200655H 00643169543362***RPA200656H 00643169543362***RPA200657H 00643169543362***RPA200658H 00643169543362***RPA200660H 00643169543362***RPA200662H 00643169543362***RPA200663H 00643169543362***RPA200665H 00643169543362***RPA200666H 00643169543362***RPA200667H 00643169543362***RPA200668H 00643169543362***RPA200669H 00643169543362***RPA200670H 00643169543362***RPA200671H 00643169543362***RPA200672H 00643169543362***RPA200673H 00643169543362***RPA200674H 00643169543362***RPA200675H 00643169543362***RPA200676H 00643169543362***RPA200680H 00643169543362***RPA200681H 00643169543362***RPA200682H 00643169543362***RPA200683H 00643169543362***RPA200685H 00643169543362***RPA200686H 00643169543362***RPA200688H 00643169543362***RPA200689H 00643169543362***RPA200690H 00643169543362***RPA200691H 00643169543362***RPA200692H 00643169543362***RPA200693H 00643169543362***RPA200694H 00643169543362***RPA200696H 00643169543362***RPA200697H 00643169543362***RPA200698H 00643169543362***RPA200700H 00643169543362***RPA200701H 00643169543362***RPA200702H 00643169543362***RPA200703H 00643169543362***RPA200704H 00643169543362***RPA200705H 00643169543362***RPA200706H 00643169543362***RPA200707H 00643169543362***RPA200708H 00643169543362***RPA200709H 00643169543362***RPA200710H 00643169543362***RPA200711H 00643169543362***RPA200712H 00643169543362***RPA200713H 00643169543362***RPA200714H 00643169543362***RPA200715H 00643169543362***RPA200716H 00643169543362***RPA200717H 00643169543362***RPA200718H 00643169543362***RPA200719H 00643169543362***RPA200720H 00643169543362***RPA200721H 00643169543362***RPA200723H 00643169543362***RPA200724H 00643169543362***RPA200725H 00643169543362***RPA200726H 00643169543362***RPA200727H 00643169543362***RPA200728H 00643169543362***RPA200730H 00643169543362***RPA200731H 00643169543362***RPA200732H 00643169543362***RPA200733H 00643169543362***RPA200734H 00643169543362***RPA200735H 00643169543362***RPA200736H 00643169543362***RPA200737H 00643169543362***RPA200741H 00643169543362***RPA200742H 00643169543362***RPA200744H 00643169543362***RPA200746H 00643169543362***RPA200747H 00643169543362***RPA200748H 00643169543362***RPA200749H 00643169543362***RPA200750H 00643169543362***RPA200751H 00643169543362***RPA200752H 00643169543362***RPA200753H 00643169543362***RPA200754H 00643169543362***RPA200755H 00643169543362***RPA200756H 00643169543362***RPA200757H 00643169543362***RPA200758H 00643169543362***RPA200759H 00643169543362***RPA200760H 00643169543362***RPA200761H 00643169543362***RPA200762H 00643169543362***RPA200763H 00643169543362***RPA200764H 00643169543362***RPA200765H 00643169543362***RPA200766H 00643169543362***RPA200767H 00643169543362***RPA200768H 00643169543362***RPA200769H 00643169543362***RPA200770H 00643169543362***RPA200771H 00643169543362***RPA200772H 00643169543362***RPA200773H 00643169543362***RPA200774H 00643169543362***RPA200775H 00643169543362***RPA200776H 00643169543362***RPA200777H 00643169543362***RPA200778H 00643169543362***RPA200779H 00643169543362***RPA200780H 00643169543362***RPA200781H 00643169543362***RPA200782H 00643169543362***RPA200783H 00643169543362***RPA200784H 00643169543362***RPA200785H 00643169543362***RPA200787H 00643169543362***RPA200788H 00643169543362***RPA200789H 00643169543362***RPA200790H 00643169543362***RPA200791H 00643169543362***RPA200792H 00643169543362***RPA200795H 00643169543362***RPA200796H 00643169543362***RPA200798H 00643169543362***RPA200800H 00643169543362***RPA200801H 00643169543362***RPA200806H 00643169543362***RPA200810H 00643169543362***RPA200827H 00643169543362***RPA200839H 00643169543362***RPA200867H 00643169543362***RPA200870H 00643169543362***RPA200913H 00643169543362***RPA200923H 00643169543379***RPB600086S 00643169543379***RPB600087S 00643169543379***RPB600088S 00643169543379***RPB600089S 00643169543379***RPB600090S 00643169543379***RPB600091S 00643169543379***RPB600092S 00643169543379***RPB600093S 00643169543379***RPB600094S 00643169543379***RPB600095S 00643169543379***RPB600096S 00643169543379***RPB600097S 00643169543379***RPB600098S 00643169543379***RPB600099S 00643169543379***RPB600100S 00643169543379***RPB600101S 00643169543379***RPB600102S 00643169543379***RPB600103S 00643169543379***RPB600104S 00643169543379***RPB600105S 00643169543379***RPB600106S 00643169543379***RPB600107S 00643169543379***RPB600108S 00643169543379***RPB600109S 00643169543379***RPB600110S 00643169543379***RPB600111S 00643169543379***RPB600112S 00643169543379***RPB600113S 00643169543379***RPB600114S 00643169543379***RPB600115S 00643169543379***RPB600116S 00643169543379***RPB600117S 00643169543379***RPB600118S 00643169543379***RPB600119S 00643169543379***RPB600120S 00643169543379***RPB600121S 00643169543379***RPB600122S 00643169543379***RPB600123S 00643169543379***RPB600124S 00643169543379***RPB600125S 00643169543379***RPB600126S 00643169543379***RPB600127S 00643169543379***RPB600128S 00643169543379***RPB600129S 00643169543379***RPB600130S 00643169543379***RPB600131S 00643169543379***RPB600132S 00643169543379***RPB600133S 00643169543379***RPB600134S 00643169543379***RPB600135S 00643169543379***RPB600136S 00643169543379***RPB600137S 00643169543379***RPB600138S 00643169543379***RPB600139S 00643169543379***RPB600140S 00643169543379***RPB600141S 00643169543379***RPB600142S 00643169543379***RPB600143S 00643169543379***RPB600144S 00643169543379***RPB600145S 00643169543379***RPB600146S 00643169543379***RPB600147S 00643169543379***RPB600148S 00643169543379***RPB600149S 00643169543379***RPB600150S 00643169543379***RPB600151S 00643169543379***RPB600152S 00643169543379***RPB600153S 00643169543379***RPB600154S 00643169543379***RPB600155S 00643169543379***RPB600156S 00643169543379***RPB600157S 00643169543379***RPB600158S 00643169543379***RPB600159S 00643169543379***RPB600160S 00643169543379***RPB600161S 00643169543379***RPB600162S 00643169543379***RPB600163S 00643169543379***RPB600164S 00643169543379***RPB600165S 00643169543379***RPB600166S 00643169543379***RPB600167S 00643169543379***RPB600168S 00643169543379***RPB600169S 00643169543379***RPB600170S 00643169543379***RPB600171S 00643169543379***RPB600172S 00643169543379***RPB600173S 00643169543379***RPB600174S 00643169543379***RPB600175S 00643169543379***RPB600176S 00643169543379***RPB600177S 00643169543379***RPB600178S 00643169543379***RPB600179S 00643169543379***RPB600180S 00643169543379***RPB600181S 00643169543379***RPB600182S 00643169543379***RPB600183S 00643169543379***RPB600184S 00643169543379***RPB600185S 00643169543379***RPB600186S 00643169543379***RPB600187S 00643169543379***RPB600188S 00643169543379***RPB600189S 00643169543379***RPB600190S 00643169543379***RPB600191S 00643169543379***RPB600192S 00643169543379***RPB600193S 00643169543379***RPB600194S 00643169543379***RPB600195S 00643169543379***RPB600196S 00643169543379***RPB600197S 00643169543379***RPB600198S 00643169543379***RPB600199S 00643169543379***RPB600200S 00643169543379***RPB600201S 00643169543379***RPB600202S 00643169543379***RPB600204S 00643169543379***RPB600205S 00643169543379***RPB600206S 00643169543379***RPB600207S 00643169543379***RPB600208S 00643169543379***RPB600209S 00643169543379***RPB600210S 00643169543379***RPB600211S 00643169543379***RPB600212S 00643169543379***RPB600213S 00643169543379***RPB600214S 00643169543379***RPB600215S 00643169543379***RPB600216S 00643169543379***RPB600217S 00643169543379***RPB600218S 00643169543379***RPB600219S 00643169543379***RPB600220S 00643169543379***RPB600221S 00643169543379***RPB600222S 00643169543379***RPB600223S 00643169543379***RPB600224S 00643169543379***RPB600225S 00643169543379***RPB600226S 00643169543379***RPB600227S 00643169543379***RPB600228S 00643169543379***RPB600229S 00643169543379***RPB600230S 00643169543379***RPB600231S 00643169543379***RPB600232S 00643169543379***RPB600233S 00643169543379***RPB600234S 00643169543379***RPB600235S 00643169543379***RPB600236S 00643169543379***RPB600237S 00643169543379***RPB600238S 00643169543379***RPB600239S 00643169543379***RPB600240S 00643169543379***RPB600241S 00643169543379***RPB600242S 00643169543379***RPB600243S 00643169543379***RPB600244S 00643169543379***RPB600245S 00643169543379***RPB600246S 00643169543379***RPB600247S 00643169543379***RPB600248S 00643169543379***RPB600249S 00643169543379***RPB600250S 00643169543379***RPB600251S 00643169543379***RPB600252S 00643169543379***RPB600253S 00643169543379***RPB600254S 00643169543379***RPB600255S 00643169543379***RPB600256S 00643169543379***RPB600257S 00643169543379***RPB600258S 00643169543379***RPB600259S 00643169543379***RPB600260S 00643169543379***RPB600261S 00643169543379***RPB600262S 00643169543379***RPB600263S 00643169543379***RPB600264S 00643169543379***RPB600265S 00643169543379***RPB600266S 00643169543379***RPB600267S 00643169543379***RPB600268S 00643169543379***RPB600269S 00643169543379***RPB600270S 00643169543379***RPB600271S 00643169543379***RPB600272S 00643169543379***RPB600273S 00643169543379***RPB600274S 00643169543379***RPB600275S 00643169543379***RPB600276S 00643169543379***RPB600277S 00643169543379***RPB600278S 00643169543379***RPB600279S 00643169543379***RPB600280S 00643169543379***RPB600282S 00643169543379***RPB600283S 00643169543379***RPB600284S 00643169543379***RPB600285S 00643169543379***RPB600286S 00643169543379***RPB600287S 00643169543379***RPB600288S 00643169543379***RPB600289S 00643169543379***RPB600290S 00643169543379***RPB600291S 00643169543379***RPB600292S 00643169543379***RPB600293S 00643169543379***RPB600294S 00643169543379***RPB600295S 00643169543379***RPB600296S 00643169543379***RPB600297S 00643169543379***RPB600298S 00643169543379***RPB600299S 00643169543379***RPB600300S 00643169543379***RPB600301S 00643169543379***RPB600302S 00643169543379***RPB600303S 00643169543379***RPB600304S 00643169543379***RPB600305S 00643169543379***RPB600306S 00643169543379***RPB600307S 00643169543379***RPB600308S 00643169543379***RPB600309S 00643169543379***RPB600310S 00643169543379***RPB600311S 00643169543379***RPB600312S 00643169543379***RPB600313S 00643169543379***RPB600314S 00643169543379***RPB600315S 00643169543379***RPB600316S 00643169543379***RPB600317S 00643169543379***RPB600318S 00643169543379***RPB600319S 00643169543379***RPB600320S 00643169543379***RPB600321S 00643169543379***RPB600322S 00643169543379***RPB600323S 00643169543379***RPB600324S 00643169543379***RPB600325S 00643169543379***RPB600326S 00643169543379***RPB600327S 00643169543379***RPB600328S 00643169543379***RPB600329S 00643169543379***RPB600330S 00643169543379***RPB600331S 00643169543379***RPB600332S 00643169543379***RPB600333S 00643169543379***RPB600334S 00643169543379***RPB600335S 00643169543379***RPB600336S 00643169543379***RPB600337S 00643169543379***RPB600338S 00643169543379***RPB600339S 00643169543379***RPB600340S 00643169543379***RPB600341S 00643169543379***RPB600342S 00643169543379***RPB600343S 00643169543379***RPB600344S 00643169543379***RPB600345S 00643169543379***RPB600346S 00643169543379***RPB600347S 00643169543379***RPB600348S 00643169543379***RPB600349S 00643169543379***RPB600350S 00643169543379***RPB600351S 00643169543379***RPB600352S 00643169543379***RPB600353S 00643169543379***RPB600354S 00643169543379***RPB600355S 00643169543379***RPB600356S 00643169543379***RPB600357S 00643169543379***RPB600358S 00643169543379***RPB600359S 00643169543379***RPB600360S 00643169543379***RPB600361S 00643169543379***RPB600362S 00643169543379***RPB600363S 00643169543379***RPB600364S 00643169543379***RPB600365S 00643169543379***RPB600366S 00643169543379***RPB600367S 00643169543379***RPB600368S 00643169543379***RPB600369S 00643169543379***RPB600370S 00643169543379***RPB600371S 00643169543379***RPB600372S 00643169543379***RPB600373S 00643169543379***RPB600374S 00643169543379***RPB600375S 00643169543379***RPB600376S 00643169543379***RPB600377S 00643169543379***RPB600378S 00643169543379***RPB600379S 00643169543379***RPB600380S 00643169543379***RPB600381S 00643169543379***RPB600382S 00643169543379***RPB600383S 00643169543379***RPB600384S 00643169543379***RPB600385S 00643169543379***RPB600386S 00643169543379***RPB600387S 00643169543379***RPB600388S 00643169543379***RPB600389S 00643169543379***RPB600390S 00643169543379***RPB600391S 00643169543379***RPB600392S 00643169543379***RPB600393S 00643169543379***RPB600394S 00643169543379***RPB600395S 00643169543379***RPB600396S 00643169543379***RPB600397S 00643169543379***RPB600398S 00643169543379***RPB600399S 00643169543379***RPB600400S 00643169543379***RPB600401S 00643169543379***RPB600402S 00643169543379***RPB600403S 00643169543379***RPB600404S 00643169543379***RPB600405S 00643169543379***RPB600406S 00643169543379***RPB600407S 00643169543379***RPB600408S 00643169543379***RPB600409S 00643169543379***RPB600410S 00643169543379***RPB600411S 00643169543379***RPB600412S 00643169543379***RPB600413S 00643169543379***RPB600414S 00643169543379***RPB600415S 00643169543379***RPB600416S 00643169543379***RPB600417S 00643169543379***RPB600418S 00643169543379***RPB600419S 00643169543379***RPB600420S 00643169543379***RPB600421S 00643169543379***RPB600422S 00643169543379***RPB600423S 00643169543379***RPB600424S 00643169543379***RPB600425S 00643169543379***RPB600426S 00643169543379***RPB600427S 00643169543379***RPB600428S 00643169543379***RPB600429S 00643169543379***RPB600430S 00643169543379***RPB600431S 00643169543379***RPB600432S 00643169543379***RPB600433S 00643169543379***RPB600434S 00643169543379***RPB600435S 00643169543379***RPB600436S 00643169543379***RPB600437S 00643169543379***RPB600438S 00643169543379***RPB600439S 00643169543379***RPB600440S 00643169543379***RPB600441S 00643169543379***RPB600442S 00643169543379***RPB600443S 00643169543379***RPB600444S 00643169543379***RPB600445S 00643169543379***RPB600446S 00643169543379***RPB600447S 00643169543379***RPB600448S 00643169543379***RPB600449S 00643169543379***RPB600450S 00643169543379***RPB600451S 00643169543379***RPB600452S 00643169543379***RPB600453S 00643169543379***RPB600454S 00643169543379***RPB600455S 00643169543379***RPB600456S 00643169543379***RPB600457S 00643169543379***RPB600458S 00643169543379***RPB600459S 00643169543379***RPB600460S 00643169543379***RPB600461S 00643169543379***RPB600462S 00643169543379***RPB600463S 00643169543379***RPB600464S 00643169543379***RPB600465S 00643169543379***RPB600466S 00643169543379***RPB600467S 00643169543379***RPB600468S 00643169543379***RPB600469S 00643169543379***RPB600470S 00643169543379***RPB600471S 00643169543379***RP |  | SoftwareDesignChange
121 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151560 | Z-0744-2017 | SOMATOM Definition Edge System, x-ray, tomography, computed***Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles | Siemens SOMATOM Definition Edge | 2 | CDRH | 2016-12-12 |  | 2018-04-27 | Siemens Medical Solutions USA, Inc | Software update that provides software and firmware bug-fixes to improve system performance | Class 2 Device Recall Siemens SOMATOM Definition Edge Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Siemens SOMATOM Definition Edge FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Siemens SOMATOM Definition Edge Date Initiated by Firm November 15, 2016 Create Date December 12, 2016 Recall Status 1 Terminated 3 on April 27, 2018 Recall Number Z-0744-2017 Recall Event ID 75812 510(K)Number K152036 Product Classification System, x-ray, tomography, computed - Product Code JAK Product SOMATOM Definition Edge System, x-ray, tomography, computed******Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles Code Information Model # 10590000 Serial Numbers: 65127 95798 73739 83277 66219 65491 65453 95738 96357 65615 73726 66467 64860 73860 65712 65304 65437 65614 65856 96305 73296 73831 83403 73840 64709 96335 73673 65639 95003 65501 96351 73778 96028 96316 83520 96288 66396 96238 96352 96324 96336 65480 96312 95768 73777 95955 73883 83399 73783 96226 96303 96047 65609 73791 73887 73071 96266 95701 74368 83537 73895 96214 95925 96347 96061 83533 74226 74297 65760 74376 65017 65516 73162 64065 65455 83524 66041 65618 73818 66066 73028 83306 65680 65734 73905 95797 65672 73886 96302 83398 73538 95717 96053 65972 95278 96326 96327 96286 95447 74200 74289 96364 95510 64140 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact Marlynne Galloway 610-448-6471 Manufacturer Reason for Recall Software update that provides software and firmware bug-fixes to improve system performance FDA Determined Cause 2 Software Design Change Action Siemens mailed a Customer Information Letter to affected customers on November 21, 2016 informing them of software updates that would provide bug-fixes to improve system performance and stability for customers with SOMATOM Definition AS, Definition Edge and Definition Flash systems using Software versions VA48A_SP1 and VA48A_SP2. The update will remotely be pushed to their systems or customers will be contacted by one of Siemens service organization to arrange a date for installation of the update on site. Customers with questions were instructed to call 1-800-888-7436. Quantity in Commerce 104 systems Distribution Nationwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens mailed a Customer Information Letter to affected customers on November 21, 2016 informing them of software updates that would provide bug-fixes to improve system performance and stability for customers with SOMATOM Definition AS, Definition Edge and Definition Flash systems using Software versions VA48A_SP1 and VA48A_SP2. The update will remotely be pushed to their systems or customers will be contacted by one of Siemens service organization to arrange a date for installation of the update on site. Customers with questions were instructed to call 1-800-888-7436. | SoftwareDesignChange
122 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151354 | Z-0743-2017 | SOMATOM Definition Flash System, x-ray, tomography, computed***Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles | Siemens SOMATOM Definition Flash | 2 | CDRH | 2016-12-12 |  | 2018-04-27 | Siemens Medical Solutions USA, Inc | Software update that provides software and firmware bug-fixes to improve system performance | Class 2 Device Recall Siemens SOMATOM Definition Flash Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Siemens SOMATOM Definition Flash FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Siemens SOMATOM Definition Flash Date Initiated by Firm November 15, 2016 Create Date December 12, 2016 Recall Status 1 Terminated 3 on April 27, 2018 Recall Number Z-0743-2017 Recall Event ID 75812 510(K)Number K143416 Product Classification System, x-ray, tomography, computed - Product Code JAK Product SOMATOM Definition Flash System, x-ray, tomography, computed******Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles Code Information Serial Numbers: 65127 95798 73739 83277 66219 65491 65453 95738 96357 65615 73726 66467 64860 73860 65712 65304 65437 65614 65856 96305 73296 73831 83403 73840 64709 96335 73673 65639 95003 65501 96351 73778 96028 96316 83520 96288 66396 96238 96352 96324 96336 65480 96312 95768 73777 95955 73883 83399 73783 96226 96303 96047 65609 73791 73887 73071 96266 95701 74368 83537 73895 96214 95925 96347 96061 83533 74226 74297 65760 74376 65017 65516 73162 64065 65455 83524 66041 65618 73818 66066 73028 83306 65680 65734 73905 95797 65672 73886 96302 83398 73538 95717 96053 65972 95278 96326 96327 96286 95447 74200 74289 96364 95510 64140 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact Marlynne Galloway 610-448-6471 Manufacturer Reason for Recall Software update that provides software and firmware bug-fixes to improve system performance FDA Determined Cause 2 Software Design Change Action Siemens mailed a Customer Information Letter to affected customers on November 21, 2016 informing them of software updates that would provide bug-fixes to improve system performance and stability for customers with SOMATOM Definition AS, Definition Edge and Definition Flash systems using Software versions VA48A_SP1 and VA48A_SP2. The update will remotely be pushed to their systems or customers will be contacted by one of Siemens service organization to arrange a date for installation of the update on site. Customers with questions were instructed to call 1-800-888-7436. Quantity in Commerce 104 systems Distribution Nationwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens mailed a Customer Information Letter to affected customers on November 21, 2016 informing them of software updates that would provide bug-fixes to improve system performance and stability for customers with SOMATOM Definition AS, Definition Edge and Definition Flash systems using Software versions VA48A_SP1 and VA48A_SP2. The update will remotely be pushed to their systems or customers will be contacted by one of Siemens service organization to arrange a date for installation of the update on site. Customers with questions were instructed to call 1-800-888-7436. | SoftwareDesignChange
123 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151353 | Z-0742-2017 | SOMATOM Definition AS System, x-ray, tomography, computed***Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles. | Siemens SOMATOM Definition AS | 2 | CDRH | 2016-12-12 |  | 2018-04-27 | Siemens Medical Solutions USA, Inc | Software update that provides software and firmware bug-fixes to improve system performance | Class 2 Device Recall Siemens SOMATOM Definition AS Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Siemens SOMATOM Definition AS FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Siemens SOMATOM Definition AS Date Initiated by Firm November 15, 2016 Create Date December 12, 2016 Recall Status 1 Terminated 3 on April 27, 2018 Recall Number Z-0742-2017 Recall Event ID 75812 510(K)Number K152036 Product Classification System, x-ray, tomography, computed - Product Code JAK Product SOMATOM Definition AS System, x-ray, tomography, computed******Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles. Code Information Serial Numbers: 65127 95798 73739 83277 66219 65491 65453 95738 96357 65615 73726 66467 64860 73860 65712 65304 65437 65614 65856 96305 73296 73831 83403 73840 64709 96335 73673 65639 95003 65501 96351 73778 96028 96316 83520 96288 66396 96238 96352 96324 96336 65480 96312 95768 73777 95955 73883 83399 73783 96226 96303 96047 65609 73791 73887 73071 96266 95701 74368 83537 73895 96214 95925 96347 96061 83533 74226 74297 65760 74376 65017 65516 73162 64065 65455 83524 66041 65618 73818 66066 73028 83306 65680 65734 73905 95797 65672 73886 96302 83398 73538 95717 96053 65972 95278 96326 96327 96286 95447 74200 74289 96364 95510 64140 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact Marlynne Galloway 610-448-6471 Manufacturer Reason for Recall Software update that provides software and firmware bug-fixes to improve system performance FDA Determined Cause 2 Software Design Change Action Siemens mailed a Customer Information Letter to affected customers on November 21, 2016 informing them of software updates that would provide bug-fixes to improve system performance and stability for customers with SOMATOM Definition AS, Definition Edge and Definition Flash systems using Software versions VA48A_SP1 and VA48A_SP2. The update will remotely be pushed to their systems or customers will be contacted by one of Siemens service organization to arrange a date for installation of the update on site. Customers with questions were instructed to call 1-800-888-7436. Quantity in Commerce 104 systems Distribution Nationwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens mailed a Customer Information Letter to affected customers on November 21, 2016 informing them of software updates that would provide bug-fixes to improve system performance and stability for customers with SOMATOM Definition AS, Definition Edge and Definition Flash systems using Software versions VA48A_SP1 and VA48A_SP2. The update will remotely be pushed to their systems or customers will be contacted by one of Siemens service organization to arrange a date for installation of the update on site. Customers with questions were instructed to call 1-800-888-7436. | SoftwareDesignChange
124 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150426 | Z-0413-2017 | The ORCHESTRA PLUS Programmer ******A portable device, equipped with a microprocessor used to program and interrogate Sorin implantable pacemakers and defibrillators. The programmer also provides measurement, ECG display and report printing functions that are essential for monitoring implant patients. | ORCHESTRA PLUS Programmer | 2 | CDRH | 2016-11-15 |  | 2020-05-15 15:34 | Sorin Group USA, Inc. | Sorin Group USA, Inc. announces a voluntary field action for the ORCHESTRA Programmer and ORCHESTRA PLUS***Programmer because the residual longevity estimate may be less accurate. | Class 2 Device Recall ORCHESTRA PLUS Programmer Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ORCHESTRA PLUS Programmer FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ORCHESTRA PLUS Programmer Date Initiated by Firm September 29, 2016 Create Date November 15, 2016 Recall Status 1 Terminated 3 on May 15, 2020 Recall Number Z-0413-2017 Recall Event ID 75456 PMA Number P980049S002 P980049S025 Product Classification Programmer, pacemaker - Product Code KRG Product The ORCHESTRA PLUS Programmer ************A portable device, equipped with a microprocessor used to program and interrogate Sorin implantable pacemakers and defibrillators. The programmer also provides measurement, ECG display and report printing functions that are essential for monitoring implant patients. Code Information Serial Numbers: NB1006418K NB1107409K NB1203456K NB1107407K NB1105419K NB0908405K NB1202447K NA0804365K NB1201447K NA0807425K NB0909407K NB1112470K NB0909402K NB1112474K NB1112444K NB1201403K NB1201402K NB1102423K NB1112471K NA0804393K NB1203431K NA0804389K NB1004403K NB1202418K NB1108403K NB1007410K NB1105438K NA0804367K NB1102406K NA0804384K NB1202458K NA0804321K NB1006415K NB1201455K NB1102402K NB1107417K NB1109421K NB1202454K NC1503413K NA0804353K NB1112482K NB1108404K NB1105436K NB0907430K NB1112485K NA0804343K NB1109418K NB1104409K NB1112414K NB0908415K NB1112407K NB1108425K NB1006402K NB1203428K NA0804426K NA0804390K NB0911402K NB0907419K NB0909417K NB1203415K NB1201442K NA0804363K NB0911414K NB0908411K NB0910412K NA0807414K NA0804387K NB1102405K NB1201409K NB1201423K NB1004416K NB1108401K NB1002406K NB1203429K NB0907433K NB1202474K NB1103415K NB1112427K NB1202462K NB1202469K NA0804319K NB1102425K NA0804339K NB1003402K NB1102404K NB0909410K NB0911404K NB1108429K NB1203444K NB1203423K NB1201434K NA0804410K NB0909430K NB1202413K NB1201454K NB1001405K NA0904419K NA0902423K NB0907425K NB1202435K NB1203417K NA0904415K NB0907413K NB1105401K NB1112455K NA0904413K NB1112423K NB1201410K NB1112404K NA0904414K NB1108428K NB1112454K NB0911401K NB1007406K NB1102428K NB1201450K NB1001409K NB1112463K NA0804433K NB1105404K NB1202423K NB1201430K NB1203453K NB1006410K NC1503412K NA0804373K NA0804378K NB1201449K NA0804358K NB1108412K NB1201406K NB0907410K NB1006403K NA0804362K NB1004406K NB0908414K NA0902406K NA0804432K NB1002401K NB0909405K NA0807429K NB0911415K NA0804313K NB1109402K NB1203454K NB1202409K NB1202453K NB1201440K NB0905015K NB1202464K NB1202442K NA0804340K NB1201439K NB1109412K NB1102407K NB1112410K NB1108418K NB1104416K NB0907414K NB1202424K NB1201445K NB1201452K NB0907417K NB1202432K NA0804411K NB1202408K NB1007402K NB1103409K NB1112459K NB1202430K NA0804383K NB1108426K NB1010412K NB0911412K NB1202440K NB1102429K NB1104405K NB1203409K NB1010408K NB1109417K NB1108410K NB1203437K NA0804324K NA0804422K NB0905017K NB1108424K NB1007412K NB1112442K NB1112438K NB1112460K NB1108422K NB1006404K NB1104406K NB1007407K NA0904416K NB1107408K NB1104413K NB1202439K NA0804391K NB1105429K NB1107401K NB1006401K NB0907409K NB1003406K NB1109420K NB1112443K NB1102411K NA0804309K NB1108416K NB1203421K NB1102419K NB1202420K NA0804323K NA0804352K NB1202438K NB1202444K NB1006405K NB1004414K NB1202459K NC1503404K NB1107414K NB1108419K NB1108423K NC1503407K NB1108409K NA0804336K NA0804341K NA0804305K NB1108432K NB1112461K NB1007404K NB1112453K NB1109416K NA0804386K NC1503408K NB0907408K NB1102415K NB1001404K NB1112420K NA0804306K NB0908417K NA0902401K NB1203402K NB0909406K NB1004404K NB0907427K NB1203451K NA0804320K NA0807424K NB1201435K NB1109422K NB0910413K NB1202477K NB1203401K NB1203452K NA0804394K NA0902422K NA0804329K NB1109414K NB1201460K NB1104404K NB1003403K NB1203448K NB0907422K NB1105431K NB1004410K NB0909401K NB1201401K NB1112475K NC1503401K NB1108406K NB1201424K NA0807416K NB1105417K NB1105422K NB0909414K NB0908420K NB1112426K NB0908401K NB1107412K NB1202467K NB1004407K NB1010413K NB1004412K NB1003405K NB1112403K NA0807415K NB1202450K NB1103407K NB1104410K NB1104411K NB1202451K NA0904408K NB1112473K NB1202473K NB1202456K NB1109407K NB0908402K NB1202468K NA0804356K NB1003404K NA0804308K NA0804359K NA0804370K NB1203441K NB1203442K NB1103417K NB1103419K NB1007411K NA0804334K NB1105406K NB1109424K NC1503414K NB1105428K NA0902411K NB1201429K NB0910414K NB1102412K NB1103412K NB0911403K NB1112434K NC1503409K NA0804409K NB0905018K NB1104402K NA0804330K NB1112448K NB1201407K NB1203411K NA0804396K NB1202436K NB0907431K NB1102409K NB1006419K NB1006409K NB1007401K NB1109419K NB1203419K NB1105403K NB1203426K NB1202455K NB1102422K NB1105407K NB1109405K NA0804360K NB1112457K NA0902409K NB0909420K NB1004405K NB1112417K NA0804427K NB1201417K NB1201404K NA0902407K NB1112480K NB1201427K NB1102420K NB1112484K NA0804397K NB0907412K NB1202437K NB1112424K NB1104418K NB1103414K NB1010416K NB1102427K NB1202463K NB1202431K NB1102410K NB1012407K NA0804369K NB1108431K NA0807428K NB1201433K NB1112441K NB1108413K NB1201412K NB1108415K NB1108420K NB1108427K NB1010409K NB1203449K NB1112462K NA0804354K NB0905011K NB0907434K NB0907423K NB1203430K NB1202429K NA0804349K NB0907428K NB1001402K NA0902405K NB1102421K NB0910401K NB1004402K NB1112419K NB1105425K NB1201457K NA0804429K NA0804368K NA0902421K NA0804423K NB0910403K NB1007413K NB1112439K NB1201416K NB0907415K NB1202448K NA0902403K NB1202445K NB1201413K NA0804317K NB1202457K NB1104414K NB1004401K NB0908406K NB0908408K NB1104403K NA0902420K NB1112433K NB1007409K NA0804392K NB1112451K NA0904410K NB1109425K NB0910408K NB1012405K NB1105420K NB1006414K NB1104419K ND1603401K NA0804366K NB1203434K NB1109423K NB1112458K NB1203407K NB1203413K NB1203406K NA0804312K NA0804310K NB1112447K NB1112477K NB1103406K NB1202446K NB1202421K NB1102401K NB1203435K NB1112406K NB1112430K NB1107404K NB1112409K NA0804315K NA0804311K NB1112422K NA0902418K NB1203422K NB1102408K NB1202478K NB1203412K NB1203425K NB1202441K NB1006413K NB1109401K NB1201408K NB1201453K NB0907429K NB1012402K NA0904407K NB1108421K NB1105432K NB0907421K NA0804385K NB1112483K NB1103404K NB1112425K NB0911410K NB0910409K NA0804318K NB1107413K NB1004409K NA0804333K NA0807418K NB1012404K NB1201415K NB1109404K NB1105427K NB1202466K NB1103416K NB1108408K NB1202434K NB0905009K NA0804348K NB1105433K NB1202414K NB1203436K NB0908413K NA0904420K NC1503405K NB1102426K NB1203447K NB0911405K NB1201441K NB1202479K NB1010414K NB1203432K NB1012401K NB1108402K NB0905013K NB1108414K NB1203427K NB1203418K NB1201431K NB1102416K NB1103410K NA0804361K NA0904409K NB1105402K NB1105440K NB1202417K NB1112429K NB1112472K NA0705031K NB1112416K NB1112440K NB0908407K NA0902415K NB1203405K NB0908419K NB1102403K NB1112446K NB1004408K NB1010417K NB1203457K NB0909416K NB1203433K NB1201411K NB1007405K NA0807426K NB0909404K NB1102424K NB1107405K NB1107406K NB1203414K NA0804374K NB1108438K NB1109403K NB1002404K NB1103403K NB1010407K NA0804342K NB1203403K NB1112476K NB1202416K NB1010402K NB1202412K NB1203445K NB1103420K NA0804382K NB1006411K NA0804398K NB1201432K NB1203446K NB1203450K NB0907405K NB1108430K NB1109415K NB1010410K NB1010411K NB1006406K NA0902424K NA0807430K NA0902413K NB0908412K NB1004411K NB1006416K NB1201451K NB1002402K NB1103418K NB1010403K NA0807412K NB1105426K NB1202465K NB1112466K NB0909403K NB1203424K NB1006418K NB1107409K NB1203456K NB1107407K NB1105419K NB0908405K NB1202447K NA0804365K NB1201447K NA0807425K NB0909407K NB1112470K NB0909402K NB1112474K NB1112444K NB1201403K NB1201402K NB1102423K NB1112471K NA0804393K NB1203431K NA0804389K NB1004403K NB1202418K NB1108403K NB1007410K NB1105438K NA0804367K NB1102406K NA0804384K NB1202458K NA0804321K NB1006415K NB1201455K NB1102402K NB1107417K NB1109421K NB1202454K NC1503413K NA0804353K NB1112482K NB1108404K NB1105436K NB0907430K NB1112485K NA0804343K NB1109418K NB1104409K NB1112414K NB0908415K NB1112407K NB1108425K NB1006402K NB1203428K NA0804426K NA0804390K NB0911402K NB0907419K NB0909417K NB1203415K NB1201442K NA0804363K NB0911414K NB0908411K NB0910412K NA0807414K NA0804387K NB1102405K NB1201409K NB1201423K NB1004416K NB1108401K NB1002406K NB1203429K NB0907433K NB1202474K NB1103415K NB1112427K NB1202462K NB1202469K NA0804319K NB1102425K NA0804339K NB1003402K NB1102404K NB0909410K NB0911404K NB1108429K NB1203444K NB1203423K NB1201434K NA0804410K NB0909430K NB1202413K NB1201454K NB1001405K NA0904419K NA0902423K NB0907425K NB1202435K NB1203417K NA0904415K NB0907413K NB1105401K NB1112455K NA0904413K NB1112423K NB1201410K NB1112404K NA0904414K NB1108428K NB1112454K NB0911401K NB1007406K NB1102428K NB1201450K NB1001409K NB1112463K NA0804433K NB1105404K NB1202423K NB1201430K NB1203453K NB1006410K NC1503412K NA0804373K NA0804378K NB1201449K NA0804358K NB1108412K NB1201406K NB0907410K NB1006403K NA0804362K NB1004406K NB0908414K NA0902406K NA0804432K NB1002401K NB0909405K NA0807429K NB0911415K NA0804313K NB1109402K NB1203454K NB1202409K NB1202453K NB1201440K NB0905015K NB1202464K NB1202442K NA0804340K NB1201439K NB1109412K NB1102407K NB1112410K NB1108418K NB1104416K NB0907414K NB1202424K NB1201445K NB1201452K NB0907417K NB1202432K NA0804411K NB1202408K NB1007402K NB1103409K NB1112459K NB1202430K NA0804383K NB1108426K NB1010412K NB0911412K NB1202440K NB1102429K NB1104405K NB1203409K NB1010408K NB1109417K NB1108410K NB1203437K NA0804324K NA0804422K NB0905017K NB1108424K NB1007412K NB1112442K NB1112438K NB1112460K NB1108422K NB1006404K NB1104406K NB1007407K NA0904416K NB1107408K NB1104413K NB1202439K NA0804391K NB1105429K NB1107401K NB1006401K NB0907409K NB1003406K NB1109420K NB1112443K NB1102411K NA0804309K NB1108416K NB1203421K NB1102419K NB1202420K NA0804323K NA0804352K NB1202438K NB1202444K NB1006405K NB1004414K NB1202459K NC1503404K NB1107414K NB1108419K NB1108423K NC1503407K NB1108409K NA0804336K NA0804341K NA0804305K NB1108432K NB1112461K NB1007404K NB1112453K NB1109416K NA0804386K NC1503408K NB0907408K NB1102415K NB1001404K NB1112420K NA0804306K NB0908417K NA0902401K NB1203402K NB0909406K NB1004404K NB0907427K NB1203451K NA0804320K NA0807424K NB1201435K NB1109422K NB0910413K NB1202477K NB1203401K NB1203452K NA0804394K NA0902422K NA0804329K NB1109414K NB1201460K NB1104404K NB1003403K NB1203448K NB0907422K NB1105431K NB1004410K NB0909401K NB1201401K NB1112475K NC1503401K NB1108406K NB1201424K NA0807416K NB1105417K NB1105422K NB0909414K NB0908420K NB1112426K NB0908401K NB1107412K NB1202467K NB1004407K NB1010413K NB1004412K NB1003405K NB1112403K NA0807415K NB1202450K NB1103407K NB1104410K NB1104411K NB1202451K NA0904408K NB1112473K NB1202473K NB1202456K NB1109407K NB0908402K NB1202468K NA0804356K NB1003404K NA0804308K NA0804359K NA0804370K NB1203441K NB1203442K NB1103417K NB1103419K NB1007411K NA0804334K NB1105406K NB1109424K NC1503414K NB1105428K NA0902411K NB1201429K NB0910414K NB1102412K NB1103412K NB0911403K NB1112434K NC1503409K NA0804409K NB0905018K NB1104402K NA0804330K NB1112448K NB1201407K NB1203411K NA0804396K NB1202436K NB0907431K NB1102409K NB1006419K NB1006409K NB1007401K NB1109419K NB1203419K NB1105403K NB1203426K NB1202455K NB1102422K NB1105407K NB1109405K NA0804360K NB1112457K NA0902409K NB0909420K NB1004405K NB1112417K NA0804427K NB1201417K NB1201404K NA0902407K NB1112480K NB1201427K NB1102420K NB1112484K NA0804397K NB0907412K NB1202437K NB1112424K NB1104418K NB1103414K NB1010416K NB1102427K NB1202463K NB1202431K NB1102410K NB1012407K NA0804369K NB1108431K NA0807428K NB1201433K NB1112441K NB1108413K NB1201412K NB1108415K NB1108420K NB1108427K NB1010409K NB1203449K NB1112462K NA0804354K NB0905011K NB0907434K NB0907423K NB1203430K NB1202429K NA0804349K NB0907428K NB1001402K NA0902405K NB1102421K NB0910401K NB1004402K NB1112419K NB1105425K NB1201457K NA0804429K NA0804368K NA0902421K NA0804423K NB0910403K NB1007413K NB1112439K NB1201416K NB0907415K NB1202448K NA0902403K NB1202445K NB1201413K NA0804317K NB1202457K NB1104414K NB1004401K NB0908406K NB0908408K NB1104403K NA0902420K NB1112433K NB1007409K NA0804392K NB1112451K NA0904410K NB1109425K NB0910408K NB1012405K NB1105420K NB1006414K NB1104419K ND1603401K NA0804366K NB1203434K NB1109423K NB1112458K NB1203407K NB1203413K NB1203406K NA0804312K NA0804310K NB1112447K NB1112477K NB1103406K NB1202446K NB1202421K NB1102401K NB1203435K NB1112406K NB1112430K NB1107404K NB1112409K NA0804315K NA0804311K NB1112422K NA0902418K NB1203422K NB1102408K NB1202478K NB1203412K NB1203425K NB1202441K NB1006413K NB1109401K NB1201408K NB1201453K NB0907429K NB1012402K NA0904407K NB1108421K NB1105432K NB0907421K NA0804385K NB1112483K NB1103404K NB1112425K NB0911410K NB0910409K NA0804318K NB1107413K NB1004409K NA0804333K NA0807418K NB1012404K NB1201415K NB1109404K NB1105427K NB1202466K NB1103416K NB1108408K NB1202434K NB0905009K NA0804348K NB1105433K NB1202414K NB1203436K NB0908413K NA0904420K NC1503405K NB1102426K NB1203447K NB0911405K NB1201441K NB1202479K NB1010414K NB1203432K NB1012401K NB1108402K NB0905013K NB1108414K NB1203427K NB1203418K NB1201431K NB1102416K NB1103410K NA0804361K NA0904409K NB1105402K NB1105440K NB1202417K NB1112429K NB1112472K NA0705031K NB1112416K NB1112440K NB0908407K NA0902415K NB1203405K NB0908419K NB1102403K NB1112446K NB1004408K NB1010417K NB1203457K NB0909416K NB1203433K NB1201411K NB1007405K NA0807426K NB0909404K NB1102424K NB1107405K NB1107406K NB1203414K NA0804374K NB1108438K NB1109403K NB1002404K NB1103403K NB1010407K NA0804342K NB1203403K NB1112476K NB1202416K NB1010402K NB1202412K NB1203445K NB1103420K NA0804382K NB1006411K NA0804398K NB1201432K NB1203446K NB1203450K NB0907405K NB1108430K NB1109415K NB1010410K NB1010411K NB1006406K NA0902424K NA0807430K NA0902413K NB0908412K NB1004411K NB1006416K NB1201451K NB1002402K NB1103418K NB1010403K NA0807412K NB1105426K NB1202465K NB1112466K NB0909403K NB1203424K NB1006418K NB1107409K NB1203456K NB1107407K NB1105419K NB0908405K NB1202447K NA0804365K NB1201447K NA0807425K NB0909407K NB1112470K NB0909402K NB1112474K NB1112444K NB1201403K NB1201402K NB1102423K NB1112471K NA0804393K NB1203431K NA0804389K NB1004403K NB1202418K NB1108403K NB1007410K NB1105438K NA0804367K NB1102406K NA0804384K NB1202458K NA0804321K NB1006415K NB1201455K NB1102402K NB1107417K NB1109421K NB1202454K NC1503413K NA0804353K NB1112482K NB1108404K NB1105436K NB0907430K NB1112485K NA0804343K NB1109418K NB1104409K NB1112414K NB0908415K NB1112407K NB1108425K NB1006402K NB1203428K NA0804426K NA0804390K NB0911402K NB0907419K NB0909417K NB1203415K NB1201442K NA0804363K NB0911414K NB0908411K NB0910412K NA0807414K NA0804387K NB1102405K NB1201409K NB1201423K NB1004416K NB1108401K NB1002406K NB1203429K NB0907433K NB1202474K NB1103415K NB1112427K NB1202462K NB1202469K NA0804319K NB1102425K NA0804339K NB1003402K NB1102404K NB0909410K NB0911404K NB1108429K NB1203444K NB1203423K NB1201434K NA0804410K NB0909430K NB1202413K NB1201454K NB1001405K NA0904419K NA0902423K NB0907425K NB1202435K NB1203417K NA0904415K NB0907413K NB1105401K NB1112455K NA0904413K NB1112423K NB1201410K NB1112404K NA0904414K NB1108428K NB1112454K NB0911401K NB1007406K NB1102428K NB1201450K NB1001409K NB1112463K NA0804433K NB1105404K NB1202423K NB1201430K NB1203453K NB1006410K NC1503412K NA0804373K NA0804378K NB1201449K NA0804358K NB1108412K NB1201406K NB0907410K NB1006403K NA0804362K NB1004406K NB0908414K NA0902406K NA0804432K NB1002401K NB0909405K NA0807429K NB0911415K NA0804313K NB1109402K NB1203454K NB1202409K NB1202453K NB1201440K NB0905015K NB1202464K NB1202442K NA0804340K NB1201439K NB1109412K NB1102407K NB1112410K NB1108418K NB1104416K NB0907414K NB1202424K NB1201445K NB1201452K NB0907417K NB1202432K NA0804411K NB1202408K NB1007402K NB1103409K NB1112459K NB1202430K NA0804383K NB1108426K NB1010412K NB0911412K NB1202440K NB1102429K NB1104405K NB1203409K NB1010408K NB1109417K NB1108410K NB1203437K NA0804324K NA0804422K NB0905017K NB1108424K NB1007412K NB1112442K NB1112438K NB1112460K NB1108422K NB1006404K NB1104406K NB1007407K NA0904416K NB1107408K NB1104413K NB1202439K NA0804391K NB1105429K NB1107401K NB1006401K NB0907409K NB1003406K NB1109420K NB1112443K NB1102411K NA0804309K NB1108416K NB1203421K NB1102419K NB1202420K NA0804323K NA0804352K NB1202438K NB1202444K NB1006405K NB1004414K NB1202459K NC1503404K NB1107414K NB1108419K NB1108423K NC1503407K NB1108409K NA0804336K NA0804341K NA0804305K NB1108432K NB1112461K NB1007404K NB1112453K NB1109416K NA0804386K NC1503408K NB0907408K NB1102415K NB1001404K NB1112420K NA0804306K NB0908417K NA0902401K NB1203402K NB0909406K NB1004404K NB0907427K NB1203451K NA0804320K NA0807424K NB1201435K NB1109422K NB0910413K NB1202477K NB1203401K NB1203452K NA0804394K NA0902422K NA0804329K NB1109414K NB1201460K NB1104404K NB1003403K NB1203448K NB0907422K NB1105431K NB1004410K NB0909401K NB1201401K NB1112475K NC1503401K NB1108406K NB1201424K NA0807416K NB1105417K NB1105422K NB0909414K NB0908420K NB1112426K NB0908401K NB1107412K NB1202467K NB1004407K NB1010413K NB1004412K NB1003405K NB1112403K NA0807415K NB1202450K NB1103407K NB1104410K NB1104411K NB1202451K NA0904408K NB1112473K NB1202473K NB1202456K NB1109407K NB0908402K NB1202468K NA0804356K NB1003404K NA0804308K NA0804359K NA0804370K NB1203441K NB1203442K NB1103417K NB1103419K NB1007411K NA0804334K NB1105406K NB1109424K NC1503414K NB1105428K NA0902411K NB1201429K NB0910414K NB1102412K NB1103412K NB0911403K NB1112434K NC1503409K NA0804409K NB0905018K NB1104402K NA0804330K NB1112448K NB1201407K NB1203411K NA0804396K NB1202436K NB0907431K NB1102409K NB1006419K NB1006409K NB1007401K NB1109419K NB1203419K NB1105403K NB1203426K NB1202455K NB1102422K NB1105407K NB1109405K NA0804360K NB1112457K NA0902409K NB0909420K NB1004405K NB1112417K NA0804427K NB1201417K NB1201404K NA0902407K NB1112480K NB1201427K NB1102420K NB1112484K NA0804397K NB0907412K NB1202437K NB1112424K NB1104418K NB1103414K NB1010416K NB1102427K NB1202463K NB1202431K NB1102410K NB1012407K NA0804369K NB1108431K NA0807428K NB1201433K NB1112441K NB1108413K NB1201412K NB1108415K NB1108420K NB1108427K NB1010409K NB1203449K NB1112462K NA0804354K NB0905011K NB0907434K NB0907423K NB1203430K NB1202429K NA0804349K NB0907428K NB1001402K NA0902405K NB1102421K NB0910401K NB1004402K NB1112419K NB1105425K NB1201457K NA0804429K NA0804368K NA0902421K NA0804423K NB0910403K NB1007413K NB1112439K NB1201416K NB0907415K NB1202448K NA0902403K NB1202445K NB1201413K NA0804317K NB1202457K NB1104414K NB1004401K NB0908406K NB0908408K NB1104403K NA0902420K NB1112433K NB1007409K NA0804392K NB1112451K NA0904410K NB1109425K NB0910408K NB1012405K NB1105420K NB1006414K NB1104419K ND1603401K NA0804366K NB1203434K NB1109423K NB1112458K NB1203407K NB1203413K NB1203406K NA0804312K NA0804310K NB1112447K NB1112477K NB1103406K NB1202446K NB1202421K NB1102401K NB1203435K NB1112406K NB1112430K NB1107404K NB1112409K NA0804315K NA0804311K NB1112422K NA0902418K NB1203422K NB1102408K NB1202478K NB1203412K NB1203425K NB1202441K NB1006413K NB1109401K NB1201408K NB1201453K NB0907429K NB1012402K NA0904407K NB1108421K NB1105432K NB0907421K NA0804385K NB1112483K NB1103404K NB1112425K NB0911410K NB0910409K NA0804318K NB1107413K NB1004409K NA0804333K NA0807418K NB1012404K NB1201415K NB1109404K NB1105427K NB1202466K NB1103416K NB1108408K NB1202434K NB0905009K NA0804348K NB1105433K NB1202414K NB1203436K NB0908413K NA0904420K NC1503405K NB1102426K NB1203447K NB0911405K NB1201441K NB1202479K NB1010414K NB1203432K NB1012401K NB1108402K NB0905013K NB1108414K NB1203427K NB1203418K NB1201431K NB1102416K NB1103410K NA0804361K NA0904409K NB1105402K NB1105440K NB1202417K NB1112429K NB1112472K NA0705031K NB1112416K NB1112440K NB0908407K NA0902415K NB1203405K NB0908419K NB1102403K NB1112446K NB1004408K NB1010417K NB1203457K NB0909416K NB1203433K NB1201411K NB1007405K NA0807426K NB0909404K NB1102424K NB1107405K NB1107406K NB1203414K NA0804374K NB1108438K NB1109403K NB1002404K NB1103403K NB1010407K NA0804342K NB1203403K NB1112476K NB1202416K NB1010402K NB1202412K NB1203445K NB1103420K NA0804382K NB1006411K NA0804398K NB1201432K NB1203446K NB1203450K NB0907405K NB1108430K NB1109415K NB1010410K NB1010411K NB1006406K NA0902424K NA0807430K NA0902413K NB0908412K NB1004411K NB1006416K NB1201451K NB1002402K NB1103418K NB1010403K NA0807412K NB1105426K NB1202465K NB1112466K NB0909403K NB1203424K Recalling Firm/ Manufacturer Sorin Group USA, Inc. 14401 W 65th Way Arvada CO 80004-3503 For Additional Information Contact Joan Ceasar 281-685-2227 Manufacturer Reason for Recall Sorin Group USA, Inc. announces a voluntary field action for the ORCHESTRA Programmer and ORCHESTRA PLUS******Programmer because the residual longevity estimate may be less accurate. FDA Determined Cause 2 Software Design Change Action The firm, LivaNova, sent an "Urgent Field Safety Notice" dated 9/29/2016 to its customers. The notice described the product, problem, actions to be taken. The notice states that a Sorin representative will contact customers to upgrade the software in the facility's programmer.************If customers have any questions, they are urged to call their LivaNova representative or contact LivaNova Technical Services at 877.663.7674. Quantity in Commerce 586 programmers Distribution Programmers were distributed nationwide, to VA/govt/military, Canadian, and other foreign consignees. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, LivaNova, sent an "Urgent Field Safety Notice" dated 9/29/2016 to its customers. The notice described the product, problem, actions to be taken. The notice states that a Sorin representative will contact customers to upgrade the software in the facility's programmer.************If customers have any questions, they are urged to call their LivaNova representative or contact LivaNova Technical Services at 877.663.7674. | SoftwareDesignChange
125 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150425 | Z-0412-2017 | The ORCHESTRA Programmer ******A portable device, equipped with a microprocessor used to program and interrogate Sorin implantable pacemakers and defibrillators. The programmer also provides measurement, ECG display and report printing functions that are essential for monitoring implant***patients. | ORCHESTRA and ORCHESTRA PLUS Programmers | 2 | CDRH | 2016-11-15 |  | 2020-05-15 15:34 | Sorin Group USA, Inc. | Sorin Group USA, Inc. announces a voluntary field action for the ORCHESTRA Programmer and ORCHESTRA PLUS***Programmer because the residual longevity estimate may be less accurate. | Class 2 Device Recall ORCHESTRA and ORCHESTRA PLUS Programmers Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ORCHESTRA and ORCHESTRA PLUS Programmers FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ORCHESTRA and ORCHESTRA PLUS Programmers Date Initiated by Firm September 29, 2016 Create Date November 15, 2016 Recall Status 1 Terminated 3 on May 15, 2020 Recall Number Z-0412-2017 Recall Event ID 75456 PMA Number P980049S002 P980049S025 Product Classification Programmer, pacemaker - Product Code KRG Product The ORCHESTRA Programmer ************A portable device, equipped with a microprocessor used to program and interrogate Sorin implantable pacemakers and defibrillators. The programmer also provides measurement, ECG display and report printing functions that are essential for monitoring implant******patients. Code Information Serial Numbers: BB0011103G BC0206105G BC0408064G BB0011111G BD0505094I BC0405016G BB0102071G BC0206096G BD0507033I BD0505110I BD0507019I BC0206076G BE0611036I BB0203034G BB0011099G BB0103048G BC0408067G BD0509044I BB0012010G BD0501100G BC0206131G BD0501096G BA0006167G BB0011169G BD0510018I BD0505029I BD0505039I BB0102074G BD0507014I BB0101021G BD0506112I BA0007003G BB0010078G BB0010213G BC0206089G BC0205079G BD0501099G BA0004162G BB0011101G BD0501094G BB0103027G BB0012068G BD0507005I BB0011003G BB0011086G BC0408057G BB0012075G BA0005050G BC0303071G BD0507060I BB0010268G BB0012154G BC0303073G BA0006171G BB0011080G BC0312045G BC0206144G BD0505016I BC0206029G BB0012184G BD0505018I BC0405050G BD0501085G BD0506136I BB0011083G BB0012056G BE0611013I BD0507053I BD0505028I BC0408078G BE0603002I BE0603012I BB0011166G BB0012150G BB0012067G BB0012171G BB0011020G BC0212074G BB0010250G BD0507025I BB0203012G BD0501061G BC0206110G BE0611023I BB0011141G BD0502041G BE0611087I BB0012017G BB0010193G BC0205082G BD0509048I BA0006178G BB0010252G BB0101092G BB0012182G BC0405037G BD0505001I BD0507065I BC0405022G BB0203059G BD0505097I BD0502013G BB0012125G BD0506123I BD0501074G BB0102069G BC0405024G BB0012151G BB0012183G BB0203052G BD0505061I BE0611062I BC0212076G BD0510028I BB0010063G BB0012019G BB0203028G BB0012062G BB0010057G BC0212080G BE0603018I BC0312059G BE0611082I BC0405025G BC0206067G BB0010273G BB0102068G BB0011032G BB0010080G BB0010271G BA0005233G BD0506111I BD0502035G BD0509046I BC0206138G BD0505002I BB0012166G BC0408072G BD0506106I BE0611041I BE0611043I BA0004161G BB0011037G BB0203037G BD0505045I BD0505052I BD0505008I BB0010294G BB0011042G BB0012070G BD0501118G BD0501116G BE0611008I BC0212068G BC0212083G BD0501109G BB0012027G BB0010207G BC0212085G BC0212086G BD0501087G BB0203050G BB0010291G BC0212079G BE0611022I BE0611009I BB0012060G BD0510024I BD0505068I BD0510022I BD0506125I BC0212048G BD0505090I BB0102059G BC0312055G BC0206146G BC0206073G BD0506124I BC0303082G BB0102075G BC0408075G BC0405047G BD0505106I BB0012131G BC0212046G BC0312066G BC0205068G BC0405052G BB0012004G BB0012013G BC0312049G BD0501066G BD0507063I BB0012063G BC0206083G BD0501092G BC0206075G BC0405023G BB0012012G BD0501071G BB0012074G BC0206071G BB0101023G BD0505053I BD0502003G BD0506101I BD0507061I BC0408052G BB0103059G BC0205089G BC0206046G BD0505080I BD0510009I BB0012160G BC0206134G BD0501081G BA0005349G BB0012146G BC0312044G BB0011120G BB0010215G BC0206027G BB0011180G BB0010082G BB0011170G BB0010297G BD0507038I BD0505077I BB0012001G BB0101020G BB0012176G BD0505074I BB0012143G BB0012141G BD0505107I BD0505033I BD0502038G BB0011036G BC0206016G BD0506103I BC0405064G BB0203046G BC0206080G BB0203020G BB0012023G BC0405045G BD0506143I BD0505040I BD0506144I BB0203018G BC0312068G BD0507069I BB0103074G BD0501075G BB0010052G BE0611069I BB0010076G BA0006181G BD0505050I BB0011105G BC0405049G BB0012008G BB0101083G BD0502031G BD0507020I BB0010216G BC0206154G BD0506098I BE0611048I BB0103022G BB0010222G BD0505048I BE0611063I BA0006183G BD0501070G BC0405006G BD0507001I BB0012181G BB0012022G BC0206042G BB0103041G BB0010254G BC0312069G BB0010269G BD0507043I BB0103047G BB0011046G BE0611024I BD0506099I BC0212061G BB0012162G BB0011081G BD0506117I BB0203013G BD0507024I BB0012158G BD0505044I BE0603014I BB0012071G BC0206116G BC0206109G BD0505066I BC0212081G BD0510014I BE0611005I BD0506110I BC0206102G BB0012002G BC0405053G BB0103039G BB0011165G BA0006179G BB0103058G BB0012119G BB0103046G BD0505036I BD0506109I BE0611073I BD0501078G BB0010285G BC0206100G BB0012018G BD0502042G BD0505101I BC0206149G BC0303090G BC0205086G BD0505085I BB0012024G BB0010105G BB0010183G BB0010051G BB0011164G BD0507044I BC0212062G BB0102077G BC0206079G BB0011038G BB0010230G BC0206018G BB0010284G BB0010049G BC0212056G BC0205071G BD0501107G BB0203011G BD0505089I BB0203022G BE0611076I BB0102064G BB0203057G BB0203041G BB0011108G BD0510020I BC0205087G BD0510031I BD0507010I BC0212078G BD0505072I BC0206148G BC0206099G BB0010241G BA0004165G BB0203035G BB0010289G BB0010184G BB0103062G BB0103043G BD0506097I BD0501089G BD0509049I BB0012136G BB0012133G BC0206098G BE0611016I BD0507057I BD0502002G BB0010214G BD0505030I BD0502015G BB0011112G BD0505047I BC0205077G BB0012135G BD0507034I BC0405005G BB0010061G BE0611021I BB0010264G BD0505058I BB0012145G BC0205062G BB0012116G BA0007063G BE0611033I BC0408069G BC0206062G BB0010279G BD0505038I BC0303080G BB0010293G BB0203044G BB0011161G BD0507058I BD0505065I BD0507002I BB0010283G BC0405032G BC0206033G BE0611006I BB0012113G BA0006170G BC0206153G BE0611070I BD0502008G BB0011102G BD0507046I BD0505042I BC0405048G BE0611010I BC0206072G BC0205074G BD0507035I BB0011104G BD0505009I BC0405008G BC0206141G BB0011118G BB0011040G BB0012132G BB0010060G BD0506121I BC0206145G BD0507072I BC0408058G BB0010071G BB0012134G BB0012033G BE0611079I BC0312053G BB0011116G BC0206024G BB0102067G BE0611011I BB0101087G BC0212088G BC0212084G BD0507062I BD0502025G BD0510017I BE0611019I BB0012085G BB0012086G BC0408066G BD0507064I BB0012200G BB0011092G BD0507031I BB0101080G BC0408073G BC0405039G BC0212087G BB0012079G BD0501064G BB0203002G BB0103038G BD0502012G BB0012054G BD0501091G BA0006174G BC0303076G BB0012155G BB0012026G BE0611030I BD0505087I BC0206107G BC0405013G BC0405028G BD0502027G BB0011029G BE0611029I BB0012163G BC0206155G BB0101027G BC0206133G BC0303088G BB0012003G BE0611046I BC0405065G BD0505049I BD0501093G BD0501108G BB0012114G BB0010242G BC0405021G BB0102060G BB0011126G BB0010062G BB0012185G BD0506093I BD0502022G BE0611028I BB0010238G BB0010070G BB0010265G BC0405062G BB0012078G BD0502006G BB0203043G BD0506091I BB0011026G BD0501068G BD0510007I BC0206115G BC0312067G BB0010256G BC0405043G BB0010288G BB0011137G BD0507039I BE0611047I BD0505067I BD0501110G BD0505041I BB0101081G BB0103052G BD0507049I BD0507013I BB0012120G BB0010075G BC0405066G BB0102062G BE0603011I BE0611003I BD0502043G BB0011167G BB0011158G BC0206081G BB0010185G BA0004159G BC0405009G BE0611015I BC0303089G BB0011131G BC0408079G BD0502024G BD0502016G BB0011002G BB0103069G BB0103024G BB0012115G BD0501082G BB0101084G BD0510016I BD0505104I BD0505019I BD0505051I BB0103061G BD0505086I BC0303086G BB0012186G BB0012178G BD0501101G BC0405042G BB0011157G BB0010206G BC0206093G BC0405067G BB0011090G BC0405030G BC0206031G BB0012180G BD0510023I BC0206152G BD0505105I BD0501095G BD0502040G BB0012011G BB0010068G BB0010300G BD0501111G BB0011179G BC0206150G BB0012077G BB0011095G BD0505079I BB0012036G BB0203008G BC0212071G BC0312046G BC0206103G BB0010190G BD0505096I BE0611018I BB0203054G BC0405020G BD0506129I BB0010224G BC0206077G BD0505059I BD0507056I BB0010228G BB0011138G BC0206151G BB0010217G BB0010239G BB0011098G BB0012159G BD0505037I BB0010245G BC0206088G BD0507073I BA0006098G BC0206091G BE0603010I BC0206125G BC0405033G BD0510013I BB0012127G BD0507036I BB0010210G BC0206084G BC0405012G BB0012081G BE0611044I BE0611050I BB0102070G BB0102079G BE0611014I BD0507030I BC0312042G BD0506127I BC0205093G BD0501073G BB0012111G BD0506126I BC0212082G BD0502034G BD0506108I BA0006175G BC0405011G BE0611002I BA0005347G BC0206143G BB0012149G BC0303087G BB0011034G BB0203029G BB0103056G BC0405040G BC0312060G BD0502029G BD0507071I BC0206044G BC0312063G BC0206119G BD0510029I BC0312065G BD0501090G BD0507070I BD0501088G BE0611080I BD0505012I BB0010081G BB0203047G BD0506137I BB0103064G BA0006100G BB0010253G BB0102021G BB0102082G BB0012059G BB0011134G BD0507041I BE0611012I BD0505083I BC0405059G BC0206104G BD0501117G BD0506142I BB0012053G BB0012069G BC0205076G BD0507029I BE0603005I BB0011121G BC0408071G BD0505007I BC0205091G BD0507022I BB0011156G BC0212055G BB0011144G BC0205069G BB0010182G BC0312054G BB0010189G BB0010275G BB0010259G BB0203048G BB0012021G BD0502004G BC0206057G BD0501084G BB0102065G BD0502018G BD0505022I BC0206069G BD0505082I BC0206043G BD0506130I BE0611032I BB0011024G BE0611034I BC0405001G BC0206142G BD0506113I BD0507015I BD0502037G BB0011001G BB0010225G BC0408060G BB0010281G BB0111034G BD0509045I BB0012080G BD0501104G BB0103042G BD0505015I BB0011114G BB0010258G BD0505011I BD0501065G BE0603007I BB0012061G BD0507004I BB0012157G BE0611059I BC0206092G BB0103050G BB0011110G BE0611026I BB0010108G BC0212064G BC0205078G BA0007064G BB0011021G BD0505035I BB0012179G BD0507037I BD0510005I BE0611038I BB0010074G BC0206124G BD0507076I BC0212077G BB0102080G BC0206021G BB0103036G BC0405041G BD0502039G BB0103035G BB0010255G BD0506105I BB0103028G BB0102081G BC0206122G BA0007001G BD0505081I BD0506095I BB0010218G BC0405056G BD0505100I BE0611020I BB0011109G BB0012137G BC0212066G BB0010110G BC0312058G BB0011087G BC0405017G BC0206053G BC0405035G BB0010208G BC0408074G BC0206106G BD0510030I BB0012165G BC0303075G BD0506114I BD0502010G BC0303081G BB0011159G BB0010229G BB0011163G BD0507011I BE0611051I BE0611040I BC0405044G BB0012082G BB0012124G BC0206019G BD0502014G BB0011171G BB0010251G BC0206032G BD0507017I BC0205070G BB0111033G BE0603013I BE0603008I BB0012015G BA0006185G BC0205061G BB0103060G BB0103067G BC0408055G BD0502009G BB0011175G BD0505017I BD0505064I BD0501083G BB0010221G BD0506128I BC0206114G BD0505021I BB0102019G BC0312050G BC0212090G BE0611075I BC0206040G BD0507003I BC0206066G BB0012126G BD0506100I BE0611060I BD0507075I BD0510032I BE0611027I BB0011100G BD0507074I BE0611017I BC0206035G BB0203005G BC0312061G BB0102066G BC0212051GEnter your input text here and click a text formatting function button below.BB0011103G BC0206105G BC0408064G BB0011111G BD0505094I BC0405016G BB0102071G BC0206096G BD0507033I BD0505110I BD0507019I BC0206076G BE0611036I BB0203034G BB0011099G BB0103048G BC0408067G BD0509044I BB0012010G BD0501100G BC0206131G BD0501096G BA0006167G BB0011169G BD0510018I BD0505029I BD0505039I BB0102074G BD0507014I BB0101021G BD0506112I BA0007003G BB0010078G BB0010213G BC0206089G BC0205079G BD0501099G BA0004162G BB0011101G BD0501094G BB0103027G BB0012068G BD0507005I BB0011003G BB0011086G BC0408057G BB0012075G BA0005050G BC0303071G BD0507060I BB0010268G BB0012154G BC0303073G BA0006171G BB0011080G BC0312045G BC0206144G BD0505016I BC0206029G BB0012184G BD0505018I BC0405050G BD0501085G BD0506136I BB0011083G BB0012056G BE0611013I BD0507053I BD0505028I BC0408078G BE0603002I BE0603012I BB0011166G BB0012150G BB0012067G BB0012171G BB0011020G BC0212074G BB0010250G BD0507025I BB0203012G BD0501061G BC0206110G BE0611023I BB0011141G BD0502041G BE0611087I BB0012017G BB0010193G BC0205082G BD0509048I BA0006178G BB0010252G BB0101092G BB0012182G BC0405037G BD0505001I BD0507065I BC0405022G BB0203059G BD0505097I BD0502013G BB0012125G BD0506123I BD0501074G BB0102069G BC0405024G BB0012151G BB0012183G BB0203052G BD0505061I BE0611062I BC0212076G BD0510028I BB0010063G BB0012019G BB0203028G BB0012062G BB0010057G BC0212080G BE0603018I BC0312059G BE0611082I BC0405025G BC0206067G BB0010273G BB0102068G BB0011032G BB0010080G BB0010271G BA0005233G BD0506111I BD0502035G BD0509046I BC0206138G BD0505002I BB0012166G BC0408072G BD0506106I BE0611041I BE0611043I BA0004161G BB0011037G BB0203037G BD0505045I BD0505052I BD0505008I BB0010294G BB0011042G BB0012070G BD0501118G BD0501116G BE0611008I BC0212068G BC0212083G BD0501109G BB0012027G BB0010207G BC0212085G BC0212086G BD0501087G BB0203050G BB0010291G BC0212079G BE0611022I BE0611009I BB0012060G BD0510024I BD0505068I BD0510022I BD0506125I BC0212048G BD0505090I BB0102059G BC0312055G BC0206146G BC0206073G BD0506124I BC0303082G BB0102075G BC0408075G BC0405047G BD0505106I BB0012131G BC0212046G BC0312066G BC0205068G BC0405052G BB0012004G BB0012013G BC0312049G BD0501066G BD0507063I BB0012063G BC0206083G BD0501092G BC0206075G BC0405023G BB0012012G BD0501071G BB0012074G BC0206071G BB0101023G BD0505053I BD0502003G BD0506101I BD0507061I BC0408052G BB0103059G BC0205089G BC0206046G BD0505080I BD0510009I BB0012160G BC0206134G BD0501081G BA0005349G BB0012146G BC0312044G BB0011120G BB0010215G BC0206027G BB0011180G BB0010082G BB0011170G BB0010297G BD0507038I BD0505077I BB0012001G BB0101020G BB0012176G BD0505074I BB0012143G BB0012141G BD0505107I BD0505033I BD0502038G BB0011036G BC0206016G BD0506103I BC0405064G BB0203046G BC0206080G BB0203020G BB0012023G BC0405045G BD0506143I BD0505040I BD0506144I BB0203018G BC0312068G BD0507069I BB0103074G BD0501075G BB0010052G BE0611069I BB0010076G BA0006181G BD0505050I BB0011105G BC0405049G BB0012008G BB0101083G BD0502031G BD0507020I BB0010216G BC0206154G BD0506098I BE0611048I BB0103022G BB0010222G BD0505048I BE0611063I BA0006183G BD0501070G BC0405006G BD0507001I BB0012181G BB0012022G BC0206042G BB0103041G BB0010254G BC0312069G BB0010269G BD0507043I BB0103047G BB0011046G BE0611024I BD0506099I BC0212061G BB0012162G BB0011081G BD0506117I BB0203013G BD0507024I BB0012158G BD0505044I BE0603014I BB0012071G BC0206116G BC0206109G BD0505066I BC0212081G BD0510014I BE0611005I BD0506110I BC0206102G BB0012002G BC0405053G BB0103039G BB0011165G BA0006179G BB0103058G BB0012119G BB0103046G BD0505036I BD0506109I BE0611073I BD0501078G BB0010285G BC0206100G BB0012018G BD0502042G BD0505101I BC0206149G BC0303090G BC0205086G BD0505085I BB0012024G BB0010105G BB0010183G BB0010051G BB0011164G BD0507044I BC0212062G BB0102077G BC0206079G BB0011038G BB0010230G BC0206018G BB0010284G BB0010049G BC0212056G BC0205071G BD0501107G BB0203011G BD0505089I BB0203022G BE0611076I BB0102064G BB0203057G BB0203041G BB0011108G BD0510020I BC0205087G BD0510031I BD0507010I BC0212078G BD0505072I BC0206148G BC0206099G BB0010241G BA0004165G BB0203035G BB0010289G BB0010184G BB0103062G BB0103043G BD0506097I BD0501089G BD0509049I BB0012136G BB0012133G BC0206098G BE0611016I BD0507057I BD0502002G BB0010214G BD0505030I BD0502015G BB0011112G BD0505047I BC0205077G BB0012135G BD0507034I BC0405005G BB0010061G BE0611021I BB0010264G BD0505058I BB0012145G BC0205062G BB0012116G BA0007063G BE0611033I BC0408069G BC0206062G BB0010279G BD0505038I BC0303080G BB0010293G BB0203044G BB0011161G BD0507058I BD0505065I BD0507002I BB0010283G BC0405032G BC0206033G BE0611006I BB0012113G BA0006170G BC0206153G BE0611070I BD0502008G BB0011102G BD0507046I BD0505042I BC0405048G BE0611010I BC0206072G BC0205074G BD0507035I BB0011104G BD0505009I BC0405008G BC0206141G BB0011118G BB0011040G BB0012132G BB0010060G BD0506121I BC0206145G BD0507072I BC0408058G BB0010071G BB0012134G BB0012033G BE0611079I BC0312053G BB0011116G BC0206024G BB0102067G BE0611011I BB0101087G BC0212088G BC0212084G BD0507062I BD0502025G BD0510017I BE0611019I BB0012085G BB0012086G BC0408066G BD0507064I BB0012200G BB0011092G BD0507031I BB0101080G BC0408073G BC0405039G BC0212087G BB0012079G BD0501064G BB0203002G BB0103038G BD0502012G BB0012054G BD0501091G BA0006174G BC0303076G BB0012155G BB0012026G BE0611030I BD0505087I BC0206107G BC0405013G BC0405028G BD0502027G BB0011029G BE0611029I BB0012163G BC0206155G BB0101027G BC0206133G BC0303088G BB0012003G BE0611046I BC0405065G BD0505049I BD0501093G BD0501108G BB0012114G BB0010242G BC0405021G BB0102060G BB0011126G BB0010062G BB0012185G BD0506093I BD0502022G BE0611028I BB0010238G BB0010070G BB0010265G BC0405062G BB0012078G BD0502006G BB0203043G BD0506091I BB0011026G BD0501068G BD0510007I BC0206115G BC0312067G BB0010256G BC0405043G BB0010288G BB0011137G BD0507039I BE0611047I BD0505067I BD0501110G BD0505041I BB0101081G BB0103052G BD0507049I BD0507013I BB0012120G BB0010075G BC0405066G BB0102062G BE0603011I BE0611003I BD0502043G BB0011167G BB0011158G BC0206081G BB0010185G BA0004159G BC0405009G BE0611015I BC0303089G BB0011131G BC0408079G BD0502024G BD0502016G BB0011002G BB0103069G BB0103024G BB0012115G BD0501082G BB0101084G BD0510016I BD0505104I BD0505019I BD0505051I BB0103061G BD0505086I BC0303086G BB0012186G BB0012178G BD0501101G BC0405042G BB0011157G BB0010206G BC0206093G BC0405067G BB0011090G BC0405030G BC0206031G BB0012180G BD0510023I BC0206152G BD0505105I BD0501095G BD0502040G BB0012011G BB0010068G BB0010300G BD0501111G BB0011179G BC0206150G BB0012077G BB0011095G BD0505079I BB0012036G BB0203008G BC0212071G BC0312046G BC0206103G BB0010190G BD0505096I BE0611018I BB0203054G BC0405020G BD0506129I BB0010224G BC0206077G BD0505059I BD0507056I BB0010228G BB0011138G BC0206151G BB0010217G BB0010239G BB0011098G BB0012159G BD0505037I BB0010245G BC0206088G BD0507073I BA0006098G BC0206091G BE0603010I BC0206125G BC0405033G BD0510013I BB0012127G BD0507036I BB0010210G BC0206084G BC0405012G BB0012081G BE0611044I BE0611050I BB0102070G BB0102079G BE0611014I BD0507030I BC0312042G BD0506127I BC0205093G BD0501073G BB0012111G BD0506126I BC0212082G BD0502034G BD0506108I BA0006175G BC0405011G BE0611002I BA0005347G BC0206143G BB0012149G BC0303087G BB0011034G BB0203029G BB0103056G BC0405040G BC0312060G BD0502029G BD0507071I BC0206044G BC0312063G BC0206119G BD0510029I BC0312065G BD0501090G BD0507070I BD0501088G BE0611080I BD0505012I BB0010081G BB0203047G BD0506137I BB0103064G BA0006100G BB0010253G BB0102021G BB0102082G BB0012059G BB0011134G BD0507041I BE0611012I BD0505083I BC0405059G BC0206104G BD0501117G BD0506142I BB0012053G BB0012069G BC0205076G BD0507029I BE0603005I BB0011121G BC0408071G BD0505007I BC0205091G BD0507022I BB0011156G BC0212055G BB0011144G BC0205069G BB0010182G BC0312054G BB0010189G BB0010275G BB0010259G BB0203048G BB0012021G BD0502004G BC0206057G BD0501084G BB0102065G BD0502018G BD0505022I BC0206069G BD0505082I BC0206043G BD0506130I BE0611032I BB0011024G BE0611034I BC0405001G BC0206142G BD0506113I BD0507015I BD0502037G BB0011001G BB0010225G BC0408060G BB0010281G BB0111034G BD0509045I BB0012080G BD0501104G BB0103042G BD0505015I BB0011114G BB0010258G BD0505011I BD0501065G BE0603007I BB0012061G BD0507004I BB0012157G BE0611059I BC0206092G BB0103050G BB0011110G BE0611026I BB0010108G BC0212064G BC0205078G BA0007064G BB0011021G BD0505035I BB0012179G BD0507037I BD0510005I BE0611038I BB0010074G BC0206124G BD0507076I BC0212077G BB0102080G BC0206021G BB0103036G BC0405041G BD0502039G BB0103035G BB0010255G BD0506105I BB0103028G BB0102081G BC0206122G BA0007001G BD0505081I BD0506095I BB0010218G BC0405056G BD0505100I BE0611020I BB0011109G BB0012137G BC0212066G BB0010110G BC0312058G BB0011087G BC0405017G BC0206053G BC0405035G BB0010208G BC0408074G BC0206106G BD0510030I BB0012165G BC0303075G BD0506114I BD0502010G BC0303081G BB0011159G BB0010229G BB0011163G BD0507011I BE0611051I BE0611040I BC0405044G BB0012082G BB0012124G BC0206019G BD0502014G BB0011171G BB0010251G BC0206032G BD0507017I BC0205070G BB0111033G BE0603013I BE0603008I BB0012015G BA0006185G BC0205061G BB0103060G BB0103067G BC0408055G BD0502009G BB0011175G BD0505017I BD0505064I BD0501083G BB0010221G BD0506128I BC0206114G BD0505021I BB0102019G BC0312050G BC0212090G BE0611075I BC0206040G BD0507003I BC0206066G BB0012126G BD0506100I BE0611060I BD0507075I BD0510032I BE0611027I BB0011100G BD0507074I BE0611017I BC0206035G BB0203005G BC0312061G BB0102066G BC0212051GCS = "Case Sensitive"BB0011103G BC0206105G BC0408064G BB0011111G BD0505094I BC0405016G BB0102071G BC0206096G BD0507033I BD0505110I BD0507019I BC0206076G BE0611036I BB0203034G BB0011099G BB0103048G BC0408067G BD0509044I BB0012010G BD0501100G BC0206131G BD0501096G BA0006167G BB0011169G BD0510018I BD0505029I BD0505039I BB0102074G BD0507014I BB0101021G BD0506112I BA0007003G BB0010078G BB0010213G BC0206089G BC0205079G BD0501099G BA0004162G BB0011101G BD0501094G BB0103027G BB0012068G BD0507005I BB0011003G BB0011086G BC0408057G BB0012075G BA0005050G BC0303071G BD0507060I BB0010268G BB0012154G BC0303073G BA0006171G BB0011080G BC0312045G BC0206144G BD0505016I BC0206029G BB0012184G BD0505018I BC0405050G BD0501085G BD0506136I BB0011083G BB0012056G BE0611013I BD0507053I BD0505028I BC0408078G BE0603002I BE0603012I BB0011166G BB0012150G BB0012067G BB0012171G BB0011020G BC0212074G BB0010250G BD0507025I BB0203012G BD0501061G BC0206110G BE0611023I BB0011141G BD0502041G BE0611087I BB0012017G BB0010193G BC0205082G BD0509048I BA0006178G BB0010252G BB0101092G BB0012182G BC0405037G BD0505001I BD0507065I BC0405022G BB0203059G BD0505097I BD0502013G BB0012125G BD0506123I BD0501074G BB0102069G BC0405024G BB0012151G BB0012183G BB0203052G BD0505061I BE0611062I BC0212076G BD0510028I BB0010063G BB0012019G BB0203028G BB0012062G BB0010057G BC0212080G BE0603018I BC0312059G BE0611082I BC0405025G BC0206067G BB0010273G BB0102068G BB0011032G BB0010080G BB0010271G BA0005233G BD0506111I BD0502035G BD0509046I BC0206138G BD0505002I BB0012166G BC0408072G BD0506106I BE0611041I BE0611043I BA0004161G BB0011037G BB0203037G BD0505045I BD0505052I BD0505008I BB0010294G BB0011042G BB0012070G BD0501118G BD0501116G BE0611008I BC0212068G BC0212083G BD0501109G BB0012027G BB0010207G BC0212085G BC0212086G BD0501087G BB0203050G BB0010291G BC0212079G BE0611022I BE0611009I BB0012060G BD0510024I BD0505068I BD0510022I BD0506125I BC0212048G BD0505090I BB0102059G BC0312055G BC0206146G BC0206073G BD0506124I BC0303082G BB0102075G BC0408075G BC0405047G BD0505106I BB0012131G BC0212046G BC0312066G BC0205068G BC0405052G BB0012004G BB0012013G BC0312049G BD0501066G BD0507063I BB0012063G BC0206083G BD0501092G BC0206075G BC0405023G BB0012012G BD0501071G BB0012074G BC0206071G BB0101023G BD0505053I BD0502003G BD0506101I BD0507061I BC0408052G BB0103059G BC0205089G BC0206046G BD0505080I BD0510009I BB0012160G BC0206134G BD0501081G BA0005349G BB0012146G BC0312044G BB0011120G BB0010215G BC0206027G BB0011180G BB0010082G BB0011170G BB0010297G BD0507038I BD0505077I BB0012001G BB0101020G BB0012176G BD0505074I BB0012143G BB0012141G BD0505107I BD0505033I BD0502038G BB0011036G BC0206016G BD0506103I BC0405064G BB0203046G BC0206080G BB0203020G BB0012023G BC0405045G BD0506143I BD0505040I BD0506144I BB0203018G BC0312068G BD0507069I BB0103074G BD0501075G BB0010052G BE0611069I BB0010076G BA0006181G BD0505050I BB0011105G BC0405049G BB0012008G BB0101083G BD0502031G BD0507020I BB0010216G BC0206154G BD0506098I BE0611048I BB0103022G BB0010222G BD0505048I BE0611063I BA0006183G BD0501070G BC0405006G BD0507001I BB0012181G BB0012022G BC0206042G BB0103041G BB0010254G BC0312069G BB0010269G BD0507043I BB0103047G BB0011046G BE0611024I BD0506099I BC0212061G BB0012162G BB0011081G BD0506117I BB0203013G BD0507024I BB0012158G BD0505044I BE0603014I BB0012071G BC0206116G BC0206109G BD0505066I BC0212081G BD0510014I BE0611005I BD0506110I BC0206102G BB0012002G BC0405053G BB0103039G BB0011165G BA0006179G BB0103058G BB0012119G BB0103046G BD0505036I BD0506109I BE0611073I BD0501078G BB0010285G BC0206100G BB0012018G BD0502042G BD0505101I BC0206149G BC0303090G BC0205086G BD0505085I BB0012024G BB0010105G BB0010183G BB0010051G BB0011164G BD0507044I BC0212062G BB0102077G BC0206079G BB0011038G BB0010230G BC0206018G BB0010284G BB0010049G BC0212056G BC0205071G BD0501107G BB0203011G BD0505089I BB0203022G BE0611076I BB0102064G BB0203057G BB0203041G BB0011108G BD0510020I BC0205087G BD0510031I BD0507010I BC0212078G BD0505072I BC0206148G BC0206099G BB0010241G BA0004165G BB0203035G BB0010289G BB0010184G BB0103062G BB0103043G BD0506097I BD0501089G BD0509049I BB0012136G BB0012133G BC0206098G BE0611016I BD0507057I BD0502002G BB0010214G BD0505030I BD0502015G BB0011112G BD0505047I BC0205077G BB0012135G BD0507034I BC0405005G BB0010061G BE0611021I BB0010264G BD0505058I BB0012145G BC0205062G BB0012116G BA0007063G BE0611033I BC0408069G BC0206062G BB0010279G BD0505038I BC0303080G BB0010293G BB0203044G BB0011161G BD0507058I BD0505065I BD0507002I BB0010283G BC0405032G BC0206033G BE0611006I BB0012113G BA0006170G BC0206153G BE0611070I BD0502008G BB0011102G BD0507046I BD0505042I BC0405048G BE0611010I BC0206072G BC0205074G BD0507035I BB0011104G BD0505009I BC0405008G BC0206141G BB0011118G BB0011040G BB0012132G BB0010060G BD0506121I BC0206145G BD0507072I BC0408058G BB0010071G BB0012134G BB0012033G BE0611079I BC0312053G BB0011116G BC0206024G BB0102067G BE0611011I BB0101087G BC0212088G BC0212084G BD0507062I BD0502025G BD0510017I BE0611019I BB0012085G BB0012086G BC0408066G BD0507064I BB0012200G BB0011092G BD0507031I BB0101080G BC0408073G BC0405039G BC0212087G BB0012079G BD0501064G BB0203002G BB0103038G BD0502012G BB0012054G BD0501091G BA0006174G BC0303076G BB0012155G BB0012026G BE0611030I BD0505087I BC0206107G BC0405013G BC0405028G BD0502027G BB0011029G BE0611029I BB0012163G BC0206155G BB0101027G BC0206133G BC0303088G BB0012003G BE0611046I BC0405065G BD0505049I BD0501093G BD0501108G BB0012114G BB0010242G BC0405021G BB0102060G BB0011126G BB0010062G BB0012185G BD0506093I BD0502022G BE0611028I BB0010238G BB0010070G BB0010265G BC0405062G BB0012078G BD0502006G BB0203043G BD0506091I BB0011026G BD0501068G BD0510007I BC0206115G BC0312067G BB0010256G BC0405043G BB0010288G BB0011137G BD0507039I BE0611047I BD0505067I BD0501110G BD0505041I BB0101081G BB0103052G BD0507049I BD0507013I BB0012120G BB0010075G BC0405066G BB0102062G BE0603011I BE0611003I BD0502043G BB0011167G BB0011158G BC0206081G BB0010185G BA0004159G BC0405009G BE0611015I BC0303089G BB0011131G BC0408079G BD0502024G BD0502016G BB0011002G BB0103069G BB0103024G BB0012115G BD0501082G BB0101084G BD0510016I BD0505104I BD0505019I BD0505051I BB0103061G BD0505086I BC0303086G BB0012186G BB0012178G BD0501101G BC0405042G BB0011157G BB0010206G BC0206093G BC0405067G BB0011090G BC0405030G BC0206031G BB0012180G BD0510023I BC0206152G BD0505105I BD0501095G BD0502040G BB0012011G BB0010068G BB0010300G BD0501111G BB0011179G BC0206150G BB0012077G BB0011095G BD0505079I BB0012036G BB0203008G BC0212071G BC0312046G BC0206103G BB0010190G BD0505096I BE0611018I BB0203054G BC0405020G BD0506129I BB0010224G BC0206077G BD0505059I BD0507056I BB0010228G BB0011138G BC0206151G BB0010217G BB0010239G BB0011098G BB0012159G BD0505037I BB0010245G BC0206088G BD0507073I BA0006098G BC0206091G BE0603010I BC0206125G BC0405033G BD0510013I BB0012127G BD0507036I BB0010210G BC0206084G BC0405012G BB0012081G BE0611044I BE0611050I BB0102070G BB0102079G BE0611014I BD0507030I BC0312042G BD0506127I BC0205093G BD0501073G BB0012111G BD0506126I BC0212082G BD0502034G BD0506108I BA0006175G BC0405011G BE0611002I BA0005347G BC0206143G BB0012149G BC0303087G BB0011034G BB0203029G BB0103056G BC0405040G BC0312060G BD0502029G BD0507071I BC0206044G BC0312063G BC0206119G BD0510029I BC0312065G BD0501090G BD0507070I BD0501088G BE0611080I BD0505012I BB0010081G BB0203047G BD0506137I BB0103064G BA0006100G BB0010253G BB0102021G BB0102082G BB0012059G BB0011134G BD0507041I BE0611012I BD0505083I BC0405059G BC0206104G BD0501117G BD0506142I BB0012053G BB0012069G BC0205076G BD0507029I BE0603005I BB0011121G BC0408071G BD0505007I BC0205091G BD0507022I BB0011156G BC0212055G BB0011144G BC0205069G BB0010182G BC0312054G BB0010189G BB0010275G BB0010259G BB0203048G BB0012021G BD0502004G BC0206057G BD0501084G BB0102065G BD0502018G BD0505022I BC0206069G BD0505082I BC0206043G BD0506130I BE0611032I BB0011024G BE0611034I BC0405001G BC0206142G BD0506113I BD0507015I BD0502037G BB0011001G BB0010225G BC0408060G BB0010281G BB0111034G BD0509045I BB0012080G BD0501104G BB0103042G BD0505015I BB0011114G BB0010258G BD0505011I BD0501065G BE0603007I BB0012061G BD0507004I BB0012157G BE0611059I BC0206092G BB0103050G BB0011110G BE0611026I BB0010108G BC0212064G BC0205078G BA0007064G BB0011021G BD0505035I BB0012179G BD0507037I BD0510005I BE0611038I BB0010074G BC0206124G BD0507076I BC0212077G BB0102080G BC0206021G BB0103036G BC0405041G BD0502039G BB0103035G BB0010255G BD0506105I BB0103028G BB0102081G BC0206122G BA0007001G BD0505081I BD0506095I BB0010218G BC0405056G BD0505100I BE0611020I BB0011109G BB0012137G BC0212066G BB0010110G BC0312058G BB0011087G BC0405017G BC0206053G BC0405035G BB0010208G BC0408074G BC0206106G BD0510030I BB0012165G BC0303075G BD0506114I BD0502010G BC0303081G BB0011159G BB0010229G BB0011163G BD0507011I BE0611051I BE0611040I BC0405044G BB0012082G BB0012124G BC0206019G BD0502014G BB0011171G BB0010251G BC0206032G BD0507017I BC0205070G BB0111033G BE0603013I BE0603008I BB0012015G BA0006185G BC0205061G BB0103060G BB0103067G BC0408055G BD0502009G BB0011175G BD0505017I BD0505064I BD0501083G BB0010221G BD0506128I BC0206114G BD0505021I BB0102019G BC0312050G BC0212090G BE0611075I BC0206040G BD0507003I BC0206066G BB0012126G BD0506100I BE0611060I BD0507075I BD0510032I BE0611027I BB0011100G BD0507074I BE0611017I BC0206035G BB0203005G BC0312061G BB0102066G BC0212051G Recalling Firm/ Manufacturer Sorin Group USA, Inc. 14401 W 65th Way Arvada CO 80004-3503 For Additional Information Contact Joan Ceasar 281-685-2227 Manufacturer Reason for Recall Sorin Group USA, Inc. announces a voluntary field action for the ORCHESTRA Programmer and ORCHESTRA PLUS******Programmer because the residual longevity estimate may be less accurate. FDA Determined Cause 2 Software Design Change Action The firm, LivaNova, sent an "Urgent Field Safety Notice" dated 9/29/2016 to its customers. The notice described the product, problem, actions to be taken. The notice states that a Sorin representative will contact customers to upgrade the software in the facility's programmer.************If customers have any questions, they are urged to call their LivaNova representative or contact LivaNova Technical Services at 877.663.7674. Quantity in Commerce 847 programmers Distribution Programmers were distributed nationwide, to VA/govt/military, Canadian, and other foreign consignees. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, LivaNova, sent an "Urgent Field Safety Notice" dated 9/29/2016 to its customers. The notice described the product, problem, actions to be taken. The notice states that a Sorin representative will contact customers to upgrade the software in the facility's programmer.************If customers have any questions, they are urged to call their LivaNova representative or contact LivaNova Technical Services at 877.663.7674. | SoftwareDesignChange
126 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150546 | Z-0410-2017 | MEVION S250 ******Product Usage: Proton Radiation Therapy System | MEVION S250 | 2 | CDRH | 2016-11-14 | 2016-11-15 21:09 | 2017-05-10 | Mevion Medical Systems, Inc. | Recovery of Incorrect Isocenter -operational change | Class 2 Device Recall MEVION S250 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall MEVION S250 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall MEVION S250 Date Initiated by Firm October 21, 2016 Date Posted November 15, 2016 Recall Status 1 Terminated 3 on May 10, 2017 Recall Number Z-0410-2017 Recall Event ID 75498 510(K)Number K120676 Product Classification System,planning,radiation therapy treatment - Product Code MUJ Product MEVION S250 ************Product Usage: Proton Radiation Therapy System Code Information Serial Number: S250-0001 through S250-0006 Recalling Firm/ Manufacturer Mevion Medical Systems, Inc. 300 Foster Street Littleton MA 01460-2017 For Additional Information Contact 978-540-1500 Manufacturer Reason for Recall Recovery of Incorrect Isocenter -operational change FDA Determined Cause 2 Software Design Change Action Mevion issued an User Notice letter on October 24, 2016 to their customers. The letter identified the affected product, problem and actions taken. Customers were informed that Mevion will be updating software when it is released. For further clarification or assistance, please contact Rob Cessac by email at rcessac@mevion.com or by phone at (978) 540-1501. Quantity in Commerce 6 units Distribution US Nationwide Distribution in the states of FL, MO, NJ, OH and OK, Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MUJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Mevion issued an User Notice letter on October 24, 2016 to their customers. The letter identified the affected product, problem and actions taken. Customers were informed that Mevion will be updating software when it is released. For further clarification or assistance, please contact Rob Cessac by email at rcessac@mevion.com or by phone at (978) 540-1501. | SoftwareDesignChange
127 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=149377 | Z-2900-2016 | AQUIOS CL Flow Cytometer, Catalog No. B30166; Intended for use with in vitro diagnostic flow cytometry applications using up to four florescent detection channels. | Aquios CL Flow Cytometer | 2 | CDRH | 2016-09-28 | 2016-09-28 21:01 | 2017-08-07 | Beckman Coulter Inc. | Beckman Coulter is recalling the AQUIOS CL System because the use of the reagent carousel can lead to misalignment of the carousel and damage to the prep probe. | Class 2 Device Recall Aquios CL Flow Cytometer Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Aquios CL Flow Cytometer FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Aquios CL Flow Cytometer Date Initiated by Firm September 06, 2016 Date Posted September 28, 2016 Recall Status 1 Terminated 3 on August 07, 2017 Recall Number Z-2900-2016 Recall Event ID 75111 510(K)Number K141932 Product Classification Flow cytometric reagents and accessories. - Product Code OYE Product AQUIOS CL Flow Cytometer, Catalog No. B30166; Intended for use with in vitro diagnostic flow cytometry applications using up to four florescent detection channels. Code Information AY20032 AZ19046 AZ12030 AY51072 AZ19047 AZ15044 AY47071 AZ15038 AY47063 AZ12033 AZ19052 AY51075 AY03004 AZ12027 AY25038 Recalling Firm/ Manufacturer Beckman Coulter Inc. 250 S Kraemer Blvd Brea CA 92821-6232 Manufacturer Reason for Recall Beckman Coulter is recalling the AQUIOS CL System because the use of the reagent carousel can lead to misalignment of the carousel and damage to the prep probe. FDA Determined Cause 2 Software Design Change Action An Urgent Medical Device Recall letter dated 9/6/16 was sent to customers to inform them that the use of Positions 9 and 10 of the reagent carousel can lead to misalignment of the carousel and damage to the prep probe. The letter informs the customers of the impact of the recall, action to be taken, and the resolution by Beckman Coulter. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding this notice, are instructed to contact Customer Support Center: http://www.beckmancoulter.com/customersupport/support; call 1-800-369-0333 in the US and Canada; outside the US and Canada, contact their local Beckman Coulter Representative. Quantity in Commerce 148 units total (12 units in US) Distribution US: IL, PA, NY, GA, WA, FL, SC.************Affected geographies: Angola, Australia, Belgium, Botswana, Canada, Chile, Columbia, Czech Republic, Denmark, Finland, France, Germany, Ghana, Israel, Italy, Malaysia, Nambia, Netherlands, Panama, Poland, Romania, Slovakia, South Africa, Spain, Sweden, Switzerland, Uganda, United Kingdom and Zambia Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = OYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action An Urgent Medical Device Recall letter dated 9/6/16 was sent to customers to inform them that the use of Positions 9 and 10 of the reagent carousel can lead to misalignment of the carousel and damage to the prep probe. The letter informs the customers of the impact of the recall, action to be taken, and the resolution by Beckman Coulter. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding this notice, are instructed to contact Customer Support Center: http://www.beckmancoulter.com/customersupport/support; call 1-800-369-0333 in the US and Canada; outside the US and Canada, contact their local Beckman Coulter Representative. | SoftwareDesignChange
128 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148524 | Z-2657-2016 | Ingenuity Core, Model number 728321******Product Usage:***Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes.*** | Philips Healthcare Ingenuity Core | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. | Class 2 Device Recall Philips Healthcare Ingenuity Core Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare Ingenuity Core FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare Ingenuity Core Date Initiated by Firm July 20, 2016 Create Date August 25, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2657-2016 Recall Event ID 74820 510(K)Number K033326 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Ingenuity Core, Model number 728321************Product Usage:******Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information 728321 -- Ingenuity Core 310004 310005***310012***310016***310020***310021***310027***310028***310033***310034***310041***310044***310045***310057***310061***310065***310072***310074***310076***310077***310078***310094***310099***310100***310103***310105***310117***310120***310122***310135 Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mrs. Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. FDA Determined Cause 2 Software Design Change Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). Quantity in Commerce 30 Distribution US Nationwide Distribution in the states of: AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,******MD,ME,MI,MN,MO,MS,NC,ND,NE,NJ,NM,******NV,NY,OH,OK,OR,PA,PR,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). | SoftwareDesignChange
129 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148523 | Z-2656-2016 | Brilliance CT 16 Power, Model number 728240******Product Usage:***Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes.*** | Philips Healthcare Brilliance CT 16 Power | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. | Class 2 Device Recall Philips Healthcare Brilliance CT 16 Power Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare Brilliance CT 16 Power FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare Brilliance CT 16 Power Date Initiated by Firm July 20, 2016 Create Date August 25, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2656-2016 Recall Event ID 74820 510(K)Number K012009 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Brilliance CT 16 Power, Model number 728240************Product Usage:******Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information 728240 -- Brilliance CT 16 Power 3100***5043***5047***5280***5345***5390***5413***5430***5495***5499***5512***6007***6009***6013***6032***6041***6045***6049***6051***6058***6064***6066***6070***6071***6077***6080***6081***6084***6089***6095***6100***6106***6109***6113***6115***6116***6123***6124***6131***6132***6134***6141***6144***6155***6161***6162***5476***6098***6120***6067 Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mrs. Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. FDA Determined Cause 2 Software Design Change Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). Quantity in Commerce 50 Distribution US Nationwide Distribution in the states of: AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,******MD,ME,MI,MN,MO,MS,NC,ND,NE,NJ,NM,******NV,NY,OH,OK,OR,PA,PR,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). | SoftwareDesignChange
130 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148525 | Z-2658-2016 | Ingenuity Core128, Model number 728323******Product Usage:***Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes.*** | Philips Healthcare Ingenuity Core128 | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. | Class 2 Device Recall Philips Healthcare Ingenuity Core128 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare Ingenuity Core128 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare Ingenuity Core128 Date Initiated by Firm July 20, 2016 Create Date August 25, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2658-2016 Recall Event ID 74820 510(K)Number K033326 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Ingenuity Core128, Model number 728323************Product Usage:******Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information 728323 -- Ingenuity Core128 320004***320053*** 320056 Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mrs. Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. FDA Determined Cause 2 Software Design Change Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). Quantity in Commerce 3 Distribution US Nationwide Distribution in the states of: AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,******MD,ME,MI,MN,MO,MS,NC,ND,NE,NJ,NM,******NV,NY,OH,OK,OR,PA,PR,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). | SoftwareDesignChange
131 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148520 | Z-2654-2016 | Brilliance BigBore Radiology CT Model number 728244******Product Usage:***Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes.*** | Philips Healthcare Brilliance BigBore Radiology CT | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. | Class 2 Device Recall Philips Healthcare Brilliance BigBore Radiology CT Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare Brilliance BigBore Radiology CT FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare Brilliance BigBore Radiology CT Date Initiated by Firm July 20, 2016 Create Date August 25, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2654-2016 Recall Event ID 74820 510(K)Number K033357 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Brilliance BigBore Radiology CT Model number 728244************Product Usage:******Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information 728244 *** Brilliance BigBore Radiology CT 7375***7781***7614***7633***7708***7789***7721***7760***7252***7380***7390***7514***7052***7244***7384***7533***7172***7200***7546***7427***7452***7407***7453***7586 Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mrs. Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. FDA Determined Cause 2 Software Design Change Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). Quantity in Commerce 24 Distribution US Nationwide Distribution in the states of: AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,******MD,ME,MI,MN,MO,MS,NC,ND,NE,NJ,NM,******NV,NY,OH,OK,OR,PA,PR,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). | SoftwareDesignChange
132 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148519 | Z-2653-2016 | Brilliance CT Big Bore Oncology, Model number 728243******Product Usage:***Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes.*** | Philips Healthcare Brilliance BigBore Oncology CT | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. | Class 2 Device Recall Philips Healthcare Brilliance BigBore Oncology CT Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare Brilliance BigBore Oncology CT FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare Brilliance BigBore Oncology CT Date Initiated by Firm July 20, 2016 Create Date August 25, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2653-2016 Recall Event ID 74820 510(K)Number K033357 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Brilliance CT Big Bore Oncology, Model number 728243************Product Usage:******Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information 728243 *** Brilliance BigBore Oncology CT 7116***7611***7354***7063***7309***7169***7205***7016***7019***7043***7069***7073***7091***7112***7132***7163***7141***7194***7226***7325***7294***7350***7316***7480***7114***7210***7310***7438***7530***7186***7119***7042 Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mrs. Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. FDA Determined Cause 2 Software Design Change Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). Quantity in Commerce 33 Distribution US Nationwide Distribution in the states of: AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,******MD,ME,MI,MN,MO,MS,NC,ND,NE,NJ,NM,******NV,NY,OH,OK,OR,PA,PR,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). | SoftwareDesignChange
133 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148521 | Z-2655-2016 | Brilliance 16, Model number 728246******Product Usage:***Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes.*** | Philips Healthcare Brilliance 16 | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. | Class 2 Device Recall Philips Healthcare Brilliance 16 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare Brilliance 16 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare Brilliance 16 Date Initiated by Firm July 20, 2016 Create Date August 25, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2655-2016 Recall Event ID 74820 510(K)Number K012009 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Brilliance 16, Model number 728246************Product Usage:******Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information 728246 -- Brilliance 16 3010***3017***3031***3040***3043***3070***3072***3085***3090***3117***3149***3216***3247***3293***3314***5006***5012***5018***5020***5024***5026***5028***5030***5031***5035***5044***5045***5052***5053***5054***5055***5058***5059***5060***5064***5069***5070***5081***5086***5088***5091***5093***5097***5099***5103***5107***5108***5110***5111***5115***5116***5118***5120***5123***5126***5127***5128***5131***5138***5140***5155***5160***5161***5163***5166***5184***5186***5188***5190***5191***5192***5193***5195***5197***5208***5209***5216***5219***5235***5236***5238A***5270***5272***5279***5282***5285***5287***5294***5295***5315***5316***5350***5354***5357***5360***5368***5394***5433***5434***5443***5453***5457***5480***5487***5501***5502***5505***5519***5523***5526***5554***5572***5573***5576***5609***5615***5633***5647***5653***5683***5723***5758***5762***5764***5771***5811***5897***5929***5937***5956***5974***5995***6021***6022***6031***6033***6048***6053***6072***6220***6229***6243***6280***6283***6298***6309***6321***6360***6375***6405***6426***6441***6465***6474***6484***6503***6511***6517***6519***6523***6549***6559***30049***32287***50011***50053***50087***50180***3123***5674***6555***32366***3465***50181***5135***5237***3179***3411***5009***5048***5077***5412***6154***30011***30110***5101***5181***3324***5405***3050***3441***6387 Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mrs. Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. FDA Determined Cause 2 Software Design Change Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). Quantity in Commerce 192 Distribution US Nationwide Distribution in the states of: AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,******MD,ME,MI,MN,MO,MS,NC,ND,NE,NJ,NM,******NV,NY,OH,OK,OR,PA,PR,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). | SoftwareDesignChange
134 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148518 | Z-2652-2016 | Brilliance 64 CT Model number 728231******Product Usage:***Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes.*** | Philips Healthcare Brilliance CT 64 | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. | Class 2 Device Recall Philips Healthcare Brilliance CT 64 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare Brilliance CT 64 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare Brilliance CT 64 Date Initiated by Firm July 20, 2016 Create Date August 25, 2016 Recall Status 1 Terminated 3 on July 02, 2018 Recall Number Z-2652-2016 Recall Event ID 74820 510(K)Number K033326 Product Classification System, x-ray, tomography, computed - Product Code JAK Product Brilliance 64 CT Model number 728231************Product Usage:******Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. Code Information 9033***9037***9043***9045***9064***9065***9070***9075***9086***9089***9104***9108***9120***9151***9156***9172***9190***9243***9259***9507***9508***9511***9526***9530***9531***9534***9547***9564***9580***9584***9598***9603***9606***9608***9613***9614***9630***9642***9655***9672***9695***9697***9690***9706***9713***9732***9735***9736***9757***9759***9812***9830***9835***9866***9875***9881***9889***9918***9936***9937***9951***9965***9968***9969***9975***9978***9992***9999***10036***10041***10046***10054***10067***10070***10080***10095***10103***10125***10141***10256***10292***10359***10363***10372***10380***10381***10386***10426***10473***10486***10505***10513***10527***10535***10561***10570***10645***10651***10655***10673***10686***10733***10743***10773***10776***29007***29016***29018***29025***29027***29037***29053***29097***90031***90047***90063***90128***95003***95058***95068***95210***95357***95811***29054***9107***9666***9700***29038***9858***9009***9029***9038***9051***9058***9067***9068***9072***9073***9092***9105***9106***9502***9503***9512***9516***9522***9525***9533***9545***9550***9551***9556***9572***9577***9591***9594***9596***9612***9618***9620***9624***9625***9626***9628***9633***9634***9635***9639***9641***9648***9650***9663***9676***9678***9685***9691***9692***9694***9705***9716***9726***9729***9740***9750***9751***9752***9753***9754***9775***9777***9780***9788***9790***9793***9809***9811***9821***9823***9826***9843***9846***9855***9857***9861***9869***9873***9878***9879***9882***9885***9894***9895***9902***9906***9917***9924***9938***9942***9948***9954***9960***9961***9966***9974***9982***9985***9994***10001***10014***10035***10037***10045***10065***10079***10085***10086***10096***10098***10111***10112***10113***10120***10148***10190***10200***10202***10224***10263***10291***10310***10332***10353***10365***10370***10371***10388***10413***10417***10424***29005***29035***29068***64015***90062***95008***95057***95348***9851T***10122***10373***10211***9139***10005***10026***10030***10031***10194***95149***95377***95392***9980***9698***9787 Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Mrs. Holly Wright Lee 440-483-2015 Manufacturer Reason for Recall After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. FDA Determined Cause 2 Software Design Change Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). Quantity in Commerce 283 Distribution US Nationwide Distribution in the states of: AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,******MD,ME,MI,MN,MO,MS,NC,ND,NE,NJ,NM,******NV,NY,OH,OK,OR,PA,PR,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.******If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For******North America and Canada, contact the Customer Care Solutions Center****** (1 -800-722-9377). | SoftwareDesignChange
135 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148151 | Z-2293-2016 | MOSAIQ Oncology Information System | MOSAIQ | 2 | CDRH | 2016-07-27 |  |  | Elekta, Inc. | It is possible that a change to an Order Set will not be saved in the current open Care Plan. This can result in a wrong chemo dose in the order. | Class 2 Device Recall MOSAIQ Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall MOSAIQ FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall MOSAIQ Date Initiated by Firm July 15, 2016 Create Date July 27, 2016 Recall Status 1 Open 3 , Classified Recall Number Z-2293-2016 Recall Event ID 74668 510(K)Number K123230 Product Classification Oncology Information System - Product Code IYE Product MOSAIQ Oncology Information System Code Information Versions 2.50.05 and higher Recalling Firm/ Manufacturer Elekta, Inc. 400 Perimeter Center Ter NE Ste 50 Atlanta GA 30346-1227 For Additional Information Contact Linda Wetsel 770-300-9725 Manufacturer Reason for Recall It is possible that a change to an Order Set will not be saved in the current open Care Plan. This can result in a wrong chemo dose in the order. FDA Determined Cause 2 Software Design Change Action Elekta sent Important Field Safety Notice 371-03-MSQ-001 to all customers on 7/15/2016. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. The letter also included a response form which is to be returned to Elekta. Quantity in Commerce 383 Distribution Worldwide distribution. US nationwide, Antigua and Barbuda, Australia, Austria, Bahamas, Bahrain, Canada, Czech Republic, Germany, Hungary, South Africa, Spain, Sweden, Turkey, and United Kingdom Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Elekta sent Important Field Safety Notice 371-03-MSQ-001 to all customers on 7/15/2016. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. The letter also included a response form which is to be returned to Elekta. | SoftwareDesignChange
136 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=143632 | Z-1172-2016 | Accu-Chek Inform II Base Unit as a part of the Accu-Chek Inform II Blood Glucose Monitoring System***Accu-Chek inform II Blood Glucose Monitoring System is intended for testing outside the body (in vitro diagnostic use) and is intended for multiple-patient use in professional healthcare settings. This system should only be use with single-use, auto-disabling lancing devices. | Accu-Chek Inform II Blood Glucose Monitoring System. | 2 | CDRH | 2016-03-14 | 2016-03-14 21:00 | 2017-05-30 | Roche Diagnostics Operations, Inc. | Accu-Chek Inform II Base Unit might produce physical transmission errors in the form of data loss in the communication between the meter and the Data Management Systems (DMS). The issue can lead to the data loss or in the worst case to an erroneous assignment of the patient data (patient mismatch). The issue will only occur at sites using POTCT1-A communication via USB. | Class 2 Device Recall AccuChek Inform II Blood Glucose Monitoring System. Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall AccuChek Inform II Blood Glucose Monitoring System. FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall AccuChek Inform II Blood Glucose Monitoring System. Date Initiated by Firm February 11, 2016 Date Posted March 14, 2016 Recall Status 1 Terminated 3 on May 30, 2017 Recall Number Z-1172-2016 Recall Event ID 73253 510(K)Number K121679 Product Classification Glucose dehydrogenase, glucose - Product Code LFR Product Accu-Chek Inform II Base Unit as a part of the Accu-Chek Inform II Blood Glucose Monitoring System******Accu-Chek inform II Blood Glucose Monitoring System is intended for testing outside the body (in vitro diagnostic use) and is intended for multiple-patient use in professional healthcare settings. This system should only be use with single-use, auto-disabling lancing devices. Code Information Model/Catalog/Part Number: 05060290001; serial numbers: UU41000000 through UU41064963 and UU41064964 through UU41122741. Recalling Firm/ Manufacturer Roche Diagnostics Operations, Inc. 9115 Hague Rd Indianapolis IN 46256-1025 For Additional Information Contact Todd Siesky 317-521-3911 Manufacturer Reason for Recall Accu-Chek Inform II Base Unit might produce physical transmission errors in the form of data loss in the communication between the meter and the Data Management Systems (DMS). The issue can lead to the data loss or in the worst case to an erroneous assignment of the patient data (patient mismatch). The issue will only occur at sites using POTCT1-A communication via USB. FDA Determined Cause 2 Software Design Change Action Roche sent an URGENT MEDICAL DEVICE RECALL CORRECTION letter dated February 11, 2016, via UPS Ground (receipt required) to the consignee. ******The Consignee instructions were as follows:****** If currently using USB connectivity, follow the instructions outlined in this Urgent Medical Device Correction (UMDC) to determine if your base unit is affected by the issue.****** Until the new software is available, do not use USB connectivity if the base unit is affected by this issue; switch the connectivity method to Ethernet (TCP/IP) or Wi-Fi.************Roche is not requesting return of affected product. Any returned product will be managed according to current procedures.******Consignees with questions were instructed to contact the ACCU-CHEK Customer Care Service Center at 1-800-440-3638.******For questions regarding this recall call 317-521-3911. Quantity in Commerce US 5,604 devices, OUS 91,925 Distribution Nationwide Distribution including AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, WI, and WV. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LFR - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Roche sent an URGENT MEDICAL DEVICE RECALL CORRECTION letter dated February 11, 2016, via UPS Ground (receipt required) to the consignee. ******The Consignee instructions were as follows:****** If currently using USB connectivity, follow the instructions outlined in this Urgent Medical Device Correction (UMDC) to determine if your base unit is affected by the issue.****** Until the new software is available, do not use USB connectivity if the base unit is affected by this issue; switch the connectivity method to Ethernet (TCP/IP) or Wi-Fi.************Roche is not requesting return of affected product. Any returned product will be managed according to current procedures.******Consignees with questions were instructed to contact the ACCU-CHEK Customer Care Service Center at 1-800-440-3638.******For questions regarding this recall call 317-521-3911. | SoftwareDesignChange
137 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142024 | Z-0856-2016 | Philips Healthcare IntelliVue Info Center iX, A.0 866023*** | Philips Healthcare IntelliVue Info Center iX, A.0 | 2 | CDRH | 2016-02-23 | 2016-02-23 21:06 | 2017-09-06 | Philips Electronics North America Corporation | Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads. | Class 2 Device Recall Philips Healthcare IntelliVue Info Center iX, A.0 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare IntelliVue Info Center iX, A.0 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare IntelliVue Info Center iX, A.0 Date Initiated by Firm November 25, 2015 Date Posted February 23, 2016 Recall Status 1 Terminated 3 on September 06, 2017 Recall Number Z-0856-2016 Recall Event ID 72720 510(K)Number K102495 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product Philips Healthcare IntelliVue Info Center iX, A.0 866023 Code Information Software A.0.X Recalling Firm/ Manufacturer Philips Electronics North America Corporation 3000 Minuteman Rd Andover MA 01810-1032 For Additional Information Contact Kelly Rector 978-659-2950 Manufacturer Reason for Recall Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads. FDA Determined Cause 2 Software Design Change Action Philips Healthcare issued a Field Safety Notice on 11/25/15 to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will provide a software correction free of charge. On 1/28/16 the Philips issued an updated letter to users.Philips has initiated a correction to address this issue. A software correction will be provided to customers with impacted devices at no charge.************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Field Safety Notice: ******Arrhythmia alarms are not affected. ST alarms, ST values and measurements are also not affected and could be used to assist with the decision to obtain a diagnostic ECG.************The ECG waves at the bedside are not affected. Printing ECG waves from the bedside is an alternative to printing from the PIIC iX.************If you need any further information or support concerning this issue, contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. Quantity in Commerce 5671 units Distribution Worldwide Distribution-US (nationwide) and the countries of Algeria Australia Austria Belgium Bulgaria Canada Denmark Estonia Finland France Germany Hong Kong Hungary Iceland India Indonesia Iraq Ireland Israel Italy Japan Kenya Korea Kuwait Luxembourg Malaysia Mongolia Morocco Netherlands New Zealand Nigeria Norway Poland Qatar Romania Russian Federation Saudi Arabia Singapore Slovenia South Africa Spain Sweden Switzerland Taiwan Turkey United Arab Emirates and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare issued a Field Safety Notice on 11/25/15 to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will provide a software correction free of charge. On 1/28/16 the Philips issued an updated letter to users.Philips has initiated a correction to address this issue. A software correction will be provided to customers with impacted devices at no charge.************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Field Safety Notice: ******Arrhythmia alarms are not affected. ST alarms, ST values and measurements are also not affected and could be used to assist with the decision to obtain a diagnostic ECG.************The ECG waves at the bedside are not affected. Printing ECG waves from the bedside is an alternative to printing from the PIIC iX.************If you need any further information or support concerning this issue, contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. | SoftwareDesignChange
138 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=143570 | Z-0853-2016 | Philips IntelliVue Measurement Module X1 ***Model: M3001A*** | Philips Healthcare | 2 | CDRH | 2016-02-23 |  | 2021-03-01 14:49 | Philips Electronics North America Corporation | The ST elevation alarm on the Patient Monitor or standalone X2 Measurement Module will not sound when indicated for all chest leads derived using Hexad 12-Lead ECG Monitoring in the Host Monitor | Class 2 Device Recall Philips Healthcare Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare Date Initiated by Firm January 04, 2016 Create Date February 23, 2016 Recall Status 1 Terminated 3 on March 01, 2021 Recall Number Z-0853-2016 Recall Event ID 73209 510(K)Number K150310 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product Philips IntelliVue Measurement Module X1 ******Model: M3001A Code Information Software Revision:K.21.54 or L.00.96 Recalling Firm/ Manufacturer Philips Electronics North America Corporation 3000 Minuteman Rd Andover MA 01810-1032 For Additional Information Contact Kelly Rector 978-659-3000 Manufacturer Reason for Recall The ST elevation alarm on the Patient Monitor or standalone X2 Measurement Module will not sound when indicated for all chest leads derived using Hexad 12-Lead ECG Monitoring in the Host Monitor FDA Determined Cause 2 Software Design Change Action Philips Healthcare issued the Field Safety Notice on January 4, 2016, and informed customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Field Safety Notice: ************Until your software is upgraded, please make sure that the ST Analysis is switched ON when using STE measurement in the Hexad 12-lead ECG Monitoring. This can be done by entering the ST Analysis Menu and select ST Analysis to ***On***. For more detailed information, please refer to the Instructions for Use (IFU) of your host monitor or to the X2 Measurement Module IFU.************Philips will provide a software correction free of charge. For further questions please call (978) 659-3000. Quantity in Commerce 13,300 Distribution Worldwide Distribution - Nationwide Distribution and to the countries of : Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iran ( Islamic Republic of ) Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Morocco, Myanmar(Burma) Netherlands Antilles, Norway, Oman, Pakistan, Philippines, Poland, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Province of China, Thailand, Turkey, United Arab Emirates, United Kingdom and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare issued the Field Safety Notice on January 4, 2016, and informed customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Field Safety Notice: ************Until your software is upgraded, please make sure that the ST Analysis is switched ON when using STE measurement in the Hexad 12-lead ECG Monitoring. This can be done by entering the ST Analysis Menu and select ST Analysis to ***On***. For more detailed information, please refer to the Instructions for Use (IFU) of your host monitor or to the X2 Measurement Module IFU.************Philips will provide a software correction free of charge. For further questions please call (978) 659-3000. | SoftwareDesignChange
139 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142026 | Z-0858-2016 | Philips Healthcare IntelliVue Info Center iX, B.0, 866389*** Physiological, Patient Monitor (With Arrhythmia Detection or Alarm) | Philips Healthcare IntelliVue Info Center iX, B.0 | 2 | CDRH | 2016-02-23 | 2016-02-23 21:06 | 2017-09-06 | Philips Electronics North America Corporation | Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads. | Class 2 Device Recall Philips Healthcare IntelliVue Info Center iX, B.0 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare IntelliVue Info Center iX, B.0 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare IntelliVue Info Center iX, B.0 Date Initiated by Firm November 25, 2015 Date Posted February 23, 2016 Recall Status 1 Terminated 3 on September 06, 2017 Recall Number Z-0858-2016 Recall Event ID 72720 510(K)Number K102495 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product Philips Healthcare IntelliVue Info Center iX, B.0, 866389****** Physiological, Patient Monitor (With Arrhythmia Detection or Alarm) Code Information Software Versions: B.0.X Recalling Firm/ Manufacturer Philips Electronics North America Corporation 3000 Minuteman Rd Andover MA 01810-1032 For Additional Information Contact Kelly Rector 978-659-2950 Manufacturer Reason for Recall Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads. FDA Determined Cause 2 Software Design Change Action Philips Healthcare issued a Field Safety Notice on 11/25/15 to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will provide a software correction free of charge. On 1/28/16 the Philips issued an updated letter to users.Philips has initiated a correction to address this issue. A software correction will be provided to customers with impacted devices at no charge.************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Field Safety Notice: ******Arrhythmia alarms are not affected. ST alarms, ST values and measurements are also not affected and could be used to assist with the decision to obtain a diagnostic ECG.************The ECG waves at the bedside are not affected. Printing ECG waves from the bedside is an alternative to printing from the PIIC iX.************If you need any further information or support concerning this issue, contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. Quantity in Commerce 1880 units Distribution Worldwide Distribution-US (nationwide) and the countries of Algeria Australia Austria Belgium Bulgaria Canada Denmark Estonia Finland France Germany Hong Kong Hungary Iceland India Indonesia Iraq Ireland Israel Italy Japan Kenya Korea Kuwait Luxembourg Malaysia Mongolia Morocco Netherlands New Zealand Nigeria Norway Poland Qatar Romania Russian Federation Saudi Arabia Singapore Slovenia South Africa Spain Sweden Switzerland Taiwan Turkey United Arab Emirates and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare issued a Field Safety Notice on 11/25/15 to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will provide a software correction free of charge. On 1/28/16 the Philips issued an updated letter to users.Philips has initiated a correction to address this issue. A software correction will be provided to customers with impacted devices at no charge.************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Field Safety Notice: ******Arrhythmia alarms are not affected. ST alarms, ST values and measurements are also not affected and could be used to assist with the decision to obtain a diagnostic ECG.************The ECG waves at the bedside are not affected. Printing ECG waves from the bedside is an alternative to printing from the PIIC iX.************If you need any further information or support concerning this issue, contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. | SoftwareDesignChange
140 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=143571 | Z-0854-2016 | Philips IntelliVue Module Measurement X2*********Model: M3002A | Philips Healthcare | 2 | CDRH | 2016-02-23 |  | 2021-03-01 14:49 | Philips Electronics North America Corporation | The ST elevation alarm on the Patient Monitor or standalone X2 Measurement Module will not sound when indicated for all chest leads derived using Hexad 12-Lead ECG Monitoring in the Host Monitor | Class 2 Device Recall Philips Healthcare Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare Date Initiated by Firm January 04, 2016 Create Date February 23, 2016 Recall Status 1 Terminated 3 on March 01, 2021 Recall Number Z-0854-2016 Recall Event ID 73209 510(K)Number K150310 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product Philips IntelliVue Module Measurement X2************Model: M3002A Code Information Software Revision: K.21.54 or L.00.96 Recalling Firm/ Manufacturer Philips Electronics North America Corporation 3000 Minuteman Rd Andover MA 01810-1032 For Additional Information Contact Kelly Rector 978-659-3000 Manufacturer Reason for Recall The ST elevation alarm on the Patient Monitor or standalone X2 Measurement Module will not sound when indicated for all chest leads derived using Hexad 12-Lead ECG Monitoring in the Host Monitor FDA Determined Cause 2 Software Design Change Action Philips Healthcare issued the Field Safety Notice on January 4, 2016, and informed customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Field Safety Notice: ************Until your software is upgraded, please make sure that the ST Analysis is switched ON when using STE measurement in the Hexad 12-lead ECG Monitoring. This can be done by entering the ST Analysis Menu and select ST Analysis to ***On***. For more detailed information, please refer to the Instructions for Use (IFU) of your host monitor or to the X2 Measurement Module IFU.************Philips will provide a software correction free of charge. For further questions please call (978) 659-3000. Quantity in Commerce 13,300 Distribution Worldwide Distribution - Nationwide Distribution and to the countries of : Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iran ( Islamic Republic of ) Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Morocco, Myanmar(Burma) Netherlands Antilles, Norway, Oman, Pakistan, Philippines, Poland, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Province of China, Thailand, Turkey, United Arab Emirates, United Kingdom and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare issued the Field Safety Notice on January 4, 2016, and informed customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Field Safety Notice: ************Until your software is upgraded, please make sure that the ST Analysis is switched ON when using STE measurement in the Hexad 12-lead ECG Monitoring. This can be done by entering the ST Analysis Menu and select ST Analysis to ***On***. For more detailed information, please refer to the Instructions for Use (IFU) of your host monitor or to the X2 Measurement Module IFU.************Philips will provide a software correction free of charge. For further questions please call (978) 659-3000. | SoftwareDesignChange
141 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142025 | Z-0857-2016 | Philips Healthcare PIIC Classic Upgrade, 866117*** Physiological, Patient Monitor (With Arrhythmia Detection or Alarm) | Philips Healthcare PIIC Classic Upgrade | 2 | CDRH | 2016-02-23 | 2016-02-23 21:06 | 2017-09-06 | Philips Electronics North America Corporation | Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads. | Class 2 Device Recall Philips Healthcare PIIC Classic Upgrade Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare PIIC Classic Upgrade FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare PIIC Classic Upgrade Date Initiated by Firm November 25, 2015 Date Posted February 23, 2016 Recall Status 1 Terminated 3 on September 06, 2017 Recall Number Z-0857-2016 Recall Event ID 72720 510(K)Number K102495 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product Philips Healthcare PIIC Classic Upgrade, 866117****** Physiological, Patient Monitor (With Arrhythmia Detection or Alarm) Code Information Software B.O Recalling Firm/ Manufacturer Philips Electronics North America Corporation 3000 Minuteman Rd Andover MA 01810-1032 For Additional Information Contact Kelly Rector 978-659-2950 Manufacturer Reason for Recall Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads. FDA Determined Cause 2 Software Design Change Action Philips Healthcare issued a Field Safety Notice on 11/25/15 to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will provide a software correction free of charge. On 1/28/16 the Philips issued an updated letter to users.Philips has initiated a correction to address this issue. A software correction will be provided to customers with impacted devices at no charge.************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Field Safety Notice: ******Arrhythmia alarms are not affected. ST alarms, ST values and measurements are also not affected and could be used to assist with the decision to obtain a diagnostic ECG.************The ECG waves at the bedside are not affected. Printing ECG waves from the bedside is an alternative to printing from the PIIC iX.************If you need any further information or support concerning this issue, contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. Quantity in Commerce 5569 units Distribution Worldwide Distribution-US (nationwide) and the countries of Algeria Australia Austria Belgium Bulgaria Canada Denmark Estonia Finland France Germany Hong Kong Hungary Iceland India Indonesia Iraq Ireland Israel Italy Japan Kenya Korea Kuwait Luxembourg Malaysia Mongolia Morocco Netherlands New Zealand Nigeria Norway Poland Qatar Romania Russian Federation Saudi Arabia Singapore Slovenia South Africa Spain Sweden Switzerland Taiwan Turkey United Arab Emirates and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare issued a Field Safety Notice on 11/25/15 to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will provide a software correction free of charge. On 1/28/16 the Philips issued an updated letter to users.Philips has initiated a correction to address this issue. A software correction will be provided to customers with impacted devices at no charge.************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Field Safety Notice: ******Arrhythmia alarms are not affected. ST alarms, ST values and measurements are also not affected and could be used to assist with the decision to obtain a diagnostic ECG.************The ECG waves at the bedside are not affected. Printing ECG waves from the bedside is an alternative to printing from the PIIC iX.************If you need any further information or support concerning this issue, contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. | SoftwareDesignChange
142 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132093 | Z-0374-2016 | Philips IntelliVue Information Center iX (release A.00, A.01, and A.02) are impacted by this issue:***866023 IntelliVue Info Center iX***866024 PIIC iX Upgrade***866117 PIIC Classic Upgrade******The intended use of the Information Center Software is to display physiologic waves, parameters, and trends, format data for strip chart recordings and printed reports, and provide the secondary annunciation of alarms from other networked medical devices at a centralized location. ******An additional intended use of the Information Center Software is to provide primary annunciation of alarms and configuration and control access for networked telemetry monitors. | Philips IntelliVue Information Center iX (release A.00, A.01, and A.02) | 2 | CDRH | 2015-12-04 | 2015-12-04 21:00 | 2020-09-21 06:06 | Philips Medical Systems, Inc. | Sp02 and/or Non Invasive Blood Pressure (NBP) alarms may become disabled without visual notification | Class 2 Device Recall Philips IntelliVue Information Center iX (release A.00, A.01, and A.02) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips IntelliVue Information Center iX (release A.00, A.01, and A.02) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips IntelliVue Information Center iX (release A.00, A.01, and A.02) Date Initiated by Firm December 19, 2014 Date Posted December 04, 2015 Recall Status 1 Terminated 3 on September 21, 2020 Recall Number Z-0374-2016 Recall Event ID 70000 510(K)Number K102495 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product Philips IntelliVue Information Center iX (release A.00, A.01, and A.02) are impacted by this issue:******866023 IntelliVue Info Center iX******866024 PIIC iX Upgrade******866117 PIIC Classic Upgrade************The intended use of the Information Center Software is to display physiologic waves, parameters, and trends, format data for strip chart recordings and printed reports, and provide the secondary annunciation of alarms from other networked medical devices at a centralized location. ************An additional intended use of the Information Center Software is to provide primary annunciation of alarms and configuration and control access for networked telemetry monitors. Code Information Software systems Revision A.02.05 or lower Recalling Firm/ Manufacturer Philips Medical Systems, Inc. 3000 Minuteman Rd Andover MA 01810-1032 For Additional Information Contact 978-687-1501 Manufacturer Reason for Recall Sp02 and/or Non Invasive Blood Pressure (NBP) alarms may become disabled without visual notification FDA Determined Cause 2 Software Design Change Action The firm, Philips Healthcare, issued a "URGENT-Field Safety Notice" on 12/19/14 to inform customers of the issue, identify details of the units affected, instructs on actions to be taken by the customer and identify what action Philips plans to take to remedy the issue. Philips will provide a software correction free of charge.******Philips is asking customers to follow the Action to be Taken by Customer/User*** section of the Field Safety Notice-To prevent this issue from occurring, customers/users should:******1) Disable the Change Bed Label field in Patient Management. This is accomplished by changing the feature from ***Optional to ***Hidden*********2) All patient transfers must be performed using the Transfer key in the Manage Patient application.******Remember that the most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment.************If you need any further information or support concerning this issue, please contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. Quantity in Commerce 10,013 units Distribution Worldwide Distribution: US (nationwide) and countries: AUSTRALIA, AUSTRIA, BELGIUM, BULGARIA, CANADA, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, HUNGARY, ICELAND, INDIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, LUXEMBOURG, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, Philips Healthcare, issued a "URGENT-Field Safety Notice" on 12/19/14 to inform customers of the issue, identify details of the units affected, instructs on actions to be taken by the customer and identify what action Philips plans to take to remedy the issue. Philips will provide a software correction free of charge.******Philips is asking customers to follow the Action to be Taken by Customer/User*** section of the Field Safety Notice-To prevent this issue from occurring, customers/users should:******1) Disable the Change Bed Label field in Patient Management. This is accomplished by changing the feature from ***Optional to ***Hidden*********2) All patient transfers must be performed using the Transfer key in the Manage Patient application.******Remember that the most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment.************If you need any further information or support concerning this issue, please contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. | SoftwareDesignChange
143 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=129687 | Z-0325-2016 | Ingenia, Intera, Achieva and Achieve dStream MR systems on Software versions R5.1.7 and R5.1.8 with the Mobi View software option. Diagnostic imaging system. | Philips Healthcare | 2 | CDRH | 2015-11-23 | 2015-11-23 21:00 | 2021-05-12 15:10 | Philips Medical Systems, Inc. | When a fused series of a sagittal, coronal or radial multi station scan is generated in MobiView, the resultant image order is reversed. This occurs with software release R5.1.7/R5.1.8. | Class 2 Device Recall Philips Healthcare Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare Date Initiated by Firm August 21, 2014 Date Posted November 23, 2015 Recall Status 1 Terminated 3 on May 12, 2021 Recall Number Z-0325-2016 Recall Event ID 69127 510(K)Number K063559 K110151 Product Classification System, nuclear magnetic resonance imaging - Product Code LNH Product Ingenia, Intera, Achieva and Achieve dStream MR systems on Software versions R5.1.7 and R5.1.8 with the Mobi View software option. Diagnostic imaging system. Code Information Software versions R5.1.7 and R5.1.8 Recalling Firm/ Manufacturer Philips Medical Systems, Inc. 3000 Minuteman Rd Andover MA 01810-1032 For Additional Information Contact 978-687-1501 Manufacturer Reason for Recall When a fused series of a sagittal, coronal or radial multi station scan is generated in MobiView, the resultant image order is reversed. This occurs with software release R5.1.7/R5.1.8. FDA Determined Cause 2 Software Design Change Action Philips Healthcare sent an Customer Information Letter dated July 2014, on August 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips will issue a Field Change Order to install updated software in a Service Pack that makes the slice numbering consistent for fused and unfused images.************If you need any further information or support concerning this issue, please contact your local Philips representative: 1-800-722-9377 Quantity in Commerce 307 units Distribution Worldwide Distribution - USA (nationwide) and to the countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, Denmark Estonia, Ethiopia. France, Germany, Hong Kong, India, Indonesia, Iraq, Ireland, Italy,******Japan, Jordan, Korea, Republic of, Libya, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, Oman, Palestinian territory, Peru, Poland, Russian Federation, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LNH 510(K)s with Product Code = LNH - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare sent an Customer Information Letter dated July 2014, on August 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips will issue a Field Change Order to install updated software in a Service Pack that makes the slice numbering consistent for fused and unfused images.************If you need any further information or support concerning this issue, please contact your local Philips representative: 1-800-722-9377 | SoftwareDesignChange
144 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=141491 | Z-0327-2016 | NxStage System One S Cycler (High Permeability Hemodialysis System)***Model no. NX1000-5-A | NxStage System One | 2 | CDRH | 2015-11-23 |  | 2019-04-29 08:20 | NxStage Medical, Inc. | Ultrafiltration (UF) Volume software error inaccurate fluid removal | Class 2 Device Recall NxStage System One Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall NxStage System One FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall NxStage System One Date Initiated by Firm October 29, 2015 Create Date November 23, 2015 Recall Status 1 Terminated 3 on April 29, 2019 Recall Number Z-0327-2016 Recall Event ID 72543 510(K)Number K133547 Product Classification Dialyzer, high permeability with or without sealed dialysate system - Product Code KDI Product NxStage System One S Cycler (High Permeability Hemodialysis System)******Model no. NX1000-5-A Code Information Software versions 4.9 and 4.10 Recalling Firm/ Manufacturer NxStage Medical, Inc. 350 Merrimack Street Lawrence MA 01843-1748 For Additional Information Contact 978-687-4700 Manufacturer Reason for Recall Ultrafiltration (UF) Volume software error inaccurate fluid removal FDA Determined Cause 2 Software Design Change Action NxStage issued issued a Field Safety Notice (FSN) on 10/29/15 to consignees who have NX1000-5 or a NX1000-5-A cycler. The FSN included Instructions for steps to be taken so that the cycler and treatments are not affected by the software error.******NxStage will be sending a representative to the facility to attach detailed instructions on each cycler as a user reference. Return the attached reply form to acknowledge receipt of this letter.******Questions or comments, please contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243). Quantity in Commerce 16 units Distribution Worldwide Distribution. US Nationwide, Netherlands, and Sweden. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = KDI - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action NxStage issued issued a Field Safety Notice (FSN) on 10/29/15 to consignees who have NX1000-5 or a NX1000-5-A cycler. The FSN included Instructions for steps to be taken so that the cycler and treatments are not affected by the software error.******NxStage will be sending a representative to the facility to attach detailed instructions on each cycler as a user reference. Return the attached reply form to acknowledge receipt of this letter.******Questions or comments, please contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243). | SoftwareDesignChange
145 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=141482 | Z-0326-2016 | NxStage System One S Cycler (High Permeability Hemodialysis System)***Model no. NX1000-5 | NxStage System One | 2 | CDRH | 2015-11-23 |  | 2019-04-29 08:20 | NxStage Medical, Inc. | Ultrafiltration (UF) Volume software error inaccurate fluid removal | Class 2 Device Recall NxStage System One Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall NxStage System One FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall NxStage System One Date Initiated by Firm October 29, 2015 Create Date November 23, 2015 Recall Status 1 Terminated 3 on April 29, 2019 Recall Number Z-0326-2016 Recall Event ID 72543 510(K)Number K133547 Product Classification Dialyzer, high permeability with or without sealed dialysate system - Product Code KDI Product NxStage System One S Cycler (High Permeability Hemodialysis System)******Model no. NX1000-5 Code Information Software versions 4.9 and 4.10 Recalling Firm/ Manufacturer NxStage Medical, Inc. 350 Merrimack Street Lawrence MA 01843-1748 For Additional Information Contact 978-687-4700 Manufacturer Reason for Recall Ultrafiltration (UF) Volume software error inaccurate fluid removal FDA Determined Cause 2 Software Design Change Action NxStage issued issued a Field Safety Notice (FSN) on 10/29/15 to consignees who have NX1000-5 or a NX1000-5-A cycler. The FSN included Instructions for steps to be taken so that the cycler and treatments are not affected by the software error.******NxStage will be sending a representative to the facility to attach detailed instructions on each cycler as a user reference. Return the attached reply form to acknowledge receipt of this letter.******Questions or comments, please contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243). Quantity in Commerce 323 units Distribution Worldwide Distribution. US Nationwide, Netherlands, and Sweden. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = KDI - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action NxStage issued issued a Field Safety Notice (FSN) on 10/29/15 to consignees who have NX1000-5 or a NX1000-5-A cycler. The FSN included Instructions for steps to be taken so that the cycler and treatments are not affected by the software error.******NxStage will be sending a representative to the facility to attach detailed instructions on each cycler as a user reference. Return the attached reply form to acknowledge receipt of this letter.******Questions or comments, please contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243). | SoftwareDesignChange
146 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=140304 | Z-0023-2016 | Centricity PACS IW by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources. Images and data can be stored, communicated, processed and displayed within the system or across computer networks at distributed locations. | GE Healthcare Centricity PACS IW | 2 | CDRH | 2015-10-05 | 2015-10-05 21:00 | 2017-01-05 | GE Healthcare | Images may be missing when a system parameter MapRoute is set to a value greater than 1. | Class 2 Device Recall GE Healthcare Centricity PACS IW Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GE Healthcare Centricity PACS IW FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GE Healthcare Centricity PACS IW Date Initiated by Firm September 16, 2015 Date Posted October 05, 2015 Recall Status 1 Terminated 3 on January 05, 2017 Recall Number Z-0023-2016 Recall Event ID 72212 510(K)Number K082318 K121387 Product Classification System, image processing, radiological - Product Code LLZ Product Centricity PACS IW by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources. Images and data can be stored, communicated, processed and displayed within the system or across computer networks at distributed locations. Code Information GE Centricity PACS-IW versions 3.5.0 through 3.7.3.9 Sp2, 3.7.3 Spa10, and 4.0.1 Recalling Firm/ Manufacturer GE Healthcare 540 W Northwest Hwy Barrington IL 60010-3051 For Additional Information Contact Deb Lahr 847-277-4472 Manufacturer Reason for Recall Images may be missing when a system parameter ***MapRoute*** is set to a value greater than 1. FDA Determined Cause 2 Software Design Change Action The firm, GE Healthcare, an "Urgent Medical Device Correction" by mail or hand delivered to all sites. The letter instructs the customer to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. Until GE Healthcare conducts remote inspection of the systems "MapRoute" parameter value. If the value is found to be greater than the recommended setting of 1, GE will reset this system parameter to a value of 1. The letter informs the user facilities not to make any changes to this parameter value.************Additionally, until a product modification is available to correct this issue, the following actions are recommended. The firm is recommending user facilities to: ************1. Utilize the image count within the QC process to alert the user of a discrepancy in the number of transmitted images from the modality to the number of images available in the PACS IW viewer.******a. If a discrepancy is identified, attempt to retransmit the exam to PACS.******b. If retransmission is unsuccessful, contact your GE Healthcare Service representative for assistance and instructions as to the impacted exam. Urgent cases impacted by this issue should be interpreted at the modality.******2. If DICOM storage commit is configured and in use, no commit notification is sent to the modality for images that were impacted by this issue.************GE Healthcare will correct all affected systems by providing a software update at no cost to the user facilities. ************If the user facilities have questions, they are encouraged to communicate questions or concerns regarding this notification to GE Healthcare Service at 1-800-437-1171 or their local Service Representative. Quantity in Commerce 921 systems are impacted. Distribution Worldwide Distribution: US (nationwide) including states of: AL AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, ON, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and Puerto Rico; and countries of: Argentina , Austria, Bahrain, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile , China, Colombia, Costa Rica, Ecuador, Egypt, France, Germany, Great Britain, Hong Kong, Hungary, India, Israel, Italy, Japan,Korea, Kuwait, Lebanon, Libya, Lithuania, Malaysia, Mexico, Netherlands, New Caledonia, Nigeria, Panam¿, Peru, Philippines, Poland, Qatar, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, Venezuela, and Vietnam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, GE Healthcare, an "Urgent Medical Device Correction" by mail or hand delivered to all sites. The letter instructs the customer to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. Until GE Healthcare conducts remote inspection of the systems "MapRoute" parameter value. If the value is found to be greater than the recommended setting of 1, GE will reset this system parameter to a value of 1. The letter informs the user facilities not to make any changes to this parameter value.************Additionally, until a product modification is available to correct this issue, the following actions are recommended. The firm is recommending user facilities to: ************1. Utilize the image count within the QC process to alert the user of a discrepancy in the number of transmitted images from the modality to the number of images available in the PACS IW viewer.******a. If a discrepancy is identified, attempt to retransmit the exam to PACS.******b. If retransmission is unsuccessful, contact your GE Healthcare Service representative for assistance and instructions as to the impacted exam. Urgent cases impacted by this issue should be interpreted at the modality.******2. If DICOM storage commit is configured and in use, no commit notification is sent to the modality for images that were impacted by this issue.************GE Healthcare will correct all affected systems by providing a software update at no cost to the user facilities. ************If the user facilities have questions, they are encouraged to communicate questions or concerns regarding this notification to GE Healthcare Service at 1-800-437-1171 or their local Service Representative. | SoftwareDesignChange
147 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138750 | Z-2321-2015 | ACUSON S2000 and ACUSON S3000 ultrasound systems with the Virtual Touch IQ option;***Model numbers:***10041461, 10440017 -S2000 system***10441730 -S3000 system***VTIQ Option- 10439521 and 10439522******Radiology: The ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications.*** | ACUSON S2000 and ACUSON S3000 ultrasound systems | 2 | CDRH | 2015-07-30 | 2015-07-30 21:00 | 2017-10-19 | Siemens Medical Solutions USA, Inc. | Potential measurement error on ACUSON S Family ultrasound system. When repositioning the Virtual Touch IQ region of interest from the original (default) location, the lateral position of the measurement tool may not align with the lateral position of the shear velocity data. | Class 2 Device Recall ACUSON S2000 and ACUSON S3000 ultrasound systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ACUSON S2000 and ACUSON S3000 ultrasound systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ACUSON S2000 and ACUSON S3000 ultrasound systems Date Initiated by Firm June 25, 2015 Date Posted July 30, 2015 Recall Status 1 Terminated 3 on October 19, 2017 Recall Number Z-2321-2015 Recall Event ID 71688 510(K)Number K132804 Product Classification System, imaging, pulsed doppler, ultrasonic - Product Code IYN Product ACUSON S2000 and ACUSON S3000 ultrasound systems with the Virtual Touch IQ option;******Model numbers:******10041461, 10440017 -S2000 system******10441730 -S3000 system******VTIQ Option- 10439521 and 10439522************Radiology: The ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. Code Information Model numbers: 10041461, 10440017 - S2000 system; 10441730 - S3000 system; VTIQ Option - 10439521 and 10439522. Serial numbers for US: 210104 211147 211148 210257 211071 201387 201389 206511 206574 206592 206619 206732 206832 208718 208772 210931 211483 212156 212157 202555 210055 210061 210423 210426 210427 210629 206049 210032 210128 210334 210361 211742 211334 211156 211165 211215 211180 211183 208220 210666 211158 211338 211425 211611 212064 211091 210343 203158 211278 210692 210748 210777 210010 210479 211550 211551 205213 211123 211448 211709 210993 210164 210839 210954 210961 210672 200113 200140 202137 210034 210215 211777 207390 210505 210521 210861 200695 210759 210761 210764 210768 210769 201673 201979 206346 210265 210488 210573 210631 210667 210783 210825 210843 211124 211167 211174 211260 211316 211319 211320 211499 211513 211627 211635 211636 211637 211639 211643 211647 211649 212070 212133 212135 206901 210044 210169 210203 210213 210217 210506 210517 210708 210853 211168 211221 211273 211291 211295 211322 211401 211405 211467 211538 211558 211653 210278 207050 210844 211151 211092 211648 210035 212235 206969 202197 211217 207411 207438 207715 208219 208482 208645 210276 210674 211218 211267 211305 211321 211449 211620 211654 212131 201279 201281 210842 210855 210992 211066 211088. Serial numbers for ROW: 211768 210671 211803 212186 206153 206181 210340 210752 210763 211060 211182 211351 211705 210494 210935 207083 207298 207553 208176 203090 203106 206480 210105 210725 210799 211059 211262 210064 202134 203148 203282 203549 205242 205260 205381 206678 206756 206819 206857 207988 209048 209088 209120 211085 211397 206341 210085 211432 211445 210144 210225 210620 210793 211178 207878 210218 210707 210718 211041 208127 200974 202859 206917 210475 210691 210106 210073 210267 210280 210281 210288 210303 210306 210377 210400 210438 210490 210514 210749 210851 210860 210901 211095 211117 211624 211662 211775 212065 212162 212239 212287 200323 200734 206304 206369 207458 207855 207999 208046 208539 208578 210375 210583 211793 203107 206451 210021 210065 210066 210069 210094 210114 210141 210142 210145 210160 210161 210210 210243 210284 210338 210499 210523 210624 210625 210626 210627 210628 210632 210634 210758 210854 210874 211046 211072 211074 211122 211143 211146 211150 211152 211224 211225 211227 211264 211276 211457 211486 211490 211498 211544 211545 211548 211593 211598 212359 202330 206161 206537 207222 207226 208016 208249 208255 208264 210075 210076 210077 210083 210084 210088 210108 210109 210113 210123 210125 210126 210127 210198 210201 210211 210251 210927 211211 211619 210060 210151 210272 210574 211226 210880 210907 210948 210951 210966 211065 211087 211105 210049 210526 210538 210784 211083 211198 211560 208316 210387 211019 200291 200819 210881 207060 210279 210819 211726 207383 210831 211170 211186 200365 206042 206403 210117 210436 210437 210751 210922 210940 210963 210990 211129 211220 211281 211301 211324 211326 211346 211424 211464 211580 211697 211698 211701 211806 212149 212233 206044 206345 206388 206415 206481 206925 207267 207353 207563 207746 207943 208068 208070 208152 208193 208563 208576 208698 208878 210293 210315 210341 210374 210434 210457 210676 210845 210847 210915 211009 211063 211290 211345 210031 210608 205238 210442 211005 211011 211118 211258 211369 211729 212173 210914 211112 211125 210780 206761 206812 211272 200641 206450 211529 211530 211531 211532 211534 208557 208881 208945 208993 209065 209172 209213 209219 209221 209276 209315 209322 209340 209370 209535 209537 209597 209612 209614 209615 209625 209640 209648 209654 209702 209705 211242 200225 211623 210045 210342 211349 211712 207245 200853 203197 205424 207078 208682 210464 211187 206405 206679 207104 207800 208103 210165 210516 210773 202847 210190 210939 209627 211149 211048 210406 201937 202992 210942 211234 212139 207218 210047 210216 210294 210307 210395 210322 211808 205602 212158 206784 210871 210917 210919 210997 211153 211163 206666 206733 206734 208001 208056 208155 211420 200259 200708 202568 206406 207919 210090 211664 211764 210685 207589 208932 210082 210801 210955 211663 207621 210091 210586 210850 210989 211025 211068 211089 211127 211229 211494 208553 210452 210471 210736 211484 207023 210591 210594 210596 210599 210877 211220 211314 210333 210394 210417 210446 210450 210493 210512 210551 210598 210661 210929 211002 211012 211128 211130 211131 211135 211268 211365 211437 212165 212221 206740 207196 208146 208311 208628 208656 208787 208879 210766 211656. Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc. 685 E Middlefield Rd Mountain View CA 94043-4045 For Additional Information Contact Sheila Pickering, Ph.D. 650-694-9398 Manufacturer Reason for Recall Potential measurement error on ACUSON S Family ultrasound system. When repositioning the Virtual Touch IQ region of interest from the original (default) location, the lateral position of the measurement tool may not align with the lateral position of the shear velocity data. FDA Determined Cause 2 Software Design Change Action On June 25, 2015 Customer Safety Advisory Notification letters were sent to all affected customers by Certified mail. The letter identified the affected product, problem and actions to be taken. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. For questions contact your Siemens Service for information. Quantity in Commerce 629 units Distribution Worldwide Distribution - US Nationwide and the countries of Algeria, Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Croatia, Czech Republic, Denmark, Dominican Rep., France, Germany, Greece, Hong Kong, India, Indonesia, Israel, Italy, Japan, Kenya, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, P.R. China, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., Ukraine, United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On June 25, 2015 Customer Safety Advisory Notification letters were sent to all affected customers by Certified mail. The letter identified the affected product, problem and actions to be taken. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. For questions contact your Siemens Service for information. | SoftwareDesignChange
148 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138433 | Z-2085-2015 | Patient Handling System (Motion Control Software), ******Product Usage:***Indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body when radiation treatment is indicated.*** | Patient Handling System | 2 | CDRH | 2015-07-16 | 2015-07-16 21:24 | 2016-02-04 | Viewray Incorporated | ViewRay received a report that the couch moved unexpectedly into the bore after performing a RTCS reboot. | Class 2 Device Recall Patient Handling System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Patient Handling System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Patient Handling System Date Initiated by Firm April 01, 2015 Date Posted July 16, 2015 Recall Status 1 Terminated 3 on February 04, 2016 Recall Number Z-2085-2015 Recall Event ID 71555 510(K)Number K111862 Product Classification Accelerator, linear, medical - Product Code IYE Product Patient Handling System (Motion Control Software), ************Product Usage:******Indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body when radiation treatment is indicated. Code Information Model #10000, Serial #'s: 100, 101, and 102 Recalling Firm/ Manufacturer Viewray Incorporated 2 Thermo Fisher Way Oakwood Village OH 44146-6536 For Additional Information Contact Mr. James Talbot 650-252-0949 Ext. 455 Manufacturer Reason for Recall ViewRay received a report that the couch moved unexpectedly into the bore after performing a RTCS reboot. FDA Determined Cause 2 Software Design Change Action ViewRay sent a Customer Advisory Notification Letters dated April 1, 2015 to thier customers. The letter identified the affected product, problem and actions to be taken. For questions or concerns regarding this issue, contact ViewRaylM Customer Support at support@viewray.com or call the ViewRaylM Customer Support center at 855-286-8875. Quantity in Commerce 3 Distribution US Nationwide Distribution in the states of CA, MO & WI Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action ViewRay sent a Customer Advisory Notification Letters dated April 1, 2015 to thier customers. The letter identified the affected product, problem and actions to be taken. For questions or concerns regarding this issue, contact ViewRaylM Customer Support at support@viewray.com or call the ViewRaylM Customer Support center at 855-286-8875. | SoftwareDesignChange
149 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=137840 | Z-1849-2015 | The VariSource iX series afterloader systems are computer controlled remote electro/mechanical systems used for medical purposes, for placing a cable incorporating an irradiated iridium seed internally or close by a malignant tumor or tumor bed in a practice known as brachytherapy. | VariSource iX series | 2 | CDRH | 2015-06-26 | 2015-06-26 23:25 | 2016-09-12 | Varian Medical Systems Inc | After an application freeze and restart of the VariSource iX series by power cycling, the Partial Fraction generated by the***system will not be correct. The application freeze issue affects only the systems equipped with magnetic Hard Disk Drives (HDD). | Class 2 Device Recall VariSource iX series Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall VariSource iX series FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall VariSource iX series Date Initiated by Firm May 18, 2015 Date Posted June 26, 2015 Recall Status 1 Terminated 3 on September 12, 2016 Recall Number Z-1849-2015 Recall Event ID 71451 510(K)Number K122236 Product Classification System, applicator, radionuclide, remote-controlled - Product Code JAQ Product The VariSource iX series afterloader systems are computer controlled remote electro/mechanical systems used for medical purposes, for placing a cable incorporating an irradiated iridium seed internally or close by a malignant tumor or tumor bed in a practice known as brachytherapy. Code Information JAQ Recalling Firm/ Manufacturer Varian Medical Systems Inc 501 Locust Ave Charlottesville VA 22902-4869 For Additional Information Contact Rachel Forsberg 434-951-8635 Manufacturer Reason for Recall After an application freeze and restart of the VariSource iX series by power cycling, the Partial Fraction generated by the******system will not be correct. The application freeze issue affects only the systems equipped with magnetic Hard Disk Drives (HDD). FDA Determined Cause 2 Software Design Change Action All the consignees affected by this recall will be notified by:************The Urgent Medical Device Correction Letter/ Field Safety Notification will be******distributed to all affected users, with a description of the problem.************The Letter will also be distributed to the Varian Sales, Marketing, and Service******organizations, informing them of the issue.************Additionally, the letter will be posted to the customer support site:******http:/ /www.MyVarian.com. Quantity in Commerce 55 Distribution Worldwide Distribution. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JAQ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action All the consignees affected by this recall will be notified by:************The Urgent Medical Device Correction Letter/ Field Safety Notification will be******distributed to all affected users, with a description of the problem.************The Letter will also be distributed to the Varian Sales, Marketing, and Service******organizations, informing them of the issue.************Additionally, the letter will be posted to the customer support site:******http:/ /www.MyVarian.com. | SoftwareDesignChange
150 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=137242 | Z-1679-2015 | RaySearch RayStation 2.5, 3.0, 3.5 and 4.0; Radiation Therapy Treatment ******Product Usage:***RayStation is a software system designed for treatment planning and analysis of radiation therapy. The treatment plans provide treatment unit set-up parameters and estimates of dose distributions expected during the proposed treatment, and may be used to administer treatments after review and approval by the intended user. The system functionality can be configured based on user needs. The intended users of RayStation shall be clinically qualified radiation therapy staff trained in using the system.*** | RayStation | 2 | CDRH | 2015-05-27 | 2015-05-27 21:00 | 2017-04-11 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories AB became aware of an issue with the dose calculation for some imported VMAT plans where controls points are defined with wide gantry angle spacing. | Class 2 Device Recall RayStation Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall RayStation FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall RayStation Date Initiated by Firm May 08, 2015 Date Posted May 27, 2015 Recall Status 1 Terminated 3 on April 11, 2017 Recall Number Z-1679-2015 Recall Event ID 71289 510(K)Number K120387 K130617 K140187 Product Classification System,planning,radiation therapy treatment - Product Code MUJ Product RaySearch RayStation 2.5, 3.0, 3.5 and 4.0; Radiation Therapy Treatment ************Product Usage:******RayStation is a software system designed for treatment planning and analysis of radiation therapy. The treatment plans provide treatment unit set-up parameters and estimates of dose distributions expected during the proposed treatment, and may be used to administer treatments after review and approval by the intended user. The system functionality can be configured based on user needs. The intended users of RayStation shall be clinically qualified radiation therapy staff trained in using the system. Code Information Software Version (Build numbers) 2.5.1.89, 3.0.0.251, 3.5.0.16, 3.5.1.6, 4.0.0.14, 4.0.1.4, 4.0.2.9, or 4.0.3.4 Recalling Firm/ Manufacturer RAYSEARCH LABORATORIES AB Sveavaegen 9 Stockholm Sweden Manufacturer Reason for Recall RaySearch Laboratories AB became aware of an issue with the dose calculation for some imported VMAT plans where controls points are defined with wide gantry angle spacing. FDA Determined Cause 2 Software Design Change Action RaySearch sent a Device Correction letter dated May 8, 2015 via e-mailed to their customers. The letter identified the affected product, problem, actions to be taken by users and contact information. Quantity in Commerce Domestic: 539 units Distribution US Nationwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MUJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action RaySearch sent a Device Correction letter dated May 8, 2015 via e-mailed to their customers. The letter identified the affected product, problem, actions to be taken by users and contact information. | SoftwareDesignChange
151 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135930 | Z-1580-2015 | ViewRay System, Radiation Therapy System | ViewRay Radiation Therapy System | 2 | CDRH | 2015-05-05 | 2015-05-05 21:03 | 2015-09-23 | Viewray Incorporated | The software was not correctly using the RT (Radiation Therapy) to MR (Magnetic Resonance image) coordinate correction for non HFS (Head First Supine) patient orientations, resulting in slice mismatch error. | Class 2 Device Recall ViewRay Radiation Therapy System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ViewRay Radiation Therapy System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ViewRay Radiation Therapy System Date Initiated by Firm January 15, 2015 Date Posted May 05, 2015 Recall Status 1 Terminated 3 on September 23, 2015 Recall Number Z-1580-2015 Recall Event ID 71050 510(K)Number K111862 Product Classification Accelerator, linear, medical - Product Code IYE Product ViewRay System, Radiation Therapy System Code Information M/N 10000, S/N 100 only Recalling Firm/ Manufacturer Viewray Incorporated 2 Thermo Fisher Way Oakwood Village OH 44146-6536 For Additional Information Contact 440-703-3210 Ext. 455 Manufacturer Reason for Recall The software was not correctly using the RT (Radiation Therapy) to MR (Magnetic Resonance image) coordinate correction for non HFS (Head First Supine) patient orientations, resulting in slice mismatch error. FDA Determined Cause 2 Software Design Change Action On January 15, 2014, ViewRay personnel discussed the investigation of the issue (slice mismatch error) with the customer. The firm advised the customer that until the issue was corrected, to not treat patients in non HFS orientations, or to use the MR acquired from the ViewRay system as the primary plan series for non HFS scans. ViewRay installed software version 3.4.5.28 on January 20 and 21, 2014, in order to prevent this slice mismatch error from occurring in any patient position. Quantity in Commerce 1 Distribution US distribution to MO. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On January 15, 2014, ViewRay personnel discussed the investigation of the issue (slice mismatch error) with the customer. The firm advised the customer that until the issue was corrected, to not treat patients in non HFS orientations, or to use the MR acquired from the ViewRay system as the primary plan series for non HFS scans. ViewRay installed software version 3.4.5.28 on January 20 and 21, 2014, in order to prevent this slice mismatch error from occurring in any patient position. | SoftwareDesignChange
152 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132485 | Z-1013-2015 | Syngo RT Therapist: The intended use of the SIEMENS branded ARTISTE, ONCOR and PRIMUS family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer | Syngo RT Therapist | 2 | CDRH | 2015-01-26 | 2015-01-26 21:00 | 2015-09-10 | Siemens Medical Solutions USA, Inc | Combination of CTVision with syngo RT Therapist / syngo RT Oncologist 4.3.SP1 automatic registration in Adaptive Targeting might result in wrong offset calculations. Applying this offset can lead to patient mistreatment. Cone Beam imaging is not affected by this problem. | Class 2 Device Recall Syngo RT Therapist Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Syngo RT Therapist FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Syngo RT Therapist Date Initiated by Firm January 02, 2015 Date Posted January 26, 2015 Recall Status 1 Terminated 3 on September 10, 2015 Recall Number Z-1013-2015 Recall Event ID 70200 510(K)Number K142434 Product Classification Accelerator, linear, medical - Product Code IYE Product Syngo RT Therapist: The intended use of the SIEMENS branded ARTISTE, ONCOR and PRIMUS family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer Code Information model # 08162815, serial #s 10010, 10046, 10586 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 40 Liberty Blvd Malvern PA 19355-1418 For Additional Information Contact Customer Support 610-219-6300 Manufacturer Reason for Recall Combination of CTVision with syngo RT Therapist / syngo RT Oncologist 4.3.SP1 automatic registration in Adaptive Targeting might result in wrong offset calculations. Applying this offset can lead to patient mistreatment. Cone Beam imaging is not affected by this problem. FDA Determined Cause 2 Software Design Change Action A customer advisory notice, dated January 2, 2015, was sent to end users which identified the product, problem, and action to be taken. It was recommended that customers use only manual registration with the 6 degrees of freedom option disabled, which provides the correct offset values. A software update will be deployed once a solution is available. Quantity in Commerce 3 Distribution UT, WI, NY Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action A customer advisory notice, dated January 2, 2015, was sent to end users which identified the product, problem, and action to be taken. It was recommended that customers use only manual registration with the 6 degrees of freedom option disabled, which provides the correct offset values. A software update will be deployed once a solution is available. | SoftwareDesignChange
153 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132078 | Z-0930-2015 | Harmony¿ vLED Surgical Lighting System, one unit packaged per box***The Harmony¿ vLED Surgical Lighting System is a variable pattern, variable intensity surgical lighting fixture designed to provide visible illumination of the surgical field or the patient for the operating room staff. | Harmony¿ vLED Surgical Lighting System | 2 | CDRH | 2015-01-06 | 2015-01-07 21:00 | 2015-07-11 | Steris Corporation | STERIS has learned that groups of LED lights (modules) contained within the vLED light head may illuminate inconsistently. | Class 2 Device Recall Harmony vLED Surgical Lighting System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Harmony vLED Surgical Lighting System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Harmony vLED Surgical Lighting System Date Initiated by Firm November 26, 2014 Date Posted January 07, 2015 Recall Status 1 Terminated 3 on July 11, 2015 Recall Number Z-0930-2015 Recall Event ID 69987 510(K)Number K100395 Product Classification Light, surgical, fiberoptic - Product Code FST Product Harmony vLED Surgical Lighting System, one unit packaged per box******The Harmony vLED Surgical Lighting System is a variable pattern, variable intensity surgical lighting fixture designed to provide visible illumination of the surgical field or the patient for the operating room staff. Code Information The following item numbers are included in this recall: B136824175, B136824176, B136824177, B136824178, B136824179, B136824427, B136824430, B136824431, LK01N, and LK05N. Recalling Firm/ Manufacturer Steris Corporation 5960 Heisley Rd Mentor OH 44060-1834 For Additional Information Contact Kathryn E. Cadorette 440-392-7231 Manufacturer Reason for Recall STERIS has learned that groups of LED lights (modules) contained within the vLED light head may illuminate inconsistently. FDA Determined Cause 2 Software Design Change Action STERIS sent an Urgent Voluntary Field Correctionn Notice to all affected consignees via FedEx. The Consignee Notification mailing was initiated the week of December 15, 2014. Customers were instructed that their local STERIS representative would contact their facility to arrange for the installation of updated software. Customers with questions were instructed to call 1-800-548-4873.******For questions regarding this recall call 440-392-7231. Quantity in Commerce 140 units Distribution Worldwide Distribution - USA including Alabama, Alaska, Arkansas, California, District of Columbia, Florida, Illinois, Indiana, Kansas, Louisiana, Mississippi, Nevada, New Jersey, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, Puerto Rico, South Carolina, Tennessee, Texas, Utah, and Virginia and Internationally to Canada, China, Greece, Indonesia, Republic of Korea, Romania, Russian Federation, Taiwan, and Viet Nam. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = FST - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action STERIS sent an Urgent Voluntary Field Correctionn Notice to all affected consignees via FedEx. The Consignee Notification mailing was initiated the week of December 15, 2014. Customers were instructed that their local STERIS representative would contact their facility to arrange for the installation of updated software. Customers with questions were instructed to call 1-800-548-4873.******For questions regarding this recall call 440-392-7231. | SoftwareDesignChange
154 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131262 | Z-0509-2015 | Total Temporomandibular Joint Replacement System- 45mm Lft Narrow Ti Mand | Total Tempomandubilar Joint Prosthesis | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Tempomandubilar Joint Prosthesis Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Tempomandubilar Joint Prosthesis FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Tempomandubilar Joint Prosthesis Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0509-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 45mm Lft Narrow Ti Mand Code Information Model 01-6546TI, Lot #s: 531880B, 535890B, 531870A, 531870B, 531880A, 535890D, 535890A, 535890C. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
155 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132001 | Z-0522-2015 | Total Temporomandibular Joint Replacement System- 45mm Rt Standard TI Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0522-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 45mm Rt Standard TI Mand******Reconstruction of the temporomandibular joint Code Information Model 24-6545TI, Lot #s: 520220B, 520220A, 520210A, 520210B Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
156 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132000 | Z-0521-2015 | Total Temporomandibular Joint Replacement System- 45mm Lft Standard Ti Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0521-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 45mm Lft Standard Ti Mand******Reconstruction of the temporomandibular joint Code Information Model 24-6546TI, Lot #s: 526040B, 526050A, 526040A, 526050B. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
157 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131271 | Z-0515-2015 | Total Temporomandibular Joint Replacement System- 50mm Rt Narrow Ti Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0515-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 50mm Rt Narrow Ti Mand******Reconstruction of the temporomandibular joint Code Information Model 01-6550TI, Lot #s: 531890A, 531890B, 531900A, 531900B, 544220B, 544230A, 544230B, 544200A, 544200B, 544210A, 544210B, 544220A. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
158 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131268 | Z-0512-2015 | Total Temporomandibular Joint Replacement System- 45mm Rt Stand Off TI Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0512-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 45mm Rt Stand Off TI Mand******Reconstruction of the temporomandibular joint Code Information Model 24-6645TI, Lot #s: 472580A, 472590B, 472580B, 472590A, 526030A, 526020A, 526020B, 526030B. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
159 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131277 | Z-0520-2015 | Total Temporomandibular Joint Replacement System- 55mm Rt Standard Ti Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0520-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 55mm Rt Standard Ti Mand******Reconstruction of the temporomandibular joint Code Information Model 24-6555TI, Lot #s: 525210B, 525200A, 525200B, 525210A. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
160 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131276 | Z-0519-2015 | Total Temporomandibular Joint Replacement System- 55mm Lft Standard Ti Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0519-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 55mm Lft Standard Ti Mand******Reconstruction of the temporomandibular joint Code Information Model 24-6556TI, Lot #s: 520240A, 520230A, 550400A, 550390B, 520230B, 520240B, 535400B, 550400B, 535400A, 535410A, 535410B, 550370A, 550370B, 550380A, 550380B, 550390A. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
161 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131273 | Z-0517-2015 | Total Temporomandibular Joint Replacement System- 50mm Rt Standard Ti Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0517-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 50mm Rt Standard Ti Mand******Reconstruction of the temporomandibular joint Code Information Model 24-6650TI, Lot #s: 518480B, 518480A, 525970A, 525970B, 525980B, 518490A, 526070A, 525990A, 525990B, 526060A, 526060B, 526070B, 525980A, 550330A, 550330B, 550340A, 550340B, 550350A, 550360A, 550360B, 518490B, 525960A, 525960B. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
162 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131272 | Z-0516-2015 | Total Temporomandibular Joint Replacement System- 50mm Rt Stand Off Ti Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0516-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 50mm Rt Stand Off Ti Mand******Reconstruction of the temporomandibular joint Code Information Model 24-6650TI, Lot #s: 531960A, 531960B. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
163 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131270 | Z-0514-2015 | Total Temporomandibular Joint Replacement System- 50mm Lft Stand Off Ti Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0514-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 50mm Lft Stand Off Ti Mand******Reconstruction of the temporomandibular joint Code Information Model 24-6651TI, Lot #s: 472560A, 472560B, 472570A, 518520A, 518520B, 472570B. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
164 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131267 | Z-0511-2015 | Total Temporomandibular Joint Replacement System- 45mm Rt Narrow Ti Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0511-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 45mm Rt Narrow Ti Mand******Reconstruction of the temporomandibular joint Code Information Model 01-6545TI, Lot #s: 535380D, 531850A, 531850B, 531860A, 531860B, 535380B, 535380C, 535380A. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
165 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131269 | Z-0513-2015 | Total Temporomandibular Joint Replacement System- 50mm Lft Narrow Ti Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0513-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 50mm Lft Narrow Ti Mand******Reconstruction of the temporomandibular joint Code Information Model 01-6551TI, Lot #s: 535370A, 535370B, 535370C, 535370D. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
166 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131265 | Z-0510-2015 | Total Temporomandibular Joint Replacement System- 45mm Lft Stand Off Ti Mand***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0510-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 45mm Lft Stand Off Ti Mand******Reconstruction of the temporomandibular joint Code Information Models 24-6646TI, Lot #s: 526000A, 526000B, 526010B, 526010A, 535390A, 535390B, 535390C, 535390D, 538720A, 538720B, 538720C, 538720D Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
167 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131274 | Z-0518-2015 | Total Temporomandibular Joint Replacement System- 50mm Lft Standard Ti Mandi***Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated. | Class 2 Device Recall Total Temporomandibular Joint Replacement System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Total Temporomandibular Joint Replacement System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Total Temporomandibular Joint Replacement System Date Initiated by Firm October 17, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on January 04, 2017 Recall Number Z-0518-2015 Recall Event ID 69644 PMA Number P020016 Product Classification Joint, temporomandibular, implant - Product Code LZD Product Total Temporomandibular Joint Replacement System- 50mm Lft Standard Ti Mandi******Reconstruction of the temporomandibular joint Code Information Model 24-6551TI, Lot #s: 544160A, 525190A, 525190B, 525180B, 544170A, 544180B, 544160B, 525180A, 544180A. Recalling Firm/ Manufacturer Biomet Microfixation, LLC 1520 Tradeport Dr Jacksonville FL 32218-2480 For Additional Information Contact Rachel Osbeck 904-741-4400 Manufacturer Reason for Recall Laser etching on the parts is wider and deeper than the conditions previously validated. FDA Determined Cause 2 Software Design Change Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. Quantity in Commerce 42 Distribution Worldwide Distribution - USA including PA, MN, CA and Internationally to Australia, Netherlands, France, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . PMA Database PMAs with Product Code = LZD - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action BIOMET sent an Urgent Medical Device Recall letter dated October 17, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps: 1) Immediately locate and remove from circulation the items invoiced/consigned/loaned to your account. 2) Carefully follow the instructions on the enclosed "Response Form." and 3)If you have further distributed these items to medical facilities, you must notify them of this action. For questions customers were instructed to call 1-800-874-7711 or 904-741-4400, extension 9448, Monday through Friday, 8 a.m. to 4 p.m. EST.******For questions regarding this recall call 904-741-9448. | SoftwareDesignChange
168 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130939 | Z-0211-2015 | Cardiohelp-I System****** Product Usage: A blood oxygenation and carbon dioxide removal system used to pump blood thorough the extracorporeal bypass circuit for circulatory and/or pulmonary support during procedures requiring cardiopulmonary bypass (for periods up to six hours). It is also intended to provide circulatory and/or pulmonary support during procedures not requiring cardiopulmonary bypass (for periods up to six hours). | Maquet Inc. | 2 | CDRH | 2014-11-13 | 2014-11-13 21:00 | 2016-04-06 | Maquet Cardiovascular Us Sales, Llc | Cardiohelp- i System may have a software issue that can potentially result in an erroneous display of a &quot;Battery Needs Service&quot; message after startup of the units when using either AC or DC power | Class 2 Device Recall Maquet Inc. Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Maquet Inc. FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Maquet Inc. Date Initiated by Firm October 17, 2014 Date Posted November 13, 2014 Recall Status 1 Terminated 3 on April 06, 2016 Recall Number Z-0211-2015 Recall Event ID 69557 510(K)Number K102726 Product Classification Console, heart-lung machine, cardiopulmonary bypass - Product Code DTQ Product Cardiohelp-I System************ Product Usage: A blood oxygenation and carbon dioxide removal system used to pump blood thorough the extracorporeal bypass circuit for circulatory and/or pulmonary support during procedures requiring cardiopulmonary bypass (for periods up to six hours). It is also intended to provide circulatory and/or pulmonary support during procedures not requiring cardiopulmonary bypass (for periods up to six hours). Code Information 90410803 90410984 90410456 90410609 90410610 90410740 90410737 90410738 90410733 90410865 90411005 90410552 90410572 90410849 90410864 90410867 90410796 90410481 90410319 90410328 90411176 90410988 90410703 90410431 90411186 90411188 90411191 90410800 90410611 90410669 90410517 90410692 90410499 90410500 90410739 90410612 90410714 90410759 90410771 90410430 90410276 90410576 90410482 90410749 90410774 90410859 90410860 90410863 90411088 90411082 90411087 90411091 90411089 90410866 90411144 90410970 90411023 90411033 90411038 90411032 90410974 90410962 90410455 90410448 90410449 90411175 90410524 90411007 90410353 90410433 90410435 90410770 90410785 90410440 90410932 90410968 90410732 90411035 90411028 90410736 90410249 90410410 90410411 90411169 90411157 90410963 90411078 90410237 90410284 90410786 90410772 90410922 90411093 90411192 90411193 90410582 90410516 90410662 90410666 90410664 90410665 90410235 90411189 90411187 90411154 90411179 90411178 90410580 90410596 90410599 90410651 90410646 90410735 90410787 90410839 90410851 90410848 90410852 90410845 90411105 90411143 90411145 90411141 90410530 90410584 90410607 90410575 90410581 90410498 90410756 90410936 90410663 90410909 90410746 90410744 90410450 90410606 90410979 90411004 90411009 90411006 90411010 90410398 90410400 90410399 90410713 90410973 90410351 90410352 90410840 90410242 90410252 90410239 90410426 90410427 90410754 90410503 90410518 90411008 90410573 90410577 90410585 90410668 90410408 90410409 90410653 90410961 90410982 90411174 90410568 90410406 90410640 90410644 90410534 90410537 90410780 90411146 90411181 90411184 90411108 90411107 90411171 90410986 90411003 90410716 90410638 90410717 90411036 90411037 90410847 90410231 90410458 90410779 90410783 90411168 90411170 90410975 90410782 90410846 90410838 90410650 90410608 90410844 90410745 90410980 90410777 90410930 90410278 90410355 90410389 90410989 90410277 90410520 90410964 90410734 90410929 90410927 90410832 90410595 90410532 90410535 90410529 90410969 90411026 90410514 90410519 90410778 90410933 90410928 90411194 90411196 90410453 90411086 90411104 90410842 90410758 90410452 90410985 90410422 90410423 90410670 90410407 90410600 90410850 90411034 90411002 90410256 90410983 90410931 90410752 90410776 90410321 90410413 90410702 90410976 90410523 90410700 90410405 90410454 90410457 90410521 90410551 90410711 90410706 90410742 90410743 90410921 90410925 90411182 90411183 90411185 90411090 90411106 90411155 90411156 90410571 90410446 90410480 90410784 90411092 90410526 90410501 90410533 90410550 90410698 90410841 90410987 90410253 90410234 90411011 90411001 90411012 90411024 90410857 90410858 90411180 90410350 90410981 90410656 90410788 90410795 90411148 90411149 90410232 90410271 90410536 90410924 90410336 90410340 90410459 90410926 90410934 90411190 90410522 90410760 90410251 90410484 90411147 90411197 90411198 90410977 90411025 90411027 90410515 90410843 90410527 90410086 90410747 90410748 90410923 90410978 Recalling Firm/ Manufacturer Maquet Cardiovascular Us Sales, Llc 45 Barbour Pond Drive Wayne NJ 07470 For Additional Information Contact Ms. Karen LeFevre 973-709-7652 Manufacturer Reason for Recall Cardiohelp- i System may have a software issue that can potentially result in an erroneous display of a "Battery Needs Service" message after startup of the units when using either AC or DC power FDA Determined Cause 2 Software Design Change Action Maquet Inc. sent a Product Recall Urgent Medicall Device Field Correction letter to all of their customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter stated that a software upgrade will be performed by an authorized Maquet Service Technician. A Maquet representative will contact customer to schedule an on-site service of the device. Any questions can be directed to the local Maquet representative or Maquet Customer Service at 1-888-627-8383 (option 2, option 2) Monday through Friday, between the hours of 8:00 am and 6:00 pm (EST). Quantity in Commerce 331 units Distribution Worldwide Distribution - US Nationwide Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = DTQ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Maquet Inc. sent a Product Recall Urgent Medicall Device Field Correction letter to all of their customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter stated that a software upgrade will be performed by an authorized Maquet Service Technician. A Maquet representative will contact customer to schedule an on-site service of the device. Any questions can be directed to the local Maquet representative or Maquet Customer Service at 1-888-627-8383 (option 2, option 2) Monday through Friday, between the hours of 8:00 am and 6:00 pm (EST). | SoftwareDesignChange
169 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130737 | Z-0190-2015 | MX 16-slice CT Scanner System and MX 16-slice CT Scanner System Phase 11 can be used as a whole body computed tomography X-ray system featuring a continuously rotating X-ray tube and detector array with multi-slice capability up to 16 slices simultaneously. The acquired X-RAY transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. The system is suitable for all patients. Access Dual CT system can be used as a Head and Whole Body Computed Tomography X-ray System featuring a continuously rotating X-ray tube and detector array with multi-slice capability. The acquired X-ray transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. | Philips Healthcare MX 16-slice CT Scanner System and MX 16-slice CT Scanner System Phase 11 | 2 | CDRH | 2014-11-07 | 2014-12-16 12:46 | 2015-07-02 | Philips Medical Systems (Cleveland) Inc | &quot;The orientation of images is displayed incorrectly. When operator chooses gantry on the right side from operation console, and the patient position is HFDL, HFDR, FFDL or FFDR, the patient position picture and text is not consistent. &quot;L&quot; and &quot;R&quot; of scan images are reversed and this is not consistent with the patient actual position. ***&quot;When a surview which is longe | Class 2 Device Recall Philips Healthcare MX 16slice CT Scanner System and MX 16slice CT Scanner System Phase 11 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare MX 16slice CT Scanner System and MX 16slice CT Scanner System Phase 11 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare MX 16slice CT Scanner System and MX 16slice CT Scanner System Phase 11 Date Initiated by Firm October 09, 2014 Date Posted December 16, 2014 Recall Status 1 Terminated 3 on July 02, 2015 Recall Number Z-0190-2015 Recall Event ID 69473 510(K)Number K083498 K091195 Product Classification System, x-ray, tomography, computed - Product Code JAK Product MX 16-slice CT Scanner System and MX 16-slice CT Scanner System Phase 11 can be used as a whole body computed tomography X-ray system featuring a continuously rotating X-ray tube and detector array with multi-slice capability up to 16 slices simultaneously. The acquired X-RAY transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. The system is suitable for all patients. Access Dual CT system can be used as a Head and Whole Body Computed Tomography X-ray System featuring a continuously rotating X-ray tube and detector array with multi-slice capability. The acquired X-ray transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. Code Information Model Number: 728131, Serial Number(s): EP16E090001,EP16E090002,EP16E090003,EP16E090004,EP16E090006,EP16E090007,EP16E090008, EP16E090009,EP16E090010,EP16E090011,EP16E090012,EP16E090013,EP16E090015,EP16E100001, EP16E100004,EP16E100005,EP16E100006,EP16E100007,EP16E100008,EP16E100010,EP16E100011, EP16E100012,EP16E100013,EP16E100015,EP16E100016,EP16E100017,EP16E100018, EP16E100019,EP16E100020,EP16E100021,EP16E100022,EP16E100023,EP16E100024,EP16E100025, EP16E100026,EP16E100027,EP16E100028,EP16E100029,EP16E100030,EP16E100031,EP16E100032, EP16E100033,EP16E100034,EP16E100035,EP16E100036,EP16E100037,EP16E100038,EP16E100039, EP16E100041,EP16E100043,EP16E100044,EP16E100045,EP16E100046,EP16E100047,EP16E100048, EP16E100049,EP16E100051,EP16E100052,EP16E100055,EP16E100056,EP16E100057,EP16E100058, EP16E100059,EP16E100060,EP16E100061,EP16E100062,EP16E100063,EP16E100067,EP16E100069, EP16E100070,EP16E100071,EP16E100072,EP16E100073,EP16E100074,EP16E100075,EP16E100076, EP16E100077,EP16E100078,EP16E100079,EP16E100080,EP16E100081,EP16E100082,EP16E100083, EP16E100084,EP16E100085,EP16E100086,EP16E100087,EP16E100088,EP16E100090,EP16E100091, EP16E100092,EP16E100093,EP16E100094,EP16E100095,EP16E100096,EP16E100099,EP16E100100, EP16E100101,EP16E100102,EP16E100103,EP16E100104,EP16E100105,EP16E100106,EP16E100107, EP16E100108,EP16E100109,EP16E100110,EP16E100111,EP16E100112,EP16E100113,EP16E110003, EP16E110005,EP16E110006,EP16E110007,EP16E110009,EP16E110011,EP16E110012,EP16E110013, EP16E110014,EP16E110015,EP16E110017,EP16E110018,EP16E110019,EP16E110020,EP16E110021, EP16E110023,EP16E110024,EP16E110025,EP16E110026,EP16E110027,EP16E110028,EP16E110029, EP16E110030,EP16E110031,EP16E110032,EP16E110033,EP16E110034,EP16E110035,EP16E110037, EP16E110038,EP16E110039,EP16E110042,EP16E110043,EP16E110044,EP16E110045,EP16E110046, EP16E110047,EP16E110048,EP16E110049,EP16E110050,EP16E110051,EP16E110052,EP16E110054, EP16E110055,EP16E110056,EP16E110057,EP16E110058,EP16E110060,EP16E110062,EP16E110063, EP16E110064,EP16E110066,EP16E110067,EP16E110069,EP16E110070,EP16E110074,EP16E110075, EP16E110077,EP16E110078,EP16E110083,EP16E110085,EP16E110089,EP16E110090,EP16E110091, EP16E110092,EP16E110093,EP16E110095,EP16E110096,EP16E110097,EP16E110099,EP16E110100, EP16E110101,EP16E110102,EP16E110104,EP16E110105,EP16E110106,EP16E110108,EP16E110109, EP16E110110,EP16E110111,EP16E110115,EP16E110117,EP16E110118,EP16E110119,EP16E120003, EP16E120004,EP16E120005,EP16E120006,EP16E120007,EP16E120008,EP16E120009,EP16E120010, EP16E120011,EP16E120014,EP16E120015,EP16E120016,EP16E120017,EP16E120019,EP16E120021, EP16E120023,EP16E120024,EP16E120025,EP16E120026,EP16E120027,EP16E120029,EP16E120030, EP16E120031,EP16E120032,EP16E120033,EP16E120034,EP16E120035,EP16E120036,EP16E120037, EP16E120039,EP16E120041,EP16E120042,EP16E120043,EP16E120044,EP16E120045,EP16E120046, EP16E120047,EP16E120048,EP16E120050,EP16E120051,EP16E120052,EP16E120054,EP16E120055, EP16E120056,EP16E120058,EP16E120059,EP16E120060,EP16E120061,EP16E120063,EP16E120065, EP16E120067,EP16E120070,EP16E120071,EP16E120072,EP16E120074,EP16E120075,EP16E120076, EP16E120078,EP16E120081,EP16E120083,EP16E120088,EP16E120090,EP16E120091,EP16E120092, EP16E120093,EP16E120095,EP16E120096,EP16E120097,EP16E120098,EP16E120099,EP16E120100, EP16E120101,EP16E120102,EP16E120104,EP16E120105,EP16E120109,EP16E120110,EP16E120111, EP16E120113,EP16E120117,EP16E120118,EP16E120120,EP16E120121,EP16E120122,EP16E120123, EP16E120124,EP16E120125,EP16E120127,EP16E120128,EP16E120129,EP16E120130,EP16E130002, EP16E130003,EP16E130004,EP16E130005,EP16E130006,EP16E130007,EP16E130009,EP16E130010, EP16E130012,EP16E130013,EP16E130014,EP16E130015,EP16E130016,EP16E130017,EP16E130020, EP16E130021,EP16E130022,EP16E130023,EP16E130025,EP16E130026,EP16E130027,EP16E130029, EP16E130030,EP16E130031,EP16E130033,EP16E130034,EP16E130035,EP16E130036,EP16E130037, EP16E130040,EP16E130041,EP16E130042,EP16E130043,EP16E130044,EP16E130045,EP16E130046, EP16E130047,EP16E130048,EP16E130049,EP16E130050,EP16E130051,EP16E130052,EP16E130053, EP16E130055,EP16E130056,EP16E130057,EP16E130058,EP16E130062,EP16E130063,EP16E130064, EP16E130065,EP16E130066,EP16E130067,EP16E130068,EP16E130069,EP16E130070,EP16E130071, EP16E130072,EP16E130073,EP16E130074,EP16E130075,EP16E130076,EP16E130077,EP16E130078, EP16E130079,EP16E130080,EP16E130081,EP16E130082,EP16E130083,EP16E130086,EP16E130087, EP16E130088,EP16E130089,EP16E130091,EP16E130092,EP16E130093,EP16E130094,EP16E130095, EP16E130097,EP16E130098,EP16E130099,EP16E130100,EP16E130101,EP16E130102,EP16E130103, EP16E130105,EP16E130106,EP16E130107,EP16E130109,EP16E130110,EP16E140001,EP16E140002, EP16E140003,EP16E140006,EP16E140007,EP16E140008,EP16E140009,EP16E140010,EP16E140011, EP16E140012,EP16E140013,EP16E140014,EP16E140015,EP16E140016,EP16E140017,EP16E140018, EP16E140019,EP16E140021,EP16E140022,EP16E140023,EP16E140024,EP16E140025,EP16E140026, EP16E140027,EP16E140028,EP16E140029,EP16E140030,EP16E140031,EP16E140032,EP16E140033, EP16E140041,EP16EB110003,EP16EB110004,EP16EB110005,EP16EB110006,EP16EB120006, EP16EB120007, EP16EB120008, EP16EB120009, EP16EB130001, EP16EB130002, EP16EB130003, EP16EB130004, EP16EB130005, EP16EB130006, EP16EB130007, EP16EB130008, EP16EB130009, EP16EB130010, EP16EB130011, EP16EB130012, EP16EB130013, EP16EB130014, EP16EB130015, EP16EB130016, EP16EB130017, EP16EB130018, EP16EB130019, EP16EB130020, EP16EB130021, EP16EB140001, EP16EB140002, P16C080001, P16C080002, P16C100001, P16C100002, P16C100003, P16C100004, P16C100005, P16C100007, P16C100008, P16C100009, P16C100010, P16C100011, P16C100012, P16C100013, P16C100014, P16C100015, P16C100017, P16C100018, P16C110001, P16C110002, P16C110003, P16C110004, P16C110005, P16C110006, P16C110007, P16C110008, P16C110011, P16C110012, P16C110013, P16C110014, P16C110015, P16C110016, P16C110017, P16C110018, P16C110019, P16C110020, P16C110021, P16C110022, P16C110023, P16C110024, P16C110025, P16C110026, P16C110027, P16C110028, P16C110029, P16C110030, P16C110031, P16C110032, P16C110033, P16C110034, P16C110035, P16C110036, P16C110037, P16C120001, P16C120002, P16C120003, P16C120006, P16C120007, P16C120008, P16C120009, P16C120010, P16C120011, P16C120012, P16C120014, P16C120015, P16C120016, P16C120017, P16C120018, P16C120019, P16C120020, P16C120021, P16C120022, P16C120023, P16C120024, P16C120025, P16C120026, P16C120027, P16C120028, P16C120029, P16C120030, P16C120031, P16C120032, P16C120033, P16C120034, P16C120035, P16C120036, P16C120037, P16C120038, P16C120039, P16C120040, P16C120041, P16C120042, P16C120043, P16C120044, P16C120045, P16C120046, P16C120047, P16C120048, P16C120049, P16C120050, P16C120051, P16C130001, P16C130002, P16C130003, P16C130004, P16C130005, P16C130006, P16C130007, P16C130008, P16C130009, P16C130010, P16C130011, P16C130012, P16C130013, P16C130014, P16C130015, P16C130016, P16C130017, P16C130018, P16C130019, P16C130020, P16C130021, P16C130023, P16C130024, P16C130025, P16C130026, P16C130027, P16C130028, P16C130029, P16C130030, P16C130031, P16C130032, P16C130033, P16C130034, P16C130035, P16C130036, P16C130037, P16C130038, P16C130039, P16C130041, P16C130043, P16C130044, P16C130045, P16C130046, P16C130047, P16C130048, P16C130049, P16C130050, P16C130051, P16C130052, P16C130053, P16E080005, P16E080007, P16E080008, P16E090002, P16E090003, P16E090004, P16E090005, P16E090006, P16E090007, P16E090008, P16E090009, P16E090010, P16E090011, P16E090012, P16E090013, P16E090014, P16E090015, P16E090016, P16E090017, P16E090018, P16E090019, P16E090020, P16E090021, P16E090024, P16E090025, P16E090026, P16E090027, P16E090028, P16EF80010, P16EM08003, P16EM09002, P16EM09003, P16EM09004, P16EVO130002, P16EVO130003, P16EVO140001, P16EVO140002, P16EVO140003, P16EVO140005, P16EVO140006, P16EVO140007, P16EVO140008 Model Number: 728132, Serial Number(s): EP16ER100003,EP16ER100004,EP16ER100005,EP16ER100007,EP16ER100008,EP16ER110003, EP16ER110004,EP16ER110005,EP16ER110006,EP16ER110007,EP16ER110008,EP16ER110009, EP16ER110010,EP16ER110011,EP16ER110012,EP16ER110013,EP16ER110014,EP16ER110015, EP16ER110016,EP16ER110017,EP16ER110018,EP16ER110019,EP16ER110020,EP16ER110021, EP16ER110022,EP16ER110023,EP16ER110024,EP16ER120001,EP16ER120002,EP16ER120003, EP16ER120004,EP16ER120005,EP16ER120006,EP16ER120007,EP16ER120008,EP16ER120009, EP16ER120010,EP16ER120011,EP16ER120012,EP16ER120013,EP16ER120015,EP16ER120016, EP16ER120017,EP16ER120018,EP16ER130001,EP16ER130002,EP16ER130003,EP16ER130004, EP16ER130005,EP16ER130006,EP16ER130007,EP16ER130008,EP16ER130009,EP16ER130010, EP16ER130011,EP16ER130012,EP16ER130013,EP16ER130014,EP16ER130015,EP16ER130016, EP16ER130017,EP16ER130018,EP16ER130019,EP16ER130020,EP16ER130021,EP16ER130022, EP16ER130023,EP16ER130024,EP16ER130025,EP16ER130026,EP16ER130027,EP16ER130028, EP16ER130029,EP16ER130030,EP16ER130031,EP16ER130032,EP16ER130033,EP16ER130034 Recalling Firm/ Manufacturer Philips Medical Systems (Cleveland) Inc 595 Miner Rd Cleveland OH 44143-2131 For Additional Information Contact Daniel Brown 440-483-7600 Manufacturer Reason for Recall "The orientation of images is displayed incorrectly. When operator chooses ***gantry on the right side*** from operation console, and the patient position is HFDL, HFDR, FFDL or FFDR, the patient position picture and text is not consistent. "L" and "R" of scan images are reversed and this is not consistent with the patient actual position. ******"When a surview which is longe FDA Determined Cause 2 Software Design Change Action Philips is informing the consignees of this issue through a Field Safety Notice (FSN) 72800615 on October 09, 2014. The US consignees will receive this notice via certified mail and Philips Key Market representatives will distribute the FSN internationally. Additionally, A Field Service Engineer will contact each site to schedule a time to install the software update. Quantity in Commerce 708 Distribution Domestic Distribution: AR, CA, FL, GA, IL, IN, KY, MI, MS, NC, ND, NJ, NV, NY, OH, PA, PR, SC, TN, and VA. Foreign distribution: Albania, Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Chile, China, Colombia, Congo, Ecuador, Egypt, France, Germany, Ghana, Greece, Gutamala, Hungray, India, Indonesia, Ireland, Italy, Japan, Jordan, Latvia, Lybyan Arab JHamahiri, Lithuania, Malawi, Mexico, Myanmar, Namibia, Nepal, Netherlands, Oman, Palestine, Paraguay, Philippines, Poland, Portugal, Romania,Russia, Rwanda, Saudi Arabia, Senegal, South Agrica, South ZKorea, Spain, Sri Lanka, Switzerland, Tunisia, Turkey, Uganda, UAE, Venezuela, Vietnam, Zambia. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 The manufacturer has initiated the recall and not all products have been corrected or removed. This record will be updated as the status changes. 510(K) Database 510(K)s with Product Code = JAK - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips is informing the consignees of this issue through a Field Safety Notice (FSN) 72800615 on October 09, 2014. The US consignees will receive this notice via certified mail and Philips Key Market representatives will distribute the FSN internationally. Additionally, A Field Service Engineer will contact each site to schedule a time to install the software update. | SoftwareDesignChange
170 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130271 | Z-0105-2015 | Philips Avalon Monitors with software revision J.30.58: ******Model******Product***FM20******M2702A ;***FM30******M2703A ;***FM50******M2705A******Intended for: " Monitoring the physiological parameters of pregnant women " Non-invasive monitoring of fetal heart rates and movements. | Philips Avalon Monitors | 2 | CDRH | 2014-10-21 | 2014-10-21 21:00 | 2017-01-31 | Philips Medical Systems, Inc. | Philips Intellivue and Avalon Fetal Monitors in time-synchronized automatic/sequence mode, the NBP automatic measurement series is stopped | Class 2 Device Recall Philips Avalon Monitors Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Avalon Monitors FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Avalon Monitors Date Initiated by Firm October 03, 2014 Date Posted October 21, 2014 Recall Status 1 Terminated 3 on January 31, 2017 Recall Number Z-0105-2015 Recall Event ID 69348 510(K)Number K111083 Product Classification System, monitoring, perinatal - Product Code HGM Product Philips Avalon Monitors with software revision J.30.58: ************Model******Product******FM20******M2702A ;******FM30******M2703A ;******FM50******M2705A************Intended for: " Monitoring the physiological parameters of pregnant women " Non-invasive monitoring of fetal heart rates and movements. Code Information software revision J.30.58: Recalling Firm/ Manufacturer Philips Medical Systems, Inc. 3000 Minuteman Rd Andover MA 01810-1032 For Additional Information Contact Same 978-687-1501 Manufacturer Reason for Recall Philips Intellivue and Avalon Fetal Monitors in time-synchronized automatic/sequence mode, the NBP automatic measurement series is stopped FDA Determined Cause 2 Software Design Change Action Philips Healthcare issued on 10/3/14 the Urgent Medical Device Correction notification/Field Safety Notice. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.******The correction will consist of a software upgrade free of charge.******A Philips Healthcare representative will contact customers with affected devices to arrange for service. ************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Urgent Medical Device Correction notification/Field Safety Notice:************During the interim period until the SW is upgraded please make sure, if the NBP repetition interval needs to be increased (e.g. from 5 to 15 minutes), to restart the automatic NBP measurement.************Please review this information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication. Quantity in Commerce 118 units Distribution Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA AUSTRIA BAHRAIN BELGIUM CHINA COLOMBIA CZECH REPUBLIC DENMARK FINLAND FRANCE GERMANY HONG KONG INDIA IRELAND ISRAEL ITALY JAPAN KAZAKHSTAN NETHERLANDS NEW ZEALAND NICARAGUA NIGERIA NORWAY OMAN POLAND ROMANIA RUSSIA SINGAPORE SLOVENIA SOUTH AFRICA SOUTH KOREA SPAIN SWEDEN SWITZERLAND TAIWAN THAILAND TURKEY UKRAINE and UNITED KINGDOM. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = HGM - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare issued on 10/3/14 the Urgent Medical Device Correction notification/Field Safety Notice. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.******The correction will consist of a software upgrade free of charge.******A Philips Healthcare representative will contact customers with affected devices to arrange for service. ************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Urgent Medical Device Correction notification/Field Safety Notice:************During the interim period until the SW is upgraded please make sure, if the NBP repetition interval needs to be increased (e.g. from 5 to 15 minutes), to restart the automatic NBP measurement.************Please review this information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication. | SoftwareDesignChange
171 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130270 | Z-0104-2015 | Philips IntelliVue Monitors with software revisions J.21.03, J.21.19. ******Model******Product***MP5******M8105A;***MP5SC******M8105AS;***MP5T ******M8105AT;***MP5 Upgrade***M8105AU ;***MP5SC Upgrade***866327******The monitors are intended to be used for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults, pediatrics, and neonates. | Philips IntelliVue Monitors | 2 | CDRH | 2014-10-21 | 2014-10-21 21:00 | 2017-01-31 | Philips Medical Systems, Inc. | Philips Intellivue and Avalon Fetal Monitors in time-synchronized automatic/sequence mode, the NBP automatic measurement series is stopped | Class 2 Device Recall Philips IntelliVue Monitors Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips IntelliVue Monitors FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips IntelliVue Monitors Date Initiated by Firm October 03, 2014 Date Posted October 21, 2014 Recall Status 1 Terminated 3 on January 31, 2017 Recall Number Z-0104-2015 Recall Event ID 69348 510(K)Number K122439 K131913 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product Philips IntelliVue Monitors with software revisions J.21.03, J.21.19. ************Model******Product******MP5******M8105A;******MP5SC******M8105AS;******MP5T ******M8105AT;******MP5 Upgrade***M8105AU ;******MP5SC Upgrade***866327************The monitors are intended to be used for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults, pediatrics, and neonates. Code Information software revisions J.21.03, J.21.19 Recalling Firm/ Manufacturer Philips Medical Systems, Inc. 3000 Minuteman Rd Andover MA 01810-1032 For Additional Information Contact Same 978-687-1501 Manufacturer Reason for Recall Philips Intellivue and Avalon Fetal Monitors in time-synchronized automatic/sequence mode, the NBP automatic measurement series is stopped FDA Determined Cause 2 Software Design Change Action Philips Healthcare issued on 10/3/14 the Urgent Medical Device Correction notification/Field Safety Notice. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.******The correction will consist of a software upgrade free of charge.******A Philips Healthcare representative will contact customers with affected devices to arrange for service. ************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Urgent Medical Device Correction notification/Field Safety Notice:************During the interim period until the SW is upgraded please make sure, if the NBP repetition interval needs to be increased (e.g. from 5 to 15 minutes), to restart the automatic NBP measurement.************Please review this information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication. Quantity in Commerce 7587 units Distribution Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA AUSTRIA BAHRAIN BELGIUM CHINA COLOMBIA CZECH REPUBLIC DENMARK FINLAND FRANCE GERMANY HONG KONG INDIA IRELAND ISRAEL ITALY JAPAN KAZAKHSTAN NETHERLANDS NEW ZEALAND NICARAGUA NIGERIA NORWAY OMAN POLAND ROMANIA RUSSIA SINGAPORE SLOVENIA SOUTH AFRICA SOUTH KOREA SPAIN SWEDEN SWITZERLAND TAIWAN THAILAND TURKEY UKRAINE and UNITED KINGDOM. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips Healthcare issued on 10/3/14 the Urgent Medical Device Correction notification/Field Safety Notice. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.******The correction will consist of a software upgrade free of charge.******A Philips Healthcare representative will contact customers with affected devices to arrange for service. ************Philips is asking customers to follow the ***Action to be Taken by Customer/User*** section of the Urgent Medical Device Correction notification/Field Safety Notice:************During the interim period until the SW is upgraded please make sure, if the NBP repetition interval needs to be increased (e.g. from 5 to 15 minutes), to restart the automatic NBP measurement.************Please review this information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication. | SoftwareDesignChange
172 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=127243 | Z-1741-2014 | enGen (TM) Laboratory Automation System Configured with TCAutomation" Software v3.5 -- For in vitro quantitative measurement of a variety of analytes of clinical interest ---******The VITROS 5,1 FS Chemistry System with enGenTM Laboratory Automation System is intended for use in the in vitro quantitative measurement of a variety of analytes of clinical interest, using both VITROS Chemistry Products Slides (colorimetric endpoint, rate, ion-selective electrode, and immunorate methods) and VITROS Chemistry Products MicroTip liquid reagents (spectrophotometric and spectrophotometric immunoassay methods.). ---The enGen (TM) Lab Automation System is a modification to VITROS Analyzers that includes the addition of an automated track sub-system (TCAutomation) for pre-analytical and post-analytical sample and data management. In addition to an Automated Track Sub-system, the enGen (TM) System includes one or more VITROS Chemistry, Immunodiagnostic and/or Integrated Systems (clinical analyzers) as well as a Data Management Subsystem (Middleware), and a Software Configuration (gsb) File. The enGen System is typically interfaced with a Laboratory Information System (LIS) via connectivity facilitated at the Data Management Subsystem level. By using the enGen (TM) Laboratory Automation System as a communication interface between the LIS and the clinical analyzers in the laboratory, the user can manage the workflow to the analyzers and other devices as an integrated system. This eliminates the need for the user to program tests for each sample uniquely on each analyzer. Within the enGen (TM) Laboratory Automation System, the physical routing of samples to the correct location is also automatic. The Buffer Module functions either as a location for samples waiting for results or as a storage area for empty sample tube carriers. When a sample is presented to a clinical analyzer via the Bypass module and is metered for testing, the sample remains on the track system in the Buffer module while waiting for the test results to be generated. If all test results are generated as expected and no additional actions are required, the sample is routed to a designated Exit module where it can be removed from the track system by the user. However, there are certain situations in which a follow-up action (retest of the sample) may be necessary. In these situations, the sample is routed back to a clinical analyzer via the Bypass module to complete the follow-up action. Examples of situations requiring a retest include, but are not limited to: repeat testing when a "No Result"is generated during the initial test event, sample dilution when an out of range result is obtained during the initial test event, and routing to an alternate analyzer when the first analyzer the sample is sent to is unavailable for testing. Once the follow-up sample metering actions are complete, the sample is routed to Buffer module where it remains until testing is complete. Once all test results are obtained the sample is routed to the designated Exit module. | enGen (TM) Laboratory Automation System | 2 | CDRH | 2014-06-10 | 2014-06-10 21:00 | 2017-04-11 | Ortho-Clinical Diagnostics | Software anomaly: potential buffer timeout issues when using enGen (TM) Laboratory Automation Systems Configured with TCAutomation (TM) Software v3.5. | Class 2 Device Recall enGen (TM) Laboratory Automation System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall enGen (TM) Laboratory Automation System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall enGen (TM) Laboratory Automation System Date Initiated by Firm April 17, 2014 Date Posted June 10, 2014 Recall Status 1 Terminated 3 on April 11, 2017 Recall Number Z-1741-2014 Recall Event ID 68269 510(K)Number K063144 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product enGen (TM) Laboratory Automation System Configured with TCAutomation" Software v3.5 -- For in vitro quantitative measurement of a variety of analytes of clinical interest ---************The VITROS 5,1 FS Chemistry System with enGenTM Laboratory Automation System is intended for use in the in vitro quantitative measurement of a variety of analytes of clinical interest, using both VITROS Chemistry Products Slides (colorimetric endpoint, rate, ion-selective electrode, and immunorate methods) and VITROS Chemistry Products MicroTip liquid reagents (spectrophotometric and spectrophotometric immunoassay methods.). ---The enGen (TM) Lab Automation System is a modification to VITROS Analyzers that includes the addition of an automated track sub-system (TCAutomation) for pre-analytical and post-analytical sample and data management. In addition to an Automated Track Sub-system, the enGen (TM) System includes one or more VITROS Chemistry, Immunodiagnostic and/or Integrated Systems (clinical analyzers) as well as a Data Management Subsystem (Middleware), and a Software Configuration (gsb) File. The enGen System is typically interfaced with a Laboratory Information System (LIS) via connectivity facilitated at the Data Management Subsystem level. By using the enGen (TM) Laboratory Automation System as a communication interface between the LIS and the clinical analyzers in the laboratory, the user can manage the workflow to the analyzers and other devices as an integrated system. This eliminates the need for the user to program tests for each sample uniquely on each analyzer. Within the enGen (TM) Laboratory Automation System, the physical routing of samples to the correct location is also automatic. The Buffer Module functions either as a location for samples waiting for results or as a storage area for empty sample tube carriers. When a sample is presented to a clinical analyzer via the Bypass module and is metered for testing, the sample remains on the track system in the Buffer module while waiting for the test results to be generated. If all test results are generated as expected and no additional actions are required, the sample is routed to a designated Exit module where it can be removed from the track system by the user. However, there are certain situations in which a follow-up action (retest of the sample) may be necessary. In these situations, the sample is routed back to a clinical analyzer via the Bypass module to complete the follow-up action. Examples of situations requiring a retest include, but are not limited to: repeat testing when a "No Result"is generated during the initial test event, sample dilution when an out of range result is obtained during the initial test event, and routing to an alternate analyzer when the first analyzer the sample is sent to is unavailable for testing. Once the follow-up sample metering actions are complete, the sample is routed to Buffer module where it remains until testing is complete. Once all test results are obtained the sample is routed to the designated Exit module. Code Information enGen (TM) Laboratory Automation Systems configured with TCAutomation" Software version 3.5 Recalling Firm/ Manufacturer Ortho-Clinical Diagnostics 100 Indigo Creek Dr Rochester NY 14626-5101 Manufacturer Reason for Recall Software anomaly: potential buffer timeout issues when using enGen (TM) Laboratory Automation Systems Configured with TCAutomation (TM) Software v3.5. FDA Determined Cause 2 Software Design Change Action Ortho Clinical Diagnostics sent an Important Product Correction Notification dated April 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.************On April 17, 2014, foreign affiliates were informed by e-mail of the issue and instructed to notify their customers of the issue and actions.************The letters instruct customers: to ensure that samples prematurely routed to Check are detected and processed appropriately, please monitor completion of all pending tests following standard laboratory practice. Customers were informed of the issue and advised to: -- Routinely monitor pending result reports from your Laboratory Information System (LIS) to determine if results are pending for samples. -- Investigate to determine the next action identified for a sample if the sample is routed to the Check location prematurely. ************ If you have any additional questions, please call Customer Technical Services at 1-800-421-3311. Quantity in Commerce Domestic: 6 units, Foreign: 3 units Distribution Worldwide Distribution - US including the states of Arizona, Delaware, Ohio, Tennessee, and North Carolina; and the countries of Mexico, England and Italy. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Ortho Clinical Diagnostics sent an Important Product Correction Notification dated April 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.************On April 17, 2014, foreign affiliates were informed by e-mail of the issue and instructed to notify their customers of the issue and actions.************The letters instruct customers: to ensure that samples prematurely routed to Check are detected and processed appropriately, please monitor completion of all pending tests following standard laboratory practice. Customers were informed of the issue and advised to: -- Routinely monitor pending result reports from your Laboratory Information System (LIS) to determine if results are pending for samples. -- Investigate to determine the next action identified for a sample if the sample is routed to the Check location prematurely. ************ If you have any additional questions, please call Customer Technical Services at 1-800-421-3311. | SoftwareDesignChange
173 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=127138 | Z-1617-2014 | McKesson Radiology-PACS.******McKesson Radiology is a medical image and information application that is intended to receive, transmit, store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from various medical imaging systems. | Picture Archive Communication System (PACS) | 2 | CDRH | 2014-05-13 | 2014-05-13 21:00 | 2016-02-02 | Mckesson Medical Immaging | High priority alert icon was not displayed in the Study List when there are no Stat studies present. However, the High priority alert icon is visible when studies marked as Stat priority exist in the Study List. | Class 2 Device Recall Picture Archive Communication System (PACS) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Picture Archive Communication System (PACS) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Picture Archive Communication System (PACS) Date Initiated by Firm April 17, 2014 Date Posted May 13, 2014 Recall Status 1 Terminated 3 on February 02, 2016 Recall Number Z-1617-2014 Recall Event ID 68208 510(K)Number K043146 Product Classification System, image processing, radiological - Product Code LLZ Product McKesson Radiology-PACS.************McKesson Radiology is a medical image and information application that is intended to receive, transmit, store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from various medical imaging systems. Code Information Model MR 12.0 Recalling Firm/ Manufacturer Mckesson Medical Immaging Suite 130-10711 Cambie Road Richmond Canada British Columbia For Additional Information Contact Paul Sumner 404-338-3556 Manufacturer Reason for Recall High priority alert icon was not displayed in the Study List when there are no Stat studies present. However, the High priority alert icon is visible when studies marked as Stat priority exist in the Study List. FDA Determined Cause 2 Software Design Change Action On April 17, 2014, McKesson Medical Imaging Company initiated a process for applying a software solution to all McKesson Radiology customer systems that may be affected to prevent future occurrences of this issue. The systems with the potentially affected software versions will be corrected without removing the systems from their location. Quantity in Commerce 45 Distribution Worldwide Distribution-USA (nationwide) including the states of AL, CA, GA, ID, IL, IA, MD, MS, MO, NE, NV, NH, NY, NC, PA, TN, TX, VT, WA, WV, WY, and the countries of Canada and France. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On April 17, 2014, McKesson Medical Imaging Company initiated a process for applying a software solution to all McKesson Radiology customer systems that may be affected to prevent future occurrences of this issue. The systems with the potentially affected software versions will be corrected without removing the systems from their location. | SoftwareDesignChange
174 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126644 | Z-1528-2014 | INOmax DSIR, model 10007, Nitric Oxide Delivery System, with Software version 3.0.0. The INOmax DS delivery system delivers INOMAX (nitric oxide for inhalation) therapy gas and provides continuous monitoring of inspired O2, NO2, and NO, and a comprehensive alarm system. The primary targeted clinical setting is the Neonatal Intensive Care Unit (NICU) and secondary targeted clinical setting is the transport of neonates. | INOmax DSIR Delivery System | 2 | CDRH | 2014-05-02 | 2014-05-02 21:00 | 2014-11-21 | INO Therapeutics (dba Ikaria) | An anomaly impacts the monitoring system of INOmax DSIR Plus (software 3.0.0). When the integrated pneumatic backup switch is activated with low or no gas flow through the injector module, an intended NO cell failure alarm will occur. In addition, an unintended NO2 cell failure alarm may also occur. The NO/NO2 cell failure alarms are caused by an average NO/NO2 circuit concentration greater tha | Class 2 Device Recall INOmax DSIR Delivery System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall INOmax DSIR Delivery System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall INOmax DSIR Delivery System Date Initiated by Firm March 27, 2014 Date Posted May 02, 2014 Recall Status 1 Terminated 3 on November 21, 2014 Recall Number Z-1528-2014 Recall Event ID 67947 510(K)Number K131686 Product Classification Apparatus, nitric oxide delivery - Product Code MRN Product INOmax DSIR, model 10007, Nitric Oxide Delivery System, with Software version 3.0.0. The INOmax DS delivery system delivers INOMAX (nitric oxide for inhalation) therapy gas and provides continuous monitoring of inspired O2, NO2, and NO, and a comprehensive alarm system. The primary targeted clinical setting is the Neonatal Intensive Care Unit (NICU) and secondary targeted clinical setting is the transport of neonates. Code Information Serial Numbers: DS20080152, DS20090218, DS20090243, DS20090271, DS20090523, DS20090598, DS20090821, DS20090925, DS20101022, DS20101076, DS20101147, DS20101374, DS20101413, DS20101420, DS20101499, DS20101637, DS20111199, DS20120099, DS20120150, DS20120169, DS20120288, DS20120306. Recalling Firm/ Manufacturer INO Therapeutics (dba Ikaria) 2902 Dairy Dr Madison WI 53718-3809 For Additional Information Contact Customer Care 877-566-9466 Manufacturer Reason for Recall An anomaly impacts the monitoring system of INOmax DSIR Plus (software 3.0.0). When the integrated pneumatic backup switch is activated with low or no gas flow through the injector module, an intended NO cell failure alarm will occur. In addition, an unintended NO2 cell failure alarm may also occur. The NO/NO2 cell failure alarms are caused by an average NO/NO2 circuit concentration greater tha FDA Determined Cause 2 Software Design Change Action Ikaria sent an Urgent Field Corrective Action letter via FedEx on March 27, 2014. The letter identified the issue, and requested that all INOmax DSIR with software version 3.0.0 be returned to Ikaria. Customers with questions were instructed to contact Ikaria Customer Care at 1-877-566-9466.******For questions regarding this recall call 877-566-9466. Quantity in Commerce 22 Distribution Nationwide Distribution including DC and VA. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MRN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Ikaria sent an Urgent Field Corrective Action letter via FedEx on March 27, 2014. The letter identified the issue, and requested that all INOmax DSIR with software version 3.0.0 be returned to Ikaria. Customers with questions were instructed to contact Ikaria Customer Care at 1-877-566-9466.******For questions regarding this recall call 877-566-9466. | SoftwareDesignChange
175 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126814 | Z-1504-2014 | Siemens ACUSON X700 Ultrasound Systems with software version 1.0.00, 1.0.01, 1.0.02, 1.1.00 and 1.1.02. ultrasound imaging system. | ACUSON X700 Ultrasound Systems | 2 | CDRH | 2014-04-24 | 2014-04-24 21:00 | 2015-11-12 | Siemens Medical Solutions USA, Inc. | When fourSight 4D ultrasound imaging, 3-Scrape real time 3D imaging, or Anatomical M-mode are used during an Obstetric or Gynecology study, data from the previous patient could replace the current patient's date in the report. | Class 2 Device Recall ACUSON X700 Ultrasound Systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ACUSON X700 Ultrasound Systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ACUSON X700 Ultrasound Systems Date Initiated by Firm April 02, 2014 Date Posted April 24, 2014 Recall Status 1 Terminated 3 on November 12, 2015 Recall Number Z-1504-2014 Recall Event ID 68027 510(K)Number K123001 Product Classification System, imaging, pulsed doppler, ultrasonic - Product Code IYN Product Siemens ACUSON X700 Ultrasound Systems with software version 1.0.00, 1.0.01, 1.0.02, 1.1.00 and 1.1.02. ultrasound imaging system. Code Information Model numbers: 1065844 *** ACUSON X700 Ultrasound System 10658846 - ACUSON X700 Ultrasound System (Russia) 10658845 - ACUSON X700 Ultrasound System (Korea) Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc. 685 E Middlefield Rd Mountain View CA 94043-4045 For Additional Information Contact Sheila Pickering 650-694-5398 Manufacturer Reason for Recall When fourSight 4D ultrasound imaging, 3-Scrape real time 3D imaging, or Anatomical M-mode are used during an Obstetric or Gynecology study, data from the previous patient could replace the current patient's date in the report. FDA Determined Cause 2 Software Design Change Action Siemens sent a Customer Safety Advisory Notification letter on April 2, 2014 to all affected customers via certified mail. The letter identified the affected product, problem and what actions to take until software is upgraded. For questions call 1-888-826-9702. Quantity in Commerce 490 Distribution Worldwide Distribution - US Nationwide and the countries of Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia Herzeg., Brunei, Chile, Colombia, Croatia, Ecuador, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kenya, Kuwait, Latvia, Libya, Liechtenstein, Malaysia, Mauritius, Montenegro, Namibia, Nepal, Nigeria, Norway, Palestinian Ter, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Saudi Arabia, Slovakia, South Africa, Spain, Switzerland, Taiwan, Turkey, Turkmenistan, U.A.E., United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYN - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens sent a Customer Safety Advisory Notification letter on April 2, 2014 to all affected customers via certified mail. The letter identified the affected product, problem and what actions to take until software is upgraded. For questions call 1-888-826-9702. | SoftwareDesignChange
176 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125643 | Z-1308-2014 | RaySearch Ray Station, Radiation Therapy Treatment Planning System, Stand-alone software | RayStation | 2 | CDRH | 2014-03-31 | 2014-03-31 21:00 | 2017-02-17 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories has recalled RaySearch Ray Station 4.0 software due to clinical dose calculation errors during radiation therapy. | Class 2 Device Recall RayStation Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall RayStation FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall RayStation Date Initiated by Firm February 06, 2014 Date Posted March 31, 2014 Recall Status 1 Terminated 3 on February 17, 2017 Recall Number Z-1308-2014 Recall Event ID 67511 510(K)Number K130617 Product Classification Radiation Therapy Treatment Planning System - Product Code MUJ Product RaySearch Ray Station, Radiation Therapy Treatment Planning System, Stand-alone software Code Information Model#: 4.30; Version: 4.0.2.9; 3D-CRT optimization or Fallback Planning Licenses Recalling Firm/ Manufacturer RAYSEARCH LABORATORIES AB Sveavaegen 9 Stockholm Sweden For Additional Information Contact Freddie Cardel 877-778-3849 Manufacturer Reason for Recall RaySearch Laboratories has recalled RaySearch Ray Station 4.0 software due to clinical dose calculation errors during radiation therapy. FDA Determined Cause 2 Software Design Change Action On February 6, 2014, RaySearch Laboratories distributed Urgent Medical Device Correction notices to their customers via email regarding the affected product. To determine if the version you are using is affected, bring up the About box in the RayStation application and check if the number reported is "4.0.2.9". If so, the notice applies to your version. Customers are advised to inform all planning staff and all users about the instructions listed on the notice. Customers are advised to inspect their system and identify all installed units with the referenced software version numbers listed on the notice. Corrective labeling is distributed to the customers in the form of the safety notification attached to this report and the software problem will be resolved in RayStation patch 4.0.3 scheduled for market release end of February 2014 and in all future releases.****** For further information and return of the reply form, please contact: Freddie Cardel, Director of Customer Support, RaySearch Americas, (877) 778-3849 or freddie.cardel@raysearchlabs.com. Quantity in Commerce 19 units (16 units domestically & 3 units internationally) Distribution OH, MD, NJ, WA, IL, MA, FL, Spain, United Kingdom, Italy, Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MUJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action On February 6, 2014, RaySearch Laboratories distributed Urgent Medical Device Correction notices to their customers via email regarding the affected product. To determine if the version you are using is affected, bring up the About box in the RayStation application and check if the number reported is "4.0.2.9". If so, the notice applies to your version. Customers are advised to inform all planning staff and all users about the instructions listed on the notice. Customers are advised to inspect their system and identify all installed units with the referenced software version numbers listed on the notice. Corrective labeling is distributed to the customers in the form of the safety notification attached to this report and the software problem will be resolved in RayStation patch 4.0.3 scheduled for market release end of February 2014 and in all future releases.****** For further information and return of the reply form, please contact: Freddie Cardel, Director of Customer Support, RaySearch Americas, (877) 778-3849 or freddie.cardel@raysearchlabs.com. | SoftwareDesignChange
177 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125602 | Z-1307-2014 | RayStation Radiation Therapy Treatment Planning System, models 2.5, 3.0, 3.5 and 4.0; a software system designed for treatment planning and analysis of radiation therapy. | RayStation | 2 | CDRH | 2014-03-28 | 2014-03-28 21:00 | 2017-02-17 | RAYSEARCH LABORATORIES AB | Potential for dose errors due to software program errors. | Class 2 Device Recall RayStation Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall RayStation FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall RayStation Date Initiated by Firm February 06, 2014 Date Posted March 28, 2014 Recall Status 1 Terminated 3 on February 17, 2017 Recall Number Z-1307-2014 Recall Event ID 67504 510(K)Number K100552 K120387 K130617 Product Classification System,planning,radiation therapy treatment - Product Code MUJ Product RayStation Radiation Therapy Treatment Planning System, models 2.5, 3.0, 3.5 and 4.0; a software system designed for treatment planning and analysis of radiation therapy. Code Information MODEL #(s): 2.5, 3.0, 3.5 and 4.0; VERSIONS: 2.5.1.89, 3.0.0.251, 3.5.0.16, 3.5.1.6, 4.0.0.14, 4.0.1.4, 4.0.2.9 Recalling Firm/ Manufacturer RAYSEARCH LABORATORIES AB Sveavaegen 9 Stockholm Sweden For Additional Information Contact Freddie Cardel 877-778-3849 Manufacturer Reason for Recall Potential for dose errors due to software program errors. FDA Determined Cause 2 Software Design Change Action RaySearch Laboratories sent an Field Safety Notice, Medical Device Correction letter dated February 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.************ACTIONS TO BE TAKEN BY THE USER******A workaround that eliminates the described error is to select the ***discard rotations*** option in the Image Registration Module before creating a rigid registration intended to be used for hybrid deformable registration. Another option is to use the ROI based deformable registration when applicable, as this functionality is not affected by the error.************SOLUTION******This issue will be resolved in a patch release of RayStation, scheduled for market release February 28, 2014. Release schedule is subject to regulatory approval for some markets. In the meantime, this field safety notice is being distributed to all customers.************CONTACT INFORMATION******For further information, please contact******For the Americas:******Customer Support, RaySearch Americas, at +1 877 778 3849 ************For Europe, Asia and rest of the world:******RaySearch Laboratories AB, at +46 8 5450 6130 ************Thank you for your cooperation, and we apologize for any inconvenience Quantity in Commerce 151 units Total (88 units domestically & 63 units internationally) Distribution Worldwide Distribution - US including the states of NY, OH, MD, CT, CO, TN, IL, AZ, MS, NJ, PA, DE, FL, CA and WA., and the countries of New Zealand, Australia, Republic of Korea, China, Italy, Germany, Norway, Belgium, Switzerland, United Kingdom and Spain. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MUJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action RaySearch Laboratories sent an Field Safety Notice, Medical Device Correction letter dated February 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.************ACTIONS TO BE TAKEN BY THE USER******A workaround that eliminates the described error is to select the ***discard rotations*** option in the Image Registration Module before creating a rigid registration intended to be used for hybrid deformable registration. Another option is to use the ROI based deformable registration when applicable, as this functionality is not affected by the error.************SOLUTION******This issue will be resolved in a patch release of RayStation, scheduled for market release February 28, 2014. Release schedule is subject to regulatory approval for some markets. In the meantime, this field safety notice is being distributed to all customers.************CONTACT INFORMATION******For further information, please contact******For the Americas:******Customer Support, RaySearch Americas, at +1 877 778 3849 ************For Europe, Asia and rest of the world:******RaySearch Laboratories AB, at +46 8 5450 6130 ************Thank you for your cooperation, and we apologize for any inconvenience | SoftwareDesignChange
178 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125492 | Z-1194-2014 | Philips HeartStart MRx Monitor/Defibrillator***Models: M3535A and M3536A with Q-CPR Meter Option B08******The HeartStart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced cardiac support, or defibrillation. It must be used by or on the order of a physician | Phillips HeartStart MRx Monitor/Defribillator | 2 | CDRH | 2014-03-12 | 2014-03-12 21:00 | 2020-09-17 13:51 | Philips Medical Systems, Inc. | When the HeartStart MRx.is used with the Q-CPR Meter in defibrillation mode, the Q-CPR Meter may incorrectly display the Do Not Touch the Patient icon. | Class 2 Device Recall Phillips HeartStart MRx Monitor/Defribillator Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Phillips HeartStart MRx Monitor/Defribillator FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Phillips HeartStart MRx Monitor/Defribillator Date Initiated by Firm February 18, 2014 Date Posted March 12, 2014 Recall Status 1 Terminated 3 on September 17, 2020 Recall Number Z-1194-2014 Recall Event ID 67459 510(K)Number K051134 Product Classification Automated external defibrillators (non-wearable) - Product Code MKJ Product Philips HeartStart MRx Monitor/Defibrillator******Models: M3535A and M3536A with Q-CPR Meter Option B08************The HeartStart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced cardiac support, or defibrillation. It must be used by or on the order of a physician Code Information MRx units with the Q-CPR Meter option B08 and with a serial number within the range: US00100253 and US00571587 Recalling Firm/ Manufacturer Philips Medical Systems, Inc. 3000 Minuteman Rd Andover MA 01810-1032 For Additional Information Contact 978-687-1501 Manufacturer Reason for Recall When the HeartStart MRx.is used with the Q-CPR Meter in defibrillation mode, the Q-CPR Meter may incorrectly display the Do Not Touch the Patient icon. FDA Determined Cause 2 Software Design Change Action The firm, Philips Healthcare, sent an "URGENT -Medical Device Correction/Field Safety Notice" letter dated February 2014 on February 18, 2014 to its domestic and foreign customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the Action to be Taken by Customer/User section of the attached Field Safety Notice; to continue to use your HeartStart MRx without the Q-CPR Meter attached. If you choose to use the HeartStart MRx with the Q-CPR Meter attached and you observe that the Q CPR Meter is displaying the Do Not Touch the Patient icon at an unexpected time, follow the MRx Instructions for Use and remove the meter from the patient and continue CPR according to your organization's protocol.************Philips plans to take to remedy the issue. Philips will provide a software upgrade free of charge.************If you have questions regarding this notification or need any further information or support, please contact your local Philips representative or call at 1-800-722-9377. Quantity in Commerce 6583 units Distribution Worldwide distribution: US (nationwide) and to countries of: ARUBA AUSTRALIA AUSTRIA BELGIUM BRAZIL CANADA ECUADOR EGYPT FINLAND FRANCE GERMANY HONG KONG INDIA INDONESIA IRELAND ISRAEL ITALY KENYA KOREA, REPUBLIC OF KUWAIT LATVIA LIBYAN ARAB JAMAHIRIYA MALAYSIA MEXICO MOROCCO NETHERLANDS NORWAY OMAN PAKISTAN PHILIPPINES POLAND QATAR RUSSIAN FEDERATION SAUDI ARABIA SINGAPORE SOUTH AFRICA SPAIN SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA THAILAND TURKEY UKRAINE UNITED ARAB EMIRATES and UNITED KINGDOM. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MKJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, Philips Healthcare, sent an "URGENT -Medical Device Correction/Field Safety Notice" letter dated February 2014 on February 18, 2014 to its domestic and foreign customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the Action to be Taken by Customer/User section of the attached Field Safety Notice; to continue to use your HeartStart MRx without the Q-CPR Meter attached. If you choose to use the HeartStart MRx with the Q-CPR Meter attached and you observe that the Q CPR Meter is displaying the Do Not Touch the Patient icon at an unexpected time, follow the MRx Instructions for Use and remove the meter from the patient and continue CPR according to your organization's protocol.************Philips plans to take to remedy the issue. Philips will provide a software upgrade free of charge.************If you have questions regarding this notification or need any further information or support, please contact your local Philips representative or call at 1-800-722-9377. | SoftwareDesignChange
179 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124662 | Z-0780-2014 | XiO RTP System******The XiO RTP System is used to create treatment plans for any cancer patient for whom external beam radiation therapy or brachytherapy has been prescribed. The system will calculate and display, both on-screen and in hard-copy, either two or three-dimensional radiation dose distributions within a patient for a given treatment plan set-up. | XiO RTP System | 2 | CDRH | 2014-01-17 | 2014-01-17 21:00 | 2014-05-13 | Elekta, Inc. | When the number of fractions is changed for a proton spot beam, the MU are recomputed and rounded if necessary, but the dose is not recalculated as it should be even after saving and re-opening the plan. | Class 2 Device Recall XiO RTP System Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall XiO RTP System FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall XiO RTP System Date Initiated by Firm November 25, 2013 Date Posted January 17, 2014 Recall Status 1 Terminated 3 on May 13, 2014 Recall Number Z-0780-2014 Recall Event ID 67078 510(K)Number K102216 Product Classification System,planning,radiation therapy treatment - Product Code MUJ Product XiO RTP System************The XiO RTP System is used to create treatment plans for any cancer patient for whom external beam radiation therapy or brachytherapy has been prescribed. The system will calculate and display, both on-screen and in hard-copy, either two or three-dimensional radiation dose distributions within a patient for a given treatment plan set-up. Code Information Model or Catalog No. MUJ; Software Version XiO Versions 4.80 and higher Recalling Firm/ Manufacturer Elekta, Inc. 400 Perimeter Center Ter Ne Ste 50 Atlanta GA 30346-1227 For Additional Information Contact Linda Wetsel 404-993-5886 Manufacturer Reason for Recall When the number of fractions is changed for a proton spot beam, the MU are recomputed and rounded if necessary, but the dose is not recalculated as it should be even after saving and re-opening the plan. FDA Determined Cause 2 Software Design Change Action An Important Safety Notice, LINXIO0017 Proton Sport Beams, Changes in Fractionation not Updating Dose, is in the process of being distributed to affected customers. The release date for this notice is November 25, 2013. The notice included a Confirmation of Receipt which was to be returned to Elekta.************For further questions please call (404) 993-5886. Quantity in Commerce 13 Distribution Worldwide Distribution - US Distribution including the states of IL, IN, LA, NJ, TX, WA and WI., and the countries of Belgium, Germany, Italy and Netherlands. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MUJ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action An Important Safety Notice, LINXIO0017 Proton Sport Beams, Changes in Fractionation not Updating Dose, is in the process of being distributed to affected customers. The release date for this notice is November 25, 2013. The notice included a Confirmation of Receipt which was to be returned to Elekta.************For further questions please call (404) 993-5886. | SoftwareDesignChange
180 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124446 | Z-0524-2014 | InstaTrak¿ 3500 Plus, ENTrakTM Plus, InstaTrak¿ 3500, ENTrakTM Computer-Assisted Navigation Systems (USA), | InstaTrak¿ 3500 Plus, ENTrakTM Plus, InstaTrak¿ 3500, ENTrakTM | 2 | CDRH | 2013-12-17 | 2013-12-17 21:00 | 2017-04-17 | GE OEC Medical Systems, Inc | GE OEC Medical Systems Inc is recalling InstaTrak¿ 3500 Plus, ENTrakTM Plus, InstaTrak¿ 3500, ENTrakTM Computer-Assisted Navigation Systems (USA) due to several potential safety issues related to the use of these products. | Class 2 Device Recall InstaTrak 3500 Plus, ENTrakTM Plus, InstaTrak 3500, ENTrakTM Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall InstaTrak 3500 Plus, ENTrakTM Plus, InstaTrak 3500, ENTrakTM FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall InstaTrak 3500 Plus, ENTrakTM Plus, InstaTrak 3500, ENTrakTM Date Initiated by Firm November 26, 2013 Date Posted December 17, 2013 Recall Status 1 Terminated 3 on April 17, 2017 Recall Number Z-0524-2014 Recall Event ID 66992 510(K)Number K960330 Product Classification System, image processing, radiological - Product Code LLZ Product InstaTrak 3500 Plus, ENTrakTM Plus, InstaTrak 3500, ENTrakTM Computer-Assisted Navigation Systems (USA), Code Information Model Numbers IT2500,IT2500 Plus,IT3500, IT3500 Plus Recalling Firm/ Manufacturer GE OEC Medical Systems, Inc 384 Wright Brothers Dr Salt Lake City UT 84116-2862 For Additional Information Contact Deena Pease 801-536-4952 Manufacturer Reason for Recall GE OEC Medical Systems Inc is recalling InstaTrak 3500 Plus, ENTrakTM Plus, InstaTrak 3500, ENTrakTM Computer-Assisted Navigation Systems (USA) due to several potential safety issues related to the use of these products. FDA Determined Cause 2 Software Design Change Action GE Healthcare sent an An Urgent Medical Device Correction letter dated November 26, 2013, to all affected customers. The letter identified the product, the problem, and provided instructions for immediate mitigation of the issue. No product is being returned. The affected systems in the field will be corrected.******Customers with questions were instructed to contact the service team at ******800-874-7378.******For questions regarding this recall call 801-536-4952. Quantity in Commerce 654 Distribution Nationwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action GE Healthcare sent an An Urgent Medical Device Correction letter dated November 26, 2013, to all affected customers. The letter identified the product, the problem, and provided instructions for immediate mitigation of the issue. No product is being returned. The affected systems in the field will be corrected.******Customers with questions were instructed to contact the service team at ******800-874-7378.******For questions regarding this recall call 801-536-4952. | SoftwareDesignChange
181 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121816 | Z-0168-2014 | COBAS INTEGRA 800 Analyzer with software 9864.C2***Cobas. COBAS Part Numbers 28122474001, 28122474692, 04559126001, 04589459970 corresponding to the Analyzer with or without closed tube sampling accessory ******A fully automated system for clinical chemistry analysis intended for the in vitro quantitative/qualitative determination of analytes in body fluids. | COBAS INTEGRA 800 Analyzer | 2 | CDRH | 2013-11-05 | 2013-11-05 21:00 | 2015-05-12 | Roche Diagnostics Operations, Inc. | Roche COBAS INTEGRA 800 system, Software Version 9864.C2 does not perform necessary ISE Service Actions. Whenever the ISE rack is removed, the ISE counters in the software are reset to zero and the corresponding Service Actions are not conducted automatically by the system. Additionally, certain maintenance activities that have to be performed manually by the operator are not requested on the sof | Class 2 Device Recall COBAS INTEGRA 800 Analyzer Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall COBAS INTEGRA 800 Analyzer FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall COBAS INTEGRA 800 Analyzer Date Initiated by Firm September 16, 2013 Date Posted November 05, 2013 Recall Status 1 Terminated 3 on May 12, 2015 Recall Number Z-0168-2014 Recall Event ID 66239 510(K)Number K951595 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product COBAS INTEGRA 800 Analyzer with software 9864.C2******Cobas. COBAS Part Numbers 28122474001, 28122474692, 04559126001, 04589459970 corresponding to the Analyzer with or without closed tube sampling accessory ************A fully automated system for clinical chemistry analysis intended for the in vitro quantitative/qualitative determination of analytes in body fluids. Code Information software version 9864.C2 Recalling Firm/ Manufacturer Roche Diagnostics Operations, Inc. 9115 Hague Rd Indianapolis IN 46256-1025 For Additional Information Contact Todd Siesky 317-576-3911 Manufacturer Reason for Recall Roche COBAS INTEGRA 800 system, Software Version 9864.C2 does not perform necessary ISE Service Actions. Whenever the ISE rack is removed, the ISE counters in the software are reset to zero and the corresponding Service Actions are not conducted automatically by the system. Additionally, certain maintenance activities that have to be performed manually by the operator are not requested on the sof FDA Determined Cause 2 Software Design Change Action Roche sent an Urgent Medical Device Correction letter dated September 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Roche Diagnostics recommends tracking all ISE counters and manually requesting the corresponding Service Actions listed in the notification be monitored until updated software becomes available during Q4 2013. ******If the affected product was moved to another site, the notification is to be forwarded. consignees were asked to complete the attached faxback form and fax it to 1-877-266-0997 and file the Urgent Medical Device Correction (UMDC) for future reference. Question should be directed to******Roche Support Network Customer Support Center 24 hours a day, seven days a week at 1-800-428-2336******For questions regarding thisr ecall call 317-576-3911. Quantity in Commerce 527 units Distribution Nationwide Distribution and Puerto Rico Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | FDA Determined Cause 2 Software Design Change Action Roche sent an Urgent Medical Device Correction letter dated September 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Roche Diagnostics recommends tracking all ISE counters and manually requesting the corresponding Service Actions listed in the notification be monitored until updated software becomes available during Q4 2013. ******If the affected product was moved to another site, the notification is to be forwarded. consignees were asked to complete the attached faxback form and fax it to 1-877-266-0997 and file the Urgent Medical Device Correction (UMDC) for future reference. Question should be directed to******Roche Support Network Customer Support Center 24 hours a day, seven days a week at 1-800-428-2336******For questions regarding thisr ecall call 317-576-3911. | SoftwareDesignChange
182 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120684 | Z-2098-2013 | GE Ultrasound ViewPoint 6. ******Product Usage: ViewPoint is intended to be used in medical practices and in clinical departments and serves the purposes of diagnostic interpretation of images, electronic documentation of examinations in the form of text and images and generation of medical reports primarily for diagnostic ultrasound. ViewPoint provides the user the ability to including images, drawings, and charts into medical reports. ViewPoint is designed to accept, transfer, display, calculate, store and process medical images and data, and enables the user to measure and annotate the images. The medical images, which ViewPoint displays to the user, can be used for diagnostic purposes. ViewPoint is intended for professional use only. ViewPoint is not intended to be used as an automated diagnosis system. | GE Ultrasound ViewPoint 6 | 2 | CDRH | 2013-08-27 | 2013-08-27 21:00 | 2014-04-25 | GE Healthcare, LLC | GE has recently become aware of a potential safety issue due to the incorrect calculation of the pulsatility index value (PIV) for venous flow measurements associated with obstetric reporting of the GE ViewPoint 6 system. The ViewPoint calculated Doppler PIV for veins could be incorrect. An incorrect calculation of the ductus venosus PIV could lead to false high risk for trisomy 13, 18, and 21, | Class 2 Device Recall GE Ultrasound ViewPoint 6 Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall GE Ultrasound ViewPoint 6 FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall GE Ultrasound ViewPoint 6 Date Initiated by Firm August 07, 2013 Date Posted August 27, 2013 Recall Status 1 Terminated 3 on April 25, 2014 Recall Number Z-2098-2013 Recall Event ID 65880 510(K)Number K103458 Product Classification System, image processing, radiological - Product Code LLZ Product GE Ultrasound ViewPoint 6. ************Product Usage: ViewPoint is intended to be used in medical practices and in clinical departments and serves the purposes of diagnostic interpretation of images, electronic documentation of examinations in the form of text and images and generation of medical reports primarily for diagnostic ultrasound. ViewPoint provides the user the ability to including images, drawings, and charts into medical reports. ViewPoint is designed to accept, transfer, display, calculate, store and process medical images and data, and enables the user to measure and annotate the images. The medical images, which ViewPoint displays to the user, can be used for diagnostic purposes. ViewPoint is intended for professional use only. ViewPoint is not intended to be used as an automated diagnosis system. Code Information software versions 6.2, 6.2.1, 6.3, and 6.3.1 Recalling Firm/ Manufacturer GE Healthcare, LLC 3000 N Grandview Blvd Waukesha WI 53188-1615 For Additional Information Contact 877-644-3114 Manufacturer Reason for Recall GE has recently become aware of a potential safety issue due to the incorrect calculation of the pulsatility index value (PIV) for venous flow measurements associated with obstetric reporting of the GE ViewPoint 6 system. The ViewPoint calculated Doppler PIV for veins could be incorrect. An incorrect calculation of the ductus venosus PIV could lead to false high risk for trisomy 13, 18, and 21, FDA Determined Cause 2 Software Design Change Action GE Healthcare sent An Urgent Medical Device Correction letter dated August 7, 2013 to all affected customers. The letter described the affected product, problem and instructed customers to discontinue use of venous PIV for making any and all clinical decisions. The letter informed customers that a GE service engineer will install a correction at no cost. For questions contact your local GE Healthcare Representative or call 877-644-3114 (US) and 49 (0)8153-9311-11(Europe/Asia) Quantity in Commerce 225 Distribution Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of: ANDORRA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BRAZIL, BULGARIA, CANADA, CHILE, COLOMBIA, CZECH REPUBLIC, ESTONIA, FINLAND, GERMANY, HONG KONG, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, KUWAIT, MALAYSIA, Norway, OMAN, PALESTINIAN TERRITORY, OCCUPIED, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TURKEY, and UNITED KINGDOM. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = LLZ - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action GE Healthcare sent An Urgent Medical Device Correction letter dated August 7, 2013 to all affected customers. The letter described the affected product, problem and instructed customers to discontinue use of venous PIV for making any and all clinical decisions. The letter informed customers that a GE service engineer will install a correction at no cost. For questions contact your local GE Healthcare Representative or call 877-644-3114 (US) and 49 (0)8153-9311-11(Europe/Asia) | SoftwareDesignChange
183 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=119700 | Z-2088-2013 | Dimension Vista 500 and Dimension Vista 1500 System Aliquot Well Double Dispense with software versions 3.5.1 and 3.6. ******In vitro diagnostic device intended to duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids. Dimension Vista chemical and immunochemical applications use photometric, turbidimetric, chemiluminescence, nephelometric, and integrated ion-selective multisensory technology for clinical use. | Dimension Vista Intelligent Lab Systems | 2 | CDRH | 2013-08-26 | 2013-08-26 21:00 | 2014-08-27 | Siemens Healthcare Diagnostics, Inc. | Siemens is conducting a field correction for Dimension Vista 500 and Dimension Vista 1500 systems using software versions 3.5.1 and 3.6. Under specific rare conditions which require a system reset, there is a potential for a sample to be dispensed into an aliquot well that has previously been used for a QC or calibrator product from a Vista vial. If two samples are dispensed into the same aliquot | Class 2 Device Recall Dimension Vista Intelligent Lab Systems Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Dimension Vista Intelligent Lab Systems FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Dimension Vista Intelligent Lab Systems Date Initiated by Firm June 19, 2013 Date Posted August 26, 2013 Recall Status 1 Terminated 3 on August 27, 2014 Recall Number Z-2088-2013 Recall Event ID 65607 510(K)Number K051087 Product Classification Analyzer, chemistry (photometric, discrete), for clinical use - Product Code JJE Product Dimension Vista 500 and Dimension Vista 1500 System Aliquot Well Double Dispense with software versions 3.5.1 and 3.6. ************In vitro diagnostic device intended to duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids. Dimension Vista chemical and immunochemical applications use photometric, turbidimetric, chemiluminescence, nephelometric, and integrated ion-selective multisensory technology for clinical use. Code Information Dimension Vista 500 Siemens material numbers: 10284473 (domestic), and 10488224 (international), Dimension Vista 1500 Siemens material numbers: 10444801 (domestic) and 10444802 (international). Recalling Firm/ Manufacturer Siemens Healthcare Diagnostics, Inc. 500 Gbc Dr Ms 514 PO BOX 6101 Newark DE 19702-2466 For Additional Information Contact Siemens Customer Service 800-441-9250 Manufacturer Reason for Recall Siemens is conducting a field correction for Dimension Vista 500 and Dimension Vista 1500 systems using software versions 3.5.1 and 3.6. Under specific rare conditions which require a system reset, there is a potential for a sample to be dispensed into an aliquot well that has previously been used for a QC or calibrator product from a Vista vial. If two samples are dispensed into the same aliquot FDA Determined Cause 2 Software Design Change Action The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated June 2013 to its consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to take the following actions: ************1) In the event that an instrument Reset is required for any reason, review the Current Activities Screen to determine if there are any pending QC or Calibration tests.******2) If there are no pending QC of Calibration tests, Rest the instrument or troubleshoot the instrument per the Operators Guide.******3) If there are pending QC or Calibration tests, restarting the software is necessary before resuming patient processing. Go to Operation -> Shutdown, confirm Yes, then on the blue console menu screen select start Vista. ************Consignees/customers were also asked to complete the Field Effectiveness Check form attached to the letter via fax to Siemens Healthcare Diagnostics at (302) 621-8467, and retain the letter with laboratory records, and forward the letter to those who may have received the product.************If you have any questions, contact your Siemens Customer Care Center at (800-441-9250)-Technical Solutions or your local Siemens Technical Support representative. Quantity in Commerce 1937 Distribution Worldwide distribution: US (nationwide) including PR and countries of: Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Malaysia, Netherlands, New Zealand, Norway, Portugal, Republic Korea, Saudia Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland, and United Kingdom. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = JJE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated June 2013 to its consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to take the following actions: ************1) In the event that an instrument Reset is required for any reason, review the Current Activities Screen to determine if there are any pending QC or Calibration tests.******2) If there are no pending QC of Calibration tests, Rest the instrument or troubleshoot the instrument per the Operators Guide.******3) If there are pending QC or Calibration tests, restarting the software is necessary before resuming patient processing. Go to Operation -> Shutdown, confirm Yes, then on the blue console menu screen select start Vista. ************Consignees/customers were also asked to complete the Field Effectiveness Check form attached to the letter via fax to Siemens Healthcare Diagnostics at (302) 621-8467, and retain the letter with laboratory records, and forward the letter to those who may have received the product.************If you have any questions, contact your Siemens Customer Care Center at (800-441-9250)-Technical Solutions or your local Siemens Technical Support representative. | SoftwareDesignChange
184 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120158 | Z-1931-2013 | MHI-TM2000 Linear Accelerator System, Software Version 3.0.0 revision 13 or earlier; Proprietary and Trade Name: MHI-TM2000/ VERO (TM); Common Name: Medical Linear Accelerator. ******MHI-TM2000 Linear Accelerator System is intended for radiation therapy of lesions, tumors and conditions anywhere in the body where radiation treatment is indicated | MHI-TM2000/ VERO (TM) | 2 | CDRH | 2013-08-08 | 2013-08-08 21:00 | 2017-03-23 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Software anomaly: Mitsubishi Heavy Industries, Ltd. became aware of a potential safety issue associated with the VERO (TM) Linear Accelerator System. The VERO (TM) may save treatment delivery record into wrong location under certain conditions. | Class 2 Device Recall MHITM2000/ VERO (TM) Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall MHITM2000/ VERO (TM) FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall MHITM2000/ VERO (TM) Date Initiated by Firm June 21, 2013 Date Posted August 08, 2013 Recall Status 1 Terminated 3 on March 23, 2017 Recall Number Z-1931-2013 Recall Event ID 65786 510(K)Number K072047 K122450 Product Classification Accelerator, linear, medical - Product Code IYE Product MHI-TM2000 Linear Accelerator System, Software Version 3.0.0 revision 13 or earlier; Proprietary and Trade Name: MHI-TM2000/ VERO (TM); Common Name: Medical Linear Accelerator. ************MHI-TM2000 Linear Accelerator System is intended for radiation therapy of lesions, tumors and conditions anywhere in the body where radiation treatment is indicated Code Information Serial Numbers #s 201902 and 203901 Recalling Firm/ Manufacturer MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK 6-22, 4-CHOME, KAN-ON-SHIN-MACHI NISHI-KU HIROSHIMA Japan For Additional Information Contact Katsuhisa Toyama 212-3976128 Manufacturer Reason for Recall Software anomaly: Mitsubishi Heavy Industries, Ltd. became aware of a potential safety issue associated with the VERO (TM) Linear Accelerator System. The VERO (TM) may save treatment delivery record into wrong location under certain conditions. FDA Determined Cause 2 Software Design Change Action The firm, Mitsubishi Heavy Industries, Ltd. (MHI), sent an "URGENT FIELD SAFETY NOTICE" letter dated June 21, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: If any minor failure occurs, please always press [ALARM RESET] button on operator console first; and if [EXIT] button is clicked after a minor failure occurs, please click [YES] on the pop up dialog to finish treatment completely. MHI is developing a correction for this issue. Tentatively planned availability date: End of October 2013. The correction will be free of charge and a MHI Customer Service representative will contact affected customers to schedule its installation as the correction becomes available.************If you need any further information or support concerning this issue, please feel free to contact our Customer Service representative: VERO GmbH, Kapellenstrasse 12, 8566 Feldkirchen Germany, at +49-89-99-1568-44 or email: franz.gum@vero-sbrt.com. Quantity in Commerce 14 (2 in US) Distribution Worldwide distribution: US (nationwide) including states of TX and FL; and countries of: Belgium, Italy, Japan and Korea. Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IYE - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action The firm, Mitsubishi Heavy Industries, Ltd. (MHI), sent an "URGENT FIELD SAFETY NOTICE" letter dated June 21, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: If any minor failure occurs, please always press [ALARM RESET] button on operator console first; and if [EXIT] button is clicked after a minor failure occurs, please click [YES] on the pop up dialog to finish treatment completely. MHI is developing a correction for this issue. Tentatively planned availability date: End of October 2013. The correction will be free of charge and a MHI Customer Service representative will contact affected customers to schedule its installation as the correction becomes available.************If you need any further information or support concerning this issue, please feel free to contact our Customer Service representative: VERO GmbH, Kapellenstrasse 12, 8566 Feldkirchen Germany, at +49-89-99-1568-44 or email: franz.gum@vero-sbrt.com. | SoftwareDesignChange
185 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=119815 | Z-1782-2013 | ARTIS zee and zeego x-ray, angiographic system with Software Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed.******Angiographic x-ray system****** | ARTIS zee and zeego | 2 | CDRH | 2013-07-22 | 2013-07-22 21:00 | 2015-02-05 | Siemens Medical Solutions USA, Inc | Siemens issued a customer safety advisory notice and field safety corrective action for a potential issue on Artis systems with Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed. The issue affects-DSA rotational acquisition. The restriction is related to any rotational 3D-acquisition program with real-time subtraction (e.g. 3D DSA). During the fill phase of the rotational aqui | Class 2 Device Recall ARTIS zee and zeego Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall ARTIS zee and zeego FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall ARTIS zee and zeego Date Initiated by Firm June 26, 2013 Date Posted July 22, 2013 Recall Status 1 Terminated 3 on February 05, 2015 Recall Number Z-1782-2013 Recall Event ID 65651 510(K)Number K073290 K090745 Product Classification System, x-ray, angiographic - Product Code IZI Product ARTIS zee and zeego x-ray, angiographic system with Software Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed.************Angiographic x-ray system Code Information Model numbers: 10094137, 10094139, 10094141, 10280959 Recalling Firm/ Manufacturer Siemens Medical Solutions USA, Inc 51 Valley Stream Pkwy Malvern PA 19355 For Additional Information Contact Siemens Customer Service Organization 888-826-9702 Manufacturer Reason for Recall Siemens issued a customer safety advisory notice and field safety corrective action for a potential issue on Artis systems with Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed. The issue affects-DSA rotational acquisition. The restriction is related to any rotational 3D-acquisition program with real-time subtraction (e.g. 3D DSA). During the fill phase of the rotational aqui FDA Determined Cause 2 Software Design Change Action Siemens sent a Customer Safety Advisory Notice dated June 26, 2013, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. ************Customers were strongly recommend to carefully evaluate one of the following options before running 3D rotational acquisition programs that include real-time subtraction until a patch is installed on their system: ************1) position the monitor in a way such that contrast advancement can be monitored during the fill phase on the live monitor in the control room.************2) For applications not requiring actions based on the contrast flow during the fill phase, the contrast may not need to be monitored during the fill phase. Although the monitor is not showing the expected subtraction, the acquisition is working correctly. There is no need to interrupt, but continue and finish the workflow.************3) Perform 3D rotational programs that include real time subtraction from the exam room. The display on the live monitor in the exam room is not affected by this problem. ************To resolve the issue, the update AX028/13/P (VC20x/VC21B), will be made available around July 2013, and scheduling an appointment for implementation of the update with Siemens Customer Service organization was recommended. ************If this device/equipment is no longer in theirr possession, customers were asked to forward this safety notice to the new owner and inform Siemens of the new user's identity. Quantity in Commerce 29 Distribution Nationwide Distribution and Puerto Rico Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = IZI - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Siemens sent a Customer Safety Advisory Notice dated June 26, 2013, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. ************Customers were strongly recommend to carefully evaluate one of the following options before running 3D rotational acquisition programs that include real-time subtraction until a patch is installed on their system: ************1) position the monitor in a way such that contrast advancement can be monitored during the fill phase on the live monitor in the control room.************2) For applications not requiring actions based on the contrast flow during the fill phase, the contrast may not need to be monitored during the fill phase. Although the monitor is not showing the expected subtraction, the acquisition is working correctly. There is no need to interrupt, but continue and finish the workflow.************3) Perform 3D rotational programs that include real time subtraction from the exam room. The display on the live monitor in the exam room is not affected by this problem. ************To resolve the issue, the update AX028/13/P (VC20x/VC21B), will be made available around July 2013, and scheduling an appointment for implementation of the update with Siemens Customer Service organization was recommended. ************If this device/equipment is no longer in theirr possession, customers were asked to forward this safety notice to the new owner and inform Siemens of the new user's identity. | SoftwareDesignChange
186 | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=119750 | Z-1780-2013 | Radiohead Access Point 1.4 GHz Radiohead Access Point ***Firmware versions: C.00.04 & C.00.05 ***Part Number: 989803171211******Product Usage:***IntelliVue Smart-hopping 1.4 GHz Access Point provides wireless connectivity between IntelliVue devices (listed below) and the Philips IntelliVue Information Center. MX40 Indicated for use by health care professionals whenever there is a need for monitoring the physiological parameters of patients. Intended for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults and pediatrics in hospital environments and during transport inside hospitals. The MX4O is to be used primarily as a traditional telemetry medical device. It connects to the IntelliVue Information Center by way of a wireless network. | Philips Healthcare IntelliVue Smart-hopping 1.4 GHz Access Point | 2 | CDRH | 2013-07-19 | 2013-07-20 21:00 | 2017-05-17 | Philips Healthcare Inc. | IntelliVue Smart-hopping 1.4 GHz Access Point may experience unexpected signal loss of several***minutes impacting wireless connectivity | Class 2 Device Recall Philips Healthcare IntelliVue Smarthopping 1.4 GHz Access Point Quick Links: Skip to main page content Skip to Search Skip to Topics Menu Skip to Common Links U.S. Food & Drug Administration Follow FDA En Español Search FDA Home Food Drugs Medical Devices Radiation-Emitting Products Vaccines, Blood & Biologics Animal & Veterinary Cosmetics Tobacco Products Class 2 Device Recall Philips Healthcare IntelliVue Smarthopping 1.4 GHz Access Point FDA Home Medical Devices Databases - 510(k) ¦ DeNovo ¦ Registration & Listing ¦ Adverse Events ¦ Recalls ¦ PMA ¦ HDE ¦ Classification ¦ Standards CFR Title 21 ¦ Radiation-Emitting Products ¦ X-Ray Assembler ¦ Medsun Reports ¦ CLIA ¦ TPLC New Search Back to Search Results Class 2 Device Recall Philips Healthcare IntelliVue Smarthopping 1.4 GHz Access Point Date Initiated by Firm July 08, 2013 Date Posted July 20, 2013 Recall Status 1 Terminated 3 on May 17, 2017 Recall Number Z-1780-2013 Recall Event ID 65622 510(K)Number K041741 Product Classification Monitor, physiological, patient(with arrhythmia detection or alarms) - Product Code MHX Product Radiohead Access Point 1.4 GHz Radiohead Access Point ******Firmware versions: C.00.04 & C.00.05 ******Part Number: 989803171211************Product Usage:******IntelliVue Smart-hopping 1.4 GHz Access Point provides wireless connectivity between IntelliVue devices (listed below) and the Philips IntelliVue Information Center. MX40 Indicated for use by health care professionals whenever there is a need for monitoring the physiological parameters of patients. Intended for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults and pediatrics in hospital environments and during transport inside hospitals. The MX4O is to be used primarily as a traditional telemetry medical device. It connects to the IntelliVue Information Center by way of a wireless network. Code Information Serial numbers ranging from: US12700465 to US20823366 Recalling Firm/ Manufacturer Philips Healthcare Inc. 3000 Minuteman Road Andover MA 01810 For Additional Information Contact 978-687-1501 Manufacturer Reason for Recall IntelliVue Smart-hopping 1.4 GHz Access Point may experience unexpected signal loss of several******minutes impacting wireless connectivity FDA Determined Cause 2 Software Design Change Action Philips sent an Urgent Medical Device Correction letter dated July 2013 to affected customers. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer. Philips will perform a firmware upgrade (version C.00..07 for all affected customers at no cost. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice.******If you need any further information or support concerning this issue, please contact Philips Customer Support at 1-800-722-9377 Quantity in Commerce 20105 units Distribution USA Nationwide Distribution Total Product Life Cycle TPLC Device Report 1 A record in this database is created when a firm initiates a correction or removal action. The record is updated if the FDA identifies a violation and classifies the action as a recall, and it is updated for a final time when the recall is terminated. Learn more about medical device recalls . 2 Per FDA policy, recall cause determinations are subject to modification up to the point of termination of the recall. 3 For details about termination of a recall see Code of Federal Regulations (CFR) Title 21 §7.55 . 510(K) Database 510(K)s with Product Code = MHX - - Links on this page: Page Last Updated: 03/21/2025 Note: If you need help accessing information in different file formats, see Instructions for Downloading Viewers and Players . Language Assistance Available: Español ¦ 繁體中文 ¦ Tiếng Việt ¦ 한국어 ¦ Tagalog ¦ Русский ¦ العربية ¦ Kreyòl Ayisyen ¦ Français ¦ Polski ¦ Português ¦ Italiano ¦ Deutsch ¦ 日本語 ¦ فارسی ¦ English Accessibility Contact FDA Careers FDA Basics FOIA No FEAR Act Nondiscrimination Website Policies / Privacy U.S. Food and Drug Administration 10903 New Hampshire Avenue Silver Spring, MD 20993 Ph. 1-888-INFO-FDA (1-888-463-6332) Contact FDA For Government For Press Combination Products Advisory Committees Science & Research Regulatory Information Safety Emergency Preparedness International Programs News & Events Training and Continuing Education Inspections/Compliance State & Local Officials Consumers Industry Health Professionals FDA Archive Vulnerability Disclosure Policy Links on this page: | Action Philips sent an Urgent Medical Device Correction letter dated July 2013 to affected customers. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer. Philips will perform a firmware upgrade (version C.00..07 for all affected customers at no cost. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice.******If you need any further information or support concerning this issue, please contact Philips Customer Support at 1-800-722-9377 | SoftwareDesignChange