## Sheet1
WEB_ADDRESS | RECALL_NUMBER | PRODUCT_DESCRIPTION | TRADE_NAME | RECALL_CLASS | CENTER | CENTER_CLASSIFICATION_DT | POSTED_INTERNET_DT | TERMINATION_DT | FIRM_NAME | MANUFACTURER_RECALL_REASON
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=212289 | Z-1233-2025 | DxC 500 AU Clinical Chemistry Analyzer, REF: C63519, and C63520 with affected software. | DxC 500 AU Clinical Chemistry Analyzer | 2 | CDRH | 2025-02-25 17:12 | 2025-02-25 20:07 |  | BECKMAN COULTER LABORATORY SYSTEMS (SUZHOU) CO., LTD. | Beckman Coulter has identified an issue which prevents the DxC 500 AU instrument from providing the correct onboard stability (Open Expiration) dates for assays that were loaded prior to a software upgrade (e.g. 1.3 to 1.4, 1.4 to 1.4.1, etc.). As a result, assays that have expired may still be used for sample processing, leading to inaccurate or erroneous patient test results. The issue only happens on assays loaded before instruments doing upgrades and does not impact new installation. It does not impact the assays loaded after upgrade.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211506 | Z-0810-2025 | Xhibit Telemetry Receiver with Software, Model 96280. Intended to provide monitoring system with patient data. |  | 2 | CDRH | 2025-01-07 13:25 | 2025-01-07 20:07 |  | Spacelabs Healthcare, Inc. | Telemetry receivers with affected software versions may experience unsolicited system shut down and restart (Offline displayed), lasting approximately 1 to 1.5 minutes, which will cause temporary loss of monitoring. If a clinical event occurred during this period, staff may not detect until the system restarts and necessary clinical intervention could be delayed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210913 | Z-0571-2025 | Essenz HLM (Heart-Lung Machine), REF: 49-00-10, with Software | Essenz HLM | 2 | CDRH | 2024-11-27 | 2024-11-27 20:07 |  | LivaNova Deutschland GmbH | Heart-Lung Machine GUI touchscreen may go dark for 10-15 seconds, safety/performance functions remain operational before reset, control maintained through backup control panel, but second reset may occur, may lead to gas blender standby, then gas blender must be reactivated using user interface on blender, sensors deactivated until new case started; GUI reset can lead to Hypoxia and Hypoperfusion.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210328 | Z-0317-2025 | The FLUOROspeed is intended to be used as a universal diagnostic imaging system for radiographic and fluoroscopic studies. Using a flat detector, it can perform a range of applications including general R/F, angiography and pediatric examinations. FLUOROspeed may be used for emergency treatment on an outpatient basis, as well as for bedside examinations. | FluoroSpeed X1 | 2 | CDRH | 2024-11-07 | 2024-11-07 20:06 |  | SHIMADZU CORPORATION MEDICAL SYSTEMS DIVISION | X-ray R/F system software with certain versions of the control board may occasionally fail to display images when switching from Fluoroscopy mode to Radiography mode, which may lead to the possibility of delay or cancellation of the examination.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210310 | Z-0316-2025 | SONIALVISION G4 is a multi-purpose X-ray R/F system. The equipment is suitable for radiographic and fluoroscopic examinations, including general radiography and pediatric examinations, excluding mammography. | SonialVision G4 | 2 | CDRH | 2024-11-07 | 2024-11-07 20:06 |  | SHIMADZU CORPORATION MEDICAL SYSTEMS DIVISION | X-ray R/F system software with certain versions of the control board may occasionally fail to display images when switching from Fluoroscopy mode to Radiography mode, which may lead to the possibility of delay or cancellation of the examination.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209749 | Z-0009-2025 | BD FACSDuet Premium Sample Preparation System, REF: 666340, 666339; and BD FACSDuet Sample Preparation System (Base), REF: 663128, 662588, used with software V 1.4.1. Intended use: to prepare human specimens to be analyzed. |  | 2 | CDRH | 2024-10-03 08:58 | 2024-10-03 21:07 |  | Becton, Dickinson and Company, BD Bio Sciences | Sample preparation system with software may cause: 1) Multidispense feature enabled with 2-tube assay: Specimen volume dispensed into tube1 at least 3% higher than volume in tube2, or 2) Absolute counts calculated using Trucount Tubes: increased volume in tube1 may result in 20% higher absolute counts in tube 1 than tube 2; could result in additional blood collection, delayed results/patient therapy.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209671 | Z-3262-2024 | PT301US Airvo 3 Respiratory Support Device with software version 1.2.0 and/or 1.5.1 The Airvo 3 is intended to provide high flow warmed and humidified respiratory gases for administration to spontaneously breathing infant, child, adolescent and adult patients in hospitals and sub-acute facilities. | F&P Airvo 3 | 2 | CDRH | 2024-09-25 | 2024-09-25 21:07 |  | Fisher & Paykel Healthcare, Ltd. | Due to a software issue, affected devices that are set up with High Pressure Oxygen (HPO), if the flow alignment alarm occurs, the device will deliver room air only. If this happens, a patient my experience oxygen desaturation that could lead to hypoxia.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209554 | Z-3073-2024 | Medtronic Handset with Communicator, Model TH90Q01, InterStim Smart Programmer and Communicator Kits, containing the Model HH90 Handset and the TM90 Communicator. | Medtronic InterStim | 2 | CDRH | 2024-09-06 | 2024-09-06 21:30 |  | Medtronic Neuromodulation | A subset of the handsets within the kit may not be able to complete the pairing process with the communicator upon initial setup.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=208032 | Z-2374-2024 | BD Pyxis: MedStation ES (Main), REF: 323; MedStation ES Tower, REF: 352; Anesthesia Station ES, REF: 327; CII Safe ES, REF: 1116-00; CII Safe ES, Desktop PC, Wired Scanner, REF: 107-254-01; CII Safe ES, Desktop PC, Wireless Scanner, REF: 107-255-01; CII Safe ES, Desktop PC, Wired Scanner 1.7.3, REF: 107-256-01; CII Safe ES, Desktop PC, Wireless Scanner 1.7.3, REF: 107-257-01 | BD Pyxis" | 2 | CDRH | 2024-07-17 | 2024-07-17 21:07 |  | CareFusion 303, Inc. | If automated dispensing cabinets have specific software versions, and Component Manager was configured to "installed mode", and software patch KB 5033688 was installed in the production environment, then this will will lead to an error that prevents the user from accessing the dispensing software application, which may delay accessing medication.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=207225 | Z-1883-2024 | Stryker Blueprint Software, Catalog #BPUE001. | Stryker Blueprint | 2 | CDRH | 2024-05-22 | 2024-05-22 21:06 |  | Tornier S.A.S. | The software bug allows for case planning with anatomic glenoid Perform / Perform Augmented implants and anatomic humeral Tornier Flex implant configurations that are incompatible and do not have regulatory approval.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203285 | Z-0093-2024 | BD Pyxis CII Safe ES, REF: 1116-00; and BD Pyxis : CII Safe ES Desktop PC, Wired SCNR, REF: 107-254-01; CII Safe ES Desktop PC, W-LESS SCNR, BIO, REF: 107-255-01; Enterprise Server, REF: 1115-00, 134157-01, 137670-02, 134148-01, 135292-01, 136449-01, 133760-01 | BD Pyxis CII Safe ES | 2 | CDRH | 2023-10-10 | 2023-10-10 21:06 |  | CareFusion 303, Inc. | When global edit is used to update multiple formulary properties simultaneously, the following properties: 1) Require Lot Number on Recall, 2) Access Destruction Bin (Witness), 3) Add to Destruction Bin (Count/Empty), 4) CII Safe Stock Out Notice are, without user warning, overwritten to the value NO , which could result in diversion of controlled substances from the automated dispensing cabinet.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=200975 | Z-2045-2023 | wDR 2.2 Mobile Digital Diagnostic X-Ray System | Mobile Digital Diagnostic X-Ray System | 2 | CDRH | 2023-06-30 | 2023-06-30 21:07 |  | SEDECAL SA | There is a software login in issue that may prevent the user from logging in.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=200971 | Z-2042-2023 | DigitalDiagnost C90, Model No. 712034 and 712035 | DigitalDiagnost C90 | 2 | CDRH | 2023-06-29 | 2023-06-30 21:07 |  | Philips North America Llc | There is a software login in issue that may prevent the user from logging in.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=200448 | Z-1993-2023 | HeartWare Ventricular Assist Device System. a) Monitor 1521US Explore Tech, b) Monitor 1521GB Explore Tech, c) Monitor 1521IL Explore Tech, d) Monitor 1521DE Explore Tech | HeartWare" Ventricular Assist Device (HVAD") System | 2 | CDRH | 2023-06-22 | 2023-06-22 21:07 |  | Heartware, Inc. | Medtronic has identified that the logfiles downloaded from the recently updated Model 1521 monitors (Serial Numbers: MON4xxxxxx for OUS and MON5xxxxxx for US), are unable to be processed by the Autologs web portal.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197517 | Z-1207-2023 | BD Pyxis Anesthesia ES System (PAS ES), REF: 327; BP Pyxis MedStation ES (Med ES Main), REF: 323; BD Pyxis MedStation ES Tower (Med ES Tower), REF: 352; all with BD Pyxis ES version 1.6.1 drawer firmware (version 1.10.2.5). Pyxis MedStation Enterprise Server (ES) System is an automated medication dispensing system supporting decentralized medication management in a pharmaceutical environment. | BD Pyxis | 2 | CDRH | 2023-03-17 | 2023-03-17 21:07 |  | CareFusion 303, Inc. | Automated dispensing cabinet drawer firmware may have inadvertently been downgraded to an earlier version of drawer firmware during new device installations or device replacement upgrades, which may lead to an increased frequency of drawer failures, which could increase the likelihood of patients experiencing delays in receiving medications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197345 | Z-0776-2023 | Guardian iOS app (MMT-8200) and Guardian Android app (MMT-8201), part of the Guardian 4 system | Guardian App for iOS v. 1.3.0; Guardian App for Android v. 1.3.0 | 2 | CDRH | 2022-12-31 | 2023-01-01 20:06 |  | Medtronic MiniMed | An app, part of a continuous glucose monitoring system, for use with smartphone devices may automatically log out from CareLink, then the app is not able to upload data. When logged out, linked care partners will not receive SMS notifications (could result in hypoglycemia or hyperglycemia), and sensor glucose values will also not be sent to the InPen app.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194595 | Z-1403-2022 | RayStation 9B SP1. For radiation treatment planning. | RayStation | 2 | CDRH | 2022-07-19 |  | 2023-10-04 11:09 | RAYSEARCH LABORATORIES AB | An issue with propagation of treatment course information from RayStation with RayTreat to RayCare has been identified. During some workflows, information may not be propagated to RayCare.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193403 | Z-1075-2022 | Digital Medical X-ray Imaging System, Model: uDR 596i, uDR 592h | uDR 592h, uDR 596i | 2 | CDRH | 2022-05-12 |  | 2022-05-17 11:13 | Shanghai United Imaging Healthcare Co., Ltd. | X-ray imaging system positioning image and protocol label is reversed for Flexion and Extension on C-Spine and L-Spine, this will cause the image to be incorrectly labeled and may cause the patient to have a repeat exposure.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192914 | Z-0995-2022 | syngo.via RT Image Suite upgraded to VB60 running on SOMATOM go.Up Model Number 11061620. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade. | SOMATOM go.Up | 2 | CDRH | 2022-04-22 |  | 2023-12-19 10:18 | Siemens Medical Solutions USA, Inc | Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192915 | Z-0996-2022 | syngo.via RT Image Suite upgraded to VB60 running on SOMATOM go.Sim Model Number 11061660. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade. | SOMATOM go.Sim | 2 | CDRH | 2022-04-22 |  | 2023-12-19 10:18 | Siemens Medical Solutions USA, Inc | Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192916 | Z-0997-2022 | syngo.via RT Image Suite upgraded to VB60 running on SOMATOM go.Open Pro Model Number 11061670. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade. | SOMATOM go.Open Pro | 2 | CDRH | 2022-04-22 |  | 2023-12-19 10:18 | Siemens Medical Solutions USA, Inc | Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192913 | Z-0994-2022 | syngo.via RT Image Suite upgraded to VB60 running on syngo.via Model Number 10496180. Running on servers with syngo.via VB60 software after upgrade and CT scanners with syngo.CT VA40A after upgrade. | syngo.via RT Image Suite | 2 | CDRH | 2022-04-22 |  | 2023-12-19 10:18 | Siemens Medical Solutions USA, Inc | Upgraded software version makes an automatic change in laser configuration settings for "markerless workflow" which may lead to the wrong laser offset coordinates being displayed in the Patient Marking step.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=191357 | Z-0622-2022 | bellavista 1000 ventilator, Catalog #301.100.030 and bellavista 1000e ventilator, 301.100.130. The bellavista 1000/1000e ventilator is intended to provide positive pressure ventilatory support to adult and pediatric patients, and optionally, infant and neonatal patients. | bellavista | 1 | CDRH | 2022-02-16 17:27 | 2022-02-16 20:06 | 2024-08-28 09:13 | Vyaire Medical | Potential cessation in ventilation can occur under specific conditions.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189350 | Z-0165-2022 | REF/Catalogue Number GSX0020A, GORE CARDIOFORM SEPTAL OCCLUDER 20 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631018 | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189357 | Z-0172-2022 | REF/Catalogue Number GSX0030H, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132648948 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189355 | Z-0170-2022 | REF/Catalogue Number GSXE0025B, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132620111 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189354 | Z-0169-2022 | REF/Catalogue Number GSXE0030, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617661 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189352 | Z-0167-2022 | REF/Catalogue Number GSXE0020, GORE CARDIOFORM SEPTAL OCCLUDER 20 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617630 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189353 | Z-0168-2022 | REF/Catalogue Number GSXE0025, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617647 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189356 | Z-0171-2022 | REF/Catalogue Number GSXE0030B, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132620128 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189351 | Z-0166-2022 | REF/Catalogue Number GSX0025A, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631025 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189284 | Z-0164-2022 | REF/Catalogue Number GSX0030A, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631032 | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188585 | Z-2471-2021 | ACUSON Juniper Diagnostic Ultrasound System | ACUSON Juniper ultrasound imaging system | 2 | CDRH | 2021-09-14 |  | 2024-10-16 07:45 | Siemens Medical Solutions USA, Inc. | Due to intermittent failures of the power supply in the ultrasound system which renders it inoperable.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187414 | Z-1841-2021 | Medline Tissue Recovery Instrument Set, Reorder #DYNJ64347A, packaged 3/case, Sterile, Rx. Convenience kit used to recover tissue, which can be used for medical research, education, or transplantation. | Tissue Recovery Instrument Set | 2 | CDRH | 2021-06-09 |  |  | Medline Industries Inc | The devices may not meet the minimum required sterility assurance level.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=185216 | Z-0918-2021 | FUJIFILM Synapse PACS Software: versions 7.0.0 and 7.0.1 - Product Usage: intended for use as a web based application on an off-the shelf PC which meets or exceeds minimum specifications and is networked with a FUJIFILM Synapse PACS server. | Synapse PACS | 2 | CDRH | 2021-01-22 |  | 2021-11-17 15:24 | Fujifilm Medical Systems U.S.A., Inc. | The software does not update measurements and calculations in the Clinical Reporting Application (CRA) when the ventricular trace is changed in the study by a different user.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184572 | Z-0717-2021 | APEX Compounding System Control Panel Module, REF: 601224, Product Code AX1000 - Product Usage: intended for intravascular administration to a patient. | Apex Compounding System | 2 | CDRH | 2020-12-23 |  | 2023-03-14 14:38 | B Braun Medical, Inc. | There is the potential for the compounding system to not immediately interrupt compounding and alert user to the presence of an air bubble exceeding 3% of the ordered volume for an ingredient. This could result in a a compounded drug that does not meet the specified accuracy range for an individual ingredient.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184574 | Z-0719-2021 | APEX Compounding System Control Panel Module, REF: 601238, Product Code AX1000 - Product Usage: intended for intravascular administration to a patient. | Apex Compounding System | 2 | CDRH | 2020-12-23 |  | 2023-03-14 14:38 | B Braun Medical, Inc. | There is the potential for the compounding system to not immediately interrupt compounding and alert user to the presence of an air bubble exceeding 3% of the ordered volume for an ingredient. This could result in a a compounded drug that does not meet the specified accuracy range for an individual ingredient.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184573 | Z-0718-2021 | APEX Compounding System Control Panel Module, REF: 601237, Product Code AX1000 - Product Usage: intended for intravascular administration to a patient. | Apex Compounding System | 2 | CDRH | 2020-12-23 |  | 2023-03-14 14:38 | B Braun Medical, Inc. | There is the potential for the compounding system to not immediately interrupt compounding and alert user to the presence of an air bubble exceeding 3% of the ordered volume for an ingredient. This could result in a a compounded drug that does not meet the specified accuracy range for an individual ingredient.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=183827 | Z-0282-2021 | Synapse PACS Software Version 5.6.1 - Product Usage: intended for use as a web based application on an off-the shelf PC which meets or exceeds minimum specifications and is networked with a FUJIFILM Synapse PACS server. | Synapse PACS | 2 | CDRH | 2020-10-22 |  | 2021-11-17 15:26 | Fujifilm Medical Systems U.S.A., Inc. | FUJIFILM has become aware of the possibility that certain CT studies may report a much higher 3D Sphere MAX value than expected. This issue is specific to 3D Sphere - other density tools are not affected.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=181876 | Z-2551-2020 | MagNA Pure 96 Instrument, clinical sample concentrator, Material number 06541089001 - Product Usage: This protocol is intended for pipetting samples into a MagNA Pure 96 Processing Cartridge (purification rack position 1) before a purification run. | MagNA Pure 96 | 2 | CDRH | 2020-07-10 |  | 2023-10-04 09:25 | Roche Molecular Systems, Inc. | When using Sample Transfer protocol version 3.0, the drop catcher is not activated on the MagNA Pure 96 instrument, which introduces a potential risk for cross-contamination leading to erroneous results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180264 | Z-1631-2020 | Pharos Excimer Laser, Model: EX-308 - Product Usage: The Pharos EX-308 Excimer Laser System is designed to treat the skin disorders psoriasis, vitiligo, atopic dermatitis, and leukoderma using phototherapy. The excimer laser system produces collimated 308 nm radiation, within the invisible region of the electromagnetic spectrum. The UV radiation falls within the range of ultraviolet B (UVB), which is used to treat photo-responsive skin disorders. | Pharos Ex-308 Excimer Laser System | 2 | CDRH | 2020-04-01 | 2020-04-01 21:00 | 2021-05-05 15:37 | Ra Medical Systems, Inc. | At doses of less than 100 mJ, in custom and auto repeat modes, a software error in the excimer laser phototherapy system causes the device to fail to cease firing as long as the foot pedal or handpiece switch is depressed. This malfunction can cause overtreatment, which could result in strong erythema, and possibly cause blisters.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179871 | Z-1476-2020 | QIAGEN Rotor Gene Q Software version 2.3.4. Real-time PCR analyzer designed for rapid thermalcycling and real-time detection of PCR assays. | QIAGEN | 2 | CDRH | 2020-03-12 |  | 2024-09-26 07:32 | Qiagen Sciences LLC | When using the Rotor-Gene Q with Software version 2.3.4 completing a LIMs export, the .csv file reports the calculated concentration result value as a logarithmic value, and could lead to a false negative result, which could lead to serious medical consequences such as the suspension or non-initiation of treatment
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178928 | Z-1124-2020 | Prometra Programmable 20mL Pump. Catalog No. 91827 The device is an implantable infusion pump. GTIN-14 00810335020XXX (XXX = in-country language specific) | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178929 | Z-1125-2020 | Prometra II Programmable 20mL Pump. Catalog No. 93827 The device is an implantable infusion pump. GTIN-14 00810335020YYY (YYY= in-country language specific) | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178926 | Z-1122-2020 | Prometra II 40mL Pump. Catalog No. 16827 The device is an implantable infusion pump. GTIN-14 00810335020242 | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178923 | Z-1120-2020 | Prometra Programmable 20mL Pump. Catalog No. 11827 The device is an implantable infusion pump. GTIN-14 00810335020082 | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178925 | Z-1121-2020 | Prometra II 20mL Pump. Catalog No. 13827 The device is an implantable infusion pump. GTIN-14 00810335020228 | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178927 | Z-1123-2020 | US Clinical 20mL Programmable Pump. Catalog No. 01827 The device is an implantable infusion pump. GTIN-14 00810335020082 | Prometra Programmable Infusion Pump | 2 | CDRH | 2020-02-10 | 2020-02-10 20:00 |  | Flowonix Medical Inc | A pump alarm function anomaly in the pump firmware code may result in the pump not sending the expected error code notifications to the Programmer, or not emit an audible alarm tone in certain "out of tolerance" conditions associated with the pump's valve system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=176760 | Z-0354-2020 | EndoTool SubQ (Versions 1.7.1, 1.7.4, 1.7.5) | EndoTool SubQ | 2 | CDRH | 2019-11-08 |  |  | Monarch Medical Technologies | Product was distributed prior to approval or clearance from FDA.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=176305 | Z-0260-2020 | Dario Blood Glucose Monitoring System, Model No. 1015-05, running Android software versions 4.3.0-4.3.2 | Dario Blood Glucose Monitoring System | 2 | CDRH | 2019-10-31 |  | 2020-05-05 13:58 | LABSTYLE INNOVATIONS | The Dario-Blood Glucose Tracker Android App versions 4.3.0-4.3.2 may experience duplicate logging of a blood glucose level reading.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175374 | Z-0171-2020 | Plum 360 Infusion System With MedNet/ Smart Card Plug And Play Module, Model # 30010. Infusion pump. | Plum 360 Infusion System with MedNet/ Smart Card Plug n Play Module | 2 | CDRH | 2019-10-22 | 2019-10-22 21:00 | 2021-04-27 16:12 | ICU Medical Inc | A software malfunction of the infuser when used with a MedNet Custom Drug Library (CDL). When a new CDL is downloaded and installed on the infuser the infuser can become inoperable and show a blank screen on the User Interface.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172822 | Z-1786-2019 | Monarch Medical Technologies EndoTool IV, Versions 1.8, Version 1.8.5 and higher, and Version 1.9 and Version 1.10. | Monarch Medical Technologies EndoTool IV | 2 | CDRH | 2019-06-15 |  |  | Monarch Medical Technologies | Insulin dosing calculations were erroneously high.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172979 | Z-1715-2019 | Prometra II Programmable Pump, Cat. No. 13827 For intrathecal infusion. | Prometra Programmable Pump System | 2 | CDRH | 2019-06-03 |  | 2020-06-05 09:23 | Flowonix Medical Inc | The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. Error 115 can also occur if dendrites form or if the pump is dropped.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171923 | Z-1712-2019 | Prometra Clinician Programmer (Touchscreen) Software version 2.00.29, Cat. No. 12828, used with Prometra and Prometra II Programmable Pumps. For intrathecal infusion. | Prometra Programmable Infusion Pump System, Prometra Clinician Programmer | 2 | CDRH | 2019-06-03 |  | 2020-06-05 09:23 | Flowonix Medical Inc | The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172978 | Z-1714-2019 | Prometra Programmable Pump, Cat. No. 11827 For intrathecal infusion. | Prometra Programmable Pump System | 2 | CDRH | 2019-06-03 |  | 2020-06-05 09:23 | Flowonix Medical Inc | The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. Error 115 can also occur if dendrites form or if the pump is dropped.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172980 | Z-1716-2019 | Prometra II Programmable Pump (Clinical), Cat. No. 01827 For intrathecal infusion. | Prometra Programmable Pump System | 2 | CDRH | 2019-06-03 |  | 2020-06-05 09:23 | Flowonix Medical Inc | The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode. Error 115 can also occur if dendrites form or if the pump is dropped.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171924 | Z-1713-2019 | Prometra Programmer (Grand Prime) Software version 2.00.29, Cat. No. 13828, used with Prometra and Prometra II Programmable Pumps. For intrathecal infusion. | Prometra Clinician Programmer | 2 | CDRH | 2019-06-03 |  | 2020-06-05 09:23 | Flowonix Medical Inc | The pump shuts down unexpectedly if Error 115, an alarm associated with a watchdog timer timeout, is triggered. There has been an increase of Error 115 occurrences due to a software defect caused by a Healthcare Provider using a Clinician Programmer, Software Version 2.00.29, to program a Bridge or Demand bolus while the pump is set to either a Periodic Flow or Multiple Rates flow mode.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171710 | Z-1137-2019 | Biomerieux Vitek 2 Test kit VITEK 2 is an automated system consisting of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | Biomerieux Vitek 2 Test kit | 2 | CDRH | 2019-04-12 |  | 2021-08-17 17:54 | bioMerieux, Inc. | Unexpected ESBL (Extended Spectrum Beta Lactamase) phenotype had been proposed for some Escherichia coli strains in conjunction with the VITEK 2 AST-N254, AST-N330, AST-N371 (ref.413723,418674, 422024) test kits.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=169845 | Z-0821-2019 | MOSAIQ¿ Oncology Information System and Sequencer MOSAIQ is an oncology information system used to manage workflows for treatment planning and delivery. It supports information flow among healthcare facility personnel and can be used wherever radiotherapy and/or chemotherapy are prescribed. | MOSAIQ Oncology Information System and Sequencer | 2 | CDRH | 2019-02-08 |  |  | Elekta, Inc. | There is a potential that Wedge IDs were not included in the DICOM RT PLAN sent from the TPS, which will cause the field in MOSAIQ to be created with no wedge. Treatment of a field without the planned wedge will receive more dose than the treatment plan indicates.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167480 | Z-0442-2019 | 3.2mm Proximal Reamer/Cannulated Drill provided as part of the Fibulock Fibular Nail System Pack. FibuLock Nail Procedure Pack: Contains instrumentation for the implantation of the FibuLock implant. Contents include: Actuation Driver, 6.2mm Reamer, 3.2mm Reamer, 2 mm Drill, Spade Tip Guide Wire, and 1.6mm x 12 K-Wire. | 3.2mm Proximal Reamer/Cannulated Drill | 2 | CDRH | 2018-11-09 |  |  | Arthrex, Inc. | There is potential to break during use.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=165490 | Z-0353-2019 | Intelligent Peri-Analytical Workstation (HW+SW), Catalog #: A89492 (iPAW Software v4.2.1), Beckman Coulter, Inc. The iPAW ((intelligent Peri-Analytical Workstation) can act as: a software client for specimen tracking and workload statistics or a physical workstation for manual sorting and label printing. | Intelligent Peri-Analytical Workstation | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes. Sorting Drive software version 4.2 was replaced by 4.2.1 (both versions contain the same option for manual barcode printing). When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167213 | Z-0357-2019 | Software Package, iPAW 4.2.1 User Kit for Windows XP and Windows 10, Part No. B95343, Manufactured by: Beckman Coulter Biomedical GmbH Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes. | Software Package, Software Package, Sorting-Drive 4.2.1 User Kit for Windows XP and Windows 10 | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes. When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167211 | Z-0356-2019 | Software Package, iPAW 4.2 User Kit, Part No. B73489, Manufactured by: Beckman Coulter Biomedical GmbH The iPAW 4.2 User Kit (PN B73489) includes the iPAW software application CD (version 4.2). This function is primarily used to copy an existing barcode ID to replace the label of an individual sample tube because the original label is deteriorated. | Software Package, iPAW 4.2 User Kit | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes. When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167207 | Z-0354-2019 | Software Package, Sorting-Drive 4.2 User Kit, Part No. B73488, Manufactured by: Beckman Coulter Biomedical GmbH The Sorting-Drive itself is an optional software to be used with the Automate 2500 Family of products. The system automates the process of sorting, decapping, and archiving samples. The Sorting-Drive 4.2. User Kit includes the Sorting-Drive Client software application that includes an option for manual printing of barcodes for sample tubes. | Software Package, Sorting-Drive 4.2 User Kit | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option which included the Sorting-Drive Client that is used for manual printing of barcodes. When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167209 | Z-0355-2019 | Software Package, Sorting-Drive 4.2.1 User Kit for Windows XP and Windows 10, Part No. B95347, Manufactured by: Beckman Coulter Biomedical GmbH The Sorting-Drive itself is an optional software to be used with the Automate 2500 Family of products. . The Sorting-Drive 4.2.1 User Kit includes the Sorting-Drive Client software application that includes an option for manual printing of barcodes for sample tubes; version 4.2 was replaced by version 4.2.1 (they both contain the same option for manual barcode printing). This function is primarily used to copy an existing barcode ID to replace the label of an individual sample tube because the original label is deteriorated. | Software Package, Sorting-Drive 4.2.1 User Kit for Windows XP and Windows 10 | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option which included the Sorting-Drive Client that is used for manual printing of barcodes; version 4.2 was replaced by 4.2.1 (they both contain the same option for manual barcode printing). When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs.The firm reports this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167218 | Z-0359-2019 | Sorting-Drive Professional Part, Sorting-Drive Professional Kit, (Sorting Drive Software v4.2.1), Part No. A94144, Beckman Coulter, Inc. The Sorting-Drive Professional Kit (PN A94144) includes the Sorting-Drive 4.2.1 User Kit (PN B95347) which includes the Sorting-Drive 4.2.1 Software DVD (PN B95346). This software version includes an option for manual barcode printing for sample tubes. Sorting-Drive software facilitates the Laboratory Information System (LIS) connection to AutoMate 2500 Family systems. | Sorting-Drive Professional Part, Sorting-Drive Professional Kit, (Sorting Drive Software v4.2.1) | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes.Sorting Drive software version 4.2 was replaced by 4.2.1 (both versions contain the same option for manual barcode printing). When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167217 | Z-0358-2019 | Sorting-Drive Product Part, Sorting-Drive Plus Kit (Sorting Drive Software v4.2.1), Part No. A89483, Beckman Coulter, Inc. The Sorting-Drive Plus Kit includes the Sorting-Drive 4.2.1 User Kit which includes the Sorting-Drive 4.2.1 Software DVD. . The Sorting-Drive software facilitates the Laboratory Information System (LIS) connection to AutoMate 2500 Family systems. . | Sorting-Drive Product Part, Sorting-Drive Plus Kit (Sorting Drive Software v4.2.1) | 3 | CDRH | 2018-10-31 |  | 2022-12-08 16:31 | Beckman Coulter Inc. | Sorting-Drive software version 4.2 (released in October 2015) introduced a new option for iPAW and the Sorting-Drive Client that is used for manual printing of barcodes. Sorting Drive software version 4.2 was replaced by 4.2.1 (both versions contain the same option for manual barcode printing). When printing consecutively numbered tube labels using this optional Barcode Counter function, the system incorrectly converts numbers with consecutive multiple zeroes in the middle, potentially leading to wrong barcode IDs. The firm reports that this does not affect automated label creation or reprinting of individual labels and that no other system functions are affected.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=165764 | Z-0111-2019 | GE Healthcare MRI Systems, Software Version 12.x, sold under the following product names: SIGNA MR/I ECHOSPD W/ EXCITE, 1.5 HISPEED HD 8-CH 400-AP, MR GOLDSEAL 1.5T HD 8CH USED, UPGRADE 1.5 EXC-II TO EXC-III HD, SIGNA MR/I HISPD W/ EXCITE, 1.5 ECHOSPEED HD 16-CH 800-AP, 1.5 HISPEED HD 4-CH 200-AP, 1.5 TWINSPEED HD 8CH-400-AP, MR UPGRADE LX/HZ TO HD FOR NON LCC, 1.5 TWINSPEED HD 16CH-800-AP, 1-5T HDXT TWINSPEED 16 CHANNEL, SIGNA MR/I TWISP TR W/EXCITE, 3T HD 16CH-800-AP, 1.5 ECHOSPEED HD 4-CH 200-AP, UPGRADE PREEXCITE TO 1.5TS HD 8CH, MR GOLDSEAL 1.5T HDXT 16CH NEW, MR GOLDSEAL 1.5T HD ES 8CH MOB, 1-5T NON LCC LX/5_5 TO HDIUPG8CH, UPGRADE 1.5 EXC-II TO EXC-III 16HD, UPGRADE 1.5 HORIZON TO HS/ES EXCITE, UPGRADE FROM 1.5 4X TO HISPEED LX, 1.5 LX TO HD 8CH FORKLIFT UPG., MR GOLDSEAL 1.5 LX HISP PL NEW, MR UPGRADE 4X/5X TO HD FOR NON LCC, MR GOLDSEAL 1.5 ES+ LX, SIGNA HISPEED LX 1.5T FIXED, 3T HD T/R W/EXCITE, 1.5 ECHOSPEED HD 8-CH 400-A, SIGNA 1.5 ECHOSPEED HD 8-CH 400-A, SIGNA EXCITE HD 1.5T, MR EXCITE12X, MR GOLDSEAL 1.5T HD 8CH USE, 1.5T BRM ES 16CH 2 VRE FIXED F, UPG 1.5 EXC-II TO EXC-III HD, EXCITE HD - UPG 3.5 EXC, 1.5 ECHOSPEED HD 4-CH 200, SIGNA HD ADVANCED, EXCITE HD, 1.5 ECHOSPEED HD 16-CH, MR 1.5T HD, 1.5 ECHOSPEED HD 8-CH, MR 1.5 EXCITE HD, MR GOLDSEAL 1.5T HD ES 8CH NEW, 1.5T EXCITE HD, SIGNA ECHOSPD LX K4 1.5T FIXED, 1.5T LX TO 8CH HD 23 UPG, 1.5 HISPEED HD 8-C, SIGNA EXCITE 1.5T HD, MR 1.5T HD ECHOSPEED WITH EXCITE, MR 1.5T HD TWINSPE, MR 1.5T HD ECHOSPEED, MR SIGNA VH/I 3.0T, MR 3.0T HD 16CH, MR 1.5T HD TWINSPEED, MR 3.0T HD 8CH, MR UPG TO 1.5 HD, MR UPG TO 1.5T HD, MR UPG 1.5ES HD 8CH, SIGNA EXCITE HD 1.5 T, SIGNA EXCITE HD, SIGNA MR/I ECHOSPD 1.5T FIX(+75798), SIGNA ES EXCITE, SIGNA 1.5T ECHO HD, MR HD (VON FA. LINK), SIGNA 1.5T ECHOSPEED HD, SIGNA MR 1.5T ECHO HD, SIGNA EXITE-III HD, SIGNA 1.5T HD, ECHOSPEED HD 8-CH, 1.5 ECHOSPEED HD 8CH, MR 1.5T EXCITE HD, MR SIGNA EXCITE 3T HD, MR 1.5 ECHOSPEED HD 8-CH 400-AP, 1.5 ECHOSPEED HD, 1.5T HD, 1.5T HD PREOWN, 1.5T ECHO SPEED, 1.5 ECHOSPEED HD , SIGNA MR 1.5 ECHOSPEED 8CH, SIGNA HD 1.5 ES, MR ECHOSPEED 1.5 HD, SIGNA 1.5TECHOSPEED, MR 1.5 ECHOSPEED HD 8-CH, 1.5T HD EXCITE, SIGNA 1.5 T ECHOSPEED, SIGNA VH/I 3.0T, MR 1.5T HD 16CH, MR 1.5T HD 8CH USED, SIGNA EXCITE3 HD, MR SIGNA EXCITE HD (UPGR.), 1.5T EXCITE-HD, 1.5T EX-HD TWIN, SIGNA 3T, MR 1.5 ECHOSPEED HD 8-CH 400-A, 1.5T EXCITE HD GS, 3T E2, 1.5 ECHOSPEED HD 8-CH 400-, SIGNA EXCITE III HD 1.5T TWIN SPEED, MR GOLDSEAL 1.5T HD 16CH USED, EXCITE II UPG, SIGNA EXCITE III HD 3.0T, MR SIGNA EXCITE, 1.5T SIGNA EXCITE III HD, MR SIGNA HD, SIGNA EXCITE HD 1.5, 3T EXCITE HD, UPG PREEXCITE TO 1.5ES HD 8CH, MR SIGNA SMARTSPEED 1.5, 1.5T HD ECHOSPEED 8 CHANNEL, SIGNA 1.5 MOBILE EXCITE HD, SIGNA 3T HD 16CH-800-AP, SIGNA 3T HD 8CH-400-AP, 1.5T MRI EXCITE 3 UPGRADE, MR 1.5 ECHOSPEED HD 4-CH 200-AP, SIGNA 1.5 ECHOSPEED HD 8-CH 400-AP, SIGNA EXCITE 1.5T ECHOSPEED HD 4CH, MR 3T HD 8CH-400-AP, SIGNA 1.5T ECHOSPEED HD 4-CH 200-AP, 1.5 ECHOSPEED HD 16CH 800AP-2669585, 1.5 ECHOSPEED HD 8-CH 400-AP2669575, 1.5T ECHOSPEED HD16-CH800-AP2669583, UPG 1.5 EXC-II TO EXC-III HD-561936, 1.5 TWINSPEED HD 8CH-400-AP-2583394, 1.5 ECHOSPEED HD 8-CH 400-AP2700482, 1.5 ECHOSPEED HDMR, 1.5 ECHOSPEED HD 8-CH 400-AP EXCITE, MR 1.5 ECHOSPEED HD 8-C, MR SIGNA 1.5T HD 8CH, MR SIGNA 15T HD UPGRADED, MR SIGNA 3T HD 8CH-400-AP, MR SIGNA 1.5T HD UPGRADED, MR SIGNA 1.5T HD 8CH MOBIL, EXCITE HD 1.5T UPGRADES, MR SIGNA EXCITE HD 15T, MR SIGNA EXCITE HD 1.5T, MR 1.5 ECHOSPEED HD 16-CH, MR PROFILE HD PHASE ARRAY, SIGNA 3T EX HD, MR SIGNA VH/I 3.0, MR SIGNA MR/I HISPEED 1.5T, 1.5T ECHOSPEED HD, 1.5 ECHOSPEED HD 8, 1.5T ECHOSPEED HD 8, 1.5 T ECHOSPEED HD 4-CH 200-AP, SIGNA EXCITE 3T HD 8CH-400-AP, MR 1.5T ECHOSPEED HD 8-CH 400-AP, MR 1.5 ECHOSPEED HD 4-CH 200-A, MR/I 1.5 ECHOSPEED HD 4CH 200-, MR 15T ECHOSPEED HD 8-CH 400-, MR 15 ECHOSPEED HD 4-CH 200-A, MR 1.5 ECHOSPEED HD 8-CH 400-AP//, SIGNA 1.5T, SIGNA MR/I ECHOSPD W/ EXCIT, 1.5T EXCITE HD 16 CH, 1.5T HD E | GE Healthcare MRI Systems, Software Version 12.x | 2 | CDRH | 2018-10-15 |  | 2020-06-18 09:37 | GE Medical Systems, LLC | The latest on-site software version may not have been reinstalled at some sites after service activities were performed resulting in images possibly flipped left/right and/or there can be patient data mismatch.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=165729 | Z-0110-2019 | GE Healthcare MRI Systems, Software Version 11.x, sold under the following product names: 1.5T CXK4 LCC MAGNET,1.5T EXCITE, 1.5T EXCITE E/S, 1.5T EXCITE GS, 1.5T EXCITE HYBRID, 1.5T EXCITE TWIN, 1.5T EXCITE UPGRAD, 1.5T HD ES/HS TR W/EXCITE, 1.5T MR MOBILE (SERIAL # R350), 1.5T SIGNA ECHOSPEED, 1.5T SIGNA EXCITE, 1.5T SIGNA EXCITE MRI, 1.5T SIGNA VIBRANT 16CH FIXED, 3.0T MRI TWINSPEED 8CH, 3.0T SIGNA EXCITE FIXED, 3.0T SIGNA EXCITE UPGRADE, 3T EXCITE, EX ES+8CH USED EXCITE, EXCITE 1.5T, EXCITE 3.0T, EXCITE TWIN, GOLDSEAL 1.5T EXCITE 4CH, GOLDSEAL 1.5T EXCITE 8CH, GOLDSEAL 1.5T EXCITE REL GOLDSEAL EXCITE 1.5T, GOLDSEAL EXCITE 1.5T 8CH, GS 1.5 EX ES+ 4CH NEW, GS 1.5 EX ES+4CH USED, GS 1.5 EX ES+8CH NEW, GS 1.5T EXCITE MRI, GS CLASS 1.5T ES+ EX 4CHN, GS SIGNA EXCITE II 1.5T, GS SIGNA MR/I 1.5T, LCC EXCITE 1.5T, MR 1.5 EX ES 8CH NEW, MR 1.5 EX ES+ 4CH, MR 1.5 EX ES+4CH MOB, MR 1.5 EX ES+8CH, MR 1.5EX8CH, MR 1.5EX8CH USE -HUNG, MR 1.5T EX8CH, MR 1.5T EXCITE 4CH, MR 1.5T EXCITE EXPERT MR, MR 1.5T EXCITE GS, MR 1.5T EXCITE GS MR, MR 1.5T GS ES+ EXCITE 4 CHANNE, MR 1.5T SIGNA EXCITE 8 CH GOLD SEAL, MR 1.5T SIGNA EXCITE II 4CH, MR 3.0T TWINSPEED, MR 3.0T SIGNA EXCI, MR 3.0T SIGNA EXCITE, MR 3.0T SIGNA EXCITE FIXED, MR EXCITE, MR EXCITE 1.5T, MR EXCITE 3, MR EXCITE ECHOSPEE, MR EXCITE ECHOSPEE UPG, MR EXCITE ECHOSPEED, MR EXCITE HISPEED MR EXCITE TWINSPEE, MR EXCITE TWINSPEED, MR GOLD SEAL 1.5 EX ES+8CH NEW, MR GOLDSEAL 1.5 ES PL EX 4 CH, MR GOLDSEAL 1.5 EX, MR GOLDSEAL 1.5 EX, MR GOLDSEAL 1.5 EX 4CH, MR GOLDSEAL 1.5 EX 8CH, MR GOLDSEAL 1.5 EX ES+ 4CH, MR GOLDSEAL 1.5 EX ES+ 4CH NEW, MR GOLDSEAL 1.5 EX ES+4CH, MR GOLDSEAL 1.5 EX ES+4CH USD, MR GOLDSEAL 1.5 EX ES+4CH USE, MR GOLDSEAL 1.5 EX ES+4CH USED, R GOLDSEAL 1.5 EX ES+8CH, MR GOLDSEAL 1.5 EX ES+8CH NEW, MR GOLDSEAL 1.5 EX ES+8CH USED, MR GOLDSEAL 1.5 EXCITE S+ 4CH, MR GOLDSEAL 1.5 FIXED ACTIVE, MR GOLDSEAL 1.5EX4CH NEW -HUNG, MR GOLDSEAL 3.0T EXCITE HD FIX, MR GS 1.5 EX, MR GS 1.5 EX ES+8C, MR GS 1.5 EX ES+8CH, MR SIGNA 1.5 EXCIT, MR SIGNA 1.5 EXCITE EXPERT, MR SIGNA 1.5T EXCITE, MR SIGNA 1.5T EXCITE 4CH MR SIGNA 1.5T EXCITE 8CH, MR SIGNA 1.5T EXCITE 8CH MOBIL, MR SIGNA 1.5T EXCITE EXPERT, MR SIGNA 1.5T EXCITE II, MR SIGNA 15T EXCITE 8CH, MR SIGNA 15T EXCITE 8CH MOBIL, MR SIGNA EXCITE 1.5 T GS (4 CHANNEL, MR SIGNA MR/I ECHOS W/EXCITE, MR SIGNA MR/I ECHOSPD W/ EXCITE, MR SIGNA MR/I ECHOSPD W/ EXCITE , MR SIGNA MR/I ECHOSPEED W/ EXCITE, MR SIGNA MR/I HISPEED W/ EXCITE MR SIGNA MR/I TWISP TR W/EXCIT, MR TWINSPEED 1.5T W/ EXCITE, MR UPG 1.5 SIGNA LX TO EXCITE, MR UPG 3.0T94 ECLIPSE TO EXCIT, MR/I ECHOSPD W/ EXCITE, MR/I ECHOSPEED W/ EXCITE, MR/I SIGNA ECHOSPD W/ EXCITE, MRI EXCITE, MRI EXCITE 1.5T FIXED, MRI SIGNA EXCITE 1.5T, PREOWN 1.5T EXITE, SIGNA 1.5 EXCITE, SIGNA 1.5T , CHOSPEED W EXCITE, SIGNA 1.5T ECHOSPEED W/EXCITE, SIGNA 1.5T ES EXCITE, SIGNA 1.5T EXCITE, SIGNA 1.5T EXCITE 2, SIGNA 1.5T EXCITE EXPERT, SIGNA 1.5T EXCITE II, SIGNA 1.5T SMART SPEED, /EXCITE, SIGNA 1.5T W/ EXCITE, SIGNA 3.0T E, SIGNA 3.0T SIGNA EXCITE FIXED, SIGNA ECHOSPEED 1.5T FIXED, SIGNA ECHOSPEED EXCITE 1.5T, SIGNA ECHOSPEED LX 1.5T FIXED SIGNA ECHOSPEED PLUS + EXCITE, SIGNA EXCITE, SIGNA EXCITE 1.5 SIGNA EXCITE 1.5 T HISPEED, SIGNA EXCITE 1.5T, SIGNA EXCITE 1.5T RELEASE 12.X/MR, SIGNA EXCITE 1.5T SMART SPEED, SIGNA EXCITE 11X, SIGNA EXCITE 3.0T, SIGNA EXCITE ECHOSPEED 1,5 T, SIGNA EXCITE ECHOSPEEED PLUS 1.5 T, SIGNA EXCITE EXPERT 1.5T, SIGNA EXCITE GS 1.5T 4 CHANNEL,SIGNA EXCITE II, SIGNA HISPD LX K4 1.5T FIXED, SIGNA HORIZON LX 1.5T FIXED, SIGNA LX 1.5 EXCITE II., SIGNA LX 1.5T EXCITE, SIGNA LX 1.5T EXCITE II, SIGNA LX1.5T HISPEED, SIGNA MR 1.5 ES EXCITE, SIGNA MR/I 1.5T EXCITE, SIGNA MR/I ECHO W/E,SIGNA MR/I ECHOESPEED W/EXCITE, SIGNA MR/I ECHOSPD, SIGNA MR/I ECHOSPD 1.5T FIXED, SIGNA MR/I ECHOSPD W/ EXCITE 1.5, SIGNA MR/I ECHOSPEED W/ EXCITE, SIGNA MR/I ECHOSPEED W/ EXCITE 1.5T, SIGNA MR/I ECHOSPEED W/EXCITE, SIGNA MR/I ECHOSPEED WITH EXCI, SIGNA MR/I ECHOSPEED WITH EXCITE SIGNA, MR/I | GE Healthcare MRI Systems, Software Version 11.x | 2 | CDRH | 2018-10-15 |  | 2020-06-18 09:37 | GE Medical Systems, LLC | The latest on-site software version may not have been reinstalled at some sites after service activities were performed resulting in images possibly flipped left/right and/or there can be patient data mismatch.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=165871 | Z-0112-2019 | GE Healthcare MRI Systems, Software Version 15.x or HD 16 sold under the following product names: *ER*15T HDXT ECHOSPEED 16 CHANNEL, *OC*1.5T HDX ECHOSPEED 8 CHANNEL, *SC*1.5T HD TO HDX UPGRADE, *TR*15T HDXT 16CH, *TR*3.0T HDX TWINSPEED 16 CHANNEL, 1,5T HDXT, 1.5 HD TO HDX UPGRADE, 1.5 HDX, 1.5 HDX ECHOSPEED 16 CH, 1.5 HDX ECHOSPEED 8 CHANNEL, 1.5 HDXT, 1.5 HDXT-16 , 1.5 LX TO HDX 16C FORKLIFT UPG, 1.5 LX TO HDX 8C FORKLIFT UPG, 1.5 T UPGRADE TO HDX, 1.5EXCIT TO HDXT UPG, 1.5T HDX 16 CHANNEL, 1.5T 16CH UPG, 1.5T 6 CH HDX , UPGRADE, 1.5T 8CH HDXT TO 16CH HDXT UPG, 1.5T CRM LX TO HDX UPG, 1.5T ES HDX W/ONCOLOGY PACK, 1.5T EXC.TO HDX UPG, 1.5T EXC.TO HDXT UPG, 1.5T EXCITE HD TO HDX UPG, 1.5T EXCITE TO HDX UPG, 1.5T EXCITE TO HDX UPG, 1.5T EXCITE TO HDX UPG. W/ACGD, 1.5T EXCITE TO HDXT UPG, 1.5T EXCITE TO HDXT UPGRADE, 1.5T EX-HDX, 1.5T EX-HDXTWIN, 1.5T HD ES/HS MOBILE, 1.5T HD TO HDX UPG. 1.5T HD TO HDX UPGRADE, 1.5T HD TO HDXT UPG, 1.5T HD TO HDXT UPGRADE 1.5T HDI,1.5T HDI 16CH FIXED, 1.5T HDI 8CH FIXED, 1.5T HDI ECHOSPEED 1.5T HDI ECHOSPEED 8 CHANNEL, 1.5T HDI ECHOSPEED 8 CHANNEL 1.5T HDI ECHOSPEED 8 CHANNEL., 1.5T HDI ECHOSPEED 8 CHANNEL, 1.5T HDI ECHOSPEED 8 CHAN¿SO2762285, 1.5T HDI ECHOSPEED 8CH, 1.5T HDI ECHOSPEED MOBILE, 1.5T HDI ECSHOSPEED 8CH, 1.5T HDI TO HDXT UPGRADE, 1.5T HDX, 1.5T HDX - MEIRAV (BREAST), 1.5T HDX 8CM 1.5T HDX ECHOSPD 16, 1.5T HDX ECHOSPEED, 1.5T HDX ECHOSPEED 1.5T HDX ECHOSPEED 16, 1.5T HDX ECHOSPEED 16 CHANNEL, 1.5T HDX ECHOSPEED 16 CHANNEL, 1.5T HDX ECHOSPEED 16 CHANNELS, 1.5T HDX ECHOSPEED 16CH¿SO2659508, 1.5T HDX ECHOSPEED 16CH-SO2718893, 1.5T HDX ECHOSPEED 32 CHANNEL, 1.5T HDX ECHOSPEED 8 CH, 1.5T HDX ECHOSPEED 8 CHANNEL, 1.5T HDX ECHOSPEED MOBILE 1.5T HDX ECHOSPEED MRI, 1.5T HDX ES, 1.5T HDX ES 16, 1.5T HDX TWINSPEED, 1.5T HDX TWINSPEED 16 CHANNEL, 1.5T HDXT 1.5T HDXT 16CH,1.5T HDXT 16CH CXK4, 1.5T HDXT 16CH ES FIXED 1.5T HDXT 8 CH, 1.5T HDXT 8CH ES FIXED, 1.5T HDXT 8CH UP, 1.5T HDXT ECHOSPEE, 1.5T HDXT ECHOSPEED, 1.5T HDXT ECHOSPEED 16 CH 1.5T HDXT ECHOSPEED 16 CHAN, 1.5T HDXT ECHOSPEED 16 CHANNEL 1.5T HDXT ECHOSPEED 8, 1.5T HDXT ECHOSPEED 8 CHANN, 1.5T HDXT ECHOSPEED 8 CHANNEL, 1.5T HDXT ECHOSPEED 8 CHANNEL , 1.5T HDXT ECHOSPEED 8_CHANNEL, 1.5T HDXT ECHOSPEED MOBILE, 1.5T HDXT ES 8 CH, 1.5T HDXT GS, 1.5T HDXT TWIN, 1.5T HDXT TWINSPEE, 1.5T HDXT TWINSPEED 16 CHAN,1.5T HDXT UPG, 1.5T HDXTID, 1.5T LX HDX, 1.5T LX TO 16 CH HDXT UPG, 1.5T LX TO 16 CHAN, 1.5T LX TO 16 CHANNEL HDX UPG, 1.5T LX TO 16 CHANNEL HDX UPG, 1.5T LX TO 16 CHANNEL HDX UPGRADE, 1.5T LX TO 16 CHANNEL HDXT, 1.5T LX TO 16 CHANNEL HDXT UPG, 1.5T LX TO 16CHANNEL HDX U/G2777896, 1.5T LX TO 8 CH HDX UPG, 1.5T LX TO 8 CH HDXT UPG, 1.5T LX TO 8 CHANNEL HDX UP, 1.5T LX TO 8 CHANNEL HDX UPG, 1.5T LX TO 8 CHANNEL HDX UPG, 1.5T LX TO 8 CHANNEL HDX UPG.827015, 1.5T LX TO 8 CHANNEL HDXT UPG, 1.5T LX TO 8 CHANNEL HDXT UPGRADE, 1.5T LX TO HDXT 1.5T NON LCC4X5X TO HDXT UPGD, 1.5T NON-LCC LX/HRZ 5.5 TO 1.5T HDX, 1.5T SIGNA HDX UPG, 1.5T SIGNA HDXT, 1.5TEXCITE TO HDX UPG W/O A, 1.5TEXCITE TO HDX UPG W/O ACGD, 1.5TNONLCC LXHZ TO HDXT UPG, 1.5TNONLCC LXHZ TO HDXUPGW/ACD, 1.5TNONLCC4X5X TO HDXUPGW/OACD, 1.5TNONLCCLXHZ TO HDXUPGW/O ACD, 1¿5T 8CH HDXT MRI, 1¿5T HDXT, 1¿5T HDXT MRI, 1-5T CRM LX TO HDX , G.W/OACGD, 1-5T CRM LX TO HDXT UPG, 1-5T EXCITE TO 15T EXCITE TO HDX UPG, 1-5T HDI ECHOSPEED 16 CHANNEL, 15T HDI ECHOSPEED 8 CHANNEL, 1-5T HDX ECHOSPEED 16 CHANN, 15T HDX ECHOSPEED 16 CHANNEL, 15T HDX ECHOSPEED 8 CHANNEL, 15T HDX ECHOSPEED MOBILE, 1-5T HDX ECHOSPEED MOBILE, 1-5T HDX ECHOSPEED TRANSPORT, 1-5T HDX TWINSPEED 8 CHANNEL, 1-5T HDXT ECHOSPEED 16, 1-5T HDXT ECHOSPEED 16 CHAN, 1-5T HDXT ECHOSPEED 16 CHAN-HD16.0, 1-5T HDXT ECHOSPEED 16 CHANNEL, 1-5T HDXT ECHOSPEED 16 CHANNEL, 1-5T HDXT ECHOSPEED 16 UPGRADE, 1-5T HDXT ECHOSPEED 8 CHANN, 15T HDXT ECHOSPEED 8 CHANNEL, 1-5T HDXT ECHOSPEED MOBILE, 1-5T HDXT ECHOSPEED RELOCATABLE, 1-5T LX TO 16 CHANNEL HDX UPG, 1-5T LX TO 16 CHANNEL HDXT, 1-5T LX TO 8 CHANNEL HDXT U, 1-5T LX TO | GE Healthcare MRI Systems, Software Version 15.x or HD 16 | 2 | CDRH | 2018-10-15 |  | 2020-06-18 09:37 | GE Medical Systems, LLC | The latest on-site software version may not have been reinstalled at some sites after service activities were performed resulting in images possibly flipped left/right and/or there can be patient data mismatch.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164777 | Z-2266-2018 | CARESCAPE R860 ventilators with software version 10SP05 Product Usage: The CARESCAPE R860 ventilator is designed to provide mechanical ventilation or support to neonatal, pediatric, and adult patients weighing 0.25 kg and above. The CARESCAPE R860 ventilator is a microprocessor based, electronically controlled, pneumatically driven ventilator that includes integrated monitoring of FiO2, airway pressure, flow, and volume. Additional respiratory gas monitoring capabilities are supported through the use of optional GE patient monitoring modules. Not all features are available for all patient types or product configurations. The CARESCAPE R860 ventilator is not a pulmonary function calculation device. The system is designed for facility use, including within facility transport, and should only be used under the orders of a clinician. | CARESCAPE R860 ventilators with software version 10SP05 | 2 | CDRH | 2018-06-20 | 2018-06-01 21:04 | 2021-04-29 14:13 | Datex-Ohmeda, Inc. | Potential for the display processor to experience an unexpected failure.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164900 | Z-2109-2018 | 12C (AdaPTinsight) Affected component: AdaPTinsight software Product Usage: 12C is used with a charged particle or photon radiation therapy system for localization of the patient position with respect to the therapy equipment and to provide correction feedback to the radiation therapy system. | AdaPTinsight | 2 | CDRH | 2018-06-07 |  | 2019-08-09 10:37 | Ion Beam Applications S.A. | IBA is initiating this recall to address an issue identified with AdaPTinsight software and to reduce risk related to this problem.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=162647 | Z-1170-2018 | Proclaim DRG Implantable Pulse Generator, Model Number 3664 | Proclaim DRG Implantable Pulse Generator | 2 | CDRH | 2018-03-23 | 2018-03-23 21:03 | 2020-11-03 09:50 | St. Jude Medical, Inc. | The firm received complaints of error messages that occurred during routine impedance checks on Proclaim DRG IPGs, model 3664. Some complaints were also associated with transient over stimulation which created discomfort for the patients.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153395 | Z-1816-2017 | TotalCare Bed Systems, Model Numbers P1900, PR1900, and U1900 with 3-Level Patient Position Monitoring and Scale Upgrade Kits, Model Numbers P1947A and P1948A). | TotalCare | 2 | CDRH | 2017-04-12 |  | 2017-09-19 | Hill-Rom, Inc. | The bed exit alarm may not function properly if the alarm has been armed for a period longer than seventy-two (72) hours without the alarm being triggered by patient movement or removal of patient weight within that time period.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153003 | Z-1494-2017 | Final Fit Software Version 1.11 and 1.12; PC Based software installed outside Nidek EC-5000 Excimer Laser System. Ophthalmic laser system use for correction of corneal refraction and ablation of the corneal surface. The system is composed of a laser generator which produces an excimer laser radiation of wavelength 193nm, a beam delivery unit, an optical system for observation, a gas system and a computer for system control. | Nidek EC-5000 Excimer Laser System, Final Fit Software Version 1.11 and 1.12 | 2 | CDRH | 2017-03-20 |  | 2017-12-05 | Nidek Inc | During treatment planning, the procedure was programmed with an unintended (wrong) correction.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152488 | Z-1045-2017 | Ambra PACS UDI: +AMBRHEALTHSOLUTIONS0/$$+7_V3.16.13.0R Software Version Number: V3.16.13.0 Intended for use as a primary diagnostic and analysis tool for diagnostic images. | Ambra PACS | 2 | CDRH | 2017-01-18 | 2017-01-18 21:00 | 2017-04-07 | DICOM GRID INC | A software error caused the window/level to become the same in one series regardless if the image had different levels; image results have a washed-out grey appearance.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150430 | Z-0700-2017 | Model 37751 Recharger Product Usage: The Medtronic 37751 Recharger is a hand-held device used to charge the battery in a patient s neurostimulator. It includes a display to provide information on the charging system. The charging system consists of the Model 37751 Recharger, an AC power supply and power cord (Model 37761) and an antenna. | Recharger | 2 | CDRH | 2016-12-03 |  | 2022-05-04 09:36 | Medtronic Neuromodulation | Medtronic has identified an increased number of complaints from customers involving reports of Rechargers (Medtronic Model 37751 Recharger, which is included in Models 37754 and 97754 Charging Systems,Spinal Cord Stimulation)that are in an unresponsive error state, where the Recharger is non-functional with a blank display screen and is beeping every 5 seconds. Medtronic has determined all Rechargers manufactured starting in November 2014 (indicated by serial numbers beginning with NKA4 or NKU4 ) are more susceptible to this error state.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148728 | Z-2711-2016 | SPECULAR MICROSCOPE CEM-530; Software version 1.08 and 1.09. Opthalmic: The NIDEK Specular Microscope CEM-530 provides non-contact, high magnification image capture of endothelium enabling observation of the size and shape of cells. Information such as the number of endothelial cells, cell density, and cell area is analyzed through the captured images. The captured images and analysis results of endothelium are used for intraocular or corneal surgery, postoperative follow-up and corneal observation such as for endothelial disorders or the corneal state of patients who wear extended-wear contact lenses. | SPECULAR MICROSCOPE CEM-530 | 2 | CDRH | 2016-09-01 | 2016-09-01 21:00 | 2017-01-09 | Nidek Inc | Software version 1.08 and 1.09 for the Specular Microscope CEM 530 included a change of analysis results feature that was not reviewed and approved by the FDA.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144871 | Z-2370-2016 | Brilliance CT 16 Power Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Brilliance CT 16 Power Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144868 | Z-2367-2016 | Brilliance BigBore Oncology Computed Tomography X-ray systems, intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Brilliance BigBore Oncology Computed Tomography X-ray systems, | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144870 | Z-2369-2016 | Brilliance 16 Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Brilliance 16 Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144867 | Z-2366-2016 | Brilliance 64 Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Brilliance 64 Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144869 | Z-2368-2016 | Brilliance BigBore Radiology Computed Tomography X-ray systems, intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Brilliance BigBore Radiology Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144872 | Z-2371-2016 | Ingenuity Core Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Ingenuity Core Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144873 | Z-2372-2016 | Ingenuity Core 128 Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Ingenuity Core 128 Computed Tomography X-ray systems | 2 | CDRH | 2016-08-09 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144201 | Z-1394-2016 | VERION Reference Unit (Vision Planner) Product Usage: The VERION Image Guided system is an advanced surgical planning, imaging and guidance technology designed to provide greater accuracy and efficiency during cataract surgery. It allows eye surgeons to capture a high-resolution, diagnostic reference image of the patients eye pre-operatively, quickly determine an optimized surgical plan that enables surgeons to see all inclusions and alignment in real-time. | VERION Image Guided System | 2 | CDRH | 2016-04-14 | 2016-04-14 21:04 | 2017-04-13 | Alcon Research, Ltd. | Alcon is conducting a voluntary medical device correction of all VERION Reference Units (Vision Planner) that are shared with the Alcon LenSx Laser System after receiving reports concerning the inclusion of unplanned arcuates on printed, saved (.pdf) or exported surgical plans.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=141582 | Z-0319-2016 | Syngo Imaging VB36D_HF02. Radiological image processing system. | Syngo Imaging | 2 | CDRH | 2015-11-20 | 2015-11-20 21:01 | 2016-10-06 | Siemens Medical Solutions USA, Inc | To provide supplementary information regarding the release of the syngo Imaging Software Version VB36D_HF02. The software provides improvements for all syngo Imaging installations running on the software Version VB36D.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139574 | Z-0117-2016 | NicoletOne Software v5.94, Catalog/Part Number: 482-649600. The NicoletOne data acquisition and review system, either with or without synchronous digital video. The system is intended for medical purposes to record, measure, store and display full band (FbEEG) cerebral EEG and extracerebral activity for Clinical EEG, Electrocorticography (EcOG), Long Term Monitoring (LTM), Intensive Care Unit (ICU) monitoring and Polysomnography (PSG) Sleep studies. While the Nicolet Neurodiagnostic systems are capable of displaying signals, such as Sp02 and EKG, the system is NOT intended for monitoring such signals for the preservation of life, The Nicolet Neurodiagnostic systems are intended to acquire, analyze, and display data | NicoletOne Software v5.94 | 2 | CDRH | 2015-10-13 | 2015-10-13 21:00 | 2016-03-08 | Natus Neurology Inc | When using the NicoletOne v5.94 software, after exiting the impedance check function and returning to the EEG screen, the impedance check signal remains active in waveform, obscuring the EEG signals.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139556 | Z-2784-2015 | Merlin@home RF Remote Monitoring Transmitter Model EX1150. Intended to aid in the remote support of SJM implantable pacemakers and ICD's. | Merlin@home RF Remote Monitoring Transmitter Model EX1150 | 2 | CDRH | 2015-09-17 | 2015-09-28 21:00 | 2016-05-19 | St Jude Medical Cardiac Rhythm Management Division | Some devices exhibit backup VVI operation due to a backup reset.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138620 | Z-2258-2015 | Computed Tomography X-ray Systems Ingenuity Core 128. Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Computed Tomography X-ray Systems Ingenuity Core 128 | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138615 | Z-2255-2015 | Computed Tomography X-ray systems Brilliance iCT, Philips Healthcare System. Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Computed Tomography X-ray systems Brilliance iCT | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138618 | Z-2256-2015 | Computed Tomography X-ray Systems Brilliance iCT SP. Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Computed Tomography X-ray Systems Brilliance iCT SP | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=136169 | Z-2254-2015 | Computed Tomography X-ray Systems Brilliance CT 64-channel with Essence technology. Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Computed Tomography X-ray Systems Brilliance CT | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138621 | Z-2259-2015 | Computed Tomography X-ray Systems Ingenuity CT scanners. Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Computed Tomography X-ray Systems Ingenuity CT | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138619 | Z-2257-2015 | Computed Tomography X-ray Systems Ingenuity Core. Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Computed Tomography X-ray Systems Ingenuity Core | 2 | CDRH | 2015-07-24 | 2015-07-24 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=133315 | Z-1125-2015 | Alive ECG App 2.1.2 ( a medical device application for the Apple iOS operating system, intended to be used with the AliveCor Heart Monitor. The AliveCor Heart Monitor is intended to record, store and transfer single-channel electrocardiogram (ECG) rhythms. The AliveCor Heart Monitor also displays ECG rhythms and detects the presence of atrial fibrillation and normal sinus rhythm (when prescribed or used under the care of a physician). | AliveECG App Version 2.1.2 used with AliveCor Heart Monitor | 3 | CDRH | 2015-02-18 | 2015-02-18 21:00 | 2015-02-18 | Alivecor SFO | Alive ECG App version 2.1.2 (intended to be used with the AliveCor Heart Monitor) crashed upon use of the application.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131698 | Z-0856-2015 | Lite DM (Lite Data Manager) Software Version(s): Versions 2.0.0.19 and lower Product Model #: D00A00224D The Lite DM (Lite Data Manager) is a software application developed by HORIBA Medical intended to receive, store, and manage Startup, Patient, and Quality Control (QC) results from an ABX Micros 60 Hematology Analyzer or ABX Micros CRP 200 Hematology Analyzer. Only one instrument connection is allowed per each Lite DM workstation. The Lite DM application runs on a PC within a Windows Operating System platform. Data transmitted from the instrument to the Lite DM workstation is a Serial RS-232 unidirectional ABX Format communication. The Lite DM can act as a standalone data manager for data received from the instrument, or allow transmission of this data to a Host (LIS  Laboratory Information System or EMR  Electronic Medical Record system). Following the completion of each blood sample analysis performed on the connected HORIBA Medical hematology analyzer, results are then communicated (uni-directional) to the Lite DM workstation. Results received will display to the user in a specific color coding format indicating any normal or panic reference range violations. Results are then rejected or accepted by the user for final reporting - formatted to guidelines stipulated by CLIA. At no time are any results interpreted; only reported for review and appropriate action. | Lite DM (Lite Data Manager) | 2 | CDRH | 2014-12-23 | 2014-12-23 21:00 | 2015-06-22 | Horiba Instruments Inc | HORIBA Medical is informing all HORIBA Medical Lite DM customers using software version 2.0.0.19 and lower of a software bug related to the potential for results received being applied to the wrong order in the Lite DM Worklist. This bug has not been seen in software versions above V2.0.0.19 IMPACT: Potential for results received being applied to the wrong order in the Lite DM Worklist. The is
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=129827 | Z-2708-2014 | Dako Autostainer Link 48 with software version Dako Link 4.0.3, Model No. AS480 Product Usage: For in vitro diagnostic use. The Dako Autostainer Link 48 is an automated slide processing system compatible with reagents for the staining of paraffin-embedded and frozen tissue sections, cytospins, cell smears and fine needle aspirates. The system is designed to automate manual staining methods routinely used in immunohistochemistry and cytochemistry, enabling the transfer of established protocols from the bench to the Autostainer Link 48 automated workstations operating either singly or in networked configurations. | Dako Autostainer Link 48 software version 4.0.3 | 2 | CDRH | 2014-09-25 | 2014-09-26 21:00 | 2015-03-03 | Dako North America Inc. | Dako is recalling the Dako Link 4.0.3 software on the Dako Autostainer Link 48 because Dako has determined that an upgrade to this version of the software may cause premature rejection of reagent bottles or inadequate aspiration of reagent if the instrument has also had a preventative maintenance or waste pump replacement, and the baseplate is not completely aligned.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=128604 | Z-2109-2014 | QLAB 10 Core Module PN 453561704771. QLAB Quantification Software is a software application package designed to view and quantify image data acquired on Philips Medical Systems ultrasound products. | Philips Ultrasound QLAB | 2 | CDRH | 2014-07-23 | 2014-07-23 21:00 | 2015-11-06 | Philips Ultrasound, Inc. | The QLAB a2DQ and aCMQ features have a defect that may use incorrect measurement values under specific conditions.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=128605 | Z-2110-2014 | QLAB 10.1 Core Module PN 453561728001. QLAB Quantification Software is a software application package designed to view and quantify image data acquired on Philips Medical Systems ultrasound products. | Philips Ultrasound QLAB | 2 | CDRH | 2014-07-23 | 2014-07-23 21:00 | 2015-11-06 | Philips Ultrasound, Inc. | The QLAB a2DQ and aCMQ features have a defect that may use incorrect measurement values under specific conditions.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122026 | Z-0032-2014 | FSE Image Upgrade Kit, Catalog # 92019C This kit is used to install and/or upgrade software used with the EVOLIS Microplate System. | Image Upgrade Kit, Catalog # 92019C | 3 | CDRH | 2013-10-21 | 2013-10-21 21:01 | 2013-11-13 | Bio-Rad Laboratories | The FSE Image Upgrade kit (part 92019C) is recalled due to issue related to the export of data from the EVOLIS analyzer to Laboratory Information Systems (LIS) when using the MONOLISA Anti-HBs EIA Quantitative Determination in EVOLIS APF version 6.0.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121962 | Z-2318-2013 | McKesson Horizon Medical Imaging (HMI) Radiological Image Processing System, product versions 4.6.1 to 11.6., is a medical image and information management application intended to receive, transmit, store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from medical imaging systems. | McKesson Horizon Medical Imaging | 2 | CDRH | 2013-09-30 | 2013-09-30 21:00 | 2017-02-06 | Mckesson Information Solutions LLC | Software malfunction may occur resulting in a certain combination of server configurations which has the potential to cause image loss.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120033 | Z-1916-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 NO. Part number: 414565. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120031 | Z-1914-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 IT. Part number: 414563. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120032 | Z-1915-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 JA. Part number: 414564. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120037 | Z-1920-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 ZH. Part number: 414569. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120038 | Z-1921-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit V2S RP5800XPE TO 6.01 WES7. Part number: 6201773. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120028 | Z-1911-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 EN. Part number: 414560. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120027 | Z-1910-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 DE. Part number: 414539. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120025 | Z-1908-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01ML1. Part number: 414537. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120036 | Z-1919-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 EL, TR. Part number: 414568. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120026 | Z-1909-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01ML2. Part number: 414538. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120029 | Z-1912-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 ES. Part number: 414561. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120030 | Z-1913-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 FR. Part number: 414562. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120035 | Z-1918-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 DA, SV. Part number: 414567. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=119016 | Z-1907-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, V2S 6.01 RP5800-WES7 Media. Part number: 413861. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120034 | Z-1917-2013 | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, UPDT V2S 6.01 PT. Part number: 414566. The product is used with the VITEK 2 automated system which consists of instruments, software and reagent cards designed for the identification and antimicrobial susceptibility testing of bacteria and yeast. | bioMerieux Vitek 2 Systems Software version 6.01 software | 2 | CDRH | 2013-08-06 | 2013-08-06 21:00 | 2015-06-01 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=119832 | Z-1814-2013 | Picis ED Pulsecheck - EMR Software Application - 2125, Software Versions: 5.2 and 5.3. The application stores patient information in a database, and it may analyze and/or display the data in different formats for evaluation by healthcare professionals for informational purposes. | ED PulseCheck | 2 | CDRH | 2013-07-29 | 2013-07-29 21:00 | 2014-12-19 | Picis Inc. | Notes associated with prescription are not printed to the prescription or to the patient chart.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=118569 | Z-1560-2013 | Tecan Freedom EVO with Touch Tools Suite Version 3.0 The intended use of the Touch Tools Suite is to run previously defined scripts and processes on a Freedom EVO platform (Touch Tools Suite is an optional software add-on for Freedom EVOware (the base operational software for the Freedom EVO 2 instrument. | Freedom EVO 2 with TouchTools Suite Version 3.0 | 2 | CDRH | 2013-06-19 | 2013-06-19 21:00 | 2013-06-20 | Tecan US, Inc. | WIth TouchTools 3.0 after the execution of a Rich User Prompt in Display and Wait mode, door lock errors are ignored. If an operator presses "Continue" following a Rich User Prompt and the door has been left open, the script will progress (continue without warning). This could lead to the instrument moving unexpectedly and potentially cause an injury to the operator if they are physically inside t
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=117984 | Z-1419-2013 | RayStation Therapy Treatment Planning System, Model Numbers 2.0, 2.5, and 3.0; CLASSIFICATION NAME: System, Planning, Radiation Therapy Treatment. RayStation is a software system designed for treatment planning and analysis of radiation therapy. | RayStation Therapy Treatment Planning System | 2 | CDRH | 2013-05-28 | 2013-05-28 21:00 | 2017-02-17 | RAYSEARCH LABORATORIES AB | Under some circumstances, imported CT, MR and PET images are offset from their true positions by one pixel in one or two directions. Users must be aware of this information to avoid incorrect dose calculations during treatment planning.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=117878 | Z-1339-2013 | LATITUDE¿ Patient Management System, Model 6488, Version 7.3. The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. | LATITUDE¿ Patient Management System | 2 | CDRH | 2013-05-20 | 2013-05-20 21:00 | 2013-11-26 | Boston Scientific CRM Corp | Boston Scientific CRM is conducting a recall on the LATITUDE Patient Management System Model 6488 Version 7.3 because between March 4 and 27, 2013, a Siebel release caused ~5000 distributed PMR clinic faxes to display with illegible characters and not the required, Patient non-compliance information that was intended.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116845 | Z-1149-2013 | MAQUET CARDIOHELP Support System: CARDIOHELP-I Software (up to version 3.3.0.0) Product Usage: The Cardiohelp System is a blood oxygenation and carbon dioxide removal system used to pump blood through the extracorporeal bypass circuit for circulatory and/or pulmonary support during procedures requiring cardiopulmonary bypass (for periods up to six hours). It is also intended to provide circulatory and/or pulmonary support during procedures not requiring cardiopulmonary bypass (for periods up to six hours). | CARDIOHELP-i | 2 | CDRH | 2013-04-18 | 2013-04-18 21:00 | 2013-08-22 | Maquet Cardiovascular, LLC | It has come to the attention of MAQUET that there have been episodes of brief unexpected shutdown, followed by an automatic device restart of the human machine interface (touch screen) of the CARDIOHELP device. The human machine interface (HMI) is the central display of the CARDIOHELP device, where measured values, settings and alarms are displayed and adjusted/confirmed, with the exception of bl
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=112676 | Z-0123-2013 | UniCel¿ DxC 600/800 with Software v5.0.11, Part Number B12218. The UniCel¿ DxC Synchron¿ Clinical Systems are fully automated, computercontrolled clinical chemistry analyzers designed for the in vitro determination of a variety of general chemistries, therapeutic drugs, and other chemistries. Analysis can be performed on serum, plasma, urine, or cerebrospinal fluid (CSF) and whole blood (sample type is chemistry dependent). | UniCel¿ DxC 600/800 with | 2 | CDRH | 2012-10-24 | 2012-10-24 21:00 | 2014-02-14 | Beckman Coulter Inc. | Beckman Coulter has confirmed that the Shuttle to Barcode Reader and Cap Piercer alignment procedure cannot be performed on certain DxC configurations with v5.0.11 software installed. If the Shuttle to Barcode Reader becomes misaligned, the Operator will not be able to perform this alignment procedure on the DxC systems listed above. Software v5.0 .11 does not allow access to the Shuttle to Bar
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=112243 | Z-2404-2012 | The Brilliance CT 16 System Product Usage: The Brilliance 16 is a Whole Body Computed Tomography X-ray System intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from the same axial plane taken at different angles. | Brilliance 16 Slice (Air) | 3 | CDRH | 2012-09-18 | 2012-09-19 21:01 | 2013-12-24 | Philips Medical Systems (Cleveland) Inc | Philips was notified that the system logout in software version 2.3.6 is now longer and may sometimes fail to logout, requiring a forced system restart or hard shutdown with the power button.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=111260 | Z-2184-2012 | IMPAX Cardiovascular (CV) Reporting Reporting tool used for structured reporting and structured report outputs (including XML style sheet customization and PDF) | IMPAX Cardiovascular (CV) Reporting | 2 | CDRH | 2012-08-10 | 2012-08-10 21:00 | 2015-01-13 | AGFA Corp. | Content entered into the "Conclusions" free text box on the Report Writer screen was not represented on the final printed report.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=109499 | Z-1818-2012 | Philips Digital Diagnost Software, software release 2.0.2 and 2.0.2SP1 including mirror icon Stationary Fluoroscopic system | Philips Healthcare Imaging | 2 | CDRH | 2012-06-15 | 2012-06-15 21:00 | 2017-02-24 | Philips Healthcare Inc. | With Digital Diagnost (DIDI) R 2.0.2 a mirrored "R in a circle" is burnt into the image (if mirrored) in the lower right corner of the image. When such an image is mirrored back inside the PACS system, an unmirrored "R in a circle" appears in the lower left corner. This can be mistaken for a "right patient side" marker, although this can appear on the left patient side
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=109143 | Z-1685-2012 | Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial compounding feature Software Versions 3.0, 3.0A and 3.1. (VB20, VB20A, and VB21A); Model #: 10440017 - ACUSON S2000 Refurb. Mainframe; Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Mountain View, CA The ACUSON S2000 Ultrasound System is a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or Harmonic Imaging and 3D/4D Imaging on a Flat Panel Display. The ACUSON S2000 ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides the ability to measure anatomical structures {fetal, abdominal, intraoperative, intraoperative neurological, pediatric, small organ, neonatal cephalic, adult cephalic, cardiac, trans-esophageal, transrectal, transvaginal, peripheral vessel, musculoskeletal (conventional), musculo-skeletal (superficial) and neonatal cardiac} and calculation packages that provide information that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. | Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial | 2 | CDRH | 2012-05-30 | 2012-05-30 21:00 | 2013-07-16 | Siemens Medical Solutions USA, Inc. | When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=109142 | Z-1684-2012 | Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial compounding feature Software Versions 3.0, 3.0A and 3.1. (VB20, VB20A, and VB21A); Model #: 10041461 - ACUSON S2000 Mainframe; Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Mountain View, CA The ACUSON S2000 Ultrasound System is a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or Harmonic Imaging and 3D/4D Imaging on a Flat Panel Display. The ACUSON S2000 ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides the ability to measure anatomical structures {fetal, abdominal, intraoperative, intraoperative neurological, pediatric, small organ, neonatal cephalic, adult cephalic, cardiac, trans-esophageal, transrectal, transvaginal, peripheral vessel, musculoskeletal (conventional), musculo-skeletal (superficial) and neonatal cardiac} and calculation packages that provide information that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. | Siemens brand ACUSON S2000 Ultrasound Systemsand the Advanced SieClearTM spatial, | 2 | CDRH | 2012-05-30 | 2012-05-30 21:00 | 2013-07-16 | Siemens Medical Solutions USA, Inc. | When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=86301 | Z-1487-2012 | Insight rev H software. The InSIGHT Gastrointestinal Motility System is intended for use by gastroenterologists, surgeons, and medically trained personnel for esophageal, biliary, and anorectal studies as an aid in documenting and diagnosing digestive motility disorders. | Insight | 2 | CDRH | 2012-05-07 | 2012-05-07 21:00 | 2012-05-07 | Sandhill Scientific, Inc | Software error may produce incorrect esophageal contraction readings.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=107808 | Z-1474-2012 | Philips eCareManager System, Model 865325, software versions 3.7, 3.7.1 and 3.8. Product Usage: The Philips eCareManager System software is intended for use in data collection, storage, and clinical information management with independent bedside devices and ancillary systems that are connected either directly or through networks. The eCareManager system provides patient information and surveillance of monitored patients at the point of care location and at a remote, supplementary care location through wide area networking technology and dedicated phone lines and is solely intended for use in a hospital environment. It is not intended to be used in a home environment. | Philips eCareManager System | 2 | CDRH | 2012-04-27 | 2012-04-27 21:00 | 2012-07-06 | Visicu, Inc. | Three versions of software for the eCareManager failed to display anticoagulants on the medications screen. The problem arose due to an unexpected change in the National Drug Data File.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=108198 | Z-1454-2012 | Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4 The Dimension Vista(R) System is an in vitro diagnostic device intended to duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids. Dimension Vista(R) chemical and immunochemical applications use photometric, turbidimetric, chemiluminescence, nephelometric and integrated ion-selective multisensory technology for clinical use. | Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4 | 2 | CDRH | 2012-04-25 | 2012-04-25 21:01 | 2014-03-05 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed customer complaints on Vista(R) software version 3.4 for Below Manufacturer Assay Range error flags associated with QC and patient test results that are within the assay range as defined in the Instructions For Use for the respective method. This issue is isolated to Vista(R) software version 3.4. Customers who have not made any changes to the default method configurations
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=107919 | Z-1285-2012 | Whole body Computed Tomography X-Ray System (Brilliance 64 and Ingenuity CT), Brilliance 64 M/N 728231, Ingenuity CT M/N 728326, Dist. By Philips Medical Systems, Cleveland, OH | Brilliance 64 and Ingenuity CT | 2 | CDRH | 2012-03-23 | 2012-03-23 21:00 | 2013-02-22 | Philips Medical Systems (Cleveland) Inc | A number of non-conformances were identified during internal and external validation testing of the Brilliance 64 and Ingenuity CT systems with software version Beta 3 (version 4.0.0.xx368).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=107347 | Z-1134-2012 | Philips IntelliVue Patient Monitors: The monitor is intended to be used for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults, pediatrics, and neonates in health care facilities. The monitor is to be used by trained health care professionals. Physiological Monitoring, Patient Monitor , Arrhythmia detector and alarm (including ST -segment measurement and alarm. | Philips IntelliVue Patient Monitors | 2 | CDRH | 2012-03-01 | 2012-03-01 21:00 | 2017-02-15 | Philips Healthcare Inc. | Central station monitors may not alarm (either visual or audible) when an alarm is announced at the affected patient monitor
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=106994 | Z-0927-2012 | AXIOM Artis and AXIOM Artis zee modular angiographic systems Product Usage: Angiographic x-ray systems | AXIOM Artis and AXIOM Artis zee modular angiographic systems | 2 | CDRH | 2012-01-27 | 2012-01-27 21:00 | 2014-01-31 | Siemens Medical Solutions USA, Inc | There is a potential misalignment in the connection of the remote foot switch and the stationary unit. This potential malfunction may occur if either the wireless foot switch or the stationary unit is equipped with improper firmware installed in the bluetooth interface performing the communication between foot switch and system. This is not a general issue but sporadic as a vendor unsystematical
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=105388 | Z-0278-2012 | SCC Soft Computer, Softlab Laboratory information System with DORA versions: 2.0.0.4 with Hot Fix 1.2808, 1.2809, 1.2956, or 1.3476; 2.0.0.5; 2.0.0.6; 2.0.0.7 with Hot Fix 1.6595; and 2.1.0.3. | Softlab | 2 | CDRH | 2011-11-23 | 2011-11-23 21:00 | 2012-10-29 | SCC Soft Computer | Potential for patient data to be sent to the wrong patient record. A software change affected how the Softlab handles recovery from certain types of table access conflicts. When such conflicts occur, this function has the potential to cause faulty writing to the database.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98566 | Z-0146-2012 | Labeling for RIO Robot Unit is comprised of three main components: 20399 RIO SURGICAL ARM, 201251 RIO GUIDANCE MODULE, 200294 RIO CAMERA STAND ASSEMBLY***PN 203999 SN ROB 125 2010-12 V 100/120/230 A 9.6/8.0/4.2 Hz 50/60 Class I Equipment. Conforms to IEC 60601-1/A2: 1995, EN 60601-1/A2: 1995 UL 60601-1: 2003, CAN/CSA-C22.2 No. 601.1-M90.***Manufactured in USA***MAKO SURGICAL CORP. 2555 DAVIE ROAD, FT. LAUDERDALE, FL 33317. RIO System - The Tactile Guidance System v2.0 is intended to assist the surgeon in providing software defined spatial boundaries for orientation and reference information to anatomical structures during orthopedic procedures. The Tactile Guidance System v2.0 is indicated for use in surgical knee procedures, in which the use of stereotactic surgery may be approriate, and where reference to rigid anatomical bony structures can be identified relative to a CT base model of the anatomy. These procedures include unicondylar knee replacement and/or patellofemoral knee replacement. | RIO Robot Arm Interactive Orthopedic System | 2 | CDRH | 2011-11-01 | 2011-11-01 21:00 | 2011-12-13 | Mako Surgical Corporation | MAKO Surgical Corp. is recalling their RIO Robotic Arm Interactive Orthopedic System (RIO) due to software issue that exist that could potentially result in a bone resection. No adverse events reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=97478 | Z-3091-2011 | Software version 3.3 for Rapid Programmer devices, Model 3832. Software version 3.3 for Rapid Programmer devices, a personal digital assistant (PDA) based programmer supporting implantable pulse generators of the Eon, Genesis, and Renew family. | Rapid Programmer | 2 | CDRH | 2011-08-30 | 2011-08-30 21:00 | 2012-09-27 | Advanced Neuromodulation Systems | Product programmed with unapproved software was distributed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=102758 | Z-2964-2011 | Lenstar LS 900 EyeSuite Biometry version i2.000, Four (4) Protective Cardboard Containers Per Each System. Haag-Streit USA, Inc. A non-invasive, non-contact OLCR (Optical Low Coherence Reflectometer) Biometer used for obtaining ocular measurements and performing calculations to assist in determination of the appropriate power and type of IOL (Intraocular Lens) for implantation after removal of the natural crystalline lens following cataract removal. | Lenstar LS 900 EyeSuite Biometry version i2.000, | 2 | CDRH | 2011-08-09 | 2011-08-09 21:00 | 2013-03-08 | Haag-Streit USA Inc | A patient record may be linked to an incorrect/inconsistent set of calculation results under specific circumstances when using EyeSuite Software version i2.000.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=99926 | Z-2733-2011 | SJM Confirm, Implantable Cardiac Monitor (ICM) Models: DM2100, DM2102 (IDE). This product is a lead less non-therapeutic implantable cardiac monitor and is implanted to aid clinicians in diagnosing intermittent symptoms which may be due to arrhythmia | SJM Confirm, Implantable Cardiac Monitor (ICM) | 2 | CDRH | 2011-07-05 | 2011-07-05 21:00 | 2011-07-19 | St Jude Medical CRMD | The recall was initiated because St. Jude Medical made new software available for the SJM Confirm DM2100 implantable cardiac monitor (ICM) that incorporates enhancements to the devices sensing algorithm. These enhancements are intended to provide improved sensitivity and specificity for diagnosing arrhythmias. The user is presented with the option to upgrade new software into the SJM Confirm ICM
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98797 | Z-2063-2011 | Optical Guidance Platform, version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03899; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA Optical Guidance Platform is for use with a charged particle accelerator to perform precise positioning of treatment target for stereotactic radiosurgery or radiotherapy treatments on cranial or extracranial lesions. | Optical Guidance Platform | 2 | CDRH | 2011-04-25 | 2011-04-25 21:00 | 2012-06-19 | Varian Medical Systems, Inc. Oncology Systems | The optical guidance platform may not be properly enforcing a 24 hour time limit between optical camera recalibrations.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=80569 | Z-2000-2011 | Siemens LANTIS 6.1 Oncology Information System with MEDOne Module, Linear Medical Accelerators; Product Numbers: 4503178, 5835991, 5851055; Product is manufactured by IMPAC Medical, Inc., Sunnyvale, CA. Medical charged-particle radiation therapy system. | LANTIS 6.1 Oncology Information System with MEDOne Module | 2 | CDRH | 2011-04-15 | 2011-04-15 21:00 | 2011-04-15 | Siemens Medical Solutions USA, Inc | The product has a software problem; it does not calculate the Creatinine Clearance correctly for patients less than two years old.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=88565 | Z-1932-2011 | Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program, 1600 and 2000 tests Cartridge Sets. Catalog Numbers: 270-2417 and 270-2415; Distributed and Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CA The Intended Use is as follows: The Bior-Rad II TURBO Hemoglobin A1c Program is intended for the percent determination of hemoglobin Alc in human whole blood using ion-exchange high-performance liquid chromatography (HPLC). The Bio-Rad VARIANT II. TURBO Hemoglobin Alc Program is intended for Professional Use Only--For In Vitro Diagnostic Use. Reagent Kit Information: The VARIANT II TURBO Hemoglobin A1c program utilizes principles of ion¿exchange high performance liquid chromatography (HPLC). The VARIANT II Hemoglobin A1c Program is based on chromatographic separation of HbAlc on a cation exchange cartridge. System Overview: The VARIANT II TURBO Hemoglobin testing System provides an integrated method for sample preparation, separation and the percent determination of percent hemoglobin Alc in EDTA human whole blood. Technology Overview: The V ARIANT II TURBO Hemoglobin Testing is a fully automated, high-¿throughput hemoglobin analyzer. It consists of two modules- the VARIANT II Chromatographic Station (VCS) and the VARIANT II Sampling Station (VSS). In addition, a personal computer is used to control the VARIANT II TURBO System using Clinical Data Management (CDM) software. CDM is a dedicated software package of instrument control, data control, results analysis, calibration, quality control and service software. | Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program | 2 | CDRH | 2011-04-11 | 2011-04-11 21:00 | 2011-08-30 | Bio-Rad Laboratories Inc | Some customers have reported the occurrence of some ramping baselines on the chromatograms. The ramping baseline can affect hemoglobin test results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98675 | Z-1901-2011 | Prismaflex System, Catalog Number 107493, Indicated for continuous solute and/or fluid removal in patients with acute renal failure or fluid overload. | Prismaflex | 2 | CDRH | 2011-04-06 | 2011-04-06 21:00 | 2011-09-01 | Gambro Renal Products, Incorporated | Prismaflex touch screen may become unresponsive during Continuous Renal Replacement Therapy procedure.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=80551 | Z-1900-2011 | Siemens brand ACUSON S2000 Ultrasound System with software versions: VA10, VA10A, VA10B, VA10C, VA15; Model Number: 10041461; Product is manufactured and distributed by Siemens Medical Solutions USA, Inc. Ultrasound Business Unit, Mountain View, CA The ACUSON SC2000 ultrasound imaging system is intended for the following applications: Cardiac, Neo-natal and Fetal cardiac, Pediatric, Transesophageal, Adult Cephalic, Peripheral Vessel, Intraoperative Neurological, Musculo-skeletal Conventional, and Musculo-skeletal Superficial applications. | ACUSON S2000 Ultrasound System | 2 | CDRH | 2011-04-06 | 2011-04-06 21:00 | 2011-04-06 | Siemens Medical Solutions USA, Inc. | The product has a software problem in which previous patient measurement data gets associated with another patient's image.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94615 | Z-1837-2011 | ACUSON S2000 ultrasound systems (all versions) with cardiovascular options installed; ACUSON S2000: 10041461; S2000 Cardio Option: 10041853; Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA The ACUSON S2000 Ultrasound System is a multi-purpose mobile, software controlled diagnostic ultrasound system with and on-screen display for thermal and mechanical indices related to potential bio-effect mechanisms. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, or Harmonic Imaging and 3D/4D Imaging on a Flat Panel Display. The ACUSON S2000 ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides the ability to measure anatomical structures {fetal, abdominal, intraoperative, intraoperative neurological, pediatric, small organ, neonatal cephalic, adult cephalic, cardiac, trans-esophageal, transrectal, transvaginal, peripheral vessel, musculoskeletal (conventional), musculo-skeletal (superficial) and neonatal cardiac} and calculation packages that provide information that provide information to the clinician that may be used adjunctively with other medical data obtained by a physician for clinical diagnosis purposes. | ACUSON S2000 ultrasound systems | 2 | CDRH | 2011-03-28 | 2011-03-28 21:00 | 2012-06-18 | Siemens Medical Solutions USA, Inc., Mountain View, CA | There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94614 | Z-1836-2011 | ACUSON and SONOLINE Antares and Antares PE ultrasound systems at product versions 4.0 and 5.0 with cardiovascular options installed; ACUSON Antares 4.0 & 5.0: 10032747, 10037592, 10040728; ACUSON Antares PE 4.0 & 5.0: 10032746, 10037591, 10038202, 10040729; SONOLINE Antares 4.0 & 5.0: 05936518, 08653771, 10037593; Antares Cardio Options: 10036613, 10036588, 10036585, 10436250; Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA The ACUSON Antare is a general purpose Diagnostic Ultrasound system. Its function is to acquire primary or secondary harmonic ultrasound echo data and display it in: B-Mode, M-Mode, Pulsed (PW) Doppler Mode, Continuous (CW) Doppler Mode, Color Doppler Mode, Amplitude Doppler Mode, a combination of modes, Harmonic Imaging, or 3D/4D imaging on a FPD display. The Antares ultrasound imaging system is intended for the following applications: Abdominal, Intraoperative, Small Parts, Transcranial, OB/GYN, Cardiac, Transesophageal, Pelvic, Neonatal/Adult Cephalic, Vascular, Intravascular, Musculoskeletal, Superficial Musculoskeletal, Great Vessel, and Peripheral Vascular applications. The system also provides for the measurement of anatomical structures that provide information that is used for clinical diagnosis purposes. | ACUSON and SONOLINE Antares and Antares PE ultrasound systems | 2 | CDRH | 2011-03-28 | 2011-03-28 21:00 | 2012-06-18 | Siemens Medical Solutions USA, Inc., Mountain View, CA | There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94616 | Z-1838-2011 | ACUSON X300, X300 PE and SONOVISTA X300 (all versions) with cardiovascular options installed; ACUSON X300: 10037409, 10132987, 10038837; SONOVISTA X300: 10133170; ACUSON X300 PE: 10348531, 10348532, 10348533; X300 Cardiovascular Options: 10133037, 10348777, 10349032, 10427797, 10427822, 10427816, 10563788; Product is manufactured and distributed by Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA The Siemens Acuson X300 ultrasound imaging system is intended for the following applications: General Radiology, Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Neonatal/Adult Cephalic, Cardiac, Transesophageal, Pelvic, Transcranial, OB/GYN, Urology, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. The system also provides for the measurement of anatomical structures and for analysis packages that provide information that is used for clinical diagnosis purposes. | ACUSON X300, X300 PE and SONOVISTA X300 | 2 | CDRH | 2011-03-28 | 2011-03-28 21:00 | 2012-06-18 | Siemens Medical Solutions USA, Inc., Mountain View, CA | There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=86329 | Z-1495-2011 | PhD System EIA/IFAversion 2.1A software, model number 426-0246, a component of the Bio-Rad PhD System. Manufactured by Bio-Rad Laboratories, France, Distributed by Bio-Rad Laboratories, Hercules, CA. EIA and IFA processing system that consists of a PhD fluid handling station (s) linked to a network computer via ethernet hub. The computer provides worklist generation, data management, data reduction and microplate reader control functions. | PhD System EIA/IFAversion 2.1A software | 3 | CDRH | 2011-03-02 | 2011-03-02 21:00 | 2011-03-11 | Bio-Rad Laboratories Inc | Incorrect patient results are being generated from out-of-range samples tested withc ompetitive enzyme immunoassays on the PhD system with software version 2.1A.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94177 | Z-0133-2011 | Computed Tomography X-Ray System MX 8000 IDT 16; Model # 728120; Serial #'s: 2513, 2518, 3147, 3217, 3014, 3050, 3068, 3188, 3091, 3051, 3099, 3126, 3183, 3303, 2539 & 3056. Produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Computed Tomography X-Ray System MX 8000 IDT 16 | 2 | CDRH | 2010-10-22 | 2010-10-22 21:00 | 2012-09-21 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare has decided to recall the Computed Tomography X-Ray System MX 8000 IDT due to the fact that the software upgrade for the MX 8000 IDT 16 v4.2.1 may not provide the optional collimation of 4x4.5 (18mm). If the customer does not verify the increment and thickness coincide with the newly selected collimation suggested by the conversion report, there is potential for a 6mm gap to ap
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92698 | Z-2209-2010 | Medtronic Carelink Monitor Model 2490C. The Medtronic CareLink Monitor Model 2490C is designed to automatically gather information from the implanted heart device. This automatic wireless communication between the heart device and the monitor takes place at times scheduled by the doctor or clinic. The gathered heart device information is sent over a telephone line to the clinic using a toll-free number. The monitor can also send wireless heart device alerts to the clinic automatically. | Medtronic Carelink Monitor Model 2490C | 2 | CDRH | 2010-08-11 | 2010-08-11 21:00 | 2011-12-24 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic has identified that a subset of 2490C CareLink Monitors recently received an incorrect software update. Patients reported that their monitor was making a "ticking sound" and they were unable to transmit device data via CareLink. These monitors are now non-functional. No patient injuries have been reported as a result of this issue.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=91258 | Z-1974-2010 | Horizon Cardiology Hemo Component Schiller Argus Pro PB-1000 Monitor unit. Argus Pro PB 1000 firmware versions 2.41+M NIBP and 2.42.01+M. A programmable diagnostic computer intended for acquiring, digitizing, storing, displaying and reviewing hemodynamic data for use in hospital cardiac catheterization laboratories. The system is cleared for us with a component physiological data collection unit - the Schiller Argus Pro PB 1000 Monitor unit. | Horizon Cardiology Hemo Component Schiller Argus Pro PB-1000 Monitor unit | 2 | CDRH | 2010-07-09 | 2010-07-09 21:00 | 2010-07-21 | McKesson Provider Technologies - Medical Imaging Group | McKesson has recently discovered that certain Horizon Cardiology Hemo systems were running specific Schiller Argus-Pro PB 1000 firmware versions that might cause a delay in the display of physiological parameters on the Horizon Cariology Hemo Real Time Monitor Screen.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=88343 | Z-1322-2010 | Baxter 1550 Single Patient System Hemodialysis Machines; Baxter Healthcare Corporation, Deerfield, IL 60015; product codes 5M5538, 5M5538R, 5M5575, 5M5551 and 5M5551R. | Baxter 1550 Hemodialysis Machine | 2 | CDRH | 2010-04-12 | 2010-04-12 21:00 | 2011-02-03 | Baxter Healthcare Renal Div | Some of the 1550 Hemodialysis Instruments may have replacement circuit boards with incorrect versions of software.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=88119 | Z-1143-2010 | FHC microTargeting Guideline 4000 ver 1.4 (GL4K) Software, Catalog # MT-GL4K. Intended to assist in functional neurological procedures. | FHC microTargeting Guideline 4000 | 2 | CDRH | 2010-03-16 | 2010-03-16 21:00 | 2010-03-18 | FHC, Inc. | The raster tab for the GL4K version 1.4 software may assign incorrect depths for waveform recordings and if used to identify the implantation target for the DBS lead, the implant may be placed in the wrong location.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=81671 | Z-2325-2009 | ARCHITECT i System Assay CD ROM - US, list 6E58-22, Version 22.01, and ARCHITECT i System Assay CD-ROM - WW (excluding US), list 6E59-23, Version 23.0, containing the ARCHITECT LH assay file; Abbott Laboratories, Diagnostic Division, Abbott Park, IL 60064 USA; Product List Number: 6E58 (US) and 6E59 (WW) Versions 23 and lower. The Architect LH assay is a Chemiluminescent Microparticle Immunoassay for the quantitative determination of human luteinizing hormone (LH) in human serum or plasma. | ARCHITECT i System Assay CD ROM | 3 | CDRH | 2009-09-30 | 2009-09-30 21:00 | 2009-11-30 | Abbott Laboratories | There continues to be an elevated level of complaint activity for calibration error code 1227, when using ARCHITECT LH reagent lots (6C25-22 or 6C25-27) due to the use of assay disk versions that do not contain the updated ARCHITECT LH assay file contained on ARCHITECT i System Assay CD ROM versions 24 and higher (US LN 6E58, WW LN 6E59).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=83671 | Z-1808-2009 | American Optisurgical, Horizon Phacoemulsification System, Model: HZN-2030, for ophthalmic surgery. | Horizon | 2 | CDRH | 2009-09-03 | 2009-09-04 21:00 | 2009-09-04 | American Optisurgical Inc | Device operation interrupted: A software bug causes Phaco to stop when vacuum goes above 300 mmHg while in Multi Burst mode. Use of this product at the parameters specified above could lead to a delay in surgery that might present a health risk.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=81931 | Z-1478-2009 | Philips Avalon Fetal Monitor FM20; Model Number: M2702A. | Philips Avalon Fetal Monitor | 2 | CDRH | 2009-06-26 | 2009-06-26 21:00 | 2020-09-23 09:36 | Philips Healthcare Inc. | Interruption of power-up/boot sequence may cause the paper scale and recorder speed settings of the monitor to revert to factory default may lead to misinterpretation of the fetal trace by the user and to potential incorrect clinical decisions
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=81932 | Z-1479-2009 | Philips Avalon Fetal Monitor FM30; Model Number: M2703A. | Philips Avalon Fetal Monitor | 2 | CDRH | 2009-06-26 | 2009-06-26 21:00 | 2020-09-23 09:36 | Philips Healthcare Inc. | Interruption of power-up/boot sequence may cause the paper scale and recorder speed settings of the monitor to revert to factory default may lead to misinterpretation of the fetal trace by the user and to potential incorrect clinical decisions
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=74895 | Z-1067-2009 | Roche/Hitachi Modular Analytics System, clinical chemistry analyzer, Modular D Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 03284522001. Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. | Roche/Hitachi Modular Analytics System | 2 | CDRH | 2009-04-10 | 2009-04-10 21:00 | 2009-09-22 | Roche Diagnostics Corp. | If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=74894 | Z-1066-2009 | Roche/Hitachi Modular Analytics System, clinical chemistry analyzer, Modular P800 Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 04998669001. Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. | Roche/Hitachi Modular Analytics System | 2 | CDRH | 2009-04-10 | 2009-04-10 21:00 | 2009-09-22 | Roche Diagnostics Corp. | If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=74896 | Z-1068-2009 | Roche/Hitachi Modular Analytics System, clinical chemistry analyzer, Modular P Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 03284549001. Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. | Roche/Hitachi Modular Analytics System | 2 | CDRH | 2009-04-10 | 2009-04-10 21:00 | 2009-09-22 | Roche Diagnostics Corp. | If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=74893 | Z-1065-2009 | Roche/Hitachi Modular Analytics System, clinical chemistry analyzer Modular D2400 Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 04998618001. Is a fully automatic, computerized chemistry analyzer intended for in vitro quantitative and qualitative determinations of a wide range analytes. The instrument is capable of performing potentiometric and photometric assays. | Roche/Hitachi Modular Analytics System | 2 | CDRH | 2009-04-10 | 2009-04-10 21:00 | 2009-09-22 | Roche Diagnostics Corp. | If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=74885 | Z-1040-2009 | QCA and iQCA software, a video microscopy software system for quantitative estrogen receptor immunohistochemistry; Cell Analysis Inc., Evanston, IL 60201 or Highland Park, IL 60035. Intended to measure and quantitate the percentage and intensity of positively stained nuclei in formalin-fixed, paraffin-embedded tissue specimens immunohistochemically stained for estrogen receptors. Indicated for use as an aid in the management, prognosis and prediction of therapy outcomes of breast cancer when used with reagents validated for those indications. | QCA and iQCA Software | 2 | CDRH | 2009-03-24 | 2009-03-24 21:00 | 2011-03-16 | Leica Microsystems, Inc. | A change in the software to open up the software architecture to include other analytes other than Estrogen Receptor was made without FDA review and approval.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=73784 | Z-1090-2009 | BD PrepStain System, Model Number: 799-14000-00; Catalog Number: 05CR000021, 05CR00021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 799-14000-00R; Catalog Number: 05CR000021, 05CR000021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 7914000-00R. PrepStain is a liquid-based thin layer cell preparation device for preparation and staining of Prep smear slides from cervical specimens in a liquid based format (gyn use) and of cytology examination slides from non-cervical specimens such as urine and lung lavage. | BD PrepStain System | 2 | CDRH | 2009-03-19 | 2009-03-19 21:00 | 2013-01-15 | Bd Diagnostic Systems Tripath | Software Error in the non-gyn program of some devices may cause non-gyn samples to be incorrectly transferred from centrifuge tubes to slides potentially result in mismatched patient results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=71755 | Z-0297-2009 | AxSYM Digoxin II Reagent Pack, Disk Version 1. (3D53-01 or 3D53-02) The AxSYM Digoxin II assay is a Microparticle Enzyme Immunoassay (MEIA) for the quantitative measurement of digoxin, a cardiovascular drug, in serum or plasma. The measurements obtained are used in the treatment of digoxin overdose and monitoring levels of digoxin to ensure appropriate therapy. | AxSYM Digoxin II Reagent Pack. | 3 | CDRH | 2008-11-21 | 2008-11-21 21:00 | 2011-09-09 | Abbott Diagnostic International, Ltd. | AxSYM TDM/ Transplant Assay Disks 3D53-03 and higher do not contain the AxSYM Digoxin II Assay file. The instrument might require installation of the AxSYM Digoxin II assay file.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=72541 | Z-0116-2009 | Artiste MV system, Part number 8139789, medical linear accelerator equipped with COHERENCE therapist RTT4.1. Medical Linear accelerator for radiation therapy. The device utilizes a console system for selecting various treatment sequences and manages commands for all motorized movement. | Artiste MV system | 2 | CDRH | 2008-10-20 | 2008-10-20 21:00 | 2010-12-22 | Siemens Medical Solutions USA, Inc | Three issues: unexpected rotation of gantry, unexpected movement of table between beams during patient setup, and single exposure images will be overwritten by double exposure images.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=73614 | Z-0088-2009 | Envision E700 Low Airloss Therapy Surface, Made By Hill-Rom Manufacturing, Inc., Charleston, SC. The Envision E700 Low Airloss Therapy Surface helps prevent and treat stage III and stage IV pressure ulcers in patients who weigh between 70 lb and 400 lb and are between 4'11" and 6' 4" in height. | Envision E700 Low Airloss Therapy Surface | 2 | CDRH | 2008-10-07 | 2008-10-07 21:00 | 2009-05-21 | Hill-Rom Manufacturing, Inc. | A defect in the software of the device may not allow the patient bed exam alarm or the patient movement alarm to function correctly.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=71838 | Z-2203-2008 | 740 Series Multiparameter Monitor Software Version 3.0 Model Numbers: 740-1, 740M-1, 740-2T, 740M-2T, 740-2MS, 740M-2MS, 740-2NL, 740M-2NL, 740-2NN, 740M-2NN, 740-3MS, 740M-3MS, 740-3NL, 740M-3NL, 740-3NN, and 740M-3NN CAS Medical Systems, Inc., Branford, Connecticut 06405 Indicated for use for non-invasive monitoring of blood pressure, oxygen saturation, pulse and temperature of adults, pediatric and neonatal patients, in the care of health professionals. | 740 Series Multiparameter Monitor | 2 | CDRH | 2008-09-11 | 2008-09-16 05:16 | 2010-08-05 | CAS Medical Systems, Inc. | Software: Audible alarm may be silenced if changed from the factory setting
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=71878 | Z-2225-2008 | GE 1.5T and 3.0T Signa HDX MR System Model Number(s): 2226300, 2377062-61, 2395001, 2395001-2, 5107849, and 5127452. This system is indicated for use as a diagnostic imaging device to produce axial sagittula, coronal and oblique images, spectroscopic images, and or spectra, dynamic images of the internal structure or organs of the entire body, | Signa HDX MR System | 2 | CDRH | 2008-09-09 | 2008-09-11 07:31 | 2011-12-17 | GE Healthcare | Possible misdiagnosis or inappropriate treatment based on the image. Signal homogeneity may be affected when BRAVO applications are set with IR Prep off. The concern is that signal inhomogeneity can result in an erroneous diagnosis.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=67341 | Z-1243-2008 | IMPAX¿ 4.1, IMPAX ¿ 4.5, IMPAX ¿ 5.0, IMPAX ¿ 5.1 IMPAX ¿ 5.2, IMPAX ¿ 6.0, IMPAX ¿ 6.2; Picture Archiving and Communications systems (PACs), used in the acceptance, transfer, display, storage and digital processing of medical images. | IMPAX | 2 | CDRH | 2008-09-08 | 2008-09-11 07:31 | 2010-12-09 | AGFA Corp. | Java version 1.6x (JAVA 6) is not compatible with AGFA IMPAX Administration/Service Tools. Due to a Software Mismatch data can be lost.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=67348 | Z-1244-2008 | WEB1000 Versions 3.1, 4.1 and 5.1, Picture Archiving and Communications systems (PACs), used in the acceptance, transfer, display, storage and digital processing of medical images. | WEB1000 | 2 | CDRH | 2008-09-08 | 2008-09-11 07:31 | 2010-12-09 | AGFA Corp. | Java version 1.6x (JAVA 6) is not compatible with AGFA IMPAX Administration/Service Tools. Due to a Software Mismatch data can be lost.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=70841 | Z-1827-2008 | MIM 4.0 and MIM 4.1 is a software package designed for use in diagnostic imaging. It is a stand alone software package which operates on Windows 2000/XP. | MM 4.0 and MIM 4.1 | 2 | CDRH | 2008-09-02 | 2008-09-03 00:30 | 2015-04-14 | MIMvista Corp | In MIM (4.0.2-4.0.5) and MIM (4.1.0-4.1.2) versions, there is an error that may occur during the calculation of statistics using the SUV Tool or Contour Statistics. It will occur only under the following specific circumstance: A fusion of a primary volume which has a Z-Axis slice separation larger than that of the secondary volume and statistics are reported on the secondary volume.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=70261 | Z-2139-2008 | CELL-DYN Sapphire Hematology Analyzer, List number 08H00-01, LIS/ Middleware Interface Specification, Abbott Diagnostics Division, Santa Clara, CA For in vitro diagnostic use in counting and characterizing blood cells. | CELL-DYN Sapphire Hematology Analzyer | 2 | CDRH | 2008-08-20 | 2008-08-21 00:30 | 2010-03-12 | Abbott Laboratories | Table in LIS Specification depicting the association between Record ID and Numerical Result Label is incorrect. If a record ID is used to configure the system for mapping, results will come out nonsensical.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=64472 | Z-1155-2008 | GE LOGIQ 3 Expert ultrasound scanner with software versions 4.0.5, 4.1.1 and 4.1.2. GE Healthcare P.O. Box 414, Milwaukee, Wisconsin 53201 | GE LOGIQ 3 Expert | 2 | CDRH | 2008-06-11 | 2008-06-12 00:30 | 2012-05-29 | GE Healthcare Clinical Systems | Inaccurate guidelines: The BE9C is an optional, small radius Bi-planar convex probe intended for use with the LOGIQ 3 Expert system in Urology application. An attempt by the user to activate the biopsy guidelines while this probe is in use will cause incorrect guidelines to be displayed on the image. The incorrect guidelines are recognized by their orientation that is left/right reversed from t
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=68116 | Z-1462-2008 | AxSYM Drugs of Abuse/Toxicology Assay Disk Version 8.0; List Number: 3D54-08; Abbott La boratories, Diagnostics Division, Abbott Park, IL 60064 | AxSYM Drugs of Abuse/Toxicology Assay Disk Version 8.0 | 3 | CDRH | 2008-05-13 | 2008-05-14 00:30 | 2009-12-01 | Abbott Laboratories | Error Message: If the Positive and Negative Interpretation Cutoff parameters (116/117) are edited after installation of the revised assay file, VRTX error #0002 in Task 40 is generated and the instrument locks up when it attempts to report an AxSYM Amphetamine/Methamphetamine II patient result.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=66398 | Z-0526-2008 | BD FACS Sample Prep Assistant II, Part Number: 337170; Product is manufactured and distributed by BD Biosciences, San Jose, CA 95131 | BD FACS Sample Prep Assistant II | 2 | CDRH | 2008-04-09 | 2008-04-10 00:30 | 2010-12-13 | BD Biosciences | Mixing Performance: SPA II devices manufactured with the affected product motor control board may exhibit suboptimal mixing performance.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=66172 | Z-0571-2008 | Ventana Image Analysis System (VIAS) - PATHWAY anti-HER-2/neu (4B5) primary antibody with software version 2.3.0.2, Model Number: VIAS-799-10000-00, Pathology Image Analysis System, Ventana Medical Systems, Inc., Tucson, AZ 85755 | Ventana Image Analysis System (VIAS) | 2 | CDRH | 2008-03-17 | 2008-03-18 00:30 | 2012-03-26 | Ventana Medical Systems Inc | Incomplete upgrade: The software upgrade from February 1, 2007 may not have been completed. This software upgrade adds the algorithm (Slidetype) for the PATHWAY anti-HER-2/neu (4B5) antibody to the VIAS system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=64414 | Z-0124-2008 | Philips PageWriter Touch Cardiograph, Model: 860284, Philips Medical Systems, Andover, MA, 01810 | Philips PageWriter Touch | 2 | CDRH | 2008-01-29 | 2008-01-30 00:30 | 2016-08-16 | Philips Medical Systems | Incorrect display: When using the 16 lead Patient Interface Module (PIM), lead tracings of V3R and V4R are reversed in the extended lead modes of Pediatric and Balanced.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=66252 | Z-0451-2008 | Philips Medical Systems, M3290A IntelliVue Information Center Software for M3170 Intellivue Patient Link, Catalog Number: 865007 and 865015; Running Software Versions J.00.23, J.00.24, J.00.25, and J.00.26, Manufactured and Distributed by; Philips Medical Systems, Andover, MA 01810 | Philips Medical Systems IntelliVue Information Center | 2 | CDRH | 2007-12-10 | 2007-12-11 00:30 | 2016-08-15 | Philips Medical Systems | Alarm failure : If changes are made to the Caregroup Editor configuration, Care Group alarm status, including alarm text and prompt tones, may fail to be displayed or annunciated at bedside monitors within the Care Group
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=65617 | Z-0294-2008 | Medtronic 0-arm Imaging System, Product Catalog Number: B1-700-00027, Medtronic Navigation, Inc., Littleton, MA 01460 (software version 3.0) | Medtronic 0-arm Imaging System | 2 | CDRH | 2007-11-26 | 2007-11-27 00:30 | 2007-11-27 | Medtronic Navigation, Inc | Reversed Instrument Position : a defect in the software version 3.0 causes an error in the position of surgical instruments and relative motion are reversed 180' on the StealthStation image. (If the StealthStation is not used, or if the patient orientation is prone or supine, the problem does not occur)
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=64656 | Z-0146-2008 | ARCHITECT c8000 Processing Module, for In Vitro Diagnostics; List Number 106-01; Affected software: v2.10, v2.11, v2.12, v2.20, v2.20DB, v2.30, v2.60, v3.00 and v3.10; Distributed by Abbott Laboratories, Irving, TX. | ARCHITECT cSystem | 2 | CDRH | 2007-11-06 | 2007-11-07 00:30 | 2010-10-04 | Abbott Laboratories, Inc | Incorrect Sample/Test Identification: Due to a software timing defect with certain versions of ARCHITECT cSystem Software, test results may be incorrectly associated to a Sample Identification during specific sample-processing conditions of the Robotic Sample Handler.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=64657 | Z-0147-2008 | ARCHITECT c16000 Processing Module, for In Vitro Diagnostics; List Number 3L77-01; Affected software: v2.10, v2.11, v2.12, v2.20, v2.20DB, v2.30, v2.60, v3.00 and v3.10; Distributed by Abbott Laboratories, Irving, TX. | ARCHITECT cSystem | 2 | CDRH | 2007-11-06 | 2007-11-07 00:30 | 2010-10-04 | Abbott Laboratories, Inc | Incorrect Sample/Test Identification: Due to a software timing defect with certain versions of ARCHITECT cSystem Software, test results may be incorrectly associated to a Sample Identification during specific sample-processing conditions of the Robotic Sample Handler.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=49265 | Z-0223-2007 | Roche Modular Drive (rchmdlri), analytical laboratory data interface. | Roche Modular Drive | 2 | CDRH | 2006-11-28 09:59 | 2006-11-29 00:30 | 2015-06-11 | Data Innovations, Inc. | Software of modular driver may incorrectly report patient results as Quality Control Results
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=49266 | Z-0224-2007 | Roche Cobas driver (rchcob6i), analytical laboratory data interface. | Roche Cobas driver | 2 | CDRH | 2006-11-28 09:59 | 2006-11-29 00:30 | 2015-06-11 | Data Innovations, Inc. | Software of modular driver may incorrectly report patient results as Quality Control Results
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=43332 | Z-0380-06 | Cytomics FC500 with CXP Software | Cytomics FC500 with CXP Software | 2 | CDRH | 2006-01-10 12:05 | 2006-01-11 00:30 | 2012-01-26 | Beckman Coulter Inc | Under certain conditions an incorrect Tube ID (Barcode) can be displayed and printed on the Runtime Panel Report in CXP Cytometer Software Version 2.0