## Sheet1
WEB_ADDRESS | RECALL_NUMBER | PRODUCT_DESCRIPTION | TRADE_NAME | RECALL_CLASS | CENTER | CENTER_CLASSIFICATION_DT | POSTED_INTERNET_DT | TERMINATION_DT | FIRM_NAME | MANUFACTURER_RECALL_REASON
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210459 | Z-0318-2025 | GE Healthcare Centricity Universal Viewer Zero Footprint, Model Numbers: 1) 2089507-083, GDZH-JINWAN-Centricity 2) 2089507-093, 10510-3-Centricity Universal Viewer Zero Footprint / ZFP-02268286, 10510-3-Centricity Universal Viewer Zero Footprint / ZFP-02030900, 1495-1-Centricity Universal Viewer Zero Footprint / ZFP-02347996, 423834-Centricity Universal Viewer Zero Footprint / ZFP-02286658, 1407636-Centricity Universal Viewer Zero Footprint / ZFP-02304121, 651240-Centricity Universal Viewer Zero Footprint / ZFP-02173181, 268034-Centricity Universal Viewer Zero Footprint / ZFP-02402520; 3) 2089507-098, JP-310027-Centricity Universal Viewer Zero Footprint / ZFP-02033125; 4) 2089507-109, 3909-2-Centricity Universal Viewer Zero Footprint / ZFP-20095312, 3909-2-Centricity Universal Viewer Zero Footprint / ZFP-20095313, 327750-Centricity Universal Viewer Zero Footprint / ZFP-02184508; 5) 2089507-111, 105468-1-Centricity Universal Viewer Zero Footprint / ZFP-02254840, 1495-3-Centricity Universal Viewer Zero Footprint / ZFP-00096090, JP-340111-Centricity Universal Viewer Zero Footprint / ZFP-01275379, JP-275353-Centricity Universal Viewer Zero Footprint / ZFP-02042636; 6) 2089507-112, 3922-25-Centricity Universal Viewer Zero Footprint / ZFP-01290685, 3922-17-Centricity Universal Viewer Zero Footprint / ZFP-00098229, 1503586-Centricity Universal Viewer Zero Footprint / ZFP-21016603, RU0994-Centricity Universal Viewer Zero Footprint / ZFP-02331255; 7) 2089507-114, 1461-1-Centricity Universal Viewer Zero Footprint / ZFP-02360229, 1461-1-Centricity Universal Viewer Zero Footprint / ZFP-01803680, 3728-2-Centricity Universal Viewer Zero Footprint / ZFP-01225897, 3728-2-Centricity Universal Viewer Zero Footprint / ZFP-01225898, 3728-2-Centricity Universal Viewer Zero Footprint / ZFP-01225904, 3728-2-Centricity Universal Viewer Zero Footprint / ZFP-01371500, M51725-Centricity Universal Viewer Zero Footprint / ZFP-01300873, M1781854-Centricity Universal Viewer Zero Footprint / ZFP-01553772, 609334-Centricity Universal Viewer Zero Footprint / ZFP-01282257, UJJF01-Centricity Universal Viewer Zero Footprint / ZFP-22186138, CMC-Centricity Universal Viewer Zero Footprint / ZFP-01820241, JP-100014-Centricity Universal Viewer Zero Footprint / ZFP-01890709, KW1102-Centricity Universal Viewer Zero Footprint / ZFP-02191178, 59219_ERPerror-Centricity Universal Viewer Zero Footprint / ZFP-02327073, QA1114-Centricity Universal Viewer Zero Footprint / ZFP-02306397, SAU1433-Centricity Universal Viewer Zero Footprint / ZFP-01457278, TTSH-CVIS-Centricity Universal Viewer Zero Footprint / ZFP-02041523, NE4651-Centricity Universal Viewer Zero Footprint / ZFP-00038599, NE6703-Centricity Universal Viewer Zero Footprint / ZFP-00038552, NE6731-Centricity Universal Viewer Zero Footprint / ZFP-00038461, NE3006-Centricity Universal Viewer Zero Footprint / ZFP-00038362, NE2196-Centricity Universal Viewer Zero Footprint / ZFP-00038588, NPV738-Centricity Universal Viewer Zero Footprint / ZFP-00038637, NE2031-Centricity Universal Viewer Zero Footprint / ZFP-00038578, NE7405-Centricity Universal Viewer Zero Footprint / ZFP-00038372, NE2162-Centricity Universal Viewer Zero Footprint / ZFP-00038563, NE4878- | GE Healthcare Centricity | 2 | CDRH | 2024-11-07 | 2024-11-07 20:06 |  | GE Healthcare | GE HealthCare has become aware of an issue in Centricity Universal Viewer Zero Footprint Client (ZFP) where there is a potential security vulnerability which could allow a malicious actor to access the system and potentially manipulate patient data.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209481 | Z-3065-2024 | Dimension Vista¿ Total Bilirubin (TBIL) Flex Reagent Cartridge; Siemens Material Number (SMN): 10445146; | Dimension Vista¿ Total Bilirubin (TBIL) Flex Reagent Cartridge | 2 | CDRH | 2024-09-06 | 2024-09-06 21:29 |  | Siemens Healthcare Diagnostics, Inc. | During manufacturing, the last three tests from the 160 test flex of lot 23206BA (wells 8 or 10) were processed on the Dimension Vista platform. There was an issue with the filling line which led to some wells to underfill in lot 23206BA. Calibrator, quality control (QC) and patient results could have been impacted.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=206603 | Z-2419-2024 | CADD-Solis Ambulatory Infusion Pumps, Model 2100. Intended for the following uses: Indicated for the following uses: Intra-arterial, subcutaneous, intraperitoneal, in close proximity to nerves, into an intraoperative site, epidural space or subarachnoid space. Therapies that require a continuous rate of infusion, intermittent bolus, and/or with patient-controlled demand doses. | smiths medical CADD Solis | 1 | CDRH | 2024-08-02 11:16 | 2024-08-03 21:08 |  | Smiths Medical ASD Inc. | CADD-Solis Ambulatory Infusion Pumps, with software versions before v4.3, may have the following issues:1. Upstream Occlusion, 2. Stop and Power Keys Unresponsive, 3. Manual Mode Air Detector, 4. Single Bubble Air Detection, 5. Error Codes Not Displayed at Power Up, 6. Audible Alarm, 7. Low Sensitivity Air in Line Detection Threshold, 8. PharmGuard Server Password. Smiths Medical corrected many of the issues included in this notification in previous software updates and the corrections were carried forward into all subsequent software releases. Please ensure you have the most recent CADD software (v4.3 - 2023) installed on your pumps.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=206607 | Z-2420-2024 | CADD-Solis VIP ambulatory Infusion Pumps, Model 2110. Indicated for the following uses: Intra-arterial, subcutaneous, intraperitoneal, in close proximity to nerves, into an intraoperative site, epidural space or subarachnoid space. Therapies that require a continuous rate of infusion, intermittent bolus, and/or with patient-controlled demand doses. | smiths medical CADD Solis | 1 | CDRH | 2024-08-02 11:16 | 2024-08-03 21:08 |  | Smiths Medical ASD Inc. | The CADD Solis VIP Ambulatory Infusion Pump is indicated for the following uses: " For intravenous, intraarterial, subcutaneous, intraperitoneal, perineural, surgical site, epidural space, or subarachnoid space infusion.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=208401 | Z-2421-2024 | CADD-Solis VIP ambulatory Infusion Pumps, Model 2120 ( 21-2120 , 21-2125 , and 21-2127 ) Indicated for the following uses: Intra-arterial, subcutaneous, intraperitoneal, in close proximity to nerves, into an intraoperative site, epidural space or subarachnoid space. Therapies that require a continuous rate of infusion, intermittent bolus, and/or with patient-controlled demand doses. | smiths medical CADD Solis | 1 | CDRH | 2024-08-02 11:16 | 2024-08-03 21:08 |  | Smiths Medical ASD Inc. | The CADD Solis VIP Ambulatory Infusion Pump is indicated for the following uses: " For intravenous, intraarterial, subcutaneous, intraperitoneal, perineural, surgical site, epidural space, or subarachnoid space infusion.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=208571 | Z-2418-2024 | Arial Call Station, Model Numbers CSK200-1069 and CSK200-1069MR, is a wireless nurse call station. | Securitas | 2 | CDRH | 2024-07-23 17:15 | 2024-07-23 21:09 |  | Securitas Healthcare LLC | This is a risk that alarms will not be received by the Arial Wireless Nurse Call System and premature depletion of the battery.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=208563 | Z-2417-2024 | Arial Call Station, Model Numbers CSK200 and CSK200MR, is a wireless, battery-operated call device installed inside resident living areas to allow immediate staff response. | Securitas | 2 | CDRH | 2024-07-23 17:11 | 2024-07-23 21:08 |  | Securitas Healthcare LLC | There is a risk that alarms will not be received by the Arial Wireless Emergency Call System and premature depletion of the battery.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=207644 | Z-2021-2024 | Brand Name: GlideScope Core 15-inch Monitor Model/Catalog Number: 0570-0404 Product Description: Serialized Monitor | GlideScope Core 15-inch Monitor | 2 | CDRH | 2024-06-04 16:21 | 2024-06-04 21:07 |  | Verathon, Inc. | Due to software issues, their is the potential for loss of image or a degradation of the image when using the monitor with other connectable devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=207645 | Z-2022-2024 | Brand Name: GlideScope Core 15-inch FHD Model/Catalog Number: 0570-0437 Software Version: Core 15 FHD v1.7 and earlier Product Description: Serialized Monitor | GlideScope Core 15-inch FDH Monitor | 2 | CDRH | 2024-06-04 16:21 | 2024-06-04 21:07 |  | Verathon, Inc. | Due to software issues, their is the potential for loss of image or a degradation of the image when using the monitor with other connectable devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=207243 | Z-1885-2024 | 8000-021-002, Scopis¿ ENT Software with TGS¿ with Sofware Version: NOVA 3.6.0 RC16. Scopis is a next-generation solution for navigating functional endoscopic sinus surgery (FESS) offering surgeons highly advanced image guidance and visualization capabilities in a single system. The Stryker Electromagnetic Navigation Unit (8000-010-003) powered by Scopis ENT Software with TGS (8000-021-002), allows the analysis and identification of sinus cells in the complex patient anatomy and planning of the natural drainage pathways through the sinus cavity. During surgery, the planned pathways are overlaid in real-time onto the endoscopic image, providing a unique Scopis augmented reality technology. Guidance of endoscopic instruments may help perform a minimally invasive, accurate and selective surgery. | Scopis ENT Software | 2 | CDRH | 2024-05-23 | 2024-05-23 21:07 |  | Stryker Leibinger GmbH & Co. KG | Unreleased software was installed on customer systems resulting in the visual feedback on the screen to show the points to be off from the physical reference point of the pointer or suction.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204804 | Z-0631-2024 | Covidien Valleylab FT10, FT Series Energy Platform, REF: VLFT10GEN, DLVLFT10GEN, VLFT10GENZD | Covidien Valleylab FT10, FT Series Energy Platform | 2 | CDRH | 2024-01-03 | 2024-01-03 20:06 |  | Covidien | Due to software issue, New systems may exhibit an error which may prohibit use of new systems resulting in system being inoperable
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203871 | Z-0354-2024 | Model 97745 Controller used with Models 97715 and 97716 Intellis Neurostimulators | Medtronic | 2 | CDRH | 2023-11-20 | 2023-11-20 20:06 |  | Medtronic Neuromodulation | Units distributed in Korea and Turkey that did not have the correct firmware installed to support the user interface in the local language.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203256 | Z-0036-2024 | Meera Mobile Operating Table-to support and position the patient immediately before, during and after surgical interventions as well as for examination and treatment. Model Number: 7200.01F2 | Meera Mobile Operating Table | 2 | CDRH | 2023-10-04 | 2023-10-04 21:06 |  | Getinge Usa Sales Inc | Under certain conditions, an issue might prevent the device from performing as intended. A specific sequence of commands on the control unit activates the traction drive and triggers an unintended driving (autodrive) of max. 7 sec. may result in the following reasonably foreseeable injury or damage to the health of patients and/or users
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=201276 | Z-2235-2023 | vRad PACS with Mammography | vRad | 2 | CDRH | 2023-07-21 | 2023-07-21 21:07 |  | Virtual Radiologic Corp. | The error resulted in intermittent failure of current (primary) radiology study images not displaying to the interpreting teleradiologist.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=200685 | Z-2044-2023 | RelieVR REF: RVX-2002, Rx Only. Virtual reality behavioral therapy device for pain relief. | RelievRx | 2 | CDRH | 2023-06-29 23:01 | 2023-06-30 21:07 |  | Appliedvr | There is the potential for the program software to malfunction which will not allow it to move forward to the next session.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199957 | Z-1977-2023 | MiSeq Dx Instrument, Catalog # M70101, containing Universal Copy Service (UCS) software. | MiSeq Dx Instrument | 2 | CDRH | 2023-06-16 | 2023-06-16 21:07 |  | Illumina, Inc. | Cybersecurity vulnerability concerning the software used for sequencing instruments.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199954 | Z-1976-2023 | NextSeq 550Dx Instrument, Catalog # 20005715, contaning Universal Copy Service software | NextSeq 550Dx Instrument | 2 | CDRH | 2023-06-16 | 2023-06-16 21:07 |  | Illumina, Inc. | Cybersecurity vulnerability concerning the software used for sequencing instruments.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199058 | Z-1352-2023 | Flex-Neck Catheter External Repair Kit REF CE-1400 LOT CE-1400/A | Flex-Neck Catheter External Repair Kit | 2 | CDRH | 2023-04-07 | 2023-04-07 21:06 |  | Merit Medical Systems, Inc. | Product that was built for design verification testing was inadvertently distributed to customers.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=198402 | Z-1216-2023 | LATITUDE NXT Remote Patient Management System, LATITUDE NXT System Server Software Model 6460. | LATITUDE" NXT Remote Patient Management System | 2 | CDRH | 2023-03-03 | 2023-03-03 20:07 |  | Boston Scientific Corporation | Under specific circumstances, the U.S. product registration system did not send up enablement requests to U.S. LATITUDE Customer Support during an 8-month interval in 2022. For these patients, the HeartLogic Index and any associated Yellow Alerts are currently not visible in LATITUDE as may have been intended by the healthcare provider..
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197059 | Z-0722-2023 | Centricity Universal Viewer 6.0. Used to display medical images (Including mammograms) and data from various imaging sources. | Centricity Universal Viewer 6.0 | 2 | CDRH | 2022-12-16 13:16 | 2022-12-16 20:06 |  | GE Medical Systems, LLC | There is the potential that after installation of the IAS tool, some studies can remain hidden without user notification. Access to those hidden studies requires GE Healthcare assistance, which can result in a delay in reading the study.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197056 | Z-0721-2023 | Centricity PACS-IW with Universal Viewer version 5.0. Used to display medical images (Including mammograms) and data from various imaging sources. | Centricity PACS-IW with Universal Viewer | 2 | CDRH | 2022-12-16 13:16 | 2022-12-16 20:06 |  | GE Medical Systems, LLC | There is the potential that after installation of the IAS tool, some studies can remain hidden without user notification. Access to those hidden studies requires GE Healthcare assistance, which can result in a delay in reading the study.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194455 | Z-1392-2022 | GreenLight HPSEA Laser Fibers, UPN 0010-2092 | GreenLight HPSEA Laser Fibers | 3 | CDRH | 2022-07-15 15:32 |  |  | Boston Scientific Corporation | Several complaints were received for the affected lot; users received a "Fiber type does not match card type" error message on the GreenLight XPS Laser console, necessitating the replacement of the laser fiber and fiber card. This may result in prolonged procedure while the fiber is being replaced.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192653 | Z-0919-2022 | MUSE Cardiology Information System, Model Numbers: a) 3101011-040, b) 3101010-401, c) 3101010-009, d) 3101010-008 | MUSE Cardiology Information System | 2 | CDRH | 2022-04-13 10:41 |  |  | GE Healthcare, LLC | Two scenarios may cause edits to measurements and diagnosis statements can be lost after a test is Signed in the MUSE NX web client.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192037 | Z-0741-2022 | Del(5q) Deletion FISH Probe Kit: In-Vitro Diagnostic to detect deletions in chromosome location from patients with AML or MDS Catalog Number: USA-LPH024 | Cytocell | 2 | CDRH | 2022-03-09 |  | 2024-11-08 10:24 | Cytocell Ltd. | individual components have been labelled with incorrect colours. The red and green colours are opposite to those specified in the product labelling. The EGR1 probe is labelled in green rather than red and the control probe labelled in red rather than green. The expected positive signal pattern in the case of 1x deleted 5q will be 1G2R instead of 1R2G, an incorrect result
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189276 | Z-0073-2022 | AngioDynamics NanoKnife Disposable Single Electrode Probes, Irreversible Electroporation (IRE) Device, 15 cm IRE Single Electrode RFID Activation, PG. Catalog No. 20400103 | AngioDynamics NanoKnife Disposable Single Electrode Probes | 2 | CDRH | 2021-10-07 |  | 2022-09-20 12:45 | Angiodynamics, Inc. | Programming issue affected RFID function of a single lot and did not allow the NanoKnife probes to be recognized by the NanoKnife generator.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188131 | Z-2253-2021 | Simpleware ScanIP software. For transfer of imaging information from a medical scanner to an output file, and for simulating/evaluating surgical treatment options. | Simpleware ScanIP | 2 | CDRH | 2021-08-12 |  | 2022-09-29 12:32 | Simpleware Product Group, SYNOPSYS NORTHERN EUROPE | A issue (bug) has been identified with the interface and image software which could result in anatomical orientation tags/labels being displayed incorrectly. This could result in operator misunderstanding and errors during resampling or re-alignment of image data
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188132 | Z-2254-2021 | Simpleware ScanIP medical software. For transfer of imaging information from a medical scanner to an output file, and for simulating/evaluating surgical treatment options. | Simpleware ScanIP Medical; Simpleware ScanIP | 2 | CDRH | 2021-08-12 |  | 2022-09-29 12:32 | Simpleware Product Group, SYNOPSYS NORTHERN EUROPE | A issue (bug) has been identified with the interface and image software which could result in anatomical orientation tags/labels being displayed incorrectly. This could result in operator misunderstanding and errors during resampling or re-alignment of image data
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187878 | Z-2139-2021 | CARESCAPE PDM-Masimo SpO2. Physiological data monitor. | CARESCAPE PDM-Masimo SpO2 | 2 | CDRH | 2021-07-23 15:12 | 2021-07-23 21:07 |  | GE Healthcare, LLC | Masimo SpO2 Saturation Values can become frozen after an extended length of use without a power down.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187397 | Z-1853-2021 | v.1.2 Upgrade Kit or v.1.2 upgrade installed on WorkMate Claris System, Model H700123. The WorkMate Claris" System is a computer-based electrocardiographic recording and monitoring system designed for efficient capture, display, and retrieval of surface and intracardiac signals during cardiac EP studies. | WorkMate Claris System, Model H700123 | 2 | CDRH | 2021-06-14 |  | 2024-06-18 11:30 | Abbott Laboratories Inc. (St Jude Medical) | Due to a software error, the user may lose functionality of the system or the screen may turn black during operation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187399 | Z-1855-2021 | WorkMate Claris v.1.2 Upgrade Kits, Model Numbers H800007 and H800012 | WorkMate Claris v.1.2 Upgrade Kits | 2 | CDRH | 2021-06-14 |  | 2024-06-18 11:30 | Abbott Laboratories Inc. (St Jude Medical) | Due to a software error, the user may lose functionality of the system or the screen may turn black during operation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187398 | Z-1854-2021 | v.1.2 Upgrade Kit or v.1.2 upgrade installed on WorkMate Claris" System Display, Model H700124, The WorkMate Claris" System is a computer-based electrocardiographic recording and monitoring system designed for efficient capture, display, and retrieval of surface and intracardiac signals during cardiac EP studies. | WorkMate Claris" System Display, Model H700124 | 2 | CDRH | 2021-06-14 |  | 2024-06-18 11:30 | Abbott Laboratories Inc. (St Jude Medical) | Due to a software error, the user may lose functionality of the system or the screen may turn black during operation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=183711 | Z-0109-2021 | SkyVision SDS System - Video Integration Model: GC-SDS-BSCB-200 - Product Usage: is a high quality, self-contained video switching and control unit for use in the operating room. | SkyVision SDS | 2 | CDRH | 2020-10-15 |  | 2021-06-01 14:33 | Skytron, Div. The KMW Group, Inc | Intermittent failure of a signal being sent to a monitor from the touch panel display causing the displaying monitor to go Blank and the touch panel to display a Blank image when reviewing the target display. If this issue were to happen during a case, cause increase risk to patient due to hospital losing the visual image.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=183198 | Z-0108-2021 | Artis zee or Artis Q systems (listed below) - Product Usage: for single and biplane diagnostic imaging and interventional procedures. System Name/Material Number: Artis Q BIPLANE 10848282 ARTIS Q CEILING 10848281 Artis Q floor 10848280 ARTIS Q ZEEGO 10848283 ARTIS Q.ZEN BIPLANE 10848355 ARTIS Q.zen ceiling 10848354 Artis Q.zen floor 10848353 Artis zee Biplane 10094141 Artis zee ceiling 10094137 Artis zee floor 10094135 Artis zee floor 10094135 Artis zee multi-purpose 10094139 Artis zeego 10280959 | Artis Zee and Artis Q systems | 2 | CDRH | 2020-10-15 |  | 2021-02-03 11:23 | Siemens Medical Solutions USA, Inc | Potential exists for the collision supervision not to work properly following loss of the individual room configuration data settings causing danger of collision with fix mounted room equipment, walls or floor may occur. This might cause components of the system to fall or tilt which could result in crushing of patients, operators or staff, collision of system parts with patients, as well as result in delay or interruption of the clinical procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180525 | Z-2261-2020 | Xhibit Central, Model No. 96102 - Product Usage: intended use is to provide clinicians with central monitoring of adult, pediatric and neonatal patient data of patients connected to networked Spacelabs Healthcare patient monitors and telemetry transmitters. | Spacelabs Healthcare Xhibit Central Station (96102) | 2 | CDRH | 2020-06-04 |  | 2024-04-18 18:36 | Spacelabs Healthcare, Inc. | The firm received reports of loss of audio alarm after a power failure or cable disconnection. The loss of audio for a telemetry patient could delay the recognition of an alarm condition.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179968 | Z-1673-2020 | LabPro Data Management System, UDI: 15099590676872 - Product Usage: The LabPro Data Management System is a Microsoft Windows based software program and is intended to manage both microbial identification (ID) and antimicrobial agent susceptibility testing (AST) data generated from MicroScan instruments or manually entered microbiology test results, for use by trained laboratory personnel. | LabPro Data Management System | 2 | CDRH | 2020-04-06 |  | 2021-04-27 16:30 | Beckman Coulter Inc. | A security-only update was released via remote diagnostics on January 14, 2020 to customers with Windows 7 operation system. Shortly after the releases some customers reported that Data Management System computers would not restart. If the computer is unable to restart, then the system cannot download data generated from MicroScan instruments or manually entered microbiology test results, resulting in a potential of delayed results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179902 | Z-1554-2020 | NaviCare Patient Safety Software Versions 3.5.400-3.9.300. The NaviCare Patient Safety (Patient Safety) Software is compatible with the following beds: 1. VersaCare Revision A-J with upgrade kit Revision K with UCB 1.18 2. Centrella (all versions) 3. Progressa (firmware version 1.15.0.0) | NaviCare Patient Safety Software | 2 | CDRH | 2020-03-22 |  | 2021-02-11 12:26 | Hill-Rom Inc. | There is a software issue which may result in failure to monitor or control patient's bed exit mechanism.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179075 | Z-1552-2020 | Quidel Triage TOX Drug Screen Control 1 Kit Box containing 5 Control 1 vials with 5 x 0.25 mL vials per box. Quidel Cardiovascular Inc. 9975 Summers Ridge Road San Diego, CA 92121 USA Quidel.com | Triage TOX Drug Screen Control 1 | 2 | CDRH | 2020-03-20 |  | 2021-01-14 18:00 | QUIDEL CARDIOVASCULAR INC | Due to an error in the programming of the Control CODE CHIP module for this lot, a failing result is inappropriately displayed for AMP and mAMP. This issue results in the Triage Meter displaying a failing control result for the AMP and mAMP tests when the control did not fail.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=168705 | Z-0222-2020 | C3 Wave App, v. 2.0.5 | C3 Wave App | 2 | CDRH | 2019-10-25 |  | 2020-10-01 12:52 | Medical Components, Inc dba MedComp | When the iPad is updated with Apple iOS software version 12, the C3 application malfunctions. Upon launching a new procedure, the ECG waves do not appear on the screen, in surface or internal mode. If the C3 application is not operating properly, the PICC procedure must be completed without using alternate methods to confirm PICC tip placement.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=176716 | Z-0162-2020 | OptiLITE, Products for Laser Surgery, Model # CGH1080F, 800 micron single-use holmium fiber assembly with flat tip, Lot Number 21144, Sterilized by Sterrad Run 4149, This device is for single-use only, sterile, UDI: 00855430006036 | OptiLITE IX Laser Surgery Accessories | 3 | CDRH | 2019-10-18 12:40 | 2019-10-18 21:00 | 2021-04-28 08:48 | Xintec Corp | It has been determined that a lot of fiber lasers were programmed incorrectly. As a result, the fibers from this lot will produce an error and will not allow the user to use the laser when it is plugged in.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175935 | Z-2558-2019 | V-Twin, Model # 6002-800, UDI: 03661540600180 Product Usage: An in-vitro diagnostic device, which is an automated chemistry analyzer, micro type, intended for clinical use in conjunction with certain materials to measure a variety of analytes, including applications in clinical chemistry, monitoring drugs of abuse and in therapeutic drug monitoring. | V-Twin | 2 | CDRH | 2019-09-23 |  |  | Vital Scientific N.V. | Instrument stopped working due to a software lockup, and no patient results are produced.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175756 | Z-2484-2019 | HS Hi-R NEO 900A NIR Reference # 657 822 Serial # from 101 Product Usage: A wide range of surgical interventions are made possible through the modular construction and the diversity of the equipment and accessories. The operating microscope HS Hi-R NEO 900/HS Hi-R NEO 900A/HS Hi-R NEO 900A NIR and HS ALLEGRA 900/590/90 are designed for use in Ophthalmology and in other disciplines where a vertical line of viewing is required. Due to their high maneuverability, the HS ALLEGRA range of operating microscopes are qualified for further applications, in particular ENT- and reconstructive microsurgery are possible, because their line of sight can be tilted or adjusted horizontally. The HS ALLEGRA 590 is particular because of its multidisciplinary application possibilities. | HS Hi-R NEO 900A NIR | 2 | CDRH | 2019-09-05 |  | 2020-04-16 09:43 | Haag-Streit USA Inc | software error in the central control unit of the floor stand.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175754 | Z-2482-2019 | HS Hi-R NEO 900 Reference # 657 820 Serial # from 101 Product Usage: A wide range of surgical interventions are made possible through the modular construction and the diversity of the equipment and accessories. The operating microscope HS Hi-R NEO 900/HS Hi-R NEO 900A/HS Hi-R NEO 900A NIR and HS ALLEGRA 900/590/90 are designed for use in Ophthalmology and in other disciplines where a vertical line of viewing is required. Due to their high maneuverability, the HS ALLEGRA range of operating microscopes are qualified for further applications, in particular ENT- and reconstructive microsurgery are possible, because their line of sight can be tilted or adjusted horizontally. The HS ALLEGRA 590 is particular because of its multidisciplinary application possibilities. | HS Hi-R NEO 900 | 2 | CDRH | 2019-09-05 |  | 2020-04-16 09:43 | Haag-Streit USA Inc | software error in the central control unit of the floor stand.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175755 | Z-2483-2019 | HS Hi-R NEO 900A Reference # 657 821 Serial # from 101 Product Usage: A wide range of surgical interventions are made possible through the modular construction and the diversity of the equipment and accessories. The operating microscope HS Hi-R NEO 900/HS Hi-R NEO 900A/HS Hi-R NEO 900A NIR and HS ALLEGRA 900/590/90 are designed for use in Ophthalmology and in other disciplines where a vertical line of viewing is required. Due to their high maneuverability, the HS ALLEGRA range of operating microscopes are qualified for further applications, in particular ENT- and reconstructive microsurgery are possible, because their line of sight can be tilted or adjusted horizontally. The HS ALLEGRA 590 is particular because of its multidisciplinary application possibilities. | HS Hi-R NEO 900A | 2 | CDRH | 2019-09-05 |  | 2020-04-16 09:43 | Haag-Streit USA Inc | software error in the central control unit of the floor stand.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=173179 | Z-1982-2019 | Kaluza C Flow Cytometry Software, RX Only in the U.S.A., Beckman Coulter, Inc. for the following software devices: (A) Kaluza C Perpetual License (B) Kaluza C Single Use 1 Year License (C) Kaluza C 5 User Network (D) Kaluza C 10 User Network (E) Kaluza C Perpetual Educational License (F) Kaluza C Single User 1 Year educational License (G) Kaluza C 10 User Network Educational License (H) Kaluza C 5 User Network Educational License (I) Kaluza C Flow Cytometry Software | Kaluza C Flow Cytometry Software | 2 | CDRH | 2019-07-17 |  |  | Beckman Coulter Inc. | Compatibility issue between the software and Microsoft updates to Windows 7, 8, and 10. The compatibility issue causes the software to be unusable which could result in a delay of reporting results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=170620 | Z-0996-2019 | AMSCO 5000 Series Washer/Disinfector Model # 5052 | AMSCO 5000 Series Washer/Disinfector Model # 5052 | 2 | CDRH | 2019-03-08 |  | 2019-07-22 09:11 | Steris Corporation | The software in the systems may not process the cycle originally intended. This could result in damage to the various medical devices and instruments processed in the washer/disinfector or potential for inadequate cleaning or disinfection.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=170619 | Z-0995-2019 | AMSCO 3000 Series Washer/Disinfector Model # 3052 | AMSCO 3000 Series Washer/Disinfector Model # 3052 | 2 | CDRH | 2019-03-08 |  | 2019-07-22 09:11 | Steris Corporation | The software in the systems may not process the cycle originally intended. This could result in damage to the various medical devices and instruments processed in the washer/disinfector or potential for inadequate cleaning or disinfection.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=169717 | Z-0705-2019 | MEVION S250i | MEVION S250i | 2 | CDRH | 2019-01-09 |  | 2020-04-10 11:37 | Mevion Medical Systems, Inc. | Treatment is allowed to continue (via partials) in situations where dDose1 interlock has tripped
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=168970 | Z-0651-2019 | VITROS 5,1 FS Chemistry System, Version 3.0 & below For use in the in vitro quantitative measurement of a variety of analytes of clinical interest in certain biological fluids (depending on the analyte) such as serum, plasma, urine and cerebrospinal fluid. | Vitros 5,1 | 2 | CDRH | 2018-12-20 |  | 2022-01-24 16:37 | Ortho-Clinical Diagnostics | Software anomaly may incorrectly extend the on-analyzer stability time for a loaded VITROS Reagent.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=168971 | Z-0652-2019 | VITROS¿ 3600 Immunodiagnostic System, Version 3.3.1 & below For use in the in vitro quantitative measurement of a variety of analytes of clinical interest in certain biological fluids (depending on the analyte) such as serum, plasma, urine and cerebrospinal fluid. | Vitros 3600 | 2 | CDRH | 2018-12-20 |  | 2022-01-24 16:37 | Ortho-Clinical Diagnostics | Software anomaly may incorrectly extend the on-analyzer stability time for a loaded VITROS Reagent.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=168972 | Z-0653-2019 | VITROS 4600 Chemistry System, Version 3.3.1 & below For use in the in vitro quantitative measurement of a variety of analytes of clinical interest in certain biological fluids (depending on the analyte) such as serum, plasma, urine and cerebrospinal fluid. | Vitros 4600 | 2 | CDRH | 2018-12-20 |  | 2022-01-24 16:37 | Ortho-Clinical Diagnostics | Software anomaly may incorrectly extend the on-analyzer stability time for a loaded VITROS Reagent.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=168968 | Z-0650-2019 | VITROS 5600 Integrated System, Version 3.3.1 & below For use in the in vitro quantitative, semi-quantitative, and qualitative measurement of a variety of analytes of clinical interest, using VITROS Chemistry Product Slides, VITROS Chemistry Products MicroTip Reagents and VITROS Immunodiagnostic Products Reagents. | Vitros 5600 | 2 | CDRH | 2018-12-20 |  | 2022-01-24 16:37 | Ortho-Clinical Diagnostics | Software anomaly may incorrectly extend the on-analyzer stability time for a loaded VITROS Reagent.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=168938 | Z-0486-2019 | Ascom Telligence Nurse Call System | Ascom Telligence Nurse Call System | 2 | CDRH | 2018-11-20 |  |  | Ascom (US) Inc. | There is an issue that could cause the Telligence System Gateway II, NGGTWY2 H, to stop communicating to the hallway devices connected to it. When this occurs, the dome lights connected to the Gateway II Ports A and B will start flashing, indicating that they are no longer communicating to the Gateway II. If this occurs, device failures that WOULD otherwise be annunciated WILL NOT be indicated at staff consoles that are configured to display failures.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=166894 | Z-3192-2018 | Xper Flex Cardio Physiomonitoring System Model # 453564669081 Software Version 1.6.0.0388 | Xper Flex Cardio Physiomonitoring System | 2 | CDRH | 2018-09-18 |  | 2021-07-15 13:34 | Invivo Corporation | Released with an incorrect, older firmware version installed. May result in intermittent loss of ECG monitoring, inaccurate ECG amplitude display and or auxiliary output, and inaccurate heart rate (HR) displayed due to QRS detection fault.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=166147 | Z-2882-2018 | Vertebral Motion Analyzer (VMA) Version 2.3.252. VMA software is a quantitative imaging software application intended to be used to process digital image files. | Vertebral Motion Analyzer | 2 | CDRH | 2018-08-21 |  |  | Ortho Kinematics, Inc | Ortho Kinematics Inc. sent a Notice of Correction to Released Testing Results, Radiological Read Report for the Vertebral Motion Analyzer (VMA) test because it contained an error. The error occurred due to a software bug that has been corrected.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164898 | Z-2284-2018 | The Proton Therapy System - Proteus 235 (brand names: Proteus Plus and Proteus ONE) Proteus 235, Beam Management System (PBS) | The Proton Therapy System - Proteus 235 (brand names: Proteus Plus and Proteus ONE) | 2 | CDRH | 2018-06-23 |  | 2020-03-25 11:39 | Ion Beam Applications S.A. | IBA is conducting a voluntary recall to address a PTS (Proton Therapy System) software issue and to reduce the risk related to this issue.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=162928 | Z-1382-2018 | Infinity Acute Care System (IACS) Monitoring Solution, Catalog Numbers: MS18841, MS20401, MS20407, MS20724, MS22956, MS25510, MS25520, MS25643, MS26196, MS26372, MS31818; UDI Information: 4049098054454, 4049098054447, 4049098054409, 4049098009799, 4049098009751 Multi-parameter, physiologic patient monitoring of adult, pediatric and neonatal patients in environments where patient care is provided by trained healthcare professsionals. The lACS obtains the physiologic, multi-parameter data from the connection to the M540 monitor and optional medical devices and displays. The transfer of this data is accomplished by the Infinity network. The M540 is intended to monitor one patient at a time. | Infinity Acute Care System (lACS) Monitoring Solution | 2 | CDRH | 2018-04-16 |  | 2020-04-20 13:02 | Draegar Medical Systems, Inc. | The invasive pressure visual and audible alarms may not activate.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161078 | Z-1316-2018 | HomeSafe AutoAlert Pendant works in conjunction with a compatible Lifeline communicator. | HomeSafe AutoAlert Pendant, AAHB, or AutoAlert Fall Detector | 2 | CDRH | 2018-04-06 |  | 2019-05-16 15:26 | Lifeline Systems Company | A programing error in some Model FD100 HomeSafe AutoAlert Pendants will render the fall detection feature inoperable.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=162538 | Z-1261-2018 | Syngo.via software is intended to be used for viewing, manipulation, communication, and storage of medical images. It can be used as a standalone device or together with a variety of cleared and unmodified syngo based software options. | Syngo.via software | 2 | CDRH | 2018-03-28 |  | 2018-10-02 | Siemens Medical Solutions USA, Inc | When the archiving configuration is changed, data received/created after the upgrade may be flagged as "Not to be archived". This is caused by the automatic function for cleaning up temporary data being disabled by the software upgrade. Due to the disabled cleanup function, disc capacity for free space decreases faster than usual. Unless the archiving configuration and the cleanup automatic function are reverted back to the original state (prior to the software upgrade), data that was incorrectly flagged "Not to be archived" must be manually prevented from being deleted by either clinical administrator or service engineer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=162049 | Z-1216-2018 | The Stryker Thoracic Pedicle Feeler 6002-350-000 is a smart instrument for remote control of the navigation software and creation of a pathway into the pedicle. The instrument has three types of buttons to remotely control the software: the UP, DOWN and SELECT buttons. The Pedicle Feeler is laser marked in 10 mm intervals to help indicate the depth in which the feeler tip has been inserted into the pedicle. Next to the navigational information, these depth indicators on the feeler tip are also helpful in determining the appropriate pedicle screw length. | Thoracic Pedicle Feeler instrument is intended for use with the Stryker¿ Navigation System. | 2 | CDRH | 2018-03-27 |  | 2019-04-03 11:28 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments and Stryker Leibinger GmbH & Co. KG are voluntarily recalling certain serial numbers Thoracic Pedicle Feelers that are used in conjunction with the Stryker¿ Navigation System because the settings on the machine were higher than intended. This change led to increased energy directed onto the tip of the Thoracic Pedicle Feeler which caused material degradation where the laser marks exist. There is the potential for the tip of the Thoracic Pedicle Feeler to break.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161800 | Z-1183-2018 | Sofia 2 Analyzer Catalog Number 20299 Sofia 2 is a bench top analyzer intended to be used with Cassette-based immunofluorescent in vitro diagnostic assays . Sofia 2 is intended for professional and laboratory use. | Sofia 2 | 2 | CDRH | 2018-03-26 |  | 2019-10-21 11:02 | Quidel Corporation | The calibration for the affected batch of analyzers was incorrectly set during the manufacturing process which may lead to a failure when calibration is first performed. There may also be e remote possibility that incorrect qualitative test results for patient samples near the assay cutoff may be reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=159083 | Z-0670-2018 | Medtronic CareLink iPro Therapy Management Software, MMT-7340 The CareLink iPro Therapy Management Software (model, MMT-7340) is intended to be used by a healthcare professional (HCP) with the iPro2 CGM System and provides upload support for the iPro2 Recorder. The CareLink iPro software is used to upload glucose sensor data stored on a recording device, capture BG meter readings through manual entry or wired-upload, retrospectively calibrate sensor data, and provide retrospective reports of continuous glucose information. The reports generated by the CareLink iPro software are intended for use by healthcare professionals to assist with diabetes clinical assessment, patient education, and therapy optimization. | Medtronic CareLink iPro Therapy Management Software, MMT-7340 | 2 | CDRH | 2018-02-21 |  | 2021-03-26 16:26 | Medtronic Inc. | Due to a time conversion error, data uploaded from the iPro(TM)2 recorder to the CareLink iPro software using the new CareLink iPro Uploader feature may result in incorrect reports. This pertains only to reports generated between September 23, 2017 and September 27, 2017 that used the new CareLink iPro Uploader feature.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=156750 | Z-2711-2017 | APLIO 500 TUS-A500 diagnostic ultrasound system | APLIO 500 TUS-A500 | 2 | CDRH | 2017-07-10 |  | 2023-10-11 13:30 | Toshiba American Medical Systems Inc | When Continuous Trace method is used as the tracing method on Spectral Doppler in velocity trace measurement, an incorrect value may be displayed in the ultrasound systems and in the ultrasound workstations listed below.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=156747 | Z-2709-2017 | XARIO 200 TUS-X200 diagnostic ultrasound system | XARIO 200 | 2 | CDRH | 2017-07-10 |  | 2023-10-11 13:30 | Toshiba American Medical Systems Inc | When Continuous Trace method is used as the tracing method on Spectral Doppler in velocity trace measurement, an incorrect value may be displayed in the ultrasound systems and in the ultrasound workstations listed below.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=156748 | Z-2710-2017 | APLIO 300 TUS-A300 diagnostic ultrasound system | APLIO 300 TUS-A300 | 2 | CDRH | 2017-07-10 |  | 2023-10-11 13:30 | Toshiba American Medical Systems Inc | When Continuous Trace method is used as the tracing method on Spectral Doppler in velocity trace measurement, an incorrect value may be displayed in the ultrasound systems and in the ultrasound workstations listed below.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=156751 | Z-2712-2017 | APLIO i-series: a) APLIO i700 TUS AI700 b) APLIO i800 TUS AI800 diagnostic ultrasound system | APLIO i-series | 2 | CDRH | 2017-07-10 |  | 2023-10-11 13:30 | Toshiba American Medical Systems Inc | When Continuous Trace method is used as the tracing method on Spectral Doppler in velocity trace measurement, an incorrect value may be displayed in the ultrasound systems and in the ultrasound workstations listed below.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=156746 | Z-2708-2017 | XARIO 100 TUS-X100 diagnostic ultrasound system | XARIO 100 TUS-X100 | 2 | CDRH | 2017-07-10 |  | 2023-10-11 13:30 | Toshiba American Medical Systems Inc | When Continuous Trace method is used as the tracing method on Spectral Doppler in velocity trace measurement, an incorrect value may be displayed in the ultrasound systems and in the ultrasound workstations listed below.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153766 | Z-1708-2017 | eCare Coordinator Product Usage: is software intended for use in data aggregation, patient interaction facilitation, storage and clinical information management with independent physiological devices and ancillary systems that are connected either directly or through networks. The software is intended to provide patient information from the patient location through networking technology to a remote care team. eCare Coordinator does not send any real time alarms and is not intended to provide automated treatment decisions. This software is an informational tool only and is not to be used as a substitute for professional judgment of healthcare providers in diagnosing and treating patients. | eCare Coordinator | 2 | CDRH | 2017-04-04 |  | 2017-10-27 | Philips Visicu | eCareCoordinator (eCC) is intended for use in data aggregation, patient interaction facilitation, storage and clinical information management with independent physiological devices and ancillary systems that are connected either directly or through networks. The software is intended to provide patient information from the patient location through networking technology to a remote care team. A software defect was discovered in the Philips eCareCoordinator (eCC) Clinical User Interface that can at times cause missing or redundant data to be saved without notification to the user.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153219 | Z-1495-2017 | Orthosize Templating Version 1.2.6 Echo Bi-Metric Hip Stem Digital Templates. For preoperative planning of orthopedic surgery. | Orthosize | 2 | CDRH | 2017-03-20 | 2017-03-13 21:00 | 2017-11-03 | Zimmer Biomet, Inc. | Digital templates were created with the incorrect files.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150176 | Z-0301-2017 | FemTouch delivery system includes the FemTouch Kit (AC-1007415) and FemTouch" Tip (SPSA-20002590) Provide fractional treatments in gynecology. | FemTouch delivery system | 2 | CDRH | 2016-10-23 | 2016-10-23 21:00 | 2017-04-11 | Lumenis Limited | The cleaning instructions provided in FemTouch Operator's Manual might be ineffective for a fully assembled device and may lead to a potential for cross-contamination.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148350 | Z-2605-2016 | Panacea Air Pro Mattress Control Unit, Model: Platinum 5000, US Model Number 5000. Use with alternating pressure air flotation mattress. | Panacea Air Pro Mattress Control Unit | 2 | CDRH | 2016-08-19 |  | 2017-01-17 | American National Mfg Inc | The Panacea Air Pro Mattress Control Unit is being recalled due to a firmware problem, which may result in the device not having the firmness range required.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=146970 | Z-2039-2016 | CytoVision Image Analysis and Capture System, an automated cell-locating device. The Cytovision is a rapid metaphase finder, image acquisition and computer aided chromosome analysis system which assists the operator in viewing chromosomes and looking for cellular anomalies. CytoVision enables a qualified Cytogeneticist to rapidly and accurately analyze the chromosome banding pattern. | Leica Biosystems Richmond Inc | 3 | CDRH | 2016-06-20 |  | 2017-12-13 | Leica Biosystems Richmond Inc. | Systems have an improperly activated Windows 7 OS, even though a valid Windows 7 license was associated with each manufactured system. The system appears to be unlicensed. These systems cannot be re-activated using standard Windows licensing procedures and are requiring checks of each potentially affected system to confirm Windows 7 OS is properly activated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144151 | Z-1361-2016 | Neurotherm NT2000iX RF Generator Model Number: RFG-NT -2000iX | NT200iX RF Generator | 2 | CDRH | 2016-04-09 | 2016-04-09 21:00 | 2016-04-14 | Neurotherm, Inc. | NT2000iX software shipped with the international setting turned on. The international settings include access to Corodotomy, Bi-Polar, and No Temperature modes were not cleared in the United States for use.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142174 | Z-0581-2016 | Natus Quantum System with NeuroWorks Software. Catalog /Part Numbers: 013926. The Natus Quantum Amplifier is intended to be used as an electroencephalograph: to acquire, display, store and archive electrophysiological signals. The amplifier should be used in conjunction with Natus NeuroWorks /SleepWorks software to acquire scalp and intracranial electroencephalographic (EEG) signals as well as polysomnographic (PSG) signals. | Neuro Works | 2 | CDRH | 2016-01-04 |  | 2016-10-03 | Natus Neurology DBA Excel Tech., Ltd. (XLTEK) | During an internal testing, Natus Medical Incorporated, identified that in Neuro Works 8.1 with Quantum hardware, the incoming signal is displayed with reversed polarity.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139364 | Z-2727-2015 | Transonic Flow Probe, 3 mm, Catalog Number HQD3FMC -- Product Usage: to measure flow intra-operatively. | Transonic Flow Probe | 2 | CDRH | 2015-09-02 | 2015-09-03 21:56 | 2016-07-18 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139370 | Z-2729-2015 | Transonic Flow Probe, 6 mm, Catalog Number HQD6FMV -- Product Usage: to measure flow intra-operatively. | Transonic Flow Probe | 2 | CDRH | 2015-09-02 | 2015-09-03 21:56 | 2016-07-18 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139314 | Z-2722-2015 | Transonic Flow Probe, 6 mm, Catalog Number HQC6FMV -- Product Usage: to measure flow intra-operatively. | Transonic Flow Probe | 2 | CDRH | 2015-09-02 | 2015-09-03 21:56 | 2016-07-18 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139322 | Z-2725-2015 | Transonic Flow Probe, 3 mm, Catalog Number HQC3MP -- Product Usage: to measure flow intra-operatively. | Transonic Flow Probe | 2 | CDRH | 2015-09-02 | 2015-09-03 21:56 | 2016-07-18 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139316 | Z-2724-2015 | Transonic Flow Probe, 2 mm, Catalog Number HQC2MP -- Product Usage: to measure flow intra-operatively. | Transonic Flow Probe | 2 | CDRH | 2015-09-02 | 2015-09-03 21:56 | 2016-07-18 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139366 | Z-2728-2015 | Transonic Flow Probe, 4 mm, Catalog Number HQD4FMC -- Product Usage: to measure flow intra-operatively. | Transonic Flow Probe | 2 | CDRH | 2015-09-02 | 2015-09-03 21:56 | 2016-07-18 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139313 | Z-2721-2015 | Transonic Flow Probe, 3 mm, Catalog Number HQC3FMC -- Product Usage: to measure flow intra-operatively. | Transonic Flow Probe | 2 | CDRH | 2015-09-02 | 2015-09-03 21:56 | 2016-07-18 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139359 | Z-2726-2015 | Transonic Flow Probe, 2 mm, Catalog Number HQD2FMC -- Product Usage: to measure flow intra-operatively. | Transonic Flow Probe | 2 | CDRH | 2015-09-02 | 2015-09-03 21:56 | 2016-07-18 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139310 | Z-2720-2015 | Transonic Flow Probe, 2 mm, Catalog Number HQC2FMC Product Usage: to measure flow intra-operatively. | Transonic Flow Probe | 2 | CDRH | 2015-09-02 | 2015-09-03 21:56 | 2016-07-18 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139315 | Z-2723-2015 | Transonic Flow Probe, 8 mm, Catalog Number HQC8FMV -- Product Usage: to measure flow intra-operatively. | Transonic Flow Probe | 2 | CDRH | 2015-09-02 | 2015-09-03 21:56 | 2016-07-18 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139032 | Z-2419-2015 | Getinge 86 Series Washer/Disinfector with G1 control system software, Model 8666 and Model 8668. For use in healthcare facilities and laboratories where reusable items such as instruments, utensils, anesthesia sets, and glassware are handled for decontamination purposes. | Getinge 86 Series with G1 control system software | 3 | CDRH | 2015-08-19 | 2015-08-25 21:01 | 2018-05-23 | Getinge Disinfection Ab | Getinge Disinfection AB has initiated a Field Correction due to the potential for unauthenticated remote access to the product over the network.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138606 | Z-2376-2015 | PSW 1203.U/1, PSW 1301.U, PSW 1307.U, PSW 1403.U, and PSW 1501.U Pacemaker Programmer software. Software approved for use with BIOTRONIK device programmers (trade names: Renamic and ICS 3000). Allows physicians to program devices to pace exclusively in the left ventricle. | Renamic and ICS 3000 | 2 | CDRH | 2015-08-13 | 2015-08-13 21:00 | 2015-12-24 | BIOTRONIK, Inc. | Ventricular packing: LV software programming versions for BIOTRONIK CRT-P and CRT-D devices are recalled because they are not approved for use in the US.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135853 | Z-1525-2015 | WEBSTER HIS Catheter, 4 Pole Fixed Curve with Auto ID Technology, for mapping of cardiac structure; i.e., stimulation and recording only. Catalog No: 01085413 | WEBSTER Fixed Curve Catheter with Auto ID Technology | 2 | CDRH | 2015-04-27 | 2015-04-27 21:01 | 2015-12-10 | Biosense Webster, Inc. | The Webster HIS catheter, 4Pole Fixed Curve with Auto ID, is not being recognized by the CARTO 3 navigation System.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=134491 | Z-1305-2015 | GE Healthcare, SIGNA 1.5T TWINSPEED, SIGNA INFINITY MRISYSTEM, SIGNA 3.0T INFINITY WITH EXCITE, SIGNA 3.0T WITH EXCITE, SIGNA EXCITE 1.5T, SIGNA EXCITE 3.0T, GE 1.5T AND 3.0T SIGNA HDX MR SYSTEM, Discovery MR750 3.0T, Discovery MR450 1.5T, Discovery MR750w 3.0T, Optima MR450w 1.5T, SIGNA Contour, Brivo MR355, Optima MR360, GE 1.5T SIGNA HDE MR SYSTEM, Signa Openspeed 0.7T MR SYSTEM, Vectra, SIGNA 0.35T OVATION WITH EXCITE, SIGNA 0.2T PROFILE With EXCITE MRI Systems | Multiple GE MRI systems | 2 | CDRH | 2015-03-20 | 2015-03-20 21:00 | 2015-09-08 | GE Healthcare | GE Healthcare has become aware of a potential safety issue involving MRI systems due to software versions not being maintained properly at some sites.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=133015 | Z-1027-2015 | Syngo.plaza: Syngo.plaza is a Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images. It supports the physician in diagnosis and treatment planning. Syngo.plaza also supports storage and archiving of DICOM Structured reports. In a comprehensive imaging syngo.plaza integrates Hospital/Radiology Information Systems (HIS/RIS) to enable customer specific workflows. Syngo.plaza optionally uses a variety of advanced postprocessing applications | Syngo.plaza | 2 | CDRH | 2015-01-28 | 2015-01-28 21:00 | 2017-02-16 | Siemens Medical Solutions USA, Inc | In case of a system crash, images may not be written to the hard disk and this may result in inconsistencies in the database. In case of a system crash (e.g. blue screen, power outage) images may not be written from cache to the hard disk and might get lost.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131433 | Z-1004-2015 | VITROS Software Version 3.1 utilized on the following systems: 1) 4600 Chemistry Systems (VITROS 5,1 FS System family member, 2) 5600 Integrated System 1) VITROS 4600 - in vitro quantitative measurement of a variety of analytes, 2) VITROS 5600 - in vitro quantitative, semi-quantitative, and qualitative measurement of analytes | VITROS 4600 Chemistry System (VITROS 5, FS System family member)VITROS 5600 Integrated System | 3 | CDRH | 2015-01-20 | 2015-01-20 21:00 | 2017-04-03 | Ortho-Clinical Diagnostics | Calibration may not occur when using calibrator barcode labels supplied with VITROS Chemisty Products Calibrator Kit 2.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132029 | Z-0971-2015 | SOMATOM Force; the intended use of computed tomography is to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angels or spiral planes taken at different angles. | SOMATOM Force | 2 | CDRH | 2015-01-14 | 2015-01-14 21:00 | 2015-04-06 | Siemens Medical Solutions USA, Inc | Artifacts are found in acquired imaging on the SOMATOM Force.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131560 | Z-0863-2015 | Bausch & Lomb Stellaris PC w/ Laser Upgrade BL1433 (Stellaris PC w v3 Software) with BL2272 (Laser Upgrade kit) Stellaris PC Vision Enhancement System with integrated laser module. System is used for the surgical treatment/removal of cataractous lens, vitreous removal and certain treatments of the posterior portion of the eye. | Bausch & Lomb Stellaris PC w/ Laser Upgrade BL1433 | 2 | CDRH | 2014-12-23 | 2014-12-23 21:00 | 2017-03-23 | Bausch & Lomb Inc | A software anomaly was identified where the system may automatically transition from "Standby" to "Ready to Fire" state while in surgical use.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131411 | Z-0862-2015 | Bausch & Lomb Stellaris PC w/ Laser, BL14304 The responsible firm name is Bausch & Lomb Incorporated Rochester, NY | Bausch & Lomb Stellaris PC w/ Laser, BL14304 | 2 | CDRH | 2014-12-23 | 2014-12-23 21:00 | 2017-03-23 | Bausch & Lomb Inc | A software anomaly was identified where the system may automatically transition from "Standby" to "Ready to Fire" state while in surgical use.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131561 | Z-0864-2015 | Bausch & Lomb Stellaris PC w/ Laser Upgrade BL14334 (Stellaris PC w v4 Software) with BL2272 (Laser Upgrade kit) Stellaris PC Vision Enhancement System with integrated laser module. System is used for the surgical treatment/removal of cataractous lens, vitreous removal and certain treatments of the posterior portion of the eye. | Bausch & Lomb Stellaris PC w/ Laser Upgrade BL14334 | 2 | CDRH | 2014-12-23 | 2014-12-23 21:00 | 2017-03-23 | Bausch & Lomb Inc | A software anomaly was identified where the system may automatically transition from "Standby" to "Ready to Fire" state while in surgical use.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=129618 | Z-0007-2015 | VITROS Chemistry Products PHYT (Phenytoin) Slides, REF/Product Code 829 8671, IVD. For in vitro diagnostic use only. VITROS Chemistry Products PHYT Slides quantitatively measure phenytoin (PHYT) concentration in serum and plasma using VITROS 250/350/950/5,1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System. | VITROS Chemistry Products PHYT Slides | 2 | CDRH | 2014-10-02 | 2014-10-02 21:05 | 2017-04-11 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics identified a potential for biased results to be generated when using VITROS PHYT Slides, Lot 2613-0150-4913. Some positively biased results were identified during testing of a different VITROS PHYT Slide lot; the affected slides were not released for distribution. Since Lot 2613-0150-4913 was manufactured in the same event, there is the potential for slides from this lot to also be affected.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=129727 | Z-2732-2014 | LAP-BAND AP SYSTEM AP Large with Access Port, Allergan, LAP-BAND AP SYSTEM Access Port I. Product was distributed for both the AP Large and AP Standard sizes. The LAP-BAND System is indicated for weight reduction for patients with obesity, with a Body Mass Index (BMI) of at least 40 kg/m2 or a BMI of at least 30 kg/m2 with one or more obesity related comorbid conditions. | Lap-Band AP System | 2 | CDRH | 2014-09-30 | 2014-09-30 21:00 | 2016-02-01 | Apollo Endosurgery Inc | Product was distributed past its expiration date.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=129373 | Z-2644-2014 | JBAIDS Instrument System Software version 3.5.0.72. The JBAIDS instrument is a portable thermocycler and real-time fluorimeter capable of simultaneous and rapid identification of multiple biological warfare and clinical pathogens from a variety of biological specimens. The software and instrument provides an integrated environment for performing the required functions to control JBAIDS instruments and to perform PCR tests and analysis. | JBAIDS Instrument System Software version 3.5.0.72. | 3 | CDRH | 2014-09-15 | 2014-09-15 21:00 | 2016-10-24 | Biofire Defense | BioFire Defense, LLC is initiating a field action for JBAIDS Software version 3.5.0.72 because the wrong version software may have been installed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=128470 | Z-2147-2014 | Boston Scientific ENDOTAK RELIANCE¿ SG, transvenous defibrillation lead, Models 0180 & 0292, Sterile EO. Product Usage: The ENDOTAK RELIANCE leads provide pacing and rate-sensing and deliver cardioversion and defibrillation shocks for automatic implantable cardioverter defibrillator (AICD) systems. | ENDOTAK RELIANCE¿ SG | 2 | CDRH | 2014-08-04 | 2014-08-05 21:00 | 2014-12-03 | Boston Scientific CRM Corp | A review of manufacturing test records for Boston Scientific ENDOTAK RELIANCE implantable leads revealed a suspected test data recording error. Specifically, some test results were recorded as "failed" without any other indication of failure. There are no reported injuries from the devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=127505 | Z-1949-2014 | DeVilbiss Model DV5x Series CPAP units. For use in treating Obstructive Sleep Apnea (OSA) by means of application of positive air pressure. Used in home and clinical environments. | DeVilbiss | 3 | CDRH | 2014-07-01 | 2014-07-01 21:00 | 2014-12-02 | DeVilbiss Healthcare LLC | A limited number of DV5x Series CPAP units were manufactured and shipped without the serial number programmed into the unit, causing the unit serial number on LCD display to appear as "RUN-TEST".
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=127459 | Z-1705-2014 | Spacelabs Medical Ultraview SL Command Module, Model 91496. Arrhythmia Detector and Alarm. The Multi-parameter Module is used with the Patient Care Management System (PCMS) to acquire, monitor, and process various clinical parameters from a patient in a clinical environment. | Ultraview SL Command Module | 2 | CDRH | 2014-06-03 | 2014-06-03 21:00 | 2014-06-03 | Spacelabs Healthcare Inc | A software anomaly causes temperature alarm limits to change to limits which are not physiologically possible and result in nuisance false positive alarms.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126669 | Z-1660-2014 | LIFEPAK 1000 defibrillator. Affected Part Numbers: 320371500XXX The LP1000 (in AED mode) is indicated for use on patients in cardiac arrest. | Physio Control LIFEPAK 1000 Defibrillator | 2 | CDRH | 2014-05-23 | 2014-05-23 21:00 | 2018-05-29 | Physio-Control, Inc. | Physio-Control has become aware of incidents where customers have attempted to use their LIFEPAK 1000 defibrillator and the device has shut down unexpectedly due to a very low battery. A defibrillator in this scenario has the potential to fail to deliver a shock, with the potential result that therapy is not delivered and a patient is not resuscitated. A software malfunction in the LIFEPAK 100
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126243 | Z-1449-2014 | Cobas c8000 Modular Analyzer Series Cobas c502, Part Number 5964067001, a clinical chemistry analyzer intended for the in vitro quantitative / qualitative determination of analytes in body fluids. | Cobas c8000 Modular Analyzer Series - Cobasc502 | 2 | CDRH | 2014-04-10 | 2014-04-10 21:00 | 2018-01-10 | Roche Diagnostics Operations, Inc. | The Initial Cassette Volume Check (ICVC) is not activated on the cobas c 502 module of the cobas 8000 modular analyzer series when manually filled cobas c packs are loaded. This issue occurs with tests that use cobas c pack MULTI or empty pre-labeled cobas c packs. The ICVC feature allows the reagent probe to dive into the reagent cassette and confirm the reagent level in the bottles. This ensures
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125517 | Z-1154-2014 | RetCam 3, RetCam Portable and RetCam Shuttle; manufactured by Clarity Medical Systems, Inc.; Pleasanton, CA 94588. General ophthalmic imaging including retinal, corneal and external imaging. | RetCam 3, RetCam Portable and RetCam Shuttle | 2 | CDRH | 2014-03-07 | 2014-03-07 21:00 | 2014-08-11 | Clarity Medical Systems Inc | Software anomaly for RetCam 3, RetCam Shuttle and RetCam Portable with software versions 6.0.x, 6.1. x and 6.2x where the patients age is calculated incorrectly.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125293 | Z-0981-2014 | GE Healthcare, PET Discovery 610, Discovery 710, Discovery 600 and Discovery 690. The GE Discovery XR and XA Systems are intended for head and whole body attenuation corrected Positron Emission Tomography (PET) imaging and localization of emission activity in patient anatomy by means of integrated PET and CT images. The Discovery XR and XA are to be used by trained health care professionals for imaging the distribution of radiopharmaceuticals in the body for the assessment of metabolic (molecular) and physiologic functions. This can assist in the evaluation, diagnosis, staging, restaging, and follow up of lesions, disease and organ function such as (but not limited to) cancer, cardiovascular disease, and brain dysfunction. These devices can also assist in radiotherapy planning. The Discovery XR and XA can also be used as stand-alone head and whole body multislice computed tomography (CT) diagnostic imaging systems. | GE Healthcare, PET Discovery 610, Discovery 710, Discovery 600 and Discovery 690. | 2 | CDRH | 2014-02-14 | 2014-02-14 21:00 | 2014-07-29 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to scatter overcorrection associated with the List Mode Replay feature of your Discovery 600, 610, 710 and Discovery 690.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124908 | Z-0844-2014 | KODAK DirectView DR 7500 Dual Detector System, MODEL DR 7500, Manufactured by Carestream Health, Inc. Made in U.S.A. | KODAK DirectView DR 7500 Dual Detector System, MODEL DR 7500 | 2 | CDRH | 2014-01-27 | 2014-01-27 21:00 | 2017-03-10 | Carestream Health, Inc. | Carestream Health, Inc. has recalled DR 7500 Dual Detector System utilizing Version 5.X Software due to a possible patient image display error.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123994 | Z-0495-2014 | Stryker brand EXCELSIOR SL-10 PRE-SHAPED J 15OCM; Catalog Number: M0031441930; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | EXCELSIOR SL-10 PRE-SHAPED J 15OCM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124014 | Z-0511-2014 | Stryker brand SUPER COMPLIANT 4MM X 7MM, Catalog Number: M003SSC04070; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | SUPER COMPLIANT 4MM X 7MM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124004 | Z-0501-2014 | Stryker brand EXCELSIOR SL-10 PRE-SHAPED J 15OCM, Catalog Number: M0031681920; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | EXCELSIOR SL-10 PRE-SHAPED J 15OCM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124001 | Z-0498-2014 | Stryker brand EXCELSIOR SL-10 150CM 2 TIP, Catalog Number: M0031681890, Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | EXCELSIOR SL-10 150CM 2 TIP | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124000 | Z-0497-2014 | Stryker brand EXCELSIOR SL-10 PRE-SHAPED S 15OCM; Catalog Number: M0031441920; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | EXCELSIOR SL-10 PRE-SHAPED J 15OCM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124013 | Z-0510-2014 | Stryker brand COMPLIANT 5MM X 30MM, Catalog Number: M003SRC05300; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | COMPLIANT 5MM X 30MM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124011 | Z-0508-2014 | Stryker brand COMPLIANT 4MM X 30MM, Catalog Number: M003SRC04300; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | COMPLIANT 4MM X 30MM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124007 | Z-0504-2014 | Stryker brand GDC VORTX 2MM X 3MM, Catalog Number: M00335302340; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | GDC VORTX 2MM X 3MM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124006 | Z-0503-2014 | Stryker brand GDC-18 360 18MM X 30CM, Catalog Number: M00334818300; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | GDC-18 360 18MM X 30CM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124003 | Z-0500-2014 | Stryker brand EXCELSIOR SL-10 PRE-SHAPED 90 150CM, Catalog Number: M0031681910; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | EXCELSIOR SL-10 PRE-SHAPED 90 150CM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123998 | Z-0496-2014 | Stryker brand EXCELSIOR SL-10 PRE-SHAPED C 15OCM; Catalog Number: M0031441940; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | EXCELSIOR SL-10 PRE-SHAPED J 15OCM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124008 | Z-0505-2014 | Stryker brand COMPLIANT 4MM X 10MM, Catalog Number: M003SRC04100; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | COMPLIANT 4MM X 10MM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124016 | Z-0512-2014 | Stryker brand SUPER COMPLIANT 7MM X 15MM, Catalog Number: M003SSC07150; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | SUPER COMPLIANT 7MM X 15MM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124002 | Z-0499-2014 | Stryker brand EXCELSIOR SL-10 PRE-SHAPED 45 150CM, Catalog Number: M0031681900, Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | EXCELSIOR SL-10 PRE-SHAPED 45 150CM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124010 | Z-0507-2014 | Stryker brand COMPLIANT 4MM X 20MM, Catalog Number: M003SRC04200; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | COMPLIANT 4MM X 15MM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123991 | Z-0494-2014 | Stryker brand EXCELSIOR 1018 PRE-SHAPED 45 150CM, Catalog Number: M0031441900; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | EXCELSIOR 1018 PRE-SHAPED 45 150CM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124012 | Z-0509-2014 | Stryker brand COMPLIANT 5MM X 15MM, Catalog Number: M003SRC05150; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | COMPLIANT 5MM X 15MM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124009 | Z-0506-2014 | Stryker brand COMPLIANT 4MM X 15MM, Catalog Number: M003SRC04150; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | COMPLIANT 4MM X 15MM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124005 | Z-0502-2014 | Stryker brand GDC-10 ULTRA SOFT 2MM X 3CM, Catalog Number: M0033432030; Product is manufactured by Stryker Neurovascular, Cork, Ireland and distributed by Stryker Neurovascular, Fremont, CA Excelsior SL-10/Excelsior 1018 Microcatheter: Stryker Neurovascular Microcatheters are intended to assist in the delivery of diagnostic agents, such as contrast media, and therapeutic agents, such as occlusion coils, into the peripheral, coronary, and neurovasculature. GDC 10, 18 Coils: The Gugliemi Detachable Coil is intended for the endovascular embolization of: (1) Intracranial aneurysms, (2) Other neurovascular abnormalities such as arteriovenous malformations and arteriovenous fistula, (3) Arterial and venous embolizations in the peripheral vasculature GDC 18-Fibered VortX Shape Coil: The GDC 18-Fibered VortX Shape Coil is intended for embolization of vascular abnormalities such as (1) Arteriovenous malformations and arteriovenous fistulae of the neurovasculature; (2) GDC 18- Fibered VoliX Shape Coil is also intended for arterial and venous embolizations in the peripheral vasculature. Transform: The Stryker Neurovascular TransForm Occlusion Balloon Catheters are indicated for use in the neurovasculature to temporarily stop or control blood flow and for balloon assisted embolization of aneurysms. | GDC-10 ULTRASOFT 2MM X 3CM | 2 | CDRH | 2013-12-13 | 2013-12-13 21:00 | 2014-08-27 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122858 | Z-0443-2014 | Blanketrol III Data Export Software Version 1.10 The Blanketrol III Data Export Software is intended for use in the electronic transfer of medical device data from the Blanketrol III to a computer, without controlling or altering the function or parameters of the Blanketrol III. | Blanketrol III Data Export Software | 2 | CDRH | 2013-12-04 | 2013-12-04 21:00 | 2014-05-13 | Cincinnati Sub-Zero Products Inc | During a customer inquiry regarding a beta version of the Bill Data Export Software, Model USB-127, it was discovered that the wrong software version was released to a customer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122918 | Z-0273-2014 | Osseocare Pro Console & Set Osseocare Pro, Model 1600870-001, 1700470-001, and 1700471-001. Device Name: Osseocare Pro console & Set Osseocare Pro. Model: 1600870-001, 1700470-001, 1700471-001. Marketing status: Class I, under K092214. Device description: Software-based, electronically controlled tabletop device. In dentistry for dental surgery, endodontics and implantology. | Osseocare Pro Console & Set Osseocare Pro | 2 | CDRH | 2013-11-13 | 2013-11-13 21:00 | 2014-04-10 | Nobel Biocare Usa Llc | The recall was initiated because there is a potential for malfunction of the OsseoCare Pro when used with the OseeoCare Pro iPad application software release 1.2.0.7.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122435 | Z-0142-2014 | S5 Heart-lung machine, Stockert S5 System. Catalogue No: 48-30/40/50-00. The Stockert S5 System is intended to be used during cardiopulmonary bypass for procedures lasting six(6) hours or less. | Sorin S5 Perfusionn System | 2 | CDRH | 2013-11-04 | 2013-11-04 21:00 | 2014-07-10 | Sorin Group Deutschland GmbH | Sorin Group is recalling certain models of S5 Perfusion Systems due to malfunction of cardioplegia control.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121601 | Z-0108-2014 | Siemens RAPIDLab¿ 1265 system blood gas analyzer Product Usage: RAPIDLab 1200 Systems: The Rapidlab 1200 systems are intended for in vitro diagnostic use by healthcare professionals in the quantitative testing of human whole blood. The systems can determine the following parameters: RAPIDLab 1240 pH, pC02, p02 RAPIDLab 1245 pH, pC02, p02, tHb, nBili, F02Hb, FCOHb, FMetHb, FHHb RAPIDLab 1260 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, lactate RAPIDLab 1265 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, lactate, tHb, nBili, F02Hb, FCOHb, FMetHb, FHHb | Siemens Healthcare | 2 | CDRH | 2013-10-29 | 2013-10-29 21:00 | 2015-04-01 | Siemens Healthcare Diagnostics Inc | RAPIDPoint¿ 400 Series, RAPIDPoint¿ 500, or RAPIDLab¿ 1200 Series blood gas analyzer, users can inadvertently enter a wrong patient ID on the demographics screen, which ultimately could lead to a misidentification of the sample
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121594 | Z-0102-2014 | Siemens RAPIDPoint¿ 400 system blood gas analyzer Product Usage: These systems are intended for near-patient and laboratory testing of blood gases, electrolytes, metabolites, total hemoglobin, and hemoglobin derivatives in arterial, venous, and capillary whole blood samples. The systems are capable of determining the following parameters: RAPIDPoint 400 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, Hct RAPIDPoint 405 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, tHb, F02Hb, FCOHb, FMetHb, FHHb, nBili | Siemens Healthcare | 2 | CDRH | 2013-10-29 | 2013-10-29 21:00 | 2015-04-01 | Siemens Healthcare Diagnostics Inc | RAPIDPoint¿ 400 Series, RAPIDPoint¿ 500, or RAPIDLab¿ 1200 Series blood gas analyzer, users can inadvertently enter a wrong patient ID on the demographics screen, which ultimately could lead to a misidentification of the sample
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121595 | Z-0103-2014 | Siemens RAPIDPoint¿¿ 405 system blood gas analyzer Product PUsage: These systems are intended for near-patient and laboratory testing of blood gases, electrolytes, metabolites, total hemoglobin, and hemoglobin derivatives in arterial, venous, and capillary whole blood samples. The systems are capable of determining the following parameters: RAPIDPoint 400 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, Hct RAPIDPoint 405 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, tHb, F02Hb, FCOHb, FMetHb, FHHb, nBili | Siemens Healthcare | 2 | CDRH | 2013-10-29 | 2013-10-29 21:00 | 2015-04-01 | Siemens Healthcare Diagnostics Inc | RAPIDPoint¿ 400 Series, RAPIDPoint¿ 500, or RAPIDLab¿ 1200 Series blood gas analyzer, users can inadvertently enter a wrong patient ID on the demographics screen, which ultimately could lead to a misidentification of the sample
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121597 | Z-0105-2014 | Siemens RAPIDLab¿ 1240 system blood gas analyzer Product Usage: RAPIDLab 1200 Systems: The Rapidlab 1200 systems are intended for in vitro diagnostic use by healthcare professionals in the quantitative testing of human whole blood. The systems can determine the following parameters: RAPIDLab 1240 pH, pC02, p02 RAPIDLab 1245 pH, pC02, p02, tHb, nBili, F02Hb, FCOHb, FMetHb, FHHb RAPIDLab 1260 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, lactate RAPIDLab 1265 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, lactate, tHb, nBili, F02Hb, FCOHb, FMetHb, FHHb | Siemens Healthcare | 2 | CDRH | 2013-10-29 | 2013-10-29 21:00 | 2015-04-01 | Siemens Healthcare Diagnostics Inc | RAPIDPoint¿ 400 Series, RAPIDPoint¿ 500, or RAPIDLab¿ 1200 Series blood gas analyzer, users can inadvertently enter a wrong patient ID on the demographics screen, which ultimately could lead to a misidentification of the sample
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121596 | Z-0104-2014 | Siemens RAPIDPoint¿ 500 system blood gas analyzer Product Usage: RAPIDPoint 500 System: The RAPIDPoint 500 system is designed for professional use in a point-of-care or laboratory environment. This system tests blood gases, electrolytes, metabolites, total hemoglobin, and hemoglobin derivatives in arterial, venous, and capillary whole blood samples. The following parameters are tested: pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, lactate, tHb, F02Hb, FCOHb, FMetHb, FHHb, nBili | Siemens Healthcare | 2 | CDRH | 2013-10-29 | 2013-10-29 21:00 | 2015-04-01 | Siemens Healthcare Diagnostics Inc | RAPIDPoint¿ 400 Series, RAPIDPoint¿ 500, or RAPIDLab¿ 1200 Series blood gas analyzer, users can inadvertently enter a wrong patient ID on the demographics screen, which ultimately could lead to a misidentification of the sample
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121599 | Z-0107-2014 | Siemens RAPIDLab¿ 1260 system blood gas analyzer Product Usage: RAPIDLab 1200 Systems: The Rapidlab 1200 systems are intended for in vitro diagnostic use by healthcare professionals in the quantitative testing of human whole blood. The systems can determine the following parameters: RAPIDLab 1240 pH, pC02, p02 RAPIDLab 1245 pH, pC02, p02, tHb, nBili, F02Hb, FCOHb, FMetHb, FHHb RAPIDLab 1260 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, lactate RAPIDLab 1265 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, lactate, tHb, nBili, F02Hb, FCOHb, FMetHb, FHHb | Siemens Healthcare | 2 | CDRH | 2013-10-29 | 2013-10-29 21:00 | 2015-04-01 | Siemens Healthcare Diagnostics Inc | RAPIDPoint¿ 400 Series, RAPIDPoint¿ 500, or RAPIDLab¿ 1200 Series blood gas analyzer, users can inadvertently enter a wrong patient ID on the demographics screen, which ultimately could lead to a misidentification of the sample
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121598 | Z-0106-2014 | Siemens RAPIDLab¿ 1245 system blood gas analyzer Product Usage: RAPIDLab 1200 Systems: The Rapidlab 1200 systems are intended for in vitro diagnostic use by healthcare professionals in the quantitative testing of human whole blood. The systems can determine the following parameters: RAPIDLab 1240 pH, pC02, p02 RAPIDLab 1245 pH, pC02, p02, tHb, nBili, F02Hb, FCOHb, FMetHb, FHHb RAPIDLab 1260 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, lactate RAPIDLab 1265 pH, pC02, p02, Na+, K+, Ca++, Cl-, glucose, lactate, tHb, nBili, F02Hb, FCOHb, FMetHb, FHHb | Siemens Healthcare | 2 | CDRH | 2013-10-29 | 2013-10-29 21:00 | 2015-04-01 | Siemens Healthcare Diagnostics Inc | RAPIDPoint¿ 400 Series, RAPIDPoint¿ 500, or RAPIDLab¿ 1200 Series blood gas analyzer, users can inadvertently enter a wrong patient ID on the demographics screen, which ultimately could lead to a misidentification of the sample
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122075 | Z-0082-2014 | MAQUET Getinge Group CARDIOHELP Base Unit REF 70104.7999 REF 70104.8012 Manufactured for: MAQUET Cardiopulmonary AG Made in Germany. Cardiopulmonary support system. | CARDIOHELPi | 2 | CDRH | 2013-10-25 | 2013-10-25 21:00 | 2017-06-07 | Maquet Medical Systems USA | It has come to the attention of MAQUET that in certain rare instances, the CARDIOHELP-i device may display an error message with audible alarm, indicating "battery defective." The alarm will occur-upon start up or when the main power supply is unplugged from the main power source. The human machine interface touch screen of the CARDIOHELP-i device may gradually become black due to the human machine interface (HMI) back light shutting down.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120124 | Z-1874-2013 | MRidium 1145 Dose Reduction System (DERS) drug library kit. (This is an option for the MRidium 3860+MR Infusion Pump System). The MRidium 3860+MRI Infusion pump are intended for general hospital or clinical use by medical professional whenever it is required to infuse patients with intra-venous or intra-arterial fluids before, during or after magnetic Resonance imagining (MRI) scans. | MRidium 1145 Dose Reduction System (DERS) | 1 | CDRH | 2013-08-02 | 2013-08-02 21:00 | 2015-07-13 | Iradimed Corporation | The Dose Error Reduction System (DERS) can indicate an incorrect recommended value for the rate via a specific key sequence during the initial infusion setup. This can result in a risk of over-infusion or under-infusion if the infusion is started with this rate value.