## Sheet1
WEB_ADDRESS | RECALL_NUMBER | PRODUCT_DESCRIPTION | TRADE_NAME | RECALL_CLASS | CENTER | CENTER_CLASSIFICATION_DT | POSTED_INTERNET_DT | TERMINATION_DT | FIRM_NAME | MANUFACTURER_RECALL_REASON
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=212390 | Z-1250-2025 | Brand Name: TBS Osteo Product Name: TBS iNsight Model/Catalog Number: TBS iNsight v3.x / pn: OS-TBSi Software Version: V3.x: (1) V3.1.2, (2) V3.1.1, (3) V3.1.0, (4)V3.0.0 Product Description: TBS iNsight is a class II Medical Device software that is installed on bone densitometers for analysis of bone microarchitecture and bone health management Component: No | TBS iNsight | 2 | CDRH | 2025-02-28 | 2025-02-28 20:07 |  | Medimaps Group | Potential variability in calculations from fast array scans compared to array scans when operating on Hologic Horizon machines.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211967 | Z-1019-2025 | Ivenix Infusion System (IIS), LVP Software LVP-SW-0005. Indicated for use in a hospital and in outpatient care environments for the controlled administration of fluids through clinically accepted routes of administration. | Ivenix Infusion System (IIS), LVP Software; LVP-SW-0005 | 1 | CDRH | 2025-02-04 10:58 | 2025-02-04 20:07 |  | Fresenius Kabi USA, LLC | Large Volume Pump Software, version 5.9.2 and earlier has potential for the following anomalies: 1. The pump may become nonfunctional if during an alarm condition the Pause Audio option is repeated 70 time or more. 2. If a secondary infusion is started at the exact moment a primary infusion completes and VTBI reaches 0, it will switch to primary. The primary infusion will infuse at the previously programed primary rate and continue until the infusion is stopped or the bag is empty.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211818 | Z-1006-2025 | Brand Name: VITROS Product Name: VITROS XT 7600 Integrated System Model/Catalog Number: 6844461 (New), 6272222 (Certified), 6904003 (China) Software Version: version 3.8.0 or higher Product Description: In vitro diagnostic chemistry analyzer. FDA 3 letter product code: JJE Component: No, not a component | VITROS XT 7600 Integrated System | 2 | CDRH | 2025-01-21 | 2025-01-21 20:06 |  | Ortho-Clinical Diagnostics, INc. | The VITROS XT 7600 Integrated System software versions 3.8.0 may not correctly notify users if the Signal Reagent Pack exceeds the on-board and/or shelf-life expiration date. This could result in the use of expired products, which would increase the risk of no results or erroneous results from the instrument.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211828 | Z-1007-2025 | Brand Name: VITROS Product Name: VITROS 3600 Immunodiagnostic System Model/Catalog Number: 6802783 (New), 6802914 (Certified/Refurbished), 6904000 (China) Software Version: version 3.8.0 or higher Product Description: In vitro diagnostic chemistry analyzer. FDA 3 letter product code: JJE Component: no, not a component | VITROS 3600 Immunodiagnostic System | 2 | CDRH | 2025-01-21 | 2025-01-21 20:06 |  | Ortho-Clinical Diagnostics, INc. | The VITROS 3600 Immunodiagnostic System software versions 3.8.0 may not correctly notify users if the Signal Reagent Pack exceeds the on-board and/or shelf-life expiration date. This could result in the use of expired products, which would increase the risk of no results or erroneous results from the instrument.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211816 | Z-1005-2025 | Brand Name: VITROS Product Name: VITROS 5600 Integrated System Model/Catalog Number: 6802413 (New), 6802415 (Certified/Refurbished), 6904001 (China) Software Version: version 3.8.0 or higher Product Description: In vitro diagnostic chemistry analyzer. FDA 3 letter product code: JJE Component: No, not a component | VITROS 5600 Integrated System | 2 | CDRH | 2025-01-21 | 2025-01-21 20:06 |  | Ortho-Clinical Diagnostics, INc. | The VITROS 5600 Integrated System software versions 3.8.0 may not correctly notify users if the Signal Reagent Pack exceeds the on-board and/or shelf-life expiration date. This could result in the use of expired products, which would increase the risk of no results or erroneous results from the instrument.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210892 | Z-0570-2025 | Arial Server Software Versions 11.1.4 through 11.1.6, SKU 54305, in conjunction with the Arial Mobile Application or Arial Care Giver Application | Arial Server Software | 2 | CDRH | 2024-11-27 | 2024-11-27 20:07 |  | Securitas Healthcare LLC | Certain software versions will not notify end users of calls initiated by call points with the Arial Emergency and Nurse Call System.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210132 | Z-0046-2025 | Welch Allyn Life2000 Ventilation System, REF MS-01-0118, within exterior carton codes: 1) BT-20-0002, 2) BT-20-0002A, 3) BT-20-0002AP, 4) BT-20-0007, 5) BT-20-0011, and 6) RMS010118CP | Welch Allyn Life2000 Ventilation System | 1 | CDRH | 2024-10-22 14:23 | 2024-10-22 21:06 |  | Baxter Healthcare Corporation | The Life2000 ventilator may fail to initiate the Low Gas Pressure alarm if the pressure gas source (Life2000 compressor, oxygen cylinder or wall source) is not supplied to the ventilator before initiating therapy.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210142 | Z-0139-2025 | Trinias unity, Digital Angiography System with Catheterization Table, Model: KS-100 This system consists of the following main components: C arm, catherization table, image processing equipment, x-ray high voltage generator, x-ray tube. There are several types of each component. | TRINIAS | 2 | CDRH | 2024-10-22 | 2024-10-22 21:07 |  | SHIMADZU CORPORATION MEDICAL SYSTEMS DIVISION | If digital angiography system is used with a specific catheterization table and the operator directs simultaneous movement of the table and the C-arm using direct memory, then direct memory button is released, the table may continue to move, which the operator can stop by pressing the stop button.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209798 | Z-3160-2024 | GE Healthcare Centricity Universal Viewer Zero Footprint Client, Model Numbers: a) 2110344-008; b) 2110344-009; c) 2110344-010; d) 2110344-011; e) 2110344-013; f) 2110344-019; g) 2110344-025; h) 2110344-029; i) 2110344-030; j) 2110344-032 ; Radiological image processing system | GE Healthcare Centricity | 2 | CDRH | 2024-09-17 | 2024-09-17 21:07 |  | GE Healthcare | Footprint Client (ZFP) versions v6.0 SP9.x and SP10.x where the latest addended report is not shown by default to the user. The issue occurs when the addendum is created on the same day as the original report but at a different time, and only when launched on the following browsers. " Internet Explorer 10 (IE10) " Internet Explorer 11 (IE11) " Firefox.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209118 | Z-2580-2024 | VITROS Immunodiagnostic VITROS Folate Reagent Pack 1/2 | VITROS Immunodiagnostic VITROS Folate Reagent Pack 1/2 | 2 | CDRH | 2024-08-12 | 2024-08-12 21:09 |  | Ortho-Clinical Diagnostics, Inc. | It was confirmed that fibrinogen in patient plasma samples precipitates out of solution upon the addition of folate stabilizer reagent as part of the pre-treatment process causing the increase in "TM5-4MB" condition codes which lead to a delay of results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=208061 | Z-2325-2024 | Kosmos on iOS impacting iOS, Model Numbers: P007639, P007944, software is part of a general-purpose diagnostic ultrasound system. | Kosmos | 2 | CDRH | 2024-07-10 18:07 | 2024-07-10 21:07 |  | Echonous Inc | Ultrasound system has a bug in affected iOS software, which after DICOM export, can cause inaccuracy in postprocessed measurements done in 3rd party software, such as PACS or other post-processing software, which may lead to inaccurate diagnoses or treatment.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=208053 | Z-2214-2024 | Software version 5.3 in: RAPIDPoint 500e Blood Gas System (USA) Siemens Material Number: 11416751 RAPIDPoint 500e Blood Gas System (China) 11416752; RAPIDPoint 500e Blood Gas System (Japan) 11416754; RAPIDPoint 500e Blood Gas System (ROW) 11416755 ; | RAPIDPoint 500e Blood Gas System | 2 | CDRH | 2024-06-26 | 2024-06-26 21:06 |  | Siemens Healthcare Diagnostics Inc | Siemens Healthineers has identified an issue with the RAPIDPoint500e system software version 5.3. This issue affects how the sample source is identified at the LIS when using the capillary mode. With this software version, samples run in the capillary mode are incorrectly labeled as arterial when displayed on the LIS, leading to the potential for results to be flagged according to an established arterial range. Capillary sample results are correctly identified as capillary on both the RAPIDPoint 500e system display and the instrument printouts. The issue only affects the new software version 5.3 and only capillary sample mode is impacted. The inability to distinguish specimen sources for blood gas may lead to incorrect diagnoses and management decisions.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=206914 | Z-1609-2024 | t:connect mobile app used in conjunction with t:slim X2 insulin pump with Control-IQ technology | t:connect mobile app | 1 | CDRH | 2024-05-06 15:02 | 2024-05-06 21:06 |  | Tandem Diabetes Care, Inc. | During normal use, the mobile app version 2.7 may crash and be automatically relaunched by the iOS operating system. This cycle intermittently repeats, which leads to excessive Bluetooth communication that may result in pump battery drain and may lead to the pump shutting down sooner than typically expected. Pump shutdown will cause insulin delivery to suspend, which could lead to an under-delivery of insulin and may result in hyperglycemia, including severe hyperglycemia.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=206605 | Z-1484-2024 | LVP Software of the Ivenix Infusion System (IIS), Version 5.8.0. | Ivenix Infusion Pump LVP Software | 1 | CDRH | 2024-04-11 13:41 | 2024-04-11 21:06 |  | Fresenius Kabi USA, LLC | Software has anomalies that have the potential to cause underdose, overdose, or delay in therapy which could lead to serious patient harm or death.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205928 | Z-1230-2024 | VITROS 4600 Chemistry System Catalog Nos. 6802445 (regular) and 6900440 (refurbished) | VITROS 4600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:06 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205924 | Z-1228-2024 | VITROS XT 7600 Integrated System, Catalog No. 6844461 | VITROS XT 7600 Integrated System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:07 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205930 | Z-1232-2024 | VITROS XT 3400 Chemistry System, Catalog No. 6844458 | VITROS XT 3400 Chemistry System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:06 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205931 | Z-1233-2024 | VITROS XT 7600 Integrated System, Catalog No. 6844461 | VITROS XT 7600 Integrated System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:06 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205912 | Z-1224-2024 | VITROS 3600 Immunodiagnostic System Catalog Nos. 6802783 (regular) and 6802914 (refurbished) | VITROS 3600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:07 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205927 | Z-1229-2024 | VITROS 3600 Immunodiagnostic System Catalog Nos. 6802783 (regular) and 6802914 (refurbished) | VITROS 3600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:06 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205923 | Z-1227-2024 | VITROS XT 3400 Chemistry System, Catalog No. 6844458 | VITROS XT 3400 Chemistry System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:07 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205921 | Z-1225-2024 | VITROS 4600 Chemistry System Catalog Nos. 6802445 (regular) and 6900440 (refurbished) | VITROS 4600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:07 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205929 | Z-1231-2024 | VITROS 5600 Integrated System Catalog Nos. 6802413 (regular) and 6802915 (refurbished) | VITROS 5600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:06 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205922 | Z-1226-2024 | VITROS 5600 Integrated System Catalog Nos. 6802413 (regular) and 6802915 (refurbished) | VITROS 5600 Immunodiagnostic System | 2 | CDRH | 2024-03-01 | 2024-03-01 20:07 |  | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205373 | Z-0864-2024 | smiths medical Medfusion Model 4000 syringe pump | smiths medical Medfusion | 1 | CDRH | 2024-02-01 14:28 | 2024-02-01 20:06 |  | Smiths Medical ASD Inc. | Medfusion Model 4000 syringe infusion pumps, with the following software versions, may have the following issues: 1. Delivery During Motor Not Running High Priority Alarm (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 2. Re-administered Loading Dose (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 3. Incorrect Critical Data Failure Alarm, 4. Interruption of Bolus or Loading Dose Delivery (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 5. Incorrect Total Bolus/Loading Dose Displayed (v1.1.0, v1.1.1, v1.1.2), 6. Volume Limit Before Bolus/Loading Dose Complete (v1.1.2), 7. Drug Library Lower Limit Displayed Incorrectly (v1.6.0, v1.6.1), 8. Depleted Battery Alarm (All versions previous to v1.6.5), 9. Loss of Wireless Connectivity (v1.5.0, v1.5.1, v1.6.0, v1.6.1, v1.6.4), 10. PharmGuard Server Password (v2.3, v2.4, v2.5). Smiths Medical corrected all issues included in this notification in previous software updates and the corrections were carried forward into all subsequent software releases. Please ensure you have the most recent Medfusion software installed on your pumps.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205366 | Z-0863-2024 | smiths medical Medfusion Model 3500 Syringe pump | smiths medical Medfusion | 1 | CDRH | 2024-02-01 13:44 | 2024-02-01 20:07 |  | Smiths Medical ASD Inc. | Medfusion syringe pumps, with software versions before v6.0.0, may have the following issues:1. Delivery During Motor Not Running High Priority Alarm, 2. Infusion Restarted with Incorrect Parameters, 3. Screen Lock, 4. Interruption of Bolus or Loading Dose Delivery, 5. Pump Displays Incorrect Bolus/Loading Dose, 6. Loading/Bolus Dose Below the Minimum Recommended Rate, 7. Motor Rate Error, 8. Incorrect Recall Last Settings, 9. Corrupt Configuration, 10. Auto Lock, 11. Toolbox Configuration Loading Dose Time Values. Smiths Medical corrected all issues included in this notification in previous software updates and the corrections were carried forward into all subsequent software releases. Please ensure you have the most recent Medfusion software (v6.0.0) installed on your pumps.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204312 | Z-0560-2024 | Roche Cobas 8000 and Cobas PRO Interface Module (CO8 IM)- IVD FlexLab Automation System -a modular system designed to automate Pre-Analytical and Post-Analytical processing, sample handling in order to automate sample processing in the Laboratory. Hardware versions: FLX-247-00 Firmware versions: COBASa_2.2.0, COBASa_2.3.0 | Inpeco FlexLab | 2 | CDRH | 2023-12-15 | 2023-12-15 20:07 |  | Inpeco S.A. | Firmware of the Interface Module with Roche Cobas 8000 and Cobas PRO (CO8 IM) may lead to a delayed sample tube processing.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204258 | Z-0522-2024 | Medtronic SynchroMed, Model A10 | Medtronic | 2 | CDRH | 2023-12-09 | 2023-12-09 20:06 |  | Medtronic Inc. | The previous software application version (1.1.300) is missing a decimal separator (a comma) for parameter range guidance values displayed on some of the programming screens: Catheter, Reservoir, Infusion, Bolus, myPTM, and Alarm.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204260 | Z-0517-2024 | SagiPlan 2.2, CE0344, Rx Only , Brachyterapy Treatment Planning System | SagiPlan 2.2 | 2 | CDRH | 2023-12-08 15:55 | 2023-12-08 20:08 |  | Bebig Isotopentechnik Gmbh | Due to software malfunction, numerical values may be rounded resulting in inaccurate measurements.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203408 | Z-0127-2024 | TruDi Navigation System, Model: FG-2000-00, Catalog: ENS022B, when used with TruDi Curette, Models: TDC0005Z and TDC0005 | TruDi Navigation System | 2 | CDRH | 2023-10-20 13:19 | 2023-10-20 21:07 |  | Acclarent, Inc. | When using affected curette and software, there is a discrepancy between the actual curette tip location and the location displayed on navigation systems intended for use during surgical procedures of the Ear Nose and Throat (ENT) and ENT skull base surgery, which may cause delayed/prolonged surgery, cerebrospinal fluid leak, visual impairment, or skull base structural damage.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203459 | Z-0094-2024 | Sapphire Infusion Pumps, Models: Multi-Therapy (REF 15031-000-0028), Epidural (REF 15032-000-0027) and SapphirePlus (REF 15038-000-0001) | Sapphire MT Infusion Pump, Sapphire Epidural Infusion Pump, SapphirePlus | 1 | CDRH | 2023-10-19 20:33 | 2023-10-20 21:07 |  | EITAN MEDICAL LTD | Infusion Pumps with affected software revision may fail to detect air in line, which may lead to air embolism.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203324 | Z-0041-2024 | Real Intelligence CORI RI.KNEE v2.0 TKA (Total Knee Arthroplasty) software version 2.0 (2.0.0.5) Product Number: ROB10299 | REAL INTELLIGENCE CORI Welcome Box USA | 2 | CDRH | 2023-10-05 | 2023-10-05 21:06 |  | Blue Belt Technologies, Inc | CORI software was missing a planning stage that appears in cases of significant knee deformation (outside of the range of 3 degrees valgus to 7 degrees varus) when the surgeon chooses to define the femoral rotational reference via the posterior condylar axis
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=202870 | Z-2618-2023 | Omnipod 5 Automated Insulin Delivery System | Omnipod 5 Automated Insulin Delivery System | 2 | CDRH | 2023-09-20 | 2023-09-20 21:06 |  | Insulet Corporation | An error message was received when using the Omnipod 5 App on compatible smartphones that prevented phone control users from controlling the Omnipod 5 Automated Insulin Delivery System from their Omnipod 5 App which may cause a delay in therapy. To reduce the volume of inquiries Insulet was receiving from customers, an email was sent to all phone control users and the cause of the error message was resolved by reverting to prior certificates on the Insulet cloud to secure communications between the Omnipod 5 App sotware and the Insulet Cloud. This issue was limited to only certain compatible Android smartphones and did not affect the Omnipod 5 Pods or the dedicated locked-down Controller provided by Insulet. All other Omnipod 5 users were still able to manage their insulin with the device without use of the smartphone app and affected users were able to switch to the locked-down control provided to all users when they initiated the device during the time the App was unavailable to them. Further, the Omnipod itself continued to deliver insulin as per its pre-programmed settings so although affected users who did not have the locked-down controller readily available may not have been able to command insulin boluses during this time, they continued to receive basal insulin without interruption. Lastly, users of the Omnipod 5 were advised during training to always have backup supplies ready in the event of a device malfunction which would include alternative means to deliver insulin subcutaneously until the issue is resolved to minimize disruption to their insulin regimen. There were 2,168 complaints received out of 7,838 users with the Android compatible smartphones affected and 9 MDRs were submitted, however, there were no serious injuries reported or reports of erroneous results as the primary issue was a delay in therapy. This defect cannot cause false results that could negatively impact patients.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=202222 | Z-2414-2023 | FreeStyle Libre 2 Readers, REF: 71951-01 and 71953-01, used with FreeStyle Libre 2 System |  | 2 | CDRH | 2023-08-17 | 2023-08-17 21:07 |  | Abbott Diabetes Care, Inc. | If blood glucose monitoring system users attempt to start a new sensor, when the old sensor wasn't used for the full 14-day wear time, then the reader may display "Incompatible Sensor" message. If users don't have the previous sensor, or FreeStyle Libre 2 app, or the sensor's full wear time hasn't ended, users may be unable to start a new sensor, which could result in no or delayed glucose results
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199871 | Z-1676-2023 | The 4Kscore Test is an in vitro serum or plasma test that combines the results of four immunoassays (Roche Elecsys total PSA (prostate specific antigen), Roche Elecsys free PSA, intact PSA, and human kallikrein 2) into a single numerical score that also incorporates the following information: a patients age, previous biopsy, and digital rectal exam (DRE). Patients file (sample type, prior biopsy status, DRE status and age) are uploaded to the Specimen Processing Module (SPM). The results from the four assays are exported automatically to Lab Manager Application (B2 LIS) which automatically triggers the 4Kscore calculation when values of the four analytes are available. The 4Kscore test results are determined automatically by the validated laboratory information system algorithm calculation software. | The 4Kscore Test | 2 | CDRH | 2023-06-02 | 2023-06-02 21:07 |  | BioReference Health, LLC | SPM software anomalies that may lead to the generation of erroneous 4Kscore Test results
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199540 | Z-1579-2023 | BD Pyxis MedStation ES (Med ES Main), REF: 323; BD Pyxis MedStation ES (Med ES Main Tower), REF: 352; BD Pyxis Anesthesia Station ES, REF: 327; BD Pyxis CII Safe ES Tower Main, REF: 1116-00; BD Pyxis CII Safe ES, Desktop PC, W-LESS SCNR, BIO, REF: 107-255-01 | BD Pyxis MedStation ES (Med ES Main) | 2 | CDRH | 2023-05-11 | 2023-05-11 21:07 |  | CareFusion 303, Inc. | Automated dispensing cabinet software is experiencing: 1) ES device download failure resulting in partial patient information loss that leads to 2) partial loss of patient transaction data, seen as a server upload processing error. Outdated/incorrect information display may contribute to the removal of medication to which the patient is allergic, incorrect dose amount, or discontinued medication.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=198114 | Z-1280-2023 | Scan Monitor/ScanWatch, model number hwa09 Wearable, Bluetooth-connected wrist watch that records heart activity (ECG) and oxygen saturation (SpO2) | Scan Monitor/ScanWatch | 2 | CDRH | 2023-03-22 | 2023-03-22 21:06 |  | Withings | A software bug eliminated the initial ECG activation and review of the ECG under a physician. However, the device is only cleared for use under the care of a physician.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=198299 | Z-1192-2023 | Alaris PC Unit 8015 utilizing 802.11 b/g or 802.11 a/b/g wireless network cards, with IP addresses from x.x.x.224 to x.x.x.255 | Alaris PC Unit | 2 | CDRH | 2023-03-01 11:30 | 2023-03-01 20:07 |  | CareFusion 303, Inc. | Infusion pump PCs with specific software/network cards/IP addresses can have network/wireless connection failure. Could cause inability to receive new datasets and inability to transfer retrospective analytics data to connected systems. Users with interfaces to electronic medication record or admin discharge transfer, may be unable to receive/send data; infusions may need to be manually programmed
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197513 | Z-0959-2023 | Getinge Flow-i Anesthesia Systems Flow-i C20, Part no. 6677200 Flow-i C30, Part no. 6677300 Flow-i C40, Part no. 6677400 | Getinge Flow-i | 1 | CDRH | 2023-01-26 | 2023-01-26 20:07 |  | Getinge Usa Sales Inc | Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197515 | Z-0961-2023 | Getinge Flow-e Anesthesia System, Part No. 6887900 | Getinge Flow-e | 1 | CDRH | 2023-01-26 | 2023-01-26 20:07 |  | Getinge Usa Sales Inc | Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197514 | Z-0960-2023 | Getinge Flow-c Anesthesia System, Part No. 6887700 | Getinge Flow-c | 1 | CDRH | 2023-01-26 | 2023-01-26 20:07 |  | Getinge Usa Sales Inc | Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=196379 | Z-0198-2023 | Canon Aplio ultrasound system, Model: Aplio i900 TUS-AI900, Aplio i800 TUS-AI800, Aplio i700 TUS-AI700, Aplio a550 CUS-AA550, Aplio a450 CUS-AA450 | Canon Aplio ultrasound system | 2 | CDRH | 2022-11-07 | 2022-11-08 20:06 |  | Canon Medical System, USA, INC. | After completion, ultrasound system's Stress Echo may not terminate if user does not manually terminate the protocol. During next patient exam, an abnormal confirmation message will be displayed "Same view was saved in this phase. Do you use this image with StressEcho Review?" If "no", acquired images not saved, and if "yes" images assigned to the previous patient; could cause treatment delay.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=195973 | Z-0133-2023 | da Vinci X (IS4200) and Xi (IS4000) systems, which consists of finished devices Surgeon Console (SSC), Patient Cart (PSCART), Vision System (VSS) | da Vinci X (IS4200) and Xi (IS4000) systems | 2 | CDRH | 2022-10-25 10:26 | 2022-10-25 21:06 |  | Intuitive Surgical, Inc. | Inadvertent energy delivery from surgical system instrument if 1) Force bipolar and bipolar instruments installed on system 2) Force bipolar connected to Force Triad generator, bipolar not connected to generator 3) Yellow pedal associated with force bipolar is pressed, released with head in surgeon console 4) Blue pedal associated with bipolar is pressed, resulting in force bipolar energy delivery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=195576 | Z-1563-2022 | DigitalDiagnost C50 1.1 | DigitalDiagnost C50 | 2 | CDRH | 2022-08-31 | 2022-09-01 21:06 |  | Philips North America Llc | Following a system restart where the Table Height 2 (TH2) is the default setting and then switched to Wallstand VS2, operators will see an incorrect orientation of the image on the first examination due to an issue in the firmware of the Wallstand VS2 board. The system will rotate the amplimat field selection by 90 degrees. The wrong amplimat field selection may cause an incorrect dose of radiation to occur. Additionally, the anatomic position markers may become mispositioned and could potentially be associated with the opposite side of the anatomy. (Added 1/3/2023), Expansion of the root cause to include incorrect amplimat field selection by 90 degrees following the first exposure using Wallstand VS2 and a subsequent system restart. Previously, this issue only occurred following a system restart where the Table Height 2 (TH2) is the default setting.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194648 | Z-1513-2022 | Xhibit Telemetry Receiver, Model: 96280 | Xhibit Telemetry Receiver, Model: 96280 | 2 | CDRH | 2022-08-11 |  |  | Spacelabs Healthcare, Inc. | Software related alarm escalation defect occurs after approximately 25 days of continuous use. Technical alarm escalation for the following alarms, signal loss, all leads off, low battery, SpO2 sensor off, and signal interference, does not occur in the telemetry receiver as specified. Caregiver may not respond to the initial technical alarm and may not be aware of the continuing technical failure.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194707 | Z-1469-2022 | Sight OLO, Automated Hematology Analyzer, Catalog No. OLO-U1 | Sight OLO | 2 | CDRH | 2022-07-27 |  |  | SIGHT DIAGNOSTICS LTD | The default reference ranges for MCH, MCHC, and RDW for age group "Child 2-11y" are mixed up and incorrectly displayed on OLO devices running software version 2.63, 2.63.1, and 2.63.1.1. Due to the issue, results may incorrectly appear to be within or outside the reference range.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194130 | Z-1306-2022 | NaviCare Nurse Call/Voalte Nurse Call | NaviCare Nurse Call/ Voalte Nurse Call System | 2 | CDRH | 2022-06-30 09:03 |  |  | Baxter Healthcare Corporation | An issue has been identified with Phillips (Emergin) and Longleaf non relay (Connexall, Vocera, Cerner) wireless integrations used with NaviCare/Voalte Nurse Call, software versions 3.9.100 through 3.9.300. Calls placed from a push button call device, such as a push button switch, call cord, or from the push buttons on a room audio station, will be canceled on the nurse call system when the call is answered at the wireless phone, regardless of the call priority.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193661 | Z-1243-2022 | Sensis, Programmable Diagnostic Computer, Model Nos. 6623974 6634633 6634658 6648153 6648161 10140973 10764561 10765502 10910620 | Sensis recording system | 2 | CDRH | 2022-06-06 13:58 | 2022-06-06 21:07 |  | Siemens Medical Solutions USA, Inc | Siemens has become aware of three potential software issues with AXIOM Sensis or Sensis/ Sensis Lite systems. This may lead to a hazardous situation for patients if treatment cannot be continued on the system and treatment needs to be continued on an alternate system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193616 | Z-1223-2022 | NIM Vital Console 4.0 (P/N: NIM4CM01) NIM with Software version 1.3.2 | NIM Vital | 2 | CDRH | 2022-06-02 | 2022-06-02 21:07 |  | Medtronic Xomed, Inc. | Software anomaly was identified.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193652 | Z-1224-2022 | NIM Vital Patient Interface 4.0 (P/N: NIM4CPB1) NIM with Software version 1.3.2 | NIM Vital 4.0 Patient Interface | 2 | CDRH | 2022-06-02 | 2022-06-02 21:07 |  | Medtronic Xomed, Inc. | Software anomaly was identified.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193084 | Z-1074-2022 | 4 Series CX phototherapy units equipped with Daavlin's ClearLink Control System software versions v3.03h and 3.04v. M4 Ser CX 311-20 120-240V 50/60Hz, item # 807QS0020CX6. | 4 Series Phototherapy Unit | 2 | CDRH | 2022-05-12 |  |  | Daavlin Distributing Company | Software issue, resuming an interrupted treatment will result in swap of utilized calibration treatment distance 0" inches to "9. Resulting in over or under dosing.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193083 | Z-1073-2022 | 1 Series CX phototherapy units equipped with Daavlin's ClearLink Control System software versions v3.03h and 3.04v. Models: [1 Ser CX 305-4 120-240V 50/60Hz, item # 06OS0004CX6], [1 Ser CX 311-4 120-240V 50/60Hz CE, item # 807OS0004CX5], [1 Ser CX 311-4 120-240V 50/60Hz, item # 807OS0004CX6], [1 Ser CX 350-4 120-240V 50/60Hz, item # 807OS0400CX6], [1 Ser CX UVA1-4 120-240V 50/60Hz, item # 807OS0400CX6A1] | 1 Series Phototherapy Unit | 2 | CDRH | 2022-05-12 |  |  | Daavlin Distributing Company | Software issue, resuming an interrupted treatment will result in swap of utilized calibration treatment distance 0" inches to "9. Resulting in over or under dosing.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192467 | Z-0998-2022 | MEDTRONIC INTERSTIM THERAPY SYSTEM FOR URINARY CONTROL, which includes the following: a) Handset with Communicator, Model Numbers: 1) TH90Q01, 2) TH90QFA, 3) TH90PFA, 4) TH90GFA, 5) TG90G01; b) InterStim X Clinician Therapy Application Software, Model Number A51300; and c) InterStim x "My Therapy" Patient Therapy Application Software, Model Number A52300 | MEDTRONIC INTERSTIM THERAPY SYSTEM FOR URINARY CONTROL | 2 | CDRH | 2022-04-22 | 2022-04-22 21:06 |  | Medtronic Neuromodulation | There is a software anomaly with the InterStim X Clinician software application with results in the data entered in the Patient Information fields not e retained and a " Data Lost" notification being displayed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192742 | Z-0987-2022 | Artis pheno- in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). Intended use-single and biplane diagnostic imaging and interventional procedures. Model: 10849000 | ARTIS Pheno (VE2) | 2 | CDRH | 2022-04-21 |  |  | Siemens Medical Solutions USA, Inc | Five potential software issues affecting Artis pheno and Artis icono systems in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). The potential software errors are related to the block movements function and manual rotation of the flat detector. Depending on the status of the intervention, the limited functionality may not be sufficient to continue with treatment as planned. This may result in a situation where clinical treatment may be delayed until an additional restart is performed
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192739 | Z-0986-2022 | Artis icono floor in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). Intended use single and biplane diagnostic imaging and interventional procedures. Model: 11327700 | ARTIS Icono | 2 | CDRH | 2022-04-21 |  |  | Siemens Medical Solutions USA, Inc | Five potential software issues affecting Artis pheno and Artis icono systems in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). The potential software errors are related to the block movements function and manual rotation of the flat detector. Depending on the status of the intervention, the limited functionality may not be sufficient to continue with treatment as planned. This may result in a situation where clinical treatment may be delayed until an additional restart is performed
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192721 | Z-0985-2022 | Artis icono biplane- in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). Intended use single and biplane diagnostic imaging and interventional procedures. Model: 11327600 | ARTIS Icono | 2 | CDRH | 2022-04-21 |  |  | Siemens Medical Solutions USA, Inc | Five potential software issues affecting Artis pheno and Artis icono systems in combination with a Siemens Healthineers table or Trumpf/MAQUET table (OEM). The potential software errors are related to the block movements function and manual rotation of the flat detector. Depending on the status of the intervention, the limited functionality may not be sufficient to continue with treatment as planned. This may result in a situation where clinical treatment may be delayed until an additional restart is performed
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192927 | Z-0942-2022 | ARIA Radiation Therapy Management (RTM) VERSIONS 13.6, 15.1, 15.5, 15.6, 16.0, 16.1 | ARIA Radiation Therapy Management | 2 | CDRH | 2022-04-14 |  |  | Varian Medical Systems Imaging Laboratory GmbH | Software issue for treatment plan and image management application may result in mismatch values which could result in treatment in the wrong location
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190782 | Z-0475-2022 | SOMATOM Definition Edge -Computed tomography system Model 10590000 | SOMATOM Definition Edge | 2 | CDRH | 2022-01-10 |  |  | Siemens Medical Solutions USA, Inc | Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190779 | Z-0472-2022 | SOMATOM Confidence -Computed tomography systems Model 10590100 | SOMATOM Confidence | 2 | CDRH | 2022-01-10 |  |  | Siemens Medical Solutions USA, Inc | Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190781 | Z-0474-2022 | SOMATOM Edge Plus-Computed tomography system Model 1026700 | SOMATOM Edge Plus | 2 | CDRH | 2022-01-10 |  |  | Siemens Medical Solutions USA, Inc | Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190780 | Z-0473-2022 | SOMATOM Definition AS-Computed tomography system Model 8098027 | SOMATOM Definition AS | 2 | CDRH | 2022-01-10 |  |  | Siemens Medical Solutions USA, Inc | Software version VB20_SP5 my lead to a relevant degradation of head image quality, a potential risk to patients is misdiagnosis.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189437 | Z-0175-2022 | Cliquid MD version 3.4 software used in conjunction with Analyst MD software on SCIEX Citrine & 4500MD series liquid chromatography-tandem mass spectrometry (LC-MS/MS) systems. Model No. 5088288 | Cliquid MD | 2 | CDRH | 2021-10-26 |  | 2024-01-23 12:16 | AB SCIEX | The values of the Internal Standard (IS) concentrations are incorrectly derived when the user builds a customized test, then set the IS column to hidden and left the column empty. This may lead to the user making incorrect conclusions using incorrect results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188897 | Z-2394-2021 | TRUMATCH CT PIN GUIDE KIT L- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting Part Number: 420578 | TRUMATCH CT PIN GUIDE KIT L | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188901 | Z-2397-2021 | TRUMATCH CT FEM CUT TIB PIN R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting Part Number: 420907 | TRUMATCH CT FEM CUT TIB PIN R | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188899 | Z-2396-2021 | TRUMATCH CT PIN GUIDE FEM R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting Part Number: 420581 | TRUMATCH CT PIN GUIDE FEM R | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188903 | Z-2399-2021 | TRUMATCH CT CUT GUIDE KIT R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting Part Number: 420916 | TRUMATCH CT CUT GUIDE KIT R | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188898 | Z-2395-2021 | TRUMATCH CT PIN GUIDE KIT R- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting Part Number: 420579 | TRUMATCH CT PIN GUIDE KIT R | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188904 | Z-2400-2021 | TRUMATCH CT CUT GUIDE FEM L- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting Part Number: 420917 | TRUMATCH CT CUT GUIDE FEM L | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188902 | Z-2398-2021 | TRUMATCH CT CUT GUIDE KIT L- TruMatch Patient Specific Instruments are intended to be used as patient-specific surgical instrumentation to assist in the positioning of a joint replacement component intra-operatively and in guiding the marking of bone before cutting Part Number: 420915 | TRUMATCH CT CUT GUIDE KIT L | 2 | CDRH | 2021-09-01 08:32 |  | 2023-05-17 11:53 | DePuy Orthopaedics, Inc. | Software: A coding error associated with the Fast3D Segmentation software. During the scanning process the three images required of the hip, knee and ankle may not be aligned, causing the anatomic landmark locations to potentially result in limb malalignment.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187857 | Z-2093-2021 | SoftGenomics version 4.1.15.6 SoftGenomics is a laboratory workflow and information management system to be used in a clinical molecular or cytogenetics laboratory or medical research laboratory setting by knowledgeable and trained users. SoftGenomics is used to record and track activities performed in ordering, specimen processing (pre-analytical steps), testing, documentation, and reporting. | SoftGenomics | 2 | CDRH | 2021-07-19 |  | 2022-04-15 13:14 | Soft Computer Consultants, Inc. | Software showing wrong results reported, PDF not match HIS.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187544 | Z-1906-2021 | Access SARS-CoV-2 lgG II Calibrator, REF C69058, IVD, CE UDI: (01)15099590742751 | Access SARS CoV-2 lgG II Calibrator | 2 | CDRH | 2021-06-21 |  | 2023-03-27 08:30 | Beckman Coulter Inc. | SARS-CoV-2 lgG II numerical results may potentially be multiplied by a factor of 1000 on systems running with assay protocol file (APF) and access assay file (AAF) versions resulting in falsely elevated numerical values.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187543 | Z-1905-2021 | Access SARS CoV-2 lgG II Reagent, REF: C69057, IVD CE UDI: (01)15099590742744 | Access SARS Cov-2 lgG II Reagent | 2 | CDRH | 2021-06-21 |  | 2023-03-27 08:30 | Beckman Coulter Inc. | SARS-CoV-2 lgG II numerical results may potentially be multiplied by a factor of 1000 on systems running with assay protocol file (APF) and access assay file (AAF) versions resulting in falsely elevated numerical values.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=186007 | Z-1333-2021 | BIOCARE Medical , IntelliPath FLX, Automated Staining Instrument, Ref: " Model Number:IPS0001US and IPS0001INTL containing software version 3.5.3.1 | IntelliPath FLX, Automated Staining Instrument | 2 | CDRH | 2021-04-01 |  | 2023-01-31 09:21 | Biocare Medical, LLC | Their is a potential that the automated staining instrument with software version 3.5.3.1 may move with random speeds in random direction across its range of motion. This could result in the instrument hitting end stops or running into other components on the working deck. This could also result in wash buffer being dispensed on random locations.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=185877 | Z-1330-2021 | Product Label - ACUSON Sequoia *** Power Input: 100-240V~, 13.0-5.4A, 50-60Hz *** (240)11148775 Model User Manual states - ACUSON Sequoia Diagnostic Ultrasound Systems | ACUSON Sequoia ultrasound systems | 2 | CDRH | 2021-04-01 |  |  | Siemens Medical Solutions USA, Inc. | The ultrasound system averages the Mean Sac Diameter (MSD) and Gestational Sac Diameter (GSD) into the Estimated Date of Delivery (EDD) calculation. This may result in an incorrect EDD, which may influence patient management decisions regarding induction of labor and elective caesarean delivery, which may result in premature births.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184992 | Z-0785-2021 | TomoTherapy Treatment Delivery System with iDMS - Product Usage: used as an integrated system for the planning and precise delivery of radiation therapy, stereotactic radiotherapy, or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. | TomoTherapy Treatment Delivery System with iDMS | 2 | CDRH | 2021-01-07 |  | 2022-09-26 05:49 | Accuray Incorporated | "MLC tickle error" may result in the delivered dose to effectively rotate from the planned dose.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184985 | Z-0784-2021 | TomoTherapy Treatment System - Product Usage: used as an integrated system for the planning and precise delivery of radiation therapy, stereotactic radiotherapy, or stereotactic radiosurgery to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. | TomoTherapy Treatment System | 2 | CDRH | 2021-01-07 |  | 2022-09-26 05:49 | Accuray Incorporated | "MLC tickle error" may result in the delivered dose to effectively rotate from the planned dose.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180659 | Z-2956-2020 | NovaPACS Diagnostic Viewer. NovaPACS is a picture archiving and communication system software that retrieves, distributes, and display images and data from different modalities including CT, MR, US, CR/DX, NM, PT, and XA. | NovaPACS | 2 | CDRH | 2020-09-11 | 2020-09-12 18:07 | 2021-06-30 13:27 | Novarad Corporation | The firm received a report of an atypical dataset being generated. When using the cross-localization feature and images from a modality that generates asymmetrical images, the cross-localization reference line may be inaccurately placed on any of the corresponding images that are open.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=182679 | Z-2886-2020 | Match Point System The Match Point System is intended to by uses as a surgical instrument to assist in the intraoperative positioning of glenoid components used with total and reverse shoulder arthroplasty by referencing anatomic landmarks of the shoulder that are identifiable on preoperative CT-image scans. | Match Point System Patient specific instrumentation for shoulder arthroplasty | 2 | CDRH | 2020-08-26 |  | 2020-12-23 11:52 | Materialise N.V. | The procedure side indicated in the top header of the Shoulder Case Planning Report incorrectly indicated the procedure side as "Left" for surgical plans for a "right" procedure side.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=182369 | Z-2750-2020 | CARESCAPE ONE, Physiological Patient Monitor - Product Usage: is both a multi-parameter physiological patient monitor and an accessory to a multi-parameter patient monitor intended for use in multiple areas and intra-hospital transport within a professional healthcare facility. | CARESCAPE ONE | 2 | CDRH | 2020-08-04 | 2020-08-04 21:47 | 2024-08-19 16:43 | GE Healthcare, LLC | CARESCAPE ONE may not provide visual and audible alarms for Ventricular Fibrillation (V Fib), if V Fib occurs at the time CARESCAPE ONE is docked to a CARESCAPE B450/B650/B850 host monitor.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=181805 | Z-2537-2020 | VistaSolution, physiological signal transmitter. A software graphical user interface for use by healthcare professionals to display physiological data collected by the wireless remote monitoring system in healthcare settings. This is a secondary, adjunct patient monitor and is not intended to replace existing standard-of-care patient monitoring. | VistaSolution Version 2.1 | 2 | CDRH | 2020-07-08 | 2020-07-10 10:21 | 2021-05-13 09:59 | Vitalconnect Inc. | A healthcare provider (HCP) noticed a patient had a lowered historic SpO2 reading that the software should have triggered a notification to the care team, but no notification was found. A software defect was encountered causing the lack of notification. The defect could result in exposing a patient to potential hypoxia for a short window of time. The firm is correcting the software to bring the performance back into compliance with specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=181744 | Z-2424-2020 | CARESCAPE Patient Data Module (PDM) used with the following monitors: CARESCAPE B450/B650/B850, SOLAR 8000M/I and Transport Pro, Model Numbers: 2042084-001 2094504-001 2042084-001-046050 2042084-001-208621 2042084-001-045915 2042084-001-215164 2042084-001-078474 2042084-001-053034 2042084-001-215165 2042084-001-079462 2042084-001-047389 2042084-001-112198 2042084-001-041673 2042084-001-227238 2042084-001-071564 2042084-001-503824 2042084-001-119839 2042084-001-569269 2042084-001-041674 2042084-001-047549 2042084-001-514651 2042084-001-163572 2042084-001-583721 2042084-001-427592 2042084-001-527358 2042084-001-583723 2042084-001-317166 2042084-001-583724 2042084-001-583715 2042084-001-583718 2042084-001-055156 2042084-001-583722 2042084-001-041835 2042084-001-074124 2042084-001-583706 2042084-001-583707 2042084-001-519926 2042084-001-583717 2042084-001-583713 2042084-001-207863 2042084-001-460604 2042084-001-079192 2042084-001-197663 2042084-001-045872 2042084-001-053195 2042084-001-092010 2042084-001-560824 2042084-001-052372 2042084-001-071088 2042084-001-046667 2042084-001-213006 2042084-001-086202 2042084-001-301821 2042084-001-045720 2042084-001-060454 2042084-001-418294 2042084-001-088907 2042084-001-065769 2042084-001-589530 2042084-001-623792 2042084-001-187563 2042084-001-320333 2042084-001-357143 2042084-001-148302 To be provided 2042084-001-082825 2042084-001-057016 2042084-001-153601 2042084-001-161552 2042084-001-140229 2042084-001-611704 2042084-001-626688 2042084-001-389208 2042084-001-657836 2042084-001-225214 2042084-001-503162 2042084-001-477881 2042084-001-316025 2042084-001-046826 2042084-001-245255 2042084-001-196543 2042084-001-557725 2042084-001-505890 2042084-001-218596 2042084-001-336980 2042084-001-524525 2042084-001-600032 2042084-001-529732 2042084-001-195742 2042084-001-195782 2042084-001-573889 2042084-001-573890 2042084-001-573886 2042084-001-573887 2042084-001-195822 2042084-001-195862 2042084-001-048420 2042084-001-065538 2042084-001-055814 2042084-001-341260 2042084-001-149220 2042084-001-079193 2042084-001-229282 2042084-001-629486 2042084-001-058695 2042084-001-061170 2042084-001-629548 2042084-001-080803 2042084-001-596662 2042084-001-250146 2042084-001-202515 2042084-001-515709 2042084-001-579491 2042084-001-505608 2042084-001-065010 2042084-001-619171 | CARESCAPE Patient Data Module | 2 | CDRH | 2020-06-23 | 2020-06-24 21:00 |  | GE Healthcare, LLC | The CARESCAPE Patient Data Module (PDM) is used with the following monitors: CARESCAPE B450/B650/B850, SOLAR 8000M/I and Transport Pro. If Pace detection is turned ON, on the monitor, and if an Automated External Defibrillator (AED) is used to perform defibrillation, the low frequency signals deployed by the AED, cause the PDM to issue a false flat line ECG waveform and alarm Asystole to be displayed on the monitor.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=181430 | Z-2403-2020 | Merge Healthcare Merge Application Server Software Release and Merge Healthcare Merge Cardio Workstation Software Release - Product Usage: VERICIS is a system intended to be used to acquire, store, print, transfer, and archive clinical information from Camtronics and other vendors systems including images, Hemodynamic studies and reports, measurements (via import from DICOM Structured Reporting, text files or optical character recognition of measurements captured on images) and cardiology signal (waveform) data. | VERICIS CARDIOVASCULAR IMAGE AND INFORMATION SYSTEM, MODEL 4.0 | 2 | CDRH | 2020-06-19 |  | 2023-06-23 13:27 | Merge Healthcare, Inc. | An error in the calculation of the Michigan Reference Ranges can generate an inaccurate Z-Score calculation. The error may reduce the sensitivity of the Z-Score in detecting an abnormality resulting in a false negative.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180938 | Z-2251-2020 | Accelerator APS Aliquoter Module, List number: 07L40 - Product Usage: The ACCELERATOR APS Workcell is a modular system designed to automate sample handling and processing in the clinical laboratory. The system allows consolidations of multiple clinical chemistry and immunoassay analytical instruments into a unified workstation. | Accelerator APS Aliquoter Module | 2 | CDRH | 2020-06-03 | 2020-06-03 21:00 | 2021-04-28 18:12 | Abbott Laboratories | In the case of an AQM clot detection error a portion of the distilled water from the Aliquoter Module hydraulic circuit may be dispensed into the Primary Tube leading to sample dilution, which could lead to incorrect results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180837 | Z-1947-2020 | Quickie/Zippie powered wheelchair. Q700M wheelchair. A battery operated powered wheelchair, which is controlled by Power Module RNET 120A Expandable (Item No.: 250ER10, Part No.: 123966). | Quickie/Zippie Powered Wheelchair | 2 | CDRH | 2020-05-11 | 2020-05-11 21:00 | 2021-04-28 08:49 | Sunrise Medical (US) LLC | Due to programming errors in the Controller, an increase in the set motor parameters can be made by the end user beyond tested safe limits.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180537 | Z-1927-2020 | The Cogent Hemodynamic Monitoring System. REF58400-000, GTIN 00840619099459; REF 58400-000R, GTIN 0087709094314; REF 58403, GTIN 00840619079505; REF 58401, GTIN 00840619079499 - Product Usage: intended for patients for whom the monitoring of continuous cardiac output and calculated hemodynamic parameters is indicated for diagnostic and prognostic evaluation by a clinician. The target population includes Critical Care Patients, Trauma Patients, and Cardiac Surgery Patients. | Cogent Hemodynamic Monitoring System | 2 | CDRH | 2020-05-07 |  | 2023-05-03 11:33 | ICU Medical, Inc. | Firm identified software issues which leads to the patient ID and patient information not being stored in the system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180396 | Z-1899-2020 | Varian Multileaf Collimator [MLC]; Millenium Multileaf Collimator [52, 80, 120 leaf]; High Definition 120 Multileaf Collimator Model Number(s): product code: H54 and H56; MLC software version 8.5 - Product Usage: is intended to assist the clinician in the delivery of external beam radiation to defined target volumes during radiosurgery and radiotherapy while sparing surrounding normal tissue and critical organs from excess radiation. | Varian Multileaf Collimator with C-Series Linac | 2 | CDRH | 2020-05-05 |  | 2021-01-21 11:00 | Varian Medical Systems, Inc. | After a recent upgrade to the collimator software version 8.5, the firm became aware that the multi-leaf collimator leaves did not move during an arc treatment. The issue is related to a transient carriage primary- secondary interlock due to carriage fault on the MLC that immediately preceded the initiation of the arc treatment. The issue will occur only with MLC software version 8.5 and only affects conformal arc treatments, including VMAT and RapidArc. There have been no reports of adverse health consequences due to this issue.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179188 | Z-1563-2020 | VNS Therapy Programmer, Model 3000 used with Models 103, 104, 105, 106, or 8103 VNS Therapy generators. | VNS Therapy Programmer | 2 | CDRH | 2020-03-24 |  |  | LivaNova USA Inc | False positive warning may occur after: 1) VNS Generator interrogated at 0mA normal output current, 2) Generator programmed to non-0mA output current, 3) In-session re-interrogation performed. Users instructed to lower output current and widen pulse width. Only system diagnostic testing evaluates output current. Users may conclude device malfunction, could lead to medical/surgical intervention.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179419 | Z-1346-2020 | COHERENCE Oncologist, Model Nos. 07333680 & 07351898 | COHERENCE Oncologist | 2 | CDRH | 2020-02-21 |  | 2021-01-04 14:33 | Siemens Medical Solutions USA, Inc | A potential safety issue arises when an offset calculation is followed by a filter operation; correcting the image alignment after this sequence will result in incorrect offset values which could lead to incorrect repositioning of the patient and dose to wrong location.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179418 | Z-1330-2020 | DP-30 Digital Ultrasonic Diagnostic Imaging System, Part No. 1155E-PA00004 - Product Usage: The Digital Ultrasonic Diagnostic Imaging System is applicable for adults, pregnant women, pediatric patients and neonates. It is intended for use in fetal, abdominal, pediatric, small organ (breast, thyroid, testes), neonatal cephalic, adult cephalic, trans-rectal, trans-vaginal, muscular-skeletal (conventional, superficial), cardiac (adult, pediatric), peripheral vascular and urology exams. | DP-30 Digital Ultrasonic Diagnostic Imaging System | 2 | CDRH | 2020-02-20 | 2020-02-05 20:03 | 2020-08-21 09:09 | Mindray DS USA, Inc. dba Mindray North America | The DP-30 displays an incorrect needle-guide bracket when used with the 65EC10EA model transducer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178080 | Z-0670-2020 | Monaco Radiation Treatment Planning System (RTP) System | Radiation Treatment Planning System | 2 | CDRH | 2019-12-12 |  |  | Elekta Inc | Monaco is using the incorrect energy when optimizing and calculating dose.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=174163 | Z-2230-2019 | iGUIDE System with software build iGuide 2.2.0, 2.2.1, and 2.2.2 | iGUIDE System | 2 | CDRH | 2019-08-09 |  | 2024-05-06 07:53 | Elekta, Inc. | Potential for iGUIDE to incorrectly monitor the 3D position.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=174647 | Z-2063-2019 | Merge OrthoPACS versions 7.3, 7.3.1, 8.0 and 8.0.1 | Merge OrthoPACS | 2 | CDRH | 2019-07-23 |  | 2023-03-17 14:53 | Merge Healthcare, Inc. | Potential for prior thumbnails to not display in reverse chronological order and images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. Merge PACS provides image manipulation tools to enable users to view and compare images such as: linking, MPR, MIP, 3D image fusion/registration of CT, MR and PET; as well as CVR (Color Volume Rendering), measurements (linear distances, angles, areas, SUV, etc.), and annotations (for example, outline and label regions of interest, label spinal vertebrae).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=174646 | Z-2062-2019 | Merge PACS versions 7.3, 7.3.1, 8.0 and 8.0.1 | Merge PACS | 2 | CDRH | 2019-07-23 |  | 2023-03-17 14:53 | Merge Healthcare, Inc. | Potential for prior thumbnails to not display in reverse chronological order and images may not scroll in reverse chronological order when mammography stacked scrolling is enabled. Merge PACS provides image manipulation tools to enable users to view and compare images such as: linking, MPR, MIP, 3D image fusion/registration of CT, MR and PET; as well as CVR (Color Volume Rendering), measurements (linear distances, angles, areas, SUV, etc.), and annotations (for example, outline and label regions of interest, label spinal vertebrae).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172008 | Z-1489-2019 | Horizon Cardiology Hemo | Horizon Cardiology Hemo | 2 | CDRH | 2019-05-24 |  | 2025-03-04 11:07 | Change Healthcare Israel Ltd. | users are not notified of procedure medication discrepancies between the Vitals and Meds, the Procedure Notes and Patient Common Data screens in Horizon/McKesson Cardiology Hemo" (Hemo). The discrepancy may also affect administered medication data in reports generated from Hemo or the Horizon/McKesson Cardiology Physician s Report as well as customers, who have implemented an outbound interface of procedure results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172011 | Z-1490-2019 | McKesson Cardiology Hemo | McKesson Cardiology Hemo | 2 | CDRH | 2019-05-24 |  | 2025-03-04 11:07 | Change Healthcare Israel Ltd. | users are not notified of procedure medication discrepancies between the Vitals and Meds, the Procedure Notes and Patient Common Data screens in Horizon/McKesson Cardiology Hemo" (Hemo). The discrepancy may also affect administered medication data in reports generated from Hemo or the Horizon/McKesson Cardiology Physician s Report as well as customers, who have implemented an outbound interface of procedure results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=172923 | Z-1384-2019 | UniCel DxH 900 Coulter Cellular Analysis System, REF number C11478 Product Usage: Note: The DxH 900 hematology analyzer is CLIA categorized under document CR180294 and is currently marketed as a member of the instrument family of the DxH 800 cleared under K140911, per FDA s guidance, Guidance for Industry and FDA Staff, Replacement Reagent and Instrument Family Policy . DxH 800 intended use states: The UniCel DxH 800 Analyzer is a quantitative multi-parameter, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel¿ DxH 800 Analyzer identifies and enumerates the parameters indicated below on the following sample types: Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV, NE%, NE#, LY%, LY#, MO%, MO#, EO%, EO#, BA%, BA#, NRBC%, NRBC#, RET%, RET#, MRV, IRF Pre-Diluted Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV Body Fluids (cerebrospinal, serous and synovial) o TNC and RBC | UniCel DxH 900 Coulter Cellular Analysis System | 1 | CDRH | 2019-05-23 | 2019-05-23 21:00 | 2022-03-31 14:47 | Beckman Coulter Inc. | Complaints received of sporadic erroneously elevated platelet results without flags or system messages. Thrombocytopenia may go unrecognized.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171362 | Z-1383-2019 | UniCel DxH 600 Coulter Cellular Analysis System, REF number B23858. Product Usage: Note: The DxH 600 hematology analyzer is the same instrument as the DxH 800 (K140911), except that it is a benchtop unit with no instrument cart, computer stand, and no connectivity capabilities. DxH 800 intended use states: The UniCel DxH 800 Analyzer is a quantitative multi-parameter, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel¿ DxH 800 Analyzer identifies and enumerates the parameters indicated below on the following sample types: Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV, NE%, NE#, LY%, LY#, MO%, MO#, EO%, EO#, BA%, BA#, NRBC%, NRBC#, RET%, RET#, MRV, IRF Pre-Diluted Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV Body Fluids (cerebrospinal, serous and synovial) o TNC and RBC | UniCel DxH 600 Coulter Cellular Analysis System | 1 | CDRH | 2019-05-23 | 2019-05-23 21:00 | 2022-03-31 14:47 | Beckman Coulter Inc. | Complaints received of sporadic erroneously elevated platelet results without flags or system messages. Thrombocytopenia may go unrecognized.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171361 | Z-1382-2019 | UniCel DxH 800 Coulter Cellular Analysis System, REF numbers 629029, B24465, B24802, B68304, B66445, and B63322. Product Usage: The UniCel DxH 800 Analyzer is a quantitative multi-parameter, automated hematology analyzer for in vitro diagnostic use in screening patient populations found in clinical laboratories. The UniCel¿ DxH 800 Analyzer identifies and enumerates the parameters indicated below on the following sample types: Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV, NE%, NE#, LY%, LY#, MO%, MO#, EO%, EO#, BA%, BA#, NRBC%, NRBC#, RET%, RET#, MRV, IRF Pre-Diluted Whole Blood (Venous and Capillary) o WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW, RDW-SD, PLT, MPV Body Fluids (cerebrospinal, serous and synovial) o TNC and RBC | UniCel DxH 800 Coulter Cellular Analysis System | 1 | CDRH | 2019-05-23 | 2019-05-23 21:00 | 2022-03-31 14:47 | Beckman Coulter Inc. | Complaints received of sporadic erroneously elevated platelet results without flags or system messages. Thrombocytopenia may go unrecognized.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=170821 | Z-0973-2019 | Stryker Sustainability Solutions (SSS) reprocessed Arthroscopic Shavers Arthroscopic shavers can be used to abrade, cut and excise tissue and bone; remove loose fragments; and, shave away debris in arthroscopic surgeries, as well as surgeries of the jaw and sinuses. Intended for use in orthopedic surgical procedures of the joints, jaw or sinuses where the cutting and removal of soft and hard tissue or bone is needed in patients requiring orthopedic surgery. | Stryker Sustainability Solutions (SSS) | 2 | CDRH | 2019-03-05 |  | 2020-09-17 13:46 | Stryker Sustainability Solutions | Affected lots of reprocessed arthroscopic shavers may display the incorrect catalog number and speed settings when connected to the generator, require manual mode for use on CORE generators, or display an error message on Crossfire 1 and 2 generators which does not allow the device to be used.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=170402 | Z-0847-2019 | VITEK 2 AST-N351 Test Kit | VITEK 2 AST-N351 Test Kit | 2 | CDRH | 2019-02-16 |  | 2021-08-16 06:48 | bioMerieux, Inc. | False Positive ESBL Phenotype
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=170000 | Z-0784-2019 | Anti-HAV IgM test System, Material No. 11820591160 Product Usage: The Roche Elecsys Anti-HAV IgM immunoassay is used for the in vitro qualitative detection of IgM antibodies to hepatitis A virus (anti-HAV IgM) in human serum and plasma (potassium EDTA, lithium or sodium heparin, sodium citrate). The assay is intended for use as an aid in the laboratory diagnosis of an acute or recently acquired hepatitis A virus infection. Assay results, in conjunction with other laboratory results and clinical information, may be used to provide presumptive evidence of infection with hepatitis A virus in persons with signs and symptoms of hepatitis and in persons at risk for hepatitis A infection. The electrochemiluminescence immunoassay ECLIA is intended for use on Elecsys and cobas e immunoassay analyzers. | Elecsys Anti HAV IgM | 2 | CDRH | 2019-01-29 |  | 2020-04-10 13:50 | Roche Diagnostics Corporation | Potential for erroneous result messages for the Anti-HAV IgM assay when run on the cobas e 602 module. Discrepant result reporting may result.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=168911 | Z-0614-2019 | Xper Flex Cardio Physiomonitoring System, Software versions 1.5.16.0322 and 1.6.1.0390, Model Numbers 453564241901, 453564241911, 453564483321, 453564483331, 453564621791, 453564621801, 453564669081 | Xper Flex Cardio Physiomonitoring System | 2 | CDRH | 2018-12-14 |  | 2024-09-20 14:45 | Philips Electronics North America Corporation | There are 3 issues: 1) Under certain conditions, real time waveforms may not be accurately plotted or displayed. 2) The oxygen saturation (SpO2) numeric value displayed on the device may freeze. 3) In auto cycle mode, if the NIBP communication is lost, the pump would not cycle and the non-invasive blood pressure (NIBP) numeric value displayed on the device would not update.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167437 | Z-0462-2019 | Xper Flex Cardio Physiomonitoring System, Software version 1.6.0.0388 | Xper Flex Cardio Physiomonitoring System | 2 | CDRH | 2018-11-14 |  |  | Invivo Corporation | The Invasive Blood Pressure (IBP) numeric values will freeze on the display when any of the following functions are performed: change an IBP Site Label using the Site Label icon or the keyboard, activate or deactivate the mean pressure waveform display option (On/Off) using the Invasive Pressure Options Dialog or the keyboard, activate or deactivate the waveform hide option (On/Off) using the Invasive Pressure Options Dialog, activate or deactivate the pressure waveform filter option (On/Off) using the keyboard, and/or perform a Pullback using the Pullback icon or the keyboard
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=166225 | Z-2968-2018 | McKesson Cardiology Hemo, Release SW version 13.0. Used for physiological monitoring, image and data processing. McKesson Cardiology Hemo is intended for complete physiological/hemodynamic monitoring, clinical data acquisition, medical image and data processing, and analytical assessment. McKesson Cardiology Hemo is also intended for patient/procedural data management, such as documentation, logging, reporting, trending, storing, reviewing, carrying out clinical calculations and exporting various representations of the acquired data. | McKesson Cardiology Hemo | 2 | CDRH | 2018-08-31 |  | 2021-11-23 12:57 | McKesson Israel Ltd. | Change Healthcare has identified an issue where, under certain circumstances, the Real Time Monitor (RTM) may not display physiological signals.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164966 | Z-2392-2018 | Reliance 1227 Cart and Utensil Washer/Disinfector, FW03101, FW03102, FW03S003 Product Usage: The Reliance 1227 Cart and Utensil Washer/Disinfector is intended for use in the cleaning and low-level disinfection of bedpans and urinals, basins, carts, beds, theatre shoes and other miscellaneous reusable items used in the care of patients. | Reliance 1227 Cart and Utensil Washer/Disinfector | 2 | CDRH | 2018-07-09 |  | 2019-11-05 08:18 | Steris Corporation | The firm has become aware that the Reliance 1227 Cart and Utensil Washer/Disinfector s Chemical Low Level alarm, intended to prevent the user from initiating a cycle when a low chemical level situation occurs, does not operate as intended. Currently, if a low chemical level situation occurs, the alarm will only be generated at the unit s next power-up or when accessing service mode. The failure of the Reliance 1227 to identify that the chemicals used for cleaning are low or empty could result in bedpans and urinals, basins, case carts, beds, theater shoes and other miscellaneous reusable patient care items not being properly cleaned before disinfection or reuse.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164139 | Z-1597-2018 | Biograph Horizon - PET/CT, PETsyngo VJ20A Software Nuclear medicine/ xray diagnostic scanner. | Biograph Horizon PET/CT | 2 | CDRH | 2018-05-02 |  | 2018-10-12 | Siemens Medical Solutions USA, Inc. | Error introduced into PET images acquired and reconstructed with VJ20A software. Array values are indexed improperly when the norm file is created during QC. During data reconstruction, incorrect values are being applied. This can lead to a gradient in the image. The severity of the error is directly related to the positioning of the PET QC phantom relative to the center of the field of view.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161733 | Z-0829-2018 | 2008T, Hemodialysis Delivery System Product Indicated for acute and chronic dialysis therapy. | 2008T Hemodialysis Delivery System | 2 | CDRH | 2018-03-01 | 2018-02-20 21:01 | 2021-07-28 12:53 | Fresenius Medical Care Renal Therapies Group, LLC | While reviewing documentation for the next software release of the 2008T, an R&D technician identified that setting the UF goal to '0' introduces a discrepancy between the UF rate displayed and the actual UF pump rate. The software anomaly is also applicable to the 2008T Machines that contained the 2.63 Bug Fix per DCAF 17-088. 2008T Upgrade kits that contained the 2.63 Bug Fix are also affected. Additionally, SW version 2.64 (in design freeze) is also impacted by this anomaly.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161520 | Z-0808-2018 | Syngo.via systems with SW VB20A | Syngo | 2 | CDRH | 2018-03-01 |  | 2018-10-01 | Siemens Medical Solutions USA, Inc | A software update was released on August 22, 2017. Performing this software installation and upgrading syngo.via system to the software version VB20A_HF02 will resolve various bugs and inconsistent diagnosis with the old system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=160086 | Z-0724-2018 | Integrity R1.2 | Integrity | 2 | CDRH | 2018-02-26 |  | 2024-03-15 16:39 | Elekta Limited | Potential for positional errors following automatic table movement (ATM).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=160085 | Z-0723-2018 | Agility with Integrity R3.2 | Agility Accelerator, linear, medical | 2 | CDRH | 2018-02-26 |  | 2024-03-15 16:39 | Elekta Limited | Potential for positional errors following automatic table movement (ATM).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=159835 | Z-0263-2018 | Xper Flex Cardio Physiomonitoring system. Used to facilitate invasive investigations of heart and vascular disease when non-invasive indicators warrant such. | Xper Flex Cardio Physiomonitoring system | 2 | CDRH | 2017-12-19 |  | 2024-09-20 14:33 | Invivo Corporation | Intermittent communication between the host system and the FlexCardio
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=157969 | Z-0202-2018 | CARESCAPE Patient Data Module (PDM) v2.6 software used with CARESCAPE Monitors B850, B650, and B450 v2.0.7 or earlier software The PDM is intended to provide uninterrupted acquisition of physiologic parameter data on adult, pediatric and neonatal patients during non-transport/bedside and transport patient care episodes. Physiological parameter data acquired by the PDM includes EGG, invasive blood pressure, non-invasive blood pressure, pulse oximetry, temperature, cardiac output and respiration. This device acquires, processes and stores information for all aforementioned parameters and transmits this information to a transport or bedside central processing unit for viewing and alarm surveillance purposes. | CARESCAPE | 2 | CDRH | 2017-12-06 |  | 2021-04-01 10:13 | GE Medical Systems Information Technologies, Inc. | Device does not produce a visual or audible impedance respiration APN alarm when an impedance respiration apnea event occurs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=157814 | Z-0161-2018 | ROSA Surgical Device 2.5.8 It is intended to be used in the operating room for the spatial positioning and orientation of an instrument holder or tool guide. Guidance is based on a pre-operative plan developed with three-dimensional imaging software, and uses fiducial markers or optical registration. The system is intended for use by neurosurgeons to guide standard neurosurgical instruments. It is indicated for any neurosurgical condition in which the use of stereotactic surgery may be appropriate. | Computer-Assisted Surgical Device | 2 | CDRH | 2017-11-21 |  | 2020-06-16 10:13 | Zimmer Biomet, Inc. | Software corrections reactivating the cooperative endoscopy mode.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=159623 | Z-0082-2018 | Xper Flex Cardio PhysioMonitoring system The Xper Flex Cardio physiomonitoring system is used to facilitate invasive investigation of heart and vascular disease when non-invasive indicators warrant such. Its users, responsible to interpret the data made available, will be professional health care providers, such as physicians, registered nurses, cardiovascular technicians, radiology technicians, and other clinical health care providers. The Xper Flex Cardio physiomonitoring system may be used to display and analyze surface ECG (electrocardiogram), respiration, invasive pressure, Sp02 (Pulse Oximetry), End Tidal C02 (carbon dioxide) and non-invasive pressure waveforms; surface body temperature and thermal cardiac output curves. The system is capable of processing/analyzing information such as multi-channel ECG signals and displaying a graphical ST segment map. The Xper Flex Cardio physiomonitoring system makes measurements that assist physicians and other clinicians to evaluate a patient's overall cardiopulmonary performance, conductive system and general vascular status. | Flex Cardio Patient Monitoring System | 2 | CDRH | 2017-11-09 |  | 2024-09-20 14:38 | Invivo Corporation | Intermittent communication between the host system and the Flex Cardio
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153847 | Z-1711-2017 | UniCel DxC 600i SYNCHRON Access Clinical System, Part No. A25637, A25638 Product Usage: The Access 2 and UniCel DxI Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. | Unicel DxC 600i SYNCHRON Access Clinical Systems | 2 | CDRH | 2017-04-04 | 2017-03-15 21:02 | 2018-03-01 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access 2 Family of Immunoassay Systems because a circuit board that controls motor functions can affect device positioning on the pipettor and incubator belt while the instrument is running. The position errors cause the instrument to stop running and cancel tests that are in process.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153845 | Z-1709-2017 | Access 2 Immunoassay Systems, Part No. 81600N Product Usage: The Access 2 and UniCel DxI Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. | Access 2 Immunoassay System | 2 | CDRH | 2017-04-04 | 2017-03-15 21:02 | 2018-03-01 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access 2 Family of Immunoassay Systems because a circuit board that controls motor functions can affect device positioning on the pipettor and incubator belt while the instrument is running. The position errors cause the instrument to stop running and cancel tests that are in process.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153846 | Z-1710-2017 | Access 2 assay Immunoassay System, Remanufactured, Part No. 386220 Product Usage: The Access 2 and UniCel DxI Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative or qualitative determination of various analyte concentrations found in human body fluids. | Access 2 Immunoassay System. | 2 | CDRH | 2017-04-04 | 2017-03-15 21:02 | 2018-03-01 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access 2 Family of Immunoassay Systems because a circuit board that controls motor functions can affect device positioning on the pipettor and incubator belt while the instrument is running. The position errors cause the instrument to stop running and cancel tests that are in process.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152742 | Z-1681-2017 | Merge Eye Care Systems, consisting of Merge Eye Station, Merge Eye Care PACS, and Merge Eye Station Import Utility (ESIU). | Merge Eye Care Systems | 2 | CDRH | 2017-03-24 | 2017-03-13 21:00 | 2020-06-22 17:51 | Merge Healthcare, Inc. | The system will completely stop allowing the downloading or acquisition of new images due to incompatible software application versions.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152040 | Z-1055-2017 | Digital RID Plate Reader and Software Product Code: AD400 | Digital RID Plate Reader | 2 | CDRH | 2017-01-19 | 2017-01-19 21:00 | 2017-04-04 | The Binding Site Group, Ltd. | If a control ring is marked after reading, the software will not flag results that are out of the specified QC range.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151856 | Z-0991-2017 | ConFORMIS ITOTAL PS-IPOLYIMPLANT KIT-RIGHT Catalog Number:TPS-111-1111 The iTotal¿ PS Knee Replacement System is intended for use as a total knee replacement in patients with knee joint pain and disability whose conditions cannot be solely addressed by the use of a prosthetic device that treats only one or two of the three knee compartments, such as a unicondylar, patellofemoral or bicompartmental prosthesis | The iTotal¿ PS Knee Replacement System | 2 | CDRH | 2017-01-10 | 2017-01-10 21:00 | 2017-01-11 | ConforMIS, Inc. | iView report displayed femoral cut measurements in the incorrect unit of measure (inches vs. mm).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151797 | Z-0962-2017 | Amplia MRI" CRT -D SureScan", Amplia MRI" Quad CRT -D SureScan", Claria MRI" CRT -D SureScan", Claria MRI" Quad CRT -D SureScan. implantable medical devices that automatically detect and treat episodes of ventricular fibrillation, ventricular tachycardia, fast ventricular tachycardia, and bradyarrhythmia. | AdaptivCRT (aCRT) | 2 | CDRH | 2017-01-07 |  | 2019-07-16 16:26 | Medtronic Inc., Cardiac Rhythm and Heart Failure | During internal testing conducted as part of next generation product development, it was discovered that the AdaptivCRT (aCRT) feature of Amplia MRI and Claria MRI CRT -D devices may appear to be enabled when it is not active.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151560 | Z-0744-2017 | SOMATOM Definition Edge System, x-ray, tomography, computed Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles | Siemens SOMATOM Definition Edge | 2 | CDRH | 2016-12-12 |  | 2018-04-27 | Siemens Medical Solutions USA, Inc | Software update that provides software and firmware bug-fixes to improve system performance
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151354 | Z-0743-2017 | SOMATOM Definition Flash System, x-ray, tomography, computed Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles | Siemens SOMATOM Definition Flash | 2 | CDRH | 2016-12-12 |  | 2018-04-27 | Siemens Medical Solutions USA, Inc | Software update that provides software and firmware bug-fixes to improve system performance
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151353 | Z-0742-2017 | SOMATOM Definition AS System, x-ray, tomography, computed Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles. | Siemens SOMATOM Definition AS | 2 | CDRH | 2016-12-12 |  | 2018-04-27 | Siemens Medical Solutions USA, Inc | Software update that provides software and firmware bug-fixes to improve system performance
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150426 | Z-0413-2017 | The ORCHESTRA PLUS Programmer A portable device, equipped with a microprocessor used to program and interrogate Sorin implantable pacemakers and defibrillators. The programmer also provides measurement, ECG display and report printing functions that are essential for monitoring implant patients. | ORCHESTRA PLUS Programmer | 2 | CDRH | 2016-11-15 |  | 2020-05-15 15:34 | Sorin Group USA, Inc. | Sorin Group USA, Inc. announces a voluntary field action for the ORCHESTRA Programmer and ORCHESTRA PLUS Programmer because the residual longevity estimate may be less accurate.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150425 | Z-0412-2017 | The ORCHESTRA Programmer A portable device, equipped with a microprocessor used to program and interrogate Sorin implantable pacemakers and defibrillators. The programmer also provides measurement, ECG display and report printing functions that are essential for monitoring implant patients. | ORCHESTRA and ORCHESTRA PLUS Programmers | 2 | CDRH | 2016-11-15 |  | 2020-05-15 15:34 | Sorin Group USA, Inc. | Sorin Group USA, Inc. announces a voluntary field action for the ORCHESTRA Programmer and ORCHESTRA PLUS Programmer because the residual longevity estimate may be less accurate.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150546 | Z-0410-2017 | MEVION S250 Product Usage: Proton Radiation Therapy System | MEVION S250 | 2 | CDRH | 2016-11-14 | 2016-11-15 21:09 | 2017-05-10 | Mevion Medical Systems, Inc. | Recovery of Incorrect Isocenter -operational change
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=149377 | Z-2900-2016 | AQUIOS CL Flow Cytometer, Catalog No. B30166; Intended for use with in vitro diagnostic flow cytometry applications using up to four florescent detection channels. | Aquios CL Flow Cytometer | 2 | CDRH | 2016-09-28 | 2016-09-28 21:01 | 2017-08-07 | Beckman Coulter Inc. | Beckman Coulter is recalling the AQUIOS CL System because the use of the reagent carousel can lead to misalignment of the carousel and damage to the prep probe.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148524 | Z-2657-2016 | Ingenuity Core, Model number 728321 Product Usage: Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Philips Healthcare Ingenuity Core | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148523 | Z-2656-2016 | Brilliance CT 16 Power, Model number 728240 Product Usage: Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Philips Healthcare Brilliance CT 16 Power | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148525 | Z-2658-2016 | Ingenuity Core128, Model number 728323 Product Usage: Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Philips Healthcare Ingenuity Core128 | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148520 | Z-2654-2016 | Brilliance BigBore Radiology CT Model number 728244 Product Usage: Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Philips Healthcare Brilliance BigBore Radiology CT | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148519 | Z-2653-2016 | Brilliance CT Big Bore Oncology, Model number 728243 Product Usage: Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Philips Healthcare Brilliance BigBore Oncology CT | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148521 | Z-2655-2016 | Brilliance 16, Model number 728246 Product Usage: Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Philips Healthcare Brilliance 16 | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148518 | Z-2652-2016 | Brilliance 64 CT Model number 728231 Product Usage: Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. | Philips Healthcare Brilliance CT 64 | 2 | CDRH | 2016-08-25 |  | 2018-07-02 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148151 | Z-2293-2016 | MOSAIQ Oncology Information System | MOSAIQ | 2 | CDRH | 2016-07-27 |  |  | Elekta, Inc. | It is possible that a change to an Order Set will not be saved in the current open Care Plan. This can result in a wrong chemo dose in the order.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=143632 | Z-1172-2016 | Accu-Chek Inform II Base Unit as a part of the Accu-Chek Inform II Blood Glucose Monitoring System Accu-Chek inform II Blood Glucose Monitoring System is intended for testing outside the body (in vitro diagnostic use) and is intended for multiple-patient use in professional healthcare settings. This system should only be use with single-use, auto-disabling lancing devices. | Accu-Chek Inform II Blood Glucose Monitoring System. | 2 | CDRH | 2016-03-14 | 2016-03-14 21:00 | 2017-05-30 | Roche Diagnostics Operations, Inc. | Accu-Chek Inform II Base Unit might produce physical transmission errors in the form of data loss in the communication between the meter and the Data Management Systems (DMS). The issue can lead to the data loss or in the worst case to an erroneous assignment of the patient data (patient mismatch). The issue will only occur at sites using POTCT1-A communication via USB.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142024 | Z-0856-2016 | Philips Healthcare IntelliVue Info Center iX, A.0 866023 | Philips Healthcare IntelliVue Info Center iX, A.0 | 2 | CDRH | 2016-02-23 | 2016-02-23 21:06 | 2017-09-06 | Philips Electronics North America Corporation | Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=143570 | Z-0853-2016 | Philips IntelliVue Measurement Module X1 Model: M3001A | Philips Healthcare | 2 | CDRH | 2016-02-23 |  | 2021-03-01 14:49 | Philips Electronics North America Corporation | The ST elevation alarm on the Patient Monitor or standalone X2 Measurement Module will not sound when indicated for all chest leads derived using Hexad 12-Lead ECG Monitoring in the Host Monitor
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142026 | Z-0858-2016 | Philips Healthcare IntelliVue Info Center iX, B.0, 866389 Physiological, Patient Monitor (With Arrhythmia Detection or Alarm) | Philips Healthcare IntelliVue Info Center iX, B.0 | 2 | CDRH | 2016-02-23 | 2016-02-23 21:06 | 2017-09-06 | Philips Electronics North America Corporation | Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=143571 | Z-0854-2016 | Philips IntelliVue Module Measurement X2 Model: M3002A | Philips Healthcare | 2 | CDRH | 2016-02-23 |  | 2021-03-01 14:49 | Philips Electronics North America Corporation | The ST elevation alarm on the Patient Monitor or standalone X2 Measurement Module will not sound when indicated for all chest leads derived using Hexad 12-Lead ECG Monitoring in the Host Monitor
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142025 | Z-0857-2016 | Philips Healthcare PIIC Classic Upgrade, 866117 Physiological, Patient Monitor (With Arrhythmia Detection or Alarm) | Philips Healthcare PIIC Classic Upgrade | 2 | CDRH | 2016-02-23 | 2016-02-23 21:06 | 2017-09-06 | Philips Electronics North America Corporation | Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132093 | Z-0374-2016 | Philips IntelliVue Information Center iX (release A.00, A.01, and A.02) are impacted by this issue: 866023 IntelliVue Info Center iX 866024 PIIC iX Upgrade 866117 PIIC Classic Upgrade The intended use of the Information Center Software is to display physiologic waves, parameters, and trends, format data for strip chart recordings and printed reports, and provide the secondary annunciation of alarms from other networked medical devices at a centralized location. An additional intended use of the Information Center Software is to provide primary annunciation of alarms and configuration and control access for networked telemetry monitors. | Philips IntelliVue Information Center iX (release A.00, A.01, and A.02) | 2 | CDRH | 2015-12-04 | 2015-12-04 21:00 | 2020-09-21 06:06 | Philips Medical Systems, Inc. | Sp02 and/or Non Invasive Blood Pressure (NBP) alarms may become disabled without visual notification
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=129687 | Z-0325-2016 | Ingenia, Intera, Achieva and Achieve dStream MR systems on Software versions R5.1.7 and R5.1.8 with the Mobi View software option. Diagnostic imaging system. | Philips Healthcare | 2 | CDRH | 2015-11-23 | 2015-11-23 21:00 | 2021-05-12 15:10 | Philips Medical Systems, Inc. | When a fused series of a sagittal, coronal or radial multi station scan is generated in MobiView, the resultant image order is reversed. This occurs with software release R5.1.7/R5.1.8.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=141491 | Z-0327-2016 | NxStage System One S Cycler (High Permeability Hemodialysis System) Model no. NX1000-5-A | NxStage System One | 2 | CDRH | 2015-11-23 |  | 2019-04-29 08:20 | NxStage Medical, Inc. | Ultrafiltration (UF) Volume software error inaccurate fluid removal
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=141482 | Z-0326-2016 | NxStage System One S Cycler (High Permeability Hemodialysis System) Model no. NX1000-5 | NxStage System One | 2 | CDRH | 2015-11-23 |  | 2019-04-29 08:20 | NxStage Medical, Inc. | Ultrafiltration (UF) Volume software error inaccurate fluid removal
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=140304 | Z-0023-2016 | Centricity PACS IW by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources. Images and data can be stored, communicated, processed and displayed within the system or across computer networks at distributed locations. | GE Healthcare Centricity PACS IW | 2 | CDRH | 2015-10-05 | 2015-10-05 21:00 | 2017-01-05 | GE Healthcare | Images may be missing when a system parameter MapRoute is set to a value greater than 1.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138750 | Z-2321-2015 | ACUSON S2000 and ACUSON S3000 ultrasound systems with the Virtual Touch IQ option; Model numbers: 10041461, 10440017 -S2000 system 10441730 -S3000 system VTIQ Option- 10439521 and 10439522 Radiology: The ultrasound imaging systems are intended for the following applications: Fetal, Abdominal, Intraoperative, Pediatric, Small Parts, Transcranial, OB/GYN, Cardiac, Pelvic, Neonatal/Adult Cephalic, Vascular, Musculoskeletal, Superficial Musculoskeletal, and Peripheral Vascular applications. | ACUSON S2000 and ACUSON S3000 ultrasound systems | 2 | CDRH | 2015-07-30 | 2015-07-30 21:00 | 2017-10-19 | Siemens Medical Solutions USA, Inc. | Potential measurement error on ACUSON S Family ultrasound system. When repositioning the Virtual Touch IQ region of interest from the original (default) location, the lateral position of the measurement tool may not align with the lateral position of the shear velocity data.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138433 | Z-2085-2015 | Patient Handling System (Motion Control Software), Product Usage: Indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body when radiation treatment is indicated. | Patient Handling System | 2 | CDRH | 2015-07-16 | 2015-07-16 21:24 | 2016-02-04 | Viewray Incorporated | ViewRay received a report that the couch moved unexpectedly into the bore after performing a RTCS reboot.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=137840 | Z-1849-2015 | The VariSource iX series afterloader systems are computer controlled remote electro/mechanical systems used for medical purposes, for placing a cable incorporating an irradiated iridium seed internally or close by a malignant tumor or tumor bed in a practice known as brachytherapy. | VariSource iX series | 2 | CDRH | 2015-06-26 | 2015-06-26 23:25 | 2016-09-12 | Varian Medical Systems Inc | After an application freeze and restart of the VariSource iX series by power cycling, the Partial Fraction generated by the system will not be correct. The application freeze issue affects only the systems equipped with magnetic Hard Disk Drives (HDD).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=137242 | Z-1679-2015 | RaySearch RayStation 2.5, 3.0, 3.5 and 4.0; Radiation Therapy Treatment Product Usage: RayStation is a software system designed for treatment planning and analysis of radiation therapy. The treatment plans provide treatment unit set-up parameters and estimates of dose distributions expected during the proposed treatment, and may be used to administer treatments after review and approval by the intended user. The system functionality can be configured based on user needs. The intended users of RayStation shall be clinically qualified radiation therapy staff trained in using the system. | RayStation | 2 | CDRH | 2015-05-27 | 2015-05-27 21:00 | 2017-04-11 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories AB became aware of an issue with the dose calculation for some imported VMAT plans where controls points are defined with wide gantry angle spacing.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135930 | Z-1580-2015 | ViewRay System, Radiation Therapy System | ViewRay Radiation Therapy System | 2 | CDRH | 2015-05-05 | 2015-05-05 21:03 | 2015-09-23 | Viewray Incorporated | The software was not correctly using the RT (Radiation Therapy) to MR (Magnetic Resonance image) coordinate correction for non HFS (Head First Supine) patient orientations, resulting in slice mismatch error.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132485 | Z-1013-2015 | Syngo RT Therapist: The intended use of the SIEMENS branded ARTISTE, ONCOR and PRIMUS family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer | Syngo RT Therapist | 2 | CDRH | 2015-01-26 | 2015-01-26 21:00 | 2015-09-10 | Siemens Medical Solutions USA, Inc | Combination of CTVision with syngo RT Therapist / syngo RT Oncologist 4.3.SP1 automatic registration in Adaptive Targeting might result in wrong offset calculations. Applying this offset can lead to patient mistreatment. Cone Beam imaging is not affected by this problem.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132078 | Z-0930-2015 | Harmony¿ vLED Surgical Lighting System, one unit packaged per box The Harmony¿ vLED Surgical Lighting System is a variable pattern, variable intensity surgical lighting fixture designed to provide visible illumination of the surgical field or the patient for the operating room staff. | Harmony¿ vLED Surgical Lighting System | 2 | CDRH | 2015-01-06 | 2015-01-07 21:00 | 2015-07-11 | Steris Corporation | STERIS has learned that groups of LED lights (modules) contained within the vLED light head may illuminate inconsistently.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131262 | Z-0509-2015 | Total Temporomandibular Joint Replacement System- 45mm Lft Narrow Ti Mand | Total Tempomandubilar Joint Prosthesis | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132001 | Z-0522-2015 | Total Temporomandibular Joint Replacement System- 45mm Rt Standard TI Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132000 | Z-0521-2015 | Total Temporomandibular Joint Replacement System- 45mm Lft Standard Ti Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131271 | Z-0515-2015 | Total Temporomandibular Joint Replacement System- 50mm Rt Narrow Ti Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131268 | Z-0512-2015 | Total Temporomandibular Joint Replacement System- 45mm Rt Stand Off TI Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131277 | Z-0520-2015 | Total Temporomandibular Joint Replacement System- 55mm Rt Standard Ti Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131276 | Z-0519-2015 | Total Temporomandibular Joint Replacement System- 55mm Lft Standard Ti Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131273 | Z-0517-2015 | Total Temporomandibular Joint Replacement System- 50mm Rt Standard Ti Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131272 | Z-0516-2015 | Total Temporomandibular Joint Replacement System- 50mm Rt Stand Off Ti Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131270 | Z-0514-2015 | Total Temporomandibular Joint Replacement System- 50mm Lft Stand Off Ti Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131267 | Z-0511-2015 | Total Temporomandibular Joint Replacement System- 45mm Rt Narrow Ti Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131269 | Z-0513-2015 | Total Temporomandibular Joint Replacement System- 50mm Lft Narrow Ti Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131265 | Z-0510-2015 | Total Temporomandibular Joint Replacement System- 45mm Lft Stand Off Ti Mand Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131274 | Z-0518-2015 | Total Temporomandibular Joint Replacement System- 50mm Lft Standard Ti Mandi Reconstruction of the temporomandibular joint | Total Temporomandibular Joint Replacement System | 2 | CDRH | 2014-12-08 | 2014-12-16 12:46 | 2017-01-04 | Biomet Microfixation, LLC | Laser etching on the parts is wider and deeper than the conditions previously validated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130939 | Z-0211-2015 | Cardiohelp-I System Product Usage: A blood oxygenation and carbon dioxide removal system used to pump blood thorough the extracorporeal bypass circuit for circulatory and/or pulmonary support during procedures requiring cardiopulmonary bypass (for periods up to six hours). It is also intended to provide circulatory and/or pulmonary support during procedures not requiring cardiopulmonary bypass (for periods up to six hours). | Maquet Inc. | 2 | CDRH | 2014-11-13 | 2014-11-13 21:00 | 2016-04-06 | Maquet Cardiovascular Us Sales, Llc | Cardiohelp- i System may have a software issue that can potentially result in an erroneous display of a &quot;Battery Needs Service&quot; message after startup of the units when using either AC or DC power
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130737 | Z-0190-2015 | MX 16-slice CT Scanner System and MX 16-slice CT Scanner System Phase 11 can be used as a whole body computed tomography X-ray system featuring a continuously rotating X-ray tube and detector array with multi-slice capability up to 16 slices simultaneously. The acquired X-RAY transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. The system is suitable for all patients. Access Dual CT system can be used as a Head and Whole Body Computed Tomography X-ray System featuring a continuously rotating X-ray tube and detector array with multi-slice capability. The acquired X-ray transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. | Philips Healthcare MX 16-slice CT Scanner System and MX 16-slice CT Scanner System Phase 11 | 2 | CDRH | 2014-11-07 | 2014-12-16 12:46 | 2015-07-02 | Philips Medical Systems (Cleveland) Inc | &quot;The orientation of images is displayed incorrectly. When operator chooses gantry on the right side from operation console, and the patient position is HFDL, HFDR, FFDL or FFDR, the patient position picture and text is not consistent. &quot;L&quot; and &quot;R&quot; of scan images are reversed and this is not consistent with the patient actual position. &quot;When a surview which is longe
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130271 | Z-0105-2015 | Philips Avalon Monitors with software revision J.30.58: Model Product FM20 M2702A ; FM30 M2703A ; FM50 M2705A Intended for: " Monitoring the physiological parameters of pregnant women " Non-invasive monitoring of fetal heart rates and movements. | Philips Avalon Monitors | 2 | CDRH | 2014-10-21 | 2014-10-21 21:00 | 2017-01-31 | Philips Medical Systems, Inc. | Philips Intellivue and Avalon Fetal Monitors in time-synchronized automatic/sequence mode, the NBP automatic measurement series is stopped
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130270 | Z-0104-2015 | Philips IntelliVue Monitors with software revisions J.21.03, J.21.19. Model Product MP5 M8105A; MP5SC M8105AS; MP5T M8105AT; MP5 Upgrade M8105AU ; MP5SC Upgrade 866327 The monitors are intended to be used for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults, pediatrics, and neonates. | Philips IntelliVue Monitors | 2 | CDRH | 2014-10-21 | 2014-10-21 21:00 | 2017-01-31 | Philips Medical Systems, Inc. | Philips Intellivue and Avalon Fetal Monitors in time-synchronized automatic/sequence mode, the NBP automatic measurement series is stopped
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=127243 | Z-1741-2014 | enGen (TM) Laboratory Automation System Configured with TCAutomation" Software v3.5 -- For in vitro quantitative measurement of a variety of analytes of clinical interest --- The VITROS 5,1 FS Chemistry System with enGenTM Laboratory Automation System is intended for use in the in vitro quantitative measurement of a variety of analytes of clinical interest, using both VITROS Chemistry Products Slides (colorimetric endpoint, rate, ion-selective electrode, and immunorate methods) and VITROS Chemistry Products MicroTip liquid reagents (spectrophotometric and spectrophotometric immunoassay methods.). ---The enGen (TM) Lab Automation System is a modification to VITROS Analyzers that includes the addition of an automated track sub-system (TCAutomation) for pre-analytical and post-analytical sample and data management. In addition to an Automated Track Sub-system, the enGen (TM) System includes one or more VITROS Chemistry, Immunodiagnostic and/or Integrated Systems (clinical analyzers) as well as a Data Management Subsystem (Middleware), and a Software Configuration (gsb) File. The enGen System is typically interfaced with a Laboratory Information System (LIS) via connectivity facilitated at the Data Management Subsystem level. By using the enGen (TM) Laboratory Automation System as a communication interface between the LIS and the clinical analyzers in the laboratory, the user can manage the workflow to the analyzers and other devices as an integrated system. This eliminates the need for the user to program tests for each sample uniquely on each analyzer. Within the enGen (TM) Laboratory Automation System, the physical routing of samples to the correct location is also automatic. The Buffer Module functions either as a location for samples waiting for results or as a storage area for empty sample tube carriers. When a sample is presented to a clinical analyzer via the Bypass module and is metered for testing, the sample remains on the track system in the Buffer module while waiting for the test results to be generated. If all test results are generated as expected and no additional actions are required, the sample is routed to a designated Exit module where it can be removed from the track system by the user. However, there are certain situations in which a follow-up action (retest of the sample) may be necessary. In these situations, the sample is routed back to a clinical analyzer via the Bypass module to complete the follow-up action. Examples of situations requiring a retest include, but are not limited to: repeat testing when a "No Result"is generated during the initial test event, sample dilution when an out of range result is obtained during the initial test event, and routing to an alternate analyzer when the first analyzer the sample is sent to is unavailable for testing. Once the follow-up sample metering actions are complete, the sample is routed to Buffer module where it remains until testing is complete. Once all test results are obtained the sample is routed to the designated Exit module. | enGen (TM) Laboratory Automation System | 2 | CDRH | 2014-06-10 | 2014-06-10 21:00 | 2017-04-11 | Ortho-Clinical Diagnostics | Software anomaly: potential buffer timeout issues when using enGen (TM) Laboratory Automation Systems Configured with TCAutomation (TM) Software v3.5.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=127138 | Z-1617-2014 | McKesson Radiology-PACS. McKesson Radiology is a medical image and information application that is intended to receive, transmit, store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from various medical imaging systems. | Picture Archive Communication System (PACS) | 2 | CDRH | 2014-05-13 | 2014-05-13 21:00 | 2016-02-02 | Mckesson Medical Immaging | High priority alert icon was not displayed in the Study List when there are no Stat studies present. However, the High priority alert icon is visible when studies marked as Stat priority exist in the Study List.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126644 | Z-1528-2014 | INOmax DSIR, model 10007, Nitric Oxide Delivery System, with Software version 3.0.0. The INOmax DS delivery system delivers INOMAX (nitric oxide for inhalation) therapy gas and provides continuous monitoring of inspired O2, NO2, and NO, and a comprehensive alarm system. The primary targeted clinical setting is the Neonatal Intensive Care Unit (NICU) and secondary targeted clinical setting is the transport of neonates. | INOmax DSIR Delivery System | 2 | CDRH | 2014-05-02 | 2014-05-02 21:00 | 2014-11-21 | INO Therapeutics (dba Ikaria) | An anomaly impacts the monitoring system of INOmax DSIR Plus (software 3.0.0). When the integrated pneumatic backup switch is activated with low or no gas flow through the injector module, an intended NO cell failure alarm will occur. In addition, an unintended NO2 cell failure alarm may also occur. The NO/NO2 cell failure alarms are caused by an average NO/NO2 circuit concentration greater tha
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126814 | Z-1504-2014 | Siemens ACUSON X700 Ultrasound Systems with software version 1.0.00, 1.0.01, 1.0.02, 1.1.00 and 1.1.02. ultrasound imaging system. | ACUSON X700 Ultrasound Systems | 2 | CDRH | 2014-04-24 | 2014-04-24 21:00 | 2015-11-12 | Siemens Medical Solutions USA, Inc. | When fourSight 4D ultrasound imaging, 3-Scrape real time 3D imaging, or Anatomical M-mode are used during an Obstetric or Gynecology study, data from the previous patient could replace the current patient's date in the report.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125643 | Z-1308-2014 | RaySearch Ray Station, Radiation Therapy Treatment Planning System, Stand-alone software | RayStation | 2 | CDRH | 2014-03-31 | 2014-03-31 21:00 | 2017-02-17 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories has recalled RaySearch Ray Station 4.0 software due to clinical dose calculation errors during radiation therapy.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125602 | Z-1307-2014 | RayStation Radiation Therapy Treatment Planning System, models 2.5, 3.0, 3.5 and 4.0; a software system designed for treatment planning and analysis of radiation therapy. | RayStation | 2 | CDRH | 2014-03-28 | 2014-03-28 21:00 | 2017-02-17 | RAYSEARCH LABORATORIES AB | Potential for dose errors due to software program errors.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125492 | Z-1194-2014 | Philips HeartStart MRx Monitor/Defibrillator Models: M3535A and M3536A with Q-CPR Meter Option B08 The HeartStart MRx is for use for the termination of ventricular tachycardia and ventricular fibrillation. The device is for use by qualified medical personnel trained in the operation of the device and qualified by training in basic life support, advanced cardiac support, or defibrillation. It must be used by or on the order of a physician | Phillips HeartStart MRx Monitor/Defribillator | 2 | CDRH | 2014-03-12 | 2014-03-12 21:00 | 2020-09-17 13:51 | Philips Medical Systems, Inc. | When the HeartStart MRx.is used with the Q-CPR Meter in defibrillation mode, the Q-CPR Meter may incorrectly display the Do Not Touch the Patient icon.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124662 | Z-0780-2014 | XiO RTP System The XiO RTP System is used to create treatment plans for any cancer patient for whom external beam radiation therapy or brachytherapy has been prescribed. The system will calculate and display, both on-screen and in hard-copy, either two or three-dimensional radiation dose distributions within a patient for a given treatment plan set-up. | XiO RTP System | 2 | CDRH | 2014-01-17 | 2014-01-17 21:00 | 2014-05-13 | Elekta, Inc. | When the number of fractions is changed for a proton spot beam, the MU are recomputed and rounded if necessary, but the dose is not recalculated as it should be even after saving and re-opening the plan.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124446 | Z-0524-2014 | InstaTrak¿ 3500 Plus, ENTrakTM Plus, InstaTrak¿ 3500, ENTrakTM Computer-Assisted Navigation Systems (USA), | InstaTrak¿ 3500 Plus, ENTrakTM Plus, InstaTrak¿ 3500, ENTrakTM | 2 | CDRH | 2013-12-17 | 2013-12-17 21:00 | 2017-04-17 | GE OEC Medical Systems, Inc | GE OEC Medical Systems Inc is recalling InstaTrak¿ 3500 Plus, ENTrakTM Plus, InstaTrak¿ 3500, ENTrakTM Computer-Assisted Navigation Systems (USA) due to several potential safety issues related to the use of these products.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121816 | Z-0168-2014 | COBAS INTEGRA 800 Analyzer with software 9864.C2 Cobas. COBAS Part Numbers 28122474001, 28122474692, 04559126001, 04589459970 corresponding to the Analyzer with or without closed tube sampling accessory A fully automated system for clinical chemistry analysis intended for the in vitro quantitative/qualitative determination of analytes in body fluids. | COBAS INTEGRA 800 Analyzer | 2 | CDRH | 2013-11-05 | 2013-11-05 21:00 | 2015-05-12 | Roche Diagnostics Operations, Inc. | Roche COBAS INTEGRA 800 system, Software Version 9864.C2 does not perform necessary ISE Service Actions. Whenever the ISE rack is removed, the ISE counters in the software are reset to zero and the corresponding Service Actions are not conducted automatically by the system. Additionally, certain maintenance activities that have to be performed manually by the operator are not requested on the sof
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120684 | Z-2098-2013 | GE Ultrasound ViewPoint 6. Product Usage: ViewPoint is intended to be used in medical practices and in clinical departments and serves the purposes of diagnostic interpretation of images, electronic documentation of examinations in the form of text and images and generation of medical reports primarily for diagnostic ultrasound. ViewPoint provides the user the ability to including images, drawings, and charts into medical reports. ViewPoint is designed to accept, transfer, display, calculate, store and process medical images and data, and enables the user to measure and annotate the images. The medical images, which ViewPoint displays to the user, can be used for diagnostic purposes. ViewPoint is intended for professional use only. ViewPoint is not intended to be used as an automated diagnosis system. | GE Ultrasound ViewPoint 6 | 2 | CDRH | 2013-08-27 | 2013-08-27 21:00 | 2014-04-25 | GE Healthcare, LLC | GE has recently become aware of a potential safety issue due to the incorrect calculation of the pulsatility index value (PIV) for venous flow measurements associated with obstetric reporting of the GE ViewPoint 6 system. The ViewPoint calculated Doppler PIV for veins could be incorrect. An incorrect calculation of the ductus venosus PIV could lead to false high risk for trisomy 13, 18, and 21,
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=119700 | Z-2088-2013 | Dimension Vista 500 and Dimension Vista 1500 System Aliquot Well Double Dispense with software versions 3.5.1 and 3.6. In vitro diagnostic device intended to duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids. Dimension Vista chemical and immunochemical applications use photometric, turbidimetric, chemiluminescence, nephelometric, and integrated ion-selective multisensory technology for clinical use. | Dimension Vista Intelligent Lab Systems | 2 | CDRH | 2013-08-26 | 2013-08-26 21:00 | 2014-08-27 | Siemens Healthcare Diagnostics, Inc. | Siemens is conducting a field correction for Dimension Vista 500 and Dimension Vista 1500 systems using software versions 3.5.1 and 3.6. Under specific rare conditions which require a system reset, there is a potential for a sample to be dispensed into an aliquot well that has previously been used for a QC or calibrator product from a Vista vial. If two samples are dispensed into the same aliquot
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=120158 | Z-1931-2013 | MHI-TM2000 Linear Accelerator System, Software Version 3.0.0 revision 13 or earlier; Proprietary and Trade Name: MHI-TM2000/ VERO (TM); Common Name: Medical Linear Accelerator. MHI-TM2000 Linear Accelerator System is intended for radiation therapy of lesions, tumors and conditions anywhere in the body where radiation treatment is indicated | MHI-TM2000/ VERO (TM) | 2 | CDRH | 2013-08-08 | 2013-08-08 21:00 | 2017-03-23 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Software anomaly: Mitsubishi Heavy Industries, Ltd. became aware of a potential safety issue associated with the VERO (TM) Linear Accelerator System. The VERO (TM) may save treatment delivery record into wrong location under certain conditions.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=119815 | Z-1782-2013 | ARTIS zee and zeego x-ray, angiographic system with Software Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed. Angiographic x-ray system | ARTIS zee and zeego | 2 | CDRH | 2013-07-22 | 2013-07-22 21:00 | 2015-02-05 | Siemens Medical Solutions USA, Inc | Siemens issued a customer safety advisory notice and field safety corrective action for a potential issue on Artis systems with Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed. The issue affects-DSA rotational acquisition. The restriction is related to any rotational 3D-acquisition program with real-time subtraction (e.g. 3D DSA). During the fill phase of the rotational aqui
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=119750 | Z-1780-2013 | Radiohead Access Point 1.4 GHz Radiohead Access Point Firmware versions: C.00.04 & C.00.05 Part Number: 989803171211 Product Usage: IntelliVue Smart-hopping 1.4 GHz Access Point provides wireless connectivity between IntelliVue devices (listed below) and the Philips IntelliVue Information Center. MX40 Indicated for use by health care professionals whenever there is a need for monitoring the physiological parameters of patients. Intended for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults and pediatrics in hospital environments and during transport inside hospitals. The MX4O is to be used primarily as a traditional telemetry medical device. It connects to the IntelliVue Information Center by way of a wireless network. | Philips Healthcare IntelliVue Smart-hopping 1.4 GHz Access Point | 2 | CDRH | 2013-07-19 | 2013-07-20 21:00 | 2017-05-17 | Philips Healthcare Inc. | IntelliVue Smart-hopping 1.4 GHz Access Point may experience unexpected signal loss of several minutes impacting wireless connectivity