## Sheet1
WEB_ADDRESS | RECALL_NUMBER | PRODUCT_DESCRIPTION | TRADE_NAME | RECALL_CLASS | CENTER | CENTER_CLASSIFICATION_DT | POSTED_INTERNET_DT | TERMINATION_DT | FIRM_NAME | MANUFACTURER_RECALL_REASON
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211309 | Z-0743-2025 | Clarity Precision Grip Attachments, for use in orthodontic treatment. Part numbers/SKUs: 359-830 359-831 359-832 359-833 359-834 359-835 359-836 359-837 | 3M Clarity Precision Grip Attachments | 2 | CDRH | 2024-12-20 16:57 | 2024-12-23 20:07 |  | 3M Company - Health Care Business | Attachments may not mate with the tooth as intended in the digital treatment design.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210846 | Z-0523-2025 | Brand Name: Z-800 Infusion System Product Name: Z-800WF Model/Catalog Number: Z-800WF. The Z-800 Infusion system is intended to provide intravenous infusion of parenteral fluids, blood, and blood products to a patient under the direction or supervision physician or other certified healthcare professional. | Z-800 Infusion System | 3 | CDRH | 2024-11-22 17:27 | 2024-11-22 20:07 |  | Zyno Medical LLC | For the Z-800WF pumps, the WiFi software was not compatible with the pump software that had been loaded in accordance with ongoing correction Z-1183-2024.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203893 | Z-0469-2024 | EVO VISIAN Implantable Collamer Lens REF: VICMO13.2. Phakic Intraocular lens. | EVO VISIAN Implantable Collamer Lens | 2 | CDRH | 2023-12-01 16:56 | 2023-12-01 20:07 |  | Staar Surgical Company | Their is a potential that intraocular implant devices may not meet specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203894 | Z-0470-2024 | EVO+VISIAN Implantable Collamer Lens, REF: VTICM5_12.1. Phakic Toric Intraocular lens. | EVO+VISIAN Implantable Collamer Lens | 2 | CDRH | 2023-12-01 16:56 | 2023-12-01 20:07 |  | Staar Surgical Company | Their is a potential that intraocular implant devices may not meet specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203884 | Z-0464-2024 | EVO+VISIAN Implantable Collamer Lens, REF:VICM5_12.1. Phakic Intraocular lens. | EVO+VISIAN Implantable Collamer Lens | 2 | CDRH | 2023-12-01 16:56 | 2023-12-01 20:07 |  | Staar Surgical Company | Their is a potential that intraocular implant devices may not meet specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203886 | Z-0466-2024 | EVO+VISIAN Implantable Collamer Lens, REF: VICM5_13.2. Phakic Intraocular lens. | EVO+VISIAN Implantable Collamer Lens | 2 | CDRH | 2023-12-01 16:56 | 2023-12-01 20:07 |  | Staar Surgical Company | Their is a potential that intraocular implant devices may not meet specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203895 | Z-0471-2024 | EVO+VISIAN Implantable Collamer Lens, REF: VTICM5_12.6. Phakic Toric Intraocular lens. | EVO+VISIAN Implantable Collamer Lens | 2 | CDRH | 2023-12-01 16:56 | 2023-12-01 20:07 |  | Staar Surgical Company | Their is a potential that intraocular implant devices may not meet specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203898 | Z-0473-2024 | EVO+VISIAN Implantable Collamer Lens, REF:VTICM5_13.7. Phakic Toric Intraocular lens. | EVO+VISIAN Implantable Collamer Lens | 2 | CDRH | 2023-12-01 16:56 | 2023-12-01 20:07 |  | Staar Surgical Company | Their is a potential that intraocular implant devices may not meet specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203885 | Z-0465-2024 | EVO+VISIAN Implantable Collamer Lens, REF: VICM5_12.6. Phakic Intraocular lens. | EVO+VISIAN Implantable Collamer Lens | 2 | CDRH | 2023-12-01 16:56 | 2023-12-01 20:07 |  | Staar Surgical Company | Their is a potential that intraocular implant devices may not meet specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203887 | Z-0467-2024 | EVO +VISIAN Implantable Collamer Lens, REF: VICM5_13.7.Phakic Intraocular lens. | EVO+VISIAN Implantable Collamer Lens | 2 | CDRH | 2023-12-01 16:56 | 2023-12-01 20:07 |  | Staar Surgical Company | Their is a potential that intraocular implant devices may not meet specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203888 | Z-0468-2024 | EVO VISIAN Implantable Collamer Lens, REF: VICMO12.1. Phakic Intraocular lens. | EVO VISIAN Implantable Collamer Lens, | 2 | CDRH | 2023-12-01 16:56 | 2023-12-01 20:07 |  | Staar Surgical Company | Their is a potential that intraocular implant devices may not meet specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203897 | Z-0472-2024 | EVO+VISIAN Implantable Collamer Lens, REF: VTICM5_13.2. Phakic Toric Intraocular lens. | EVO+VISIAN Implantable Collamer Lens | 2 | CDRH | 2023-12-01 16:56 | 2023-12-01 20:07 |  | Staar Surgical Company | Their is a potential that intraocular implant devices may not meet specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203899 | Z-0474-2024 | EVO VISIAN Implantable Collamer Lens, REF: VICMO12.6. Phakic Intraocular lens. | EVO VISIAN Implantable Collamer Lens | 2 | CDRH | 2023-12-01 16:56 | 2023-12-01 20:07 |  | Staar Surgical Company | Their is a potential that intraocular implant devices may not meet specifications.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=198972 | Z-1322-2023 | Philips DreamStation Auto BiPAP and CPAP, Model Numbers UCDSX500S11, UDSX500S11, UDSX500S11F, UDSX700S11, UDSX700S11F, UFRX500S14. Used to deliver positive airway pressure therapy for the treatment of Obstructive Sleep Apnea. | DreamStation1 (Uno Remediated Device) | 1 | CDRH | 2023-04-06 14:21 | 2023-04-06 21:08 |  | Philips Respironics, Inc. | A limited number of remediated Philips DreamStation units may experience communication issues when connecting to the cloud-based care management application. Should this be the case, this may affect the ability to download prescription settings to the device or the device may have incorrect prescription settings for the patient.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=191283 | Z-0550-2022 | OXY-1 System, Model Number 001-0500-001. Used during cardiopulmonary bypass. OXY-1 System consists of Console, Pump Drive, E Drive, and Pole Mount. The disposable Pump Lung Unit PLU part number 001-0400-100 is sold separately and is not part of this recall. | OXY-1 System for extracorporeal circulation | 1 | CDRH | 2022-02-03 16:41 | 2022-02-03 20:06 | 2023-02-28 17:30 | Abiomed, Inc. | The firm has received complaints of the OXY-1 System Console experiencing a power interruption while in use, which disrupts therapy delivered to a patient on support. Disruption of therapy could lead to prolonged hypoxia.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190716 | Z-0478-2022 | OCT-Camera 211 01 A3; Running software iOCT-Control 3 v1.4 | Haag Streit Surgical | 2 | CDRH | 2022-01-10 |  |  | Haag-Streit USA Inc | Malfunction of the automatic laser beam shut-off, the OCT unit might not recognize whether the laser beam is safely switched off
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190715 | Z-0477-2022 | OCT-Camera 211 01 A1, Running software iOCT-Control 3 v1.4 | Haag Streit Surgical | 2 | CDRH | 2022-01-10 |  |  | Haag-Streit USA Inc | Malfunction of the automatic laser beam shut-off, the OCT unit might not recognize whether the laser beam is safely switched off
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190077 | Z-0333-2022 | Boston Scientific AUTOLITH TOUCH Bipolar Electrohydraulic Lithotripter, REF M005466800. The device is designed to be used with Northgate Technologies Inc. bipolar disposable EHL probes for the fragmentation of calculi. | AUTOLITH TOUCH Bipolar Electrohydraulic Lit | 2 | CDRH | 2021-12-03 16:52 |  |  | Northgate Technologies, Inc. | The languages for Russian, Romanian, Slovak, and Czech have the power settings for high and low switched in the device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187867 | Z-2098-2021 | SoftCytogenetics | SoftCytogenetics | 2 | CDRH | 2021-07-19 |  | 2023-02-14 10:48 | Soft Computer Consultants, Inc. | Software is given wrong results going out to HIS on test that had reruns.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184732 | Z-0860-2021 | Medtronic Intellis Spinal Cord Stimulator, Model Number A710, Clinical Programmer application version 1.3.80 - Product Usage: intended for use by clinicians in the programming of the following Medtronic neurostimulators (external and implantable) for pain therapy. | Intellis Spinal Cord Stimulator | 2 | CDRH | 2021-01-14 |  |  | Medtronic Neuromodulation | A710 Intellis Clinician Application has a software issue that can result in the inability to program the Intellis implantable neurostimulation device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180336 | Z-2479-2020 | BodyGuardian Heart Remote Monitoring Kit, consisting of two BG Heart monitors, either Models BC-1001-A or BC1001-D, and a cell phone, either Models BP1008 or BP10014, that is paired to both of the monitors. | BodyGuardian Heart Remote Monitoring Kit | 2 | CDRH | 2020-07-01 |  | 2022-11-18 16:42 | Preventice Services, LLC | The device data being collected and transferred to the monitoring center may not be accurate due to nonvalidated association between the phone software and the heart monitors, therefore, the patient's report should not be used to evaluate their condition.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180858 | Z-1888-2020 | Abbott Accelerator a3600 Centrifuge Module tube loading algorithm, Inpeco SA P/N: FLX-202, with Firmware Versions: CM_1-7-0.H86 or xCMd_2-5-0.elf and CM_1-7-1.H86 or xCMd_2-5-1-849fad9.elf | Inpeco SA | 2 | CDRH | 2020-05-05 07:00 |  |  | Inpeco S.A. | The FlexLab Centrifuge Module loading algorithm may lead to unbalanced loads which could damage the centrifuge.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180838 | Z-1887-2020 | Inpeco FlexLab Centrifuge Module tube loading algorithm, Inpeco SA P/N: FLX-202, with Firmware Versions: CM_1-7-0.H86 or xCMd_2-5-0.elf and CM_1-7-1.H86 or xCMd_2-5-1-849fad9.elf | Inpeco SA | 2 | CDRH | 2020-05-05 07:00 |  |  | Inpeco S.A. | The FlexLab Centrifuge Module loading algorithm may lead to unbalanced loads which could damage the centrifuge.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180873 | Z-1889-2020 | Siemens Aptio Automation Centrifuge Module, tube loading algorithm, Inpeco SA P/N: FLX-202, with Firmware Versions: CM_1-7-0.H86 or xCMd_2-5-0.elf and CM_1-7-1.H86 or xCMd_2-5-1-849fad9.elf | Inpeco SA | 2 | CDRH | 2020-05-05 07:00 |  |  | Inpeco S.A. | The FlexLab Centrifuge Module loading algorithm may lead to unbalanced loads which could damage the centrifuge.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175422 | Z-2553-2019 | Gold Cup Leadwire and Snap Leadwire with the following component descriptions: 1. 24" Snap Leadwire, Blue; 2. 40" Snap Leadwire, Blue; 3. 24" Snap Leadwire, Red; 4. 40" Snap Leadwire, Red; 5. 24" Snap Leadwire, Green; 6. 40" Snap Leadwire, Green; 7. 24" Snap Leadwire, Brown; 8. 40" Snap Leadwire, Brown; 9. 24" Snap Leadwire, Orange; 10. 40" Snap Leadwire, Orange; 11. 24" Snap Leadwire, White; 12. 40" Snap Leadwire, White; 13. 24" Snap Leadwire, Black; 14. 40" Snap Leadwire, Black; 15. 24" Snap Leadwire, Yellow; 16. 40" Snap Leadwire, Yellow; 17. 24" Snap Leadwire, Set of 5; 18. 24" Snap Leadwire, Set of 7; 19. 40" Snap Leadwire, Set of 5; 20. 40" Snap Leadwire, Set of 7; 21. 24" Snap Leadwire, Set of 5; 22. 40" Snap Leadwire, Set of 5; 23. 48" Gold Cup Leadwires, Blue; 24. 48" Gold Cup Leadwires, Red; 25. 48" Gold Cup Leadwires, Green; 26. 48" Gold Cup Leadwires, Brown; 27. 48" Gold Cup Leadwires, Orange; 28. 48" Gold Cup Leadwires, White; 29. 48" Gold Cup Leadwires, Black; 30. 48" Gold Cup Leadwires, Yellow; 31. 48" Gold Cup Leadwires, Set of 5; 32. 48" Gold Cup Leadwires, Set of 8; Product Usage: The snap electrode leads are used in conjunction with disposable electrodes and the Gold Cup electrodes are reusable electrodes and leads. Both types of electrodes and leads are used for evoked potential testing for audiological, neurological and/or vestibular testing. The electrode leads are connected to the main device by the electrode lead connector and the electrode is attached to the patient contact end of electrode lead. | Gold Cup Electrode and Snap Electrode Leads | 2 | CDRH | 2019-09-23 |  |  | GN Otometrics | The electrode lead connector has a protruding conductor beyond insulation of approximately 0.1 mm.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=169583 | Z-0612-2019 | Stationary Fluoroscopic X-ray System CombiDiagnost R90, Model 709030 Stationary Fluoroscopic X-ray System | CombiDiagnost R90 | 2 | CDRH | 2019-02-12 |  | 2020-09-02 10:25 | Philips Medical Systems Gmbh, DMC | kV/mA lockin not functioning as specified which may result in the patient receiving an increased radiation dose.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167920 | Z-0683-2019 | Medtronic MiniMed 630G 6025805-003_C, Rx Only Medtronic MiniMed 670G 6025805-005_B, Rx Only Medtronic MiniMed 640G (International) | Medtronic MiniMed 630G 6025805-003_C, 670G 6025805-005_B | 2 | CDRH | 2018-12-22 |  |  | Medtronic Inc. | Reports of occurrences in which insulin pumps with version 4.10 software have failed to make expected audio sounds during alerts, alarms or sirens. This could cause a user to miss system notifications, alarms or sirens associated with how the pump is working, and with high and low glucose alerts.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=165408 | Z-2572-2018 | VNS Therapy Programming System, Rx Only, Model 3000, v1.0.2.2 | VNS Therapy Programming System | 2 | CDRH | 2018-07-28 |  | 2023-11-29 14:20 | LivaNova USA Inc | Unintended warning message displayed on generators programmed with a Model 3000 v.1.0.2.2 programmer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164356 | Z-1930-2018 | Biograph Horizon, PET/CT System using VJ10A, VJ10B, VJ20A scanners that provide registration and fusion of high resolution physiologic and anatomic information. | Siemens Biograph Horizon | 2 | CDRH | 2018-05-24 |  | 2018-10-12 | Siemens Medical Solutions USA, Inc. | Possibility that the Biograph Horizon systems performing CT retrospective cardiac gating or PET cardiac gating examinations may experience a waveform sampling issue caused by a firmware change within the Universal Physiological Measurement Module.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=155157 | Z-0254-2018 | enGen Track System with TCAutomation Software Version 4.2 | enGen Laboratory Automation System | 2 | CDRH | 2017-12-14 |  | 2021-08-23 14:18 | Ortho-Clinical Diagnostics | A software anomaly with TCA Software V4.2 may potentially cause a delay in reporting of results due to patient samples being routed to an unintended location when samples are reintroduced back onto the enGEN track
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=155590 | Z-2601-2017 | Merge OrthoPACS software. The firm name on the label is Merge Healthcare, Hartland, WI. | Merge OrthoPACS | 2 | CDRH | 2017-06-16 |  | 2019-10-15 07:51 | Merge Healthcare, Inc. | Studies that have QC, PDE, or HL7 ADT updates applied may fail to propagate via telmed protocol to the destination server and may fail to import on the destination server.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=155552 | Z-2591-2017 | Merge PACS software. Merge PACS (Picture Archiving Communication System) is designed and marketed for soft copy reading, communication, and storage of studies produced by digital modalities, including digital mammography | Merge PACS | 2 | CDRH | 2017-06-15 |  | 2019-08-08 13:49 | Merge Healthcare, Inc. | When cut lines are shown on an image, the lines may go from being vertical and then shift to horizontal causing confusion to the physician. This may cause a misdiagnosis.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=155551 | Z-2590-2017 | Merge PACS software. The firm name on the label is Merge Healthcare, Hartland, WI. | Merge PACS | 2 | CDRH | 2017-06-15 |  | 2019-08-05 17:45 | Merge Healthcare, Inc. | Merge PACS did not show unviewed images when the last view was skipped. There is a potential risk to health of a physician misdiagnosis because not all images are available for viewing.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=154482 | Z-1770-2017 | Combi Dockable Neurosurgery table for the Siemens MAGNETOM Aera (1.5T) and Skyra (3T) MRI Systems | Combi Dockable Neurosurgery table for the Siemens MAGNETOM Aera (1.5T) and Skyra (3T) MRI Systems | 2 | CDRH | 2017-04-10 |  | 2017-10-24 | Siemens Medical Solutions USA, Inc | Due to a potential manufacturing error, a rare possibility exists for a sudden drop in table height of the Combi Dockable Table Neurosurgery.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152047 | Z-1499-2017 | Arial Wireless Water-Resistant Call Pendant The Arial Pendant tag is part of the Arial wireless emergency call management system. It enables residents in assisted living, skilled nursing or independent living to call staff with the press of a button. | Arial Water-Resistant Pendant, model 59350 | 2 | CDRH | 2017-03-20 |  | 2017-11-20 | Stanley Security Solutions Inc | Devices were incorrectly programmed during manufacturing therefore depressing the pendant button may result in an alarm not sounding as intended.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152049 | Z-1098-2017 | Softbank software Product Usage: Decision support software for transfusion service. It keeps track of inventory from outside sources, multi-site inventory control, records of testing of units, and allows for record keeping for transfusion preparation. | Softbank 25.5.2.0-25.5.3.1 | 2 | CDRH | 2017-01-26 |  | 2019-08-21 10:16 | Soft Computer Consultants, Inc. | Software error. Potential for incorrect results
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152048 | Z-1097-2017 | SCC Soft Computer Softbank II software Product Usage: Supports single and multi-site transfusion services in healthcare facilities. Used by healthcare personal to document, query, and view the integrated information regarding patients and products. Quality control testing, test and transfusion history, transfusion management, inventory management, product distribution, and final disposition are all monitored using the software. | Softbank II 23.1.2.8 -25.3.0.0 | 2 | CDRH | 2017-01-26 |  | 2019-08-21 10:16 | Soft Computer Consultants, Inc. | Software error. Potential for incorrect results
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=149743 | Z-0586-2017 | Version 1.2.0 of Accu-Chek Connect Diabetes Management App (iOS) released on July 11, 2016 | ACCU-CHEK Connect Diabetes Management App | 2 | CDRH | 2016-11-16 | 2016-11-16 21:00 | 2017-06-06 | Roche Diabetes Care, Inc. | Roche Diabetes Care, Inc. initiated a voluntary recall for the version 1.2.0 of Accu-Chek Connect Diabetes Management App (iOS), released on July 11, 2016, due to the software containing a programming error (bug) in the Bolus Advisor feature, which could lead to an incorrect insulin bolus recommendation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148502 | Z-2629-2016 | OEC 9800, fluoroscopic x-ray system | OEC 9800 | 2 | CDRH | 2016-08-23 |  | 2016-12-19 | GE OEC Medical Systems, Inc | A firmware issue in the supplied LCD workstation monitor(s). Under some circumstances, the monitor(s) may appear dark with no video displayed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148501 | Z-2628-2016 | OEC UroView 2800, fluoroscopic x-ray system | OEC UroView 2800 | 2 | CDRH | 2016-08-23 |  | 2016-12-19 | GE OEC Medical Systems, Inc | A firmware issue in the supplied LCD workstation monitor(s). Under some circumstances, the monitor(s) may appear dark with no video displayed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=147531 | Z-2118-2016 | Ingenuity Core 128-Computed Tomography X-ray system Product Usage: The Ingenuity Core 128 scanner is a whole body Computed Tomography X-ray system intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes.. | Ingenuity Core 128 | 2 | CDRH | 2016-06-30 |  | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field that certain Ingenuity Core128 systems running software version 3.5.4 exhibited intermittent swirl-like ring artifacts that may appear on reconstructed images. A patient rescan may be required if the images cannot be used for interpretation due to the swirl-like artifact.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=147530 | Z-2117-2016 | Ingenuity Core-Computed Tomography X-ray system Product Usage: The Ingenuity Core scanner is a whole body Computed Tomography X-ray system intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. | Ingenuity Core | 2 | CDRH | 2016-06-30 |  | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field that certain Ingenuity Core systems running software version 3.5.4 exhibited intermittent swirl-like ring artifacts that may appear on reconstructed images. A patient rescan may be required if the images cannot be used for interpretation due to the swirl-like artifact.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=147502 | Z-2116-2016 | Brilliance 64, Computed Tomography X-ray system Product Usage: The Brilliance CT 64 scanner is a whole body Computed Tomography X-ray system intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes | Brilliance 64 | 2 | CDRH | 2016-06-30 |  | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field that certain Brilliance 64 systems running software version 3.5.4 exhibited intermittent swirl-like ring artifacts that may appear on reconstructed images. A patient rescan may be required if the images cannot be used for interpretation due to the swirl-like artifact.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=145304 | Z-1688-2016 | LenSx Laser System. Indicated for use in patients undergoing cataract surgery and refractive surgery. | LenSx Laser System | 2 | CDRH | 2016-05-17 |  | 2017-10-20 | Alcon Research, Ltd. | Alcon is conducting a voluntary recall of all LenSx Lasers with Corneal Flap Capability after receiving reports concerning incomplete corneal flap creation, which manifests as an un-dissected area(s) within the side cut.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144965 | Z-1605-2016 | Medtronic CareLink" Monitor (2490C) and Medtronic CareLink Express" Monitor (2020B) Product Usage: The 2490C CareLink Home Monitors and 2020B CareLink Express Monitors are remote monitoring systems that interrogate implanted devices and transmit the data to Medtronic s CareLink Network for viewing by the physician. The data is transmitted either through an analog telephone line, a cellular connection or an internet connection. | Carelink and Carelink Express Monitor | 2 | CDRH | 2016-05-04 |  | 2017-09-25 | Medtronic Inc., Cardiac Rhythm and Heart Failure | A recent firmware update developed by Medtronic for the 2490C CareLink Monitors and 2020B CareLink Express Monitors included incorrect data on the country analog modem dial-up configuration table used to dial into the CareLink Network. This incorrect data resulted in the removal of the 0 prefix necessary for select countries as well as an incorrect alignment of phone numbers for other select cou
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=141861 | Z-0564-2016 | Radiographic Transparency for Certain and External Connection Tapered Implants. Images that assist with measurements for dental implants. | BIOMET 3i ART 1036 Dental Implant Radiographic Transparency | 2 | CDRH | 2015-12-31 | 2015-12-31 21:01 | 2016-12-14 | Biomet 3i, LLC | Images contained within radiographic transparency ART 1036 are undersized.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138104 | Z-0205-2016 | Arthroscopic Energy 50¿ Probe with Suction, Catalog number: AES-50S Peoduct Usage: For resection, ablation, and coagulation of soft tissue and hemostasis of blood vessels in arthroscopic and orthopedic surgical procedures. | AES-50S | 2 | CDRH | 2015-11-03 | 2015-11-03 21:00 | 2017-04-11 | ConMed Corporation | Probe programming error. The incorrect program can generate more heat at the surgical site, resulting in an elevated probability of tissue damage.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=140819 | Z-0170-2016 | ARIA Radiation Oncology, versions 10, 11 , 13.0 and 13.5 with Clinical Assessment License. Model number HIT Product Usage: ARIA Radiation Oncology is a radiation treatment plan and image management application. | ARIA Radiation Oncology | 2 | CDRH | 2015-10-27 | 2015-10-30 17:47 | 2017-05-12 | Varian Medical Systems, Inc. | An anomaly was identified with the ARIA for Radiation Oncology software with a Clinical Assessment license. Admin Instructions may not print on drug order prescriptions or be transmitted with e-Rx.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139329 | Z-2812-2015 | ARTISTE, ONCOR Avantgarde, ONCOR Impression Plus, ONCOR Expression, PRIMUS, PRIMUS Plus, Syngo RT Therapist; therapeutic treatment of cancer Product Usage: The intended use of the SIEMENS branded ARTISTE and ONCOR family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer. PRIMUSTM family of linear accelerator systems is to deliver x-ray radiation for therapeutic treatment of cancer. Syngo RT Therapist is a software application whose indication for use includes the viewing, processing, filming, and archiving of medical images. It also permits patient data management, patient selection/setup, patient positioning verification, treatment planning, treatment delivery/verification, and treatment recording. | ARTISTE, ONCOR Avantgarde, Impression Plus, and Expression, PRIMUS, PRIMUS Plus, Syngo RT Therapist | 2 | CDRH | 2015-09-23 | 2015-09-24 21:00 | 2016-09-01 | Siemens Medical Solutions USA, Inc | A software fix has been released to prevent automatic movement resulting in a collision safety risk for patients.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139463 | Z-2469-2015 | Centricity PACS-IW with Universal Viewer Product Usage: The Centricity PACS-IW with Universal Viewer is a picture archiving and communications system, a device that provides one or more capabilities relating to the acceptance, transfer, display, storage, and digital processing of medical images. Its hardware components may include workstations, digitizers, communications devices, computers, video monitors, magnetic, optical disk, or other digital data storage devices, and hardcopy devices. The software components may provide functions for performing operations related to image manipulation, enhancement, compression or quantification. Used in Radiology. | GE Healthcare Inc. | 2 | CDRH | 2015-08-26 | 2015-08-27 21:02 | 2017-01-05 | GE Healthcare | Images from the Centricity PACS-IW with Universal Viewer and Centricity Universal Viewer may be missing when a system parameter MapRoute is set to a value greater than 1.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139464 | Z-2470-2015 | Centricity Universal Viewer Product Usage: The Centricity Universal Viewer Version 6.0 software is a picture archiving and communications system, a device that provides one or more capabilities relating to the acceptance, transfer, display, storage, and digital processing of medical images. Its hardware components may include workstations, digitizers, communications devices, computers, video monitors, magnetic, optical disk, or other digital data storage devices, and hardcopy devices. The software components may provide functions for performing operations related to image manipulation, enhancement, compression or quantification. Used in Radiology. | GE Healthcare Inc. | 2 | CDRH | 2015-08-26 | 2015-08-27 21:02 | 2017-01-05 | GE Healthcare | Images from the Centricity PACS-IW with Universal Viewer and Centricity Universal Viewer may be missing when a system parameter MapRoute is set to a value greater than 1.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=137300 | Z-2054-2015 | Siemens Healthcare Diagnostics Dimension Vista Systems, used with the Dimension Vista Intelligent Lab System software version 3.6.1. In vitro diagnostic analyzer. | Dimension Vista Intelligent Lab System | 2 | CDRH | 2015-07-09 | 2015-07-09 21:00 | 2017-03-10 | Siemens Healthcare Diagnostics, Inc. | Two software issues may occur in Vista software version 3.6.1. Issue #1: Samples may stop processing without notification. Issue #2: Timing issue in Vista 1500 may cause wrong reagent or no reagent delivery. This may lead to unflagged, unexpected low results and results flagged with assay errors.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=137356 | Z-1918-2015 | StoneLight 30 Laser System (NS3000). For use in surgical procedures using open, laparoscopic and endoscopic incision, excision, resection, ablation, vaporization, coagulation and hemostasis of soft tissue in use in medical specialties including: Urology, Urinary Lithotripsy, and General Surgery. | StoneLight 30 Laser System (NS3000) | 2 | CDRH | 2015-06-26 | 2015-06-27 21:00 | 2016-10-28 | New Star Lasers, Inc. | During internal testing a situation was discovered where when changing the settings, a new setting could not be accepted and potentially result in a discrepancy in the energy delivered (higher or lower) from the value displayed on the Graphical User Interface display.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135515 | Z-1554-2015 | Philips Healthcare DigitalDiagnost System X-Ray | Philips Healthcare DigitalDiagnost System X-Ray | 2 | CDRH | 2015-05-21 | 2015-05-19 21:03 | 2016-06-14 | Philips Medical Systems, Inc. | The system is designed to emit a beep upon termination of an exposure. However, if the system has been powered on for over 12 hours, the system will no longer emit this signal. This is a failure to comply with 21CFR 1020.31(j).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135516 | Z-1555-2015 | Philips Healthcare DuraDiagnost X- Ray | Philips Healthcare DuraDiagnost X- Ray | 2 | CDRH | 2015-05-21 | 2015-05-19 21:03 | 2016-06-14 | Philips Medical Systems, Inc. | The system is designed to emit a beep upon termination of an exposure. However, if the system has been powered on for over 12 hours, the system will no longer emit this signal. This is a failure to comply with 21CFR 1020.31(j).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135485 | Z-1574-2015 | MHI-TM2000 Linear Accelerator System (Software Version 3.5.0 and 3.5.1) | MHI-TM2000 Linear Accelerator System | 2 | CDRH | 2015-05-01 | 2015-05-01 21:01 | 2017-04-11 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | The operator console allows users to deliver therapeutic radiation to patients even though a specific communication error relevant to imaging conditions has occurred. In addition, the treatment record (the delivered radiation record) cannot be saved.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135639 | Z-1561-2015 | AMSCO C and AMSCO 400 Steam Sterilizers, STERIS Corporation. AMSCO C Small Sterilizer, AMSCO 400 Small Sterilizer, AMSCO 400 Medium Sterilizer. Designed for sterilization of heat and moisture-stable materials used in healthcare facilities. | AMSCO C and AMSCO 400 Steam Sterilizers | 2 | CDRH | 2015-04-30 | 2015-04-30 21:00 | 2016-01-20 | Steris Corporation | STERIS has identified that the control board software in select AMSCO 400 and AMSCO C units will interrupt and cancel a processing cycle should the selected sterilizer shutdown time coincide with a processing cycle. This may result in a procedure delay.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=134007 | Z-1212-2015 | COMPASS SW Version 3.1, Catalog Number CS10-100, medical linear accelerator, radiological | COMPASS SW Version 3.1, Catalog Number CS10-100 | 2 | CDRH | 2015-02-26 | 2015-02-26 21:00 | 2015-08-18 | Iba Dosimetry Gmbh | Error in the software. During internal tests of the current development version of the Compass SW it was found that dose reconstruction for DMLC plans when the jaws move or mlc leaves reverse during beam on will not be scaled correctly for all control points.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130189 | Z-1034-2015 | Animas Vibe Insulin Infusion Pump. This product is indicated for continuous subcutaneous infusion of insulin for the treatment of diabetes and has a continuous glucose monitoring feature. | Animas Vibe Insulin Infusion Pump | 2 | CDRH | 2015-01-30 | 2015-01-30 21:00 | 2015-08-10 | Animas Corporation | Calibration factors in the pump overwritten during a programming step. The force sensor could send a lower signal value to the pump processor, with loss of prime warnings, occlusion alarms and the pump unable to detect a cartridge during the prime sequence. Field action initiated 8/29/2011.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130952 | Z-0993-2015 | TEGRIS System manufactured by MAQUET GMBH in Germany The Maquet Tegris OR Integration System is designed to be used as the central operating unit in an operating room. The integration system has two main functions: Recording and distribution of images and videos and interaction with medical and non-medical devices. | Maquet Medical Systems | 2 | CDRH | 2015-01-16 | 2015-01-16 21:00 | 2015-11-17 | Maquet Medical Systems USA | The wrong movement results on the MAGNUS OR table. The button commands on the TEGRIS touchscreen for Lower Leg Up and Upper Leg Down are switched in the software for the integration with the MAGNUS operating table system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131533 | Z-0836-2015 | adaPT insight (12C), v1.3.2, PAT.109 (US), Treatment room 4: I2C is used with a charged particle or photon radiation therapy system for localization of the patient position with respect to the therapy equipment and to provide correction feedback to the radiation therapy device. | adaPT insight (12C), v1.3.2 | 2 | CDRH | 2014-12-19 | 2014-12-19 21:00 | 2015-06-29 | Ion Beam Applications S.A. | Two issues were identified: 1. In the release of adaPT insight v1.3.2., the kV/kV Single Source acquisition workflow was linked to a wrong geometrical calibration file, resulting in an alignment offset for this specific mode of acquisition. 2. A gantry collision incident required the IBA service team to verify the gantry alignment and protons I x-ray beams colinearity, du
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130892 | Z-0463-2015 | Abbott m2000sp, an automated system for performing sample preparation for nucleic acid testing. | Abbott Molecular Abbott m2000sp | 2 | CDRH | 2014-11-25 | 2014-11-25 21:00 | 2016-03-08 | Abbott Molecular | Some versions of Application Specifications (App Spec) are incompatible with m2000sp system software version 6.0 and 7.0. This issue will cause Error Code 9000 (An unexpected error has occurred: Software error) to be generated when the Available Reagent Packs screen is selected.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130364 | Z-0042-2015 | McKesson Cardiology" Hemo intended for complete physiological/hemodynamic monitoring, clinical data acquisition, medical image and data processing, and analytical assessment. | McKesson Cardiology Hemo System | 2 | CDRH | 2014-10-08 | 2014-10-08 21:00 | 2016-05-18 | Mckesson Information Solutions LLC | Software Error: The McKesson Cardiology Hemo calculation section incorrectly converts the Hemoglobin value before it is utilized in the applicable formula calculations.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=129577 | Z-2662-2014 | ADVIA Centaur Immunoassay System, Catalog Numbers/Siemens Material Numbers (SMN) 10284980, 10286140, 10309524, 10309525, 10310210, 10313282, 10314322, 10316248, 10316372, 10316968, 10317060, 10317403, 10319111, 10319433, 10320929, 10321568, 10322149, 10322731, 10323204, 10325015, 10326217, 10327008, 10327379, 10328250, 10328647, 10329364, 10330873, 10331013, 10332617, 10334139, 10334759, 10337512, 10337526, 10339677, 10340551, 10340737, 10341051, 10341110, 10361010, 10361011, 10361012 --- For in vitro diagnostic use This immunoassay system is intended for professional use in a laboratory environment only. Tests performed using this system are intended for in vitro diagnostic use. | ADVIA Centaur Immunoassay System | 2 | CDRH | 2014-09-18 | 2014-09-18 21:00 | 2017-05-02 | Siemens Healthcare Diagnostics | The firm is conducting a field correction for the ADVIA Centaur and ADVIA Centaur XP Immunoassay systems, regarding manually loaded sample racks that are improperly loaded or pushed too far on the sample entry queue during normal operation of the ADVIA Centaur and ADVIA Centaur XP Immunoassay systems. Pushing the sample racks may cause misreads of Sample IDs (SID) if the sample rack loading instructions are not followed exactly as described in the ADVIA Centaur or ADVIA Centaur XP Operators Guide.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=129578 | Z-2663-2014 | ADVIA Centaur XP Immunoassay System, Catalog Numbers/Siemens Material Numbers (SMN) 10285219, 10316507, 10317207, 10317284, 10319668, 10320757, 10323213, 10324519, 10327135, 10327836, 10328940, 10329339, 10336292, 10338631, 10364455, 10388696, 10471899 --- For in vitro diagnostic use This immunoassay system is intended for professional use in a laboratory environment only. Tests performed using this system are intended for in vitro diagnostic use. | ADVIA Centaur XP Immunoassay System | 2 | CDRH | 2014-09-18 | 2014-09-18 21:00 | 2017-05-02 | Siemens Healthcare Diagnostics | The firm is conducting a field correction for the ADVIA Centaur and ADVIA Centaur XP Immunoassay systems, regarding manually loaded sample racks that are improperly loaded or pushed too far on the sample entry queue during normal operation of the ADVIA Centaur and ADVIA Centaur XP Immunoassay systems. Pushing the sample racks may cause misreads of Sample IDs (SID) if the sample rack loading instructions are not followed exactly as described in the ADVIA Centaur or ADVIA Centaur XP Operators Guide.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=128235 | Z-2044-2014 | Lynx TM20 Torquemaster Low Speed Handpiece MTI Dental Products | Lynx TM20 Torquemaster Low Speed Handpiece | 2 | CDRH | 2014-07-12 | 2014-07-12 21:00 | 2015-03-12 | MTI Precision Products LLC. | Sharp edge on the body/housing of the Lynx TM20 TorqueMaster Low Speed Handpiece.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124800 | Z-0910-2014 | MooreMedical Supporting Health & Care...MooreBrand Blood Glucose Monitor for Self-Testing & Multi-Patient Use TRUEtrack. Manufactured by: NIPRO DIAGNOSTICS, Fort Lauderdale, FL 33309 U.S.A. Is intended for the quantitative determination of glucose in human whole blood taken from finger or forearm. | MooreMedical Supporting Health & Care...MooreBrand Blood Glucose Monitor for Self-Testing & Multi-Pa | 2 | CDRH | 2014-02-03 | 2014-02-03 21:00 | 2016-01-28 | Nipro Diagnostics, Inc. | Manufactured with an incorrect factory set unit of measure.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124790 | Z-0909-2014 | MooreMedical Supporting Health & Care...MooreBrand Blood Glucose Monitor for Self-Testing & Multi-Patient Use TRUEbalance. Manufactured by: NIPRO DIAGNOSTICS, Fort Lauderdale, FL 33309 U.S.A. Is intended for the quantitative determination of glucose in human whole blood taken from finger or forearm. | MooreMedical Supporting Health & Care...MooreBrand Blood Glucose Monitor for Self-Testing & Multi-Pa | 2 | CDRH | 2014-02-03 | 2014-02-03 21:00 | 2016-01-28 | Nipro Diagnostics, Inc. | Manufactured with an incorrect factory set unit of measure.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123002 | Z-0544-2014 | PedFuse Reset Screw; Model: 01-80115-55 The PedFuse Reset pedicle screw implants intended to be used to help provide immobilization and stabilization of spinal segments as an adjunct to fusion of the thoracic, lumbar, and/or sacral spine. | PedFuse Reset Screw | 2 | CDRH | 2013-12-19 | 2013-12-19 21:00 | 2016-08-22 | SpineFrontier, Inc. | During routine servicing, it was discovered that one 8mmx55mm Reset screw was assembled with an 8mm tulip head and a 9mm screw. The implant should have consisted of an 8mm tulip head and an 8mm screw. A recall was initiated to remove all remaining screws of the same lot from the field to investigate the nonconformity.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124194 | Z-0520-2014 | Mega Needle Driver, 8MM assembly of the da Vinci S IS1200 and da Vinci Si, IS2000. Manufactured by Intuitive Surgical, Sunnyvale, CA. EndoWrist Instruments, including scissors, scalpels, forceps, needle drivers and electrocautery are intended for endoscopic manipulation of tissue, including: grasping, cutting, blunt and sharp dissection, approximation, ligation, electrocautery and suturing. Mega Needle Driver and Large Needle Driver instruments are part of the EndoWrist instrument family and are intended for use in endoscopic manipulation and suturing. | Mega Needle Driver, 8MM assembly of the da Vinci S IS1200 and da Vinci Si, IS2000 | 2 | CDRH | 2013-12-16 | 2013-12-16 21:00 | 2014-09-04 | Intuitive Surgical, Inc. | Firm has become aware of the potential for detachment of the jaw insert for Large Needle Drivers and Mega Needle Drivers.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124238 | Z-0521-2014 | Large Needle Driver, 8MM assembly of the da Vinci S IS1200 and da Vinci Si, IS2000. Manufactured by Intuitive Surgical, Sunnyvale, CA. EndoWrist Instruments, including scissors, scalpels, forceps, needle drivers and electrocautery are intended for endoscopic manipulation of tissue, including: grasping, cutting, blunt and sharp dissection, approximation, ligation, electrocautery and suturing. Mega Needle Driver and Large Needle Driver instruments are part of the EndoWrist instrument family and are intended for use in endoscopic manipulation and suturing. | Large Needle Driver, 8MM assembly of the da Vinci S IS1200 and da Vinci Si, IS2000 | 2 | CDRH | 2013-12-16 | 2013-12-16 21:00 | 2014-09-04 | Intuitive Surgical, Inc. | Firm has become aware of the potential for detachment of the jaw insert for Large Needle Drivers and Mega Needle Drivers.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121035 | Z-0023-2014 | GE, Optima CT660. Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data. | GE, Optima CT660 | 2 | CDRH | 2013-10-17 | 2013-10-17 21:00 | 2013-12-05 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue related with the DMPR function and Screen Saved Images on your scanner which involves Optima CT660, Brivo CT385, Discovery PET/CT 610, Discovery PET/CT 710. The DMPR and Screen Save functionality may possibly generate the same Series UID. It has been seen that some PACS may combine series under a patient using only the Series UID
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121037 | Z-0024-2014 | GE PET Discovery 610, PET Discovery 710. Intended for head and whole body attenuation corrected Positron Emission Tomography (PET) imaging. | GE PET Discovery 610, PET Discovery 710 | 2 | CDRH | 2013-10-17 | 2013-10-17 21:00 | 2013-12-05 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue related with the DMPR function and Screen Saved Images on your scanner which involves Optima CT660, Brivo CT385, Discovery PET/CT 610, Discovery PET/CT 710. The DMPR and Screen Save functionality may possibly generate the same Series UID. It has been seen that some PACS may combine series under a patient using only the Series UID
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116713 | Z-1340-2013 | Ref: VP1600 HD Still Capture System ConMed Linvatec . Captures, stores and transfers images generated by a medical endoscopic camera system during surgical procedures. | HD Still Capture | 2 | CDRH | 2013-05-20 | 2013-05-20 21:00 | 2013-11-06 | Linvatec Corp. dba ConMed Linvatec | The VP1600 HD Still Capture Image System is recalled due to an anomaly between the software code in the VP1600 console and the communications interface for various keyboard key strokes (Caps lock and Num Lock) which may lock up and requires the user to reboot to continue.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116757 | Z-1042-2013 | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; Model: Primeview 3i, Part No. 7341428, Current R610/R630; Product Usage: The intended use of the SIEMENS branded ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | COHERENCE RT Therapist | 2 | CDRH | 2013-04-01 | 2013-04-02 21:00 | 2014-08-28 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116758 | Z-1043-2013 | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; Model: Primeview 3i 2.1, Part No. 8147675, Current R640; Product Usage: The intended use of the SIEMENS branded ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | COHERENCE RT Therapist | 2 | CDRH | 2013-04-01 | 2013-04-02 21:00 | 2014-08-28 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116746 | Z-1036-2013 | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; Model: COHERENCE AG TH, Part No. 5863506, Current R630; Product Usage: The intended use of the SIEMENS branded ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | COHERENCE RT Therapist | 2 | CDRH | 2013-04-01 | 2013-04-02 21:00 | 2014-08-28 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116747 | Z-1037-2013 | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; Model: COHERENCE TH, Part No. 7339125, - Current R630; Product Usage: The intended use of the SIEMENS branded ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | COHERENCE RT Therapist | 2 | CDRH | 2013-04-01 | 2013-04-02 21:00 | 2014-08-28 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116748 | Z-1038-2013 | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; Model: COHERENCE Impression TH, Part No. 7341410, Current R630; Product Usage: The intended use of the SIEMENS branded ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | COHERENCE RT Therapist | 2 | CDRH | 2013-04-01 | 2013-04-02 21:00 | 2014-08-28 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116750 | Z-1040-2013 | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; Model: COHERENCE TH 2.1, Part No. 8147667, Current R640; Product Usage: The intended use of the SIEMENS branded ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | COHERENCE RT Therapist | 2 | CDRH | 2013-04-01 | 2013-04-02 21:00 | 2014-08-28 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116751 | Z-1041-2013 | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; Model: COHERENCE TH 2.2, Part No. 10568431, Current R640; Product Usage: The intended use of the SIEMENS branded ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | COHERENCE RT Therapist | 2 | CDRH | 2013-04-01 | 2013-04-02 21:00 | 2014-08-28 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116749 | Z-1039-2013 | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; Model: COHERENCE TH 2.0, Part No. 8139839, Current R630; Product Usage: The intended use of the SIEMENS branded ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | COHERENCE RT Therapist | 2 | CDRH | 2013-04-01 | 2013-04-02 21:00 | 2014-08-28 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116759 | Z-1044-2013 | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; Model: Primeview 3i 2.0, Part No. 8139847, Current R630; Product Usage: The intended use of the SIEMENS branded ONCOR and PRIMUS family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. | COHERENCE RT Therapist | 2 | CDRH | 2013-04-01 | 2013-04-02 21:00 | 2014-08-28 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=115819 | Z-0951-2013 | Vial Label: 1) LIATEST VWF:Ag Buffer - 5 ml Diagnostica Stago 2) LIATEST VWF:Ag Latex + 4 ml R3 Diagnostica Stago 3) LIATEST VWF:Ag Latex Diluent 4 mL Diagnostica Stago Model, catalogue, number 00518 - 58978 - 58979 Kit Label: LIATEST VWF:Ag DIAGNOSTICA STAGO S.A.S. 92600 Asni¿res sur Seine (France) Coagulation Assay | Diagnostica Stago Inc. | 2 | CDRH | 2013-03-14 | 2013-03-14 21:00 | 2013-07-19 | Diagnostica Stago, Inc. | Diagnostica Stago has confirmed that sometimes the delta OD measured can be changed leading to a lowered percentage result.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=115393 | Z-0767-2013 | Phillips Easy Diagnost Eleva DRF. Model Number: 706034 Product Usage: Universal Fluoroscopic Applications | Phillips Easy Diagnost Elva DRF | 2 | CDRH | 2013-02-01 | 2013-02-01 21:00 | 2017-01-23 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be misinterpreted as a RIGHT lead marker for side indication if the image is mirrored within a PACS system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=114521 | Z-0522-2013 | Cobas b 123 POC system cobas b 123 POC system is fully automated POC system for whole blood in vitro measurement of pH, blood gases (BG), electrolytes Na+ K+, iCa2+ (ISE), hematocrit (Hct), metabolites (Gluc, Lac), total hemoglobin (tHb), hemoglobin derivatives (O2Hb, HHB, COHB, MetHb), and oxygen saturation (SO2). In addition, the cobas b 123 POC system calculates derived parameters. It is dedicated for use in a Point-of-Care environment and laboratory. | Cobas b 123 POC system | 2 | CDRH | 2012-12-13 | 2012-12-13 21:00 | 2013-08-07 | Roche Diagnostics Operations, Inc. | It has been determined that cobas b 123 Fluid Pack COOX cuvettes for lot numbers 21426121 and 21426171, exhibit a significant increase in the out-of-box failure rate caused by a change in the thickness in the cuvette layer. Internal investigations found the altered thickness of the cuvette was due to production equipment settings. Cuvettes were produced within specifications, but within the upper
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=89071 | Z-0027-2013 | OrthoHelix Part #HCS-056-32 3.8MM Cannulated Hex Driver NON STERILE MAT: 465 SS The drivers are used to drive screws in the Mini MaxTorque Screw System, which is intended to be used to stabilize and aid in the fusion of fractures and osteotomies involving small bones of the hand, wrist, afoot, and ankle. | Mini MaxTorque HCS-056-32 Drivers | 2 | CDRH | 2012-10-09 | 2012-10-09 21:00 | 2013-02-20 | Orthohelix Surgical Designs Inc | The tips of the drivers are breaking while being used to drive the screws into patients.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=89070 | Z-0026-2013 | OrthoHelix Part #HCS-056-32 3.2MM Cannulated Hex Driver NON STERILE MAT: 465 SS The drivers are used to drive screws in the Mini MaxTorque Screw System, which is intended to be used to stabilize and aid in the fusion of fractures and osteotomies involving small bones of the hand, wrist, afoot, and ankle. | Mini MaxTorque HCS-056-32 Drivers | 2 | CDRH | 2012-10-09 | 2012-10-09 21:00 | 2013-02-20 | Orthohelix Surgical Designs Inc | The tips of the drivers are breaking while being used to drive the screws into patients.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=110419 | Z-2173-2012 | Siemens syngo(R) Lab Data Manager. The syngo(R) Lab Data Manager is classified as a data calculator and processing module for use with In-vitro diagnostic devices. It also provides work flow management as well as remote instrument control to multiple interfaced instruments | Siemens syngo(R) Lab Data Manager | 2 | CDRH | 2012-08-09 | 2012-08-09 21:00 | 2014-04-01 | Siemens Healthcare Diagnostics, Inc. | The recalling firm has discovered that when using a Sysmex(R) CA-7000 or CA-1500 coagulation analyzer with the syngo Lab Data Manager (LDM), there is a potential for the syngo LDM system, to display and transmit to the LIS an incorrect result following the syngo LDM's application of the configured result number format. The problem is encountered when the result number format configured on the Sys
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=110464 | Z-2041-2012 | Varian brand Varian Eclipse Treatment Planning System for radiotherapy; Model Number: H48; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA The Eclipse Treatment Planning System (Eclipse TPS) is used to plan radiotherapy treatments for patients with malignant or benign diseases. Eclipse TPS is used to plan external beam irradiation with photon, electron and proton beams, as well as for internal irradiation (brachytherapy) treatments. In addition, the Eclipse Proton Eye algorithm is specifically indicated for planning proton treatment of neoplasms of the eye. | Eclipse Treatment Planning System | 2 | CDRH | 2012-07-19 | 2012-07-19 21:00 | 2014-08-06 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly with the Eclipse Treatment Planning (for radiotherapy treatments) where changing the prescription in a calculated VMAT or RapidArc plan may lead to Monitor Units which do not reflect the new prescription. If the dose prescription of a VMAT or RapidArc plan is changed after the initial Dose Volume Optimization, the Monitor Units and dose distribution may not reflect
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=107894 | Z-1368-2012 | GEMINI GXL 6, 10 & 16 Computed Tomography X-Ray Systems, M/N 882390, 882400 & 882410, Distributed by: Philips Medical System, Cleveland, OH Product Usage: The Computed Tomography X-Ray Systems are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. Image processing and display workstations provide software applications to process, analyze, display, quantify and interpret medical images/data. The PET and CT images may be registered and displayed in a fused (overlaid in the same spatial orientation) format to provide combined metabolic and anatomical data at different angles. Trained professionals use the images in: The evaluation, detection and diagnosis of lesions, disease and organ function such as but not limited to cancer, cardiovascular disease, and neurological disorders. The detection, localization, and staging of tumors and diagnosing cancer patients. Treatment planning and interventional radiology procedures. The device includes software that provides a quantified analysis of regional cerebral activity from PET images. Cardiac imaging software provides functionality for the quantification of cardiology images and datasets including but not limited to myocardial perfusion for the display of wall motion and quantification of left ventricular function parameters from gated myocardial perfusion studies and for the 3D alignment of coronary artery images from CT coronary angiography onto the myocardium. Both subsystems (PET and CT) can also be operated independently as fully functional, diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | GEMINI GXL 6, 10 & 16 Computed Tomography X-Ray Systems, | 2 | CDRH | 2012-04-05 | 2012-04-05 21:00 | 2013-02-22 | Philips Medical Systems (Cleveland) Inc | TumorLOC will save the incorrect contour when the user saves the contour while in the edit mode, and the software does not force the user to exit the edit mode before saving. In TumorLOC, when saving while a contour is in edit mode, that contour will shift to the slice corresponding to the most recently drawn contour of the last organ in the Organ list.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=107897 | Z-1369-2012 | GEMINI TF 16, 64 & Big Bore Computed Tomography X-Ray Systems, M/N 882470, 882471 & 882476, Distributed by: Philips Medical System, Cleveland, OH Product Usage: The Computed Tomography X-Ray Systems are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. Image processing and display workstations provide software applications to process, analyze, display, quantify and interpret medical images/data. The PET and CT images may be registered and displayed in a fused (overlaid in the same spatial orientation) format to provide combined metabolic and anatomical data at different angles. Trained professionals use the images in: The evaluation, detection and diagnosis of lesions, disease and organ function such as but not limited to cancer, cardiovascular disease, and neurological disorders. The detection, localization, and staging of tumors and diagnosing cancer patients. Treatment planning and interventional radiology procedures. The device includes software that provides a quantified analysis of regional cerebral activity from PET images. Cardiac imaging software provides functionality for the quantification of cardiology images and datasets including but not limited to myocardial perfusion for the display of wall motion and quantification of left ventricular function parameters from gated myocardial perfusion studies and for the 3D alignment of coronary artery images from CT coronary angiography onto the myocardium. Both subsystems (PET and CT) can also be operated independently as fully functional, diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | GEMINI TF 16, 64 & Big Bore Computed Tomography X-Ray Systems | 2 | CDRH | 2012-04-05 | 2012-04-05 21:00 | 2013-02-22 | Philips Medical Systems (Cleveland) Inc | TumorLOC will save the incorrect contour when the user saves the contour while in the edit mode, and the software does not force the user to exit the edit mode before saving. In TumorLOC, when saving while a contour is in edit mode, that contour will shift to the slice corresponding to the most recently drawn contour of the last organ in the Organ list.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=107893 | Z-1367-2012 | Extended Brilliance Workstation Computed Tomography X-Ray Systems, M/N 728260, Distributed by: Philips Medical System, Cleveland, OH Product Usage: The Computed Tomography X-Ray Systems are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. Image processing and display workstations provide software applications to process, analyze, display, quantify and interpret medical images/data. The PET and CT images may be registered and displayed in a fused (overlaid in the same spatial orientation) format to provide combined metabolic and anatomical data at different angles. Trained professionals use the images in: The evaluation, detection and diagnosis of lesions, disease and organ function such as but not limited to cancer, cardiovascular disease, and neurological disorders. The detection, localization, and staging of tumors and diagnosing cancer patients. Treatment planning and interventional radiology procedures. The device includes software that provides a quantified analysis of regional cerebral activity from PET images. Cardiac imaging software provides functionality for the quantification of cardiology images and datasets including but not limited to myocardial perfusion for the display of wall motion and quantification of left ventricular function parameters from gated myocardial perfusion studies and for the 3D alignment of coronary artery images from CT coronary angiography onto the myocardium. Both subsystems (PET and CT) can also be operated independently as fully functional, diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | Extended Brilliance Workstation Computed Tomography X-Ray Systems | 2 | CDRH | 2012-04-05 | 2012-04-05 21:00 | 2013-02-22 | Philips Medical Systems (Cleveland) Inc | TumorLOC will save the incorrect contour when the user saves the contour while in the edit mode, and the software does not force the user to exit the edit mode before saving. In TumorLOC, when saving while a contour is in edit mode, that contour will shift to the slice corresponding to the most recently drawn contour of the last organ in the Organ list.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=107891 | Z-1366-2012 | Brilliance Big Bore Computed Tomography X-Ray Systems, M/N 728243, 728244, Distributed by: Philips Medical System, Cleveland, OH Product Usage: The Computed Tomography X-Ray Systems are intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. Image processing and display workstations provide software applications to process, analyze, display, quantify and interpret medical images/data. The PET and CT images may be registered and displayed in a fused (overlaid in the same spatial orientation) format to provide combined metabolic and anatomical data at different angles. Trained professionals use the images in: The evaluation, detection and diagnosis of lesions, disease and organ function such as but not limited to cancer, cardiovascular disease, and neurological disorders. The detection, localization, and staging of tumors and diagnosing cancer patients. Treatment planning and interventional radiology procedures. The device includes software that provides a quantified analysis of regional cerebral activity from PET images. Cardiac imaging software provides functionality for the quantification of cardiology images and datasets including but not limited to myocardial perfusion for the display of wall motion and quantification of left ventricular function parameters from gated myocardial perfusion studies and for the 3D alignment of coronary artery images from CT coronary angiography onto the myocardium. Both subsystems (PET and CT) can also be operated independently as fully functional, diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | Brilliance Big Bore Computed Tomography X-Ray Systems | 2 | CDRH | 2012-04-05 | 2012-04-05 21:00 | 2013-02-22 | Philips Medical Systems (Cleveland) Inc | TumorLOC will save the incorrect contour when the user saves the contour while in the edit mode, and the software does not force the user to exit the edit mode before saving. In TumorLOC, when saving while a contour is in edit mode, that contour will shift to the slice corresponding to the most recently drawn contour of the last organ in the Organ list.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=106686 | Z-0973-2012 | Volcano pcFM Software Kit ; Manufactured by: Volcano Corporation, 2870 Kilgore Road, Rancho Cardova, CA 95624. The volcano pcFM Software Kit is a research only off line tool intended for review and printout of physiological measurements recorded by ComboMap System. It is not intended for analysis or manipulation of the original measurement data. | pcFM Software Kit | 2 | CDRH | 2012-02-07 | 2012-02-07 21:00 | 2012-02-07 | Volcano Corporation | Internal review of validation of pcMF Software kit found that instructions directs user to remove the rear panel, causing the potential for electric shock.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=101368 | Z-0006-2012 | Model 104 VNS Therapy Demipulse Duo Generator Cyberonics, Inc. Houston, Texas, Model 104 Generators are provided in a sterile single use packages as part of the VNS Therapy System. Product labeling may be found at http://us.cyberonics.com/en/vns-therapy-for-epilepsy/healthcare-professionals/vns-therapy/manuals-page. Product Usage: The VNS Therapy System is indicated for use as an adjunctive therapy in reducing the frequency of seizures in adults and adolescents over 12 years of age with partial onset seizures, which are refractory to antiepileptic medications and for the adjunctive long-term treatment of chronic or recurrent depression for patients 18 years of age or older who are experiencing a major depressive episode and have not had an adequate response to four or more adequate antidepressant treatments. | Cyberonics | 2 | CDRH | 2011-10-04 | 2011-10-04 21:00 | 2012-12-04 | Cyberonics, Inc | An investigation was initiated based on a report from the field in which an Intensive Follow-up Indicator (IFI) message was unexpectedly received by a medical professional when using Model 250 version 8.0 software to interrogate a patient's Model 103 Generator.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=101351 | Z-0005-2012 | VNS Therapy Demipulse Model 103 Generator Cyberonics, Inc. Houston, Texas, Model 103 Generators are provided in a sterile single use packages as part of the VNS Therapy System. Product labeling may be found at http://us.cyberonics.com/en/vns-therapy-for-epilepsy/healthcare-professionals/vns-therapy/manuals-page. Product Usage: The VNS Therapy System is indicated for use as an adjunctive therapy in reducing the frequency of seizures in adults and adolescents over 12 years of age with partial onset seizures, which are refractory to antiepileptic medications and for the adjunctive long-term treatment of chronic or recurrent depression for patients 18 years of age or older who are experiencing a major depressive episode and have not had an adequate response to four or more adequate antidepressant treatments. | Cyberonics | 2 | CDRH | 2011-10-04 | 2011-10-04 21:00 | 2012-12-04 | Cyberonics, Inc | An investigation was initiated based on a report from the field in which an Intensive Follow-up Indicator (IFI) message was unexpectedly received by a medical professional when using Model 250 version 8.0 software to interrogate a patient's Model 103 Generator.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=102297 | Z-3178-2011 | STOCKERT 70 Radio Frequency Generator, Catalog Number 57001 Indicated for use with a Biosense Webster, Inc. diagnostic/ablation deflectable tip catheter to deliver radio frequency energy during cardiac ablation procedures. Stockert 70 in foreign countries is also known as: EP-Shuttle/Stockert J50/Stockert J70 With software versions: 1.035/1.035J/1.035J70 to 1.037/1.037J70 | EP-Shuttle/Stockert J50/Stockert J70 | 2 | CDRH | 2011-09-12 | 2011-09-12 21:00 | 2012-10-12 | Biosense Webster, Inc. | Biosense Webster has initiated a voluntary field notification to all customers that have a Stockert 70 Radio Frequency Generator, to communicate undetected software anomalies in the form of software release note in order to amend the labeling and thereby enhance the safety of the product.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=101749 | Z-2915-2011 | GEMINI TF Big Bore computer tomograph software version 3.6. Catalog #882476, Gantry S/N 9001-9023. The GEMINI TF Big Bore is a diagnostic imaging systems for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. The CT subsystem produces cross-sectional images of the body by computer reconstruction of x-ray transmission data. The PET subsystem produces images of the distribution of PET radiopharmaceuticals in the patient body (specific radiopharmaceuticals are used for whole body, brain, heart and other organ imaging). Attenuation correction is accomplished by CTAC. The device also provides for list mode, dynamic, and gated acquisitions. Both subsystems (PET and CT) can also be operated independently as fully functional, diagnostic imaging systems including application of the CT scanner as a radiation therapy simulation scanner. | GEMINI TF Big Bore | 2 | CDRH | 2011-08-01 | 2011-08-01 21:00 | 2012-09-21 | Philips Medical Systems (Cleveland) Inc | Two non-safety related defects have been found on the Philips GEMINI TF Big Bore running version 3.6; if these defects were to recur, it could impact the performance or reliability of the system. Philips is issuing a software update to all Philips GEMINI TF Big Bore customers to correct these defects.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=100877 | Z-2890-2011 | BD BACTEC(tm) MGI (tm) System Software, Catalog 445922, Lot V5.02A. Product Usage: The BACTEC(tm) MGIT (tm) System Software is used in the MGIT 960 and MGIT 320 instruments for the rapid detection of mycobacteria from clinical specimens. In addition, an Antimicrobial/Drug Susceptibility Testing (AST or DST) system is used as a rapid qualitative procedure for susceptibility testing of Mycobacterium tuberculosis, from culture, to streptomycin, isoniazid, rifampin, ethambutol, and pyrazinamide. The routine AST set consists of a Growth Control tube and a tube for each drug, as well as a barcoded tube carrier that holds the set. The carrier barcode provides set size and identification. A tube rack capable of holding up to 55 AST tubes in various set combinations is available for tube setup and transport. Large or small AST carrier sets can be placed in the rack. The MGIT AST system is based on growth of the Mycobacterium tuberculosis strain under test as measured by the instrument. AST sets are entered into the BACTEC MGIT instrument and are tested automatically. Analysis of fluorescence in a drug containing tube is compared to the fluorescence of the Growth Control tube to determine susceptibility results. | BD BACTEC(tm) MGIT (tm) 960 System | 2 | CDRH | 2011-07-28 | 2011-07-28 21:00 | 2011-07-29 | Becton Dickinson & Co. | Software contained three additional set configurations which caused the system to reinitialized the AST set carrier defaults.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98621 | Z-2662-2011 | Elekta Impac Software Sequencer Verify & Record System, for MOSAIQ The intended use of Sequencer (generically referred to as a verification system) is to assist in the process of process of patient treatment on a radiotherapy treatment machine. It gives the capability to notify users of actions that need to take place prior to treatment, to display reference images for set up purposes and to auto setup the machine to predefined settings. | Elekta Impac Software Sequencer | 2 | CDRH | 2011-06-24 | 2011-06-24 21:00 | 2011-09-16 | Impac Medical Systems Inc | There's a problem in the database conversion program that is used when upgrading from the Multi-ACC¿S to MOSAIQ or from one version of MOSAIQ to another. If the gantry start angle is incorrect prior to pressing beam-on, the patient will be mistreated.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98564 | Z-2064-2011 | Defibtech Lifeline AED and ReviveR AED, Model: DDU-100 Series AEDs Software version 2.004 or earlier The product under correction is the DDU-100 AED. The AED is indicated for use on victims of sudden cardiac arrest (SCA) when the patient is unconscious, unresponsive and not breathing. The indications for use also state that for patients under 8 years old, child/ infant electrode pads are to be used and that therapy should not be delayed to determine the patients exact age or weight. The AED must also be used by or on the order of a physician. These AEDs are sold under the brand names Lifeline AED and ReviveR AED. | Defibtech Lifeline AED and ReviveR AED | 1 | CDRH | 2011-05-16 | 2011-05-16 21:00 | 2021-03-30 13:37 | Defibtech, LLC | Software with version 2.004 software or earlier may cause an affected AED to cancel shock during the charging process
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92218 | Z-1984-2011 | VARIANT nbs Sickle Program Resin Update CD-ROM; Model Number:250-3020, included in 250-3000; Distributed and Manufactured by Bio-Rad Laboratories, Inc. The VARIANT nbs Sickle Cell Program is intended as a qualitative screen for the presence of hemoglobins F, A, S, D, C and E in eluates of neonatal blood collected on filter paper by high-performance liquid chromatography(HPLC). The VARIANT nbs Sickle Cell Program is intended for Professional use only. | VARIANT(tm) nbs Sickle Cell Prograin Reorder Pack. | 3 | CDRH | 2011-04-14 | 2011-04-14 21:00 | 2011-08-08 | Bio-Rad Laboratories Inc | The affected resin Update CD-ROMs cause all user-defined settings in the Setup/Test/Data Setup sub screen and Set/Test/Pattern Setting sub screen to be overwritten when the Update Kit procedure is performed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90074 | Z-1733-2011 | Abbott brand CELL-DYN Emerald Hematology Analyzer System; List Numbers: 09H39-01; The product is manufactured in France for Abbott Diagnostics Division of Abbott Laboratories, Abbott Park, IL The CELL-DYN Emerald is an automated hematology analyzer designed for in-vitro diagnostic use in clinical laboratories. | CELL-DYN Emerald Hematology Analyzer System | 2 | CDRH | 2011-03-18 | 2011-03-18 21:00 | 2011-03-18 | Abbott Laboratories | Abbott has determined that gain settings for the non-standard specimen types may not match the gain settings for the standard specimen type. There is the potential for impact to product performance.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=97200 | Z-1572-2011 | MOSAIQ Oncology Information System Intended Use : Electronic Patient Records Management | MOSAIQ Oncology Information System | 2 | CDRH | 2011-03-07 | 2011-03-07 21:00 | 2012-04-19 | Siemens Medical Solutions USA, Inc | The potential for injury to a patient should the treatment data be incorrectly uploaded to patient records.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94463 | Z-1004-2011 | Optovue iVue optical coherence tomography system, computer controlled opthalmic imaging system, manufactured by Optovue, Fremont, CA Computer controlled opthalmic imaging system. Device scans the patient's eye using low coherence interferometer to measure the reflectivity of the retinal tissue. The cross sectional retinal tissue structure is composed of a sequence of A-scans. It has a traditional patient and instrument interface. The computer has a graphic user interface for acquiring and analyzing the image. The iVue offers three scans: Retina, Glaucoma and Cornea. For the cornea scan, a lens must be attached to the front of the device for proper scanning. This is the Cornea Anterior Module (CAM). | Optovue iVue | 2 | CDRH | 2011-02-01 | 2011-02-01 21:00 | 2011-02-08 | Optovue Inc. | Use of the CAM attachment may result in a positioning anomaly and mistreatment in the eyes of the patient.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94510 | Z-0143-2011 | The label (which is affixed to ScanPoint Communications Cradle) states: "...VERATHON ScanPoint Docking Station...Part No. 0570-0168...Bothell, WA USA..." | Scanpoint Communication Cradles | 3 | CDRH | 2010-10-26 | 2010-10-26 21:00 | 2011-04-01 | Verathon, Inc. | Incorrect programming causes QuickPrint to fail to recognize the communication Cradles. As a result, ScanPoint Communications Cradles could not communicate with ScanPoint Online Tool while using the QuickPrint interface. This will affect the transfer of exams and device calibration through QuickPrint interface.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92717 | Z-2212-2010 | DYNASTY¿ A-CLASS¿ POLY LINER, REF DLXP-LD36, 1 EACH, Rx ONLY, STERILE EO, GROUP D, I.D. 36 mm, LIP 15o, LINER STD, Wright Medical Technology, 4677 Airline Road, Arlington, TN 38002. Product is used in total hip arthroplasty. | DYNASTY ¿ A-CLASS ¿ POLY LINER | 3 | CDRH | 2010-08-13 | 2010-08-12 21:00 | 2011-05-31 | Wright Medical Technology Inc | The inner and outer product packaging was missing the following information: shelf life, translations, manufacturing date, CE marking and also listed the incorrect sterilization method.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87992 | Z-1534-2010 | Hitachi Oasis Magnetic Resonance Scanners Software V1.0C The intended use: The MR system produces transverse, coronal, sagittal, oblique, and curved cross-sectional images that display the internal structure of the head, body, or extremities. The images produced by the MR system reflect the spatial distribution of protons (hydrogen nuclei) exhibiting magnetic resonance. | Hitachi Oasis Magnetic Resonance Scanners Software | 2 | CDRH | 2010-05-04 | 2010-05-04 21:00 | 2012-02-24 | Hitachi Medical Systems America Inc | The device has a software error that could potentially lead to a misdiagnosis.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=89730 | Z-1403-2010 | GE Centricity Web Diagnostic (WebDX) software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. Device that receives, stores and communicates medical images and data from various imaging sources. | GE Centricity Web Diagnostic (WebDX) Software | 2 | CDRH | 2010-04-16 | 2010-04-16 21:00 | 2010-11-15 | Ge Healthcare It | There is a potential safety issue associated with the use of GE Centricity PACS-IW and PACS Web Diagnostic software when dragging-and-dropping images from the Navigator causing flipping of the images that may impact patient safety.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=89729 | Z-1402-2010 | GE Centricity PACS-IW software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. Device that receives, stores and communicates medical images and data from various imaging sources. | GE Centricity PACS-IW Software | 2 | CDRH | 2010-04-16 | 2010-04-16 21:00 | 2010-11-15 | Ge Healthcare It | There is a potential safety issue associated with the use of GE Centricity PACS-IW and PACS Web Diagnostic software when dragging-and-dropping images from the Navigator causing flipping of the images that may impact patient safety.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=76985 | Z-2113-2009 | Infinity Orange Enteral Feeding Pump, REF INFORNG, Zevex Enteral Nutrition Delivery Systems, Salt Lake City, UT 84123. | Infinity Orange Enteral Feeding Pump | 2 | CDRH | 2009-09-20 | 2009-09-21 21:00 | 2011-10-27 | Zevex International, Inc. | Enteral feeding pumps may not accurately deliver fluid, potentially under-infusing or over-infusing.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=75996 | Z-2112-2009 | EnteraLite Infinity Enteral Feeding Pump, REF INFP01, Zevex Enteral Nutrition Delivery Systems, Salt Lake City, UT 84123. A rotary peristaltic pump designed to deliver programmed doses of enteral nutrition solutions at selectable rates. | EnteraLite Infinity Enteral Feeding Pump | 2 | CDRH | 2009-09-20 | 2009-09-21 21:00 | 2011-10-27 | Zevex International, Inc. | Enteral feeding pumps may not accurately deliver fluid, potentially under-infusing or over-infusing.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=71931 | Z-2241-2008 | Medtronic SPECIFY, 3998, Lead Kit for Spinal Cord Stimulation (SCS). Manufactured at Medtronic, Inc., Villalba, Puerto Rico, USA. Contents of the inner package are STERILE. Method of sterilization: Ethylene Oxide. Model 3998 lead is an implantable, permanent lead. It is used to deliver electrical pulses to the dorsal aspect of the spinal cord. The lead consists of two polyurethane lead bodies joined to one silicone rubber paddle. The lead has two parallel rows of four platinum iridium electrodes on the distal end. | Medtronic SPECIFY | 2 | CDRH | 2008-09-18 | 2008-09-18 21:00 | 2010-10-08 | Medtronic Neuromodulation | Medtronic is recalling certain lots of Model 3998 SPECIFY surgical lead kits due to a package mislabeling issue. The package labeling incorrectly states the lead length is 28 cm in length when it should state 20 cm. Recalled lots are to be returned for replacement or credit.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=71855 | Z-2204-2008 | da Vinci S Surgical System, Model number IS2000 Endoscopic Instrument Control System, manufactured by Intuitive Surgical, Sunnyvale, CA. | da Vinci S Surgical System IS2000 | 2 | CDRH | 2008-09-11 | 2008-09-16 05:16 | 2009-07-29 | Intuitive Surgical, Inc. | Defective software chip may cause the system to fail and lock up.