## Sheet1
WEB_ADDRESS | RECALL_NUMBER | PRODUCT_DESCRIPTION | TRADE_NAME | RECALL_CLASS | CENTER | CENTER_CLASSIFICATION_DT | POSTED_INTERNET_DT | TERMINATION_DT | FIRM_NAME | MANUFACTURER_RECALL_REASON
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194852 | Z-1560-2022 | Maquet Critical Care AB/Getinge Servo-u Ventilator System, Part Number 6694800 Servo-n Ventilator System, Part Number 6688600 Servo-u MR ventilator: Part Number 6888800 Servo-air: Part Number 6882000 | Maquet Critical Care AB/Getinge Servo-u, Servo-n, Servo-u MR, Servo-air | 1 | CDRH | 2022-08-22 07:07 |  |  | Getinge Usa Sales Inc | Ventilators may generate a combination of alarms which may result in loss of communication, technical alarms being triggered, and loss of ventilation. If the device stops ventilating due to this issue, it may lead to hypoventilation and consequently desaturation, with patient outcomes including hypoxemia and hypoxic injury, which might result in circulatory failure.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193698 | Z-1260-2022 | illumina Model NextSeq 550 Dx REF 20005715 | NextSeq 550 Dx Instrument | 2 | CDRH | 2022-06-13 |  |  | Illumina, Inc. | cybersecurity vulnerability
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193602 | Z-1259-2022 | illumina REF DX-410-1001 Model: MiSeq Dx illumina REF 15036706 Model: MiSeq Dx illumina REF 20014053 Model: MiSeq Dx | MiSeq Dx Instrument | 2 | CDRH | 2022-06-13 |  |  | Illumina, Inc. | cybersecurity vulnerability
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192179 | Z-0835-2022 | FORE-SIGHT ELITE Tissue Oximeter module, REF: HEMFSM10 | FORE-SIGHT ELITE Tissue Oximeter | 2 | CDRH | 2022-03-30 |  |  | Edwards Lifesciences, LLC | The StO2 values may be inaccurately low when using either the FORE-SIGHT ELITE Tissue Oximeter Module or the FORE-SIGHT ELITE Absolute Tissue Oximeter Monitor with the Fore-Sight Elite large sensor in certain somatic locations (arms and legs). While the StO2 absolute values are impacted, the directional trend remains accurate, but may have a larger magnitude change. Low StO2 values may lead to unintended or inappropriate treatment
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=192119 | Z-0834-2022 | FORE-SIGHT ELITE Absolute Tissue Oximeter, REF: 01-06-3000, Contents: Monitor 1 EA., Accessory Kit 1 EA. | FORE-SIGHT ELITE Absolute Tissue Oximeter | 2 | CDRH | 2022-03-30 |  |  | Edwards Lifesciences, LLC | The StO2 values may be inaccurately low when using either the FORE-SIGHT ELITE Tissue Oximeter Module or the FORE-SIGHT ELITE Absolute Tissue Oximeter Monitor with the Fore-Sight Elite large sensor in certain somatic locations (arms and legs). While the StO2 absolute values are impacted, the directional trend remains accurate, but may have a larger magnitude change. Low StO2 values may lead to unintended or inappropriate treatment
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=191822 | Z-0728-2022 | 1688 Camera Control Unit (CCU), Catalog number:1688010000. Used in general laparoscopy, nasopharyngoscopy, ear endoscopy, sinuscopy, and plastic surgery wherever a laparoscope/endoscope/arthroscope/sinuscope is used. | Camera Control Unit (CCU) | 2 | CDRH | 2022-03-07 | 2022-03-07 20:06 |  | Stryker Corporation | A software defect in the camera control unit (CCU) will cause the image on the monitor to flip upside-down into an incorrect orientation. The potential of harms are conversion to open procedure, additional medical intervention, or a revision surgery.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188994 | Z-2473-2021 | McKesson Lap Sponge, X-Ray Detectable 12"x 12", 5/pack, Sterile McKesson Part # 16-2112121 | McKesson Lap Sponge | 2 | CDRH | 2021-09-14 10:47 |  | 2023-07-07 10:57 | Cypress Medical Products LLC | Manufacturer issued a recall due to an EO indicator color. The indicator should change from red to blue signifying sterilization. However, on one case of product, the indicators did not change
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184970 | Z-0695-2021 | Revolution CT with Apex Edition CT System - Product Usage: intended for head, whole body, cardiac and vascular X-ray Computed Tomography applications. | Revolution CT with Apex Edition CT System | 2 | CDRH | 2020-12-22 |  | 2023-08-29 12:17 | GE Healthcare, LLC | A potential for protocol(s) with unintended patient orientation or scan setting when using Intelligent Protocoling
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184853 | Z-0694-2021 | Revolution CT ES - Product Usage: intended for head, whole body, cardiac and vascular X-ray Computed Tomography applications. | Revolution CT ES | 2 | CDRH | 2020-12-22 |  | 2023-08-29 12:17 | GE Healthcare, LLC | A potential for protocol(s) with unintended patient orientation or scan setting when using Intelligent Protocoling
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179478 | Z-1436-2020 | Atellica IM Humidity Pack (Qty 1), SMN 11313496, UDI 00630414243726, Software Version V1.21.0 and lower - Product Usage: Product Usage: The system is intended for the qualitative and quantitative analysis of various body fluids, using photometric, turbidimetric, chemiluminescent, and integrated ion selective electrode technology for clinical use. Humidity packs maintain reagent compartment humidity between 70100%. The operator loads 7 empty humidity packs initially. | Atellica IM Humidity Pack | 2 | CDRH | 2020-03-04 |  | 2022-07-11 14:53 | Siemens Healthcare Diagnostics, Inc. | A software error is causing the analyzer to incorrectly eject affected Humidity Packs as expired.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179477 | Z-1435-2020 | Atellica IM Humidity Pack (Qty 5), SMN 11313505, UDI 00630414234526, Software Version V1.21.0 and lower - Product Usage: The system is intended for the qualitative and quantitative analysis of various body fluids, using photometric, turbidimetric, chemiluminescent, and integrated ion selective electrode technology for clinical use. Humidity packs maintain reagent compartment humidity between 70100%. The operator loads 7 empty humidity packs initially. | Atellica IM Humidity Pack | 2 | CDRH | 2020-03-04 |  | 2022-07-11 14:53 | Siemens Healthcare Diagnostics, Inc. | A software error is causing the analyzer to incorrectly eject affected Humidity Packs as expired.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=179431 | Z-1441-2020 | SoundWave Professional Suite Software 3.2, Ref CI-6055-014, SPN 058-023-32085, UDI: (01)07630016853227, Logical SoundWave Professional Suite 3.2, CE 0123 - Product Usage: Soundwave 3.2 is used by healthcare professionals to program Advanced Bionics sound processors for use with Advanced Bionics cochlear implants, a device is intended to restore a level of auditory sensation to children with profound sensorineural deafness via electrical stimulation of the auditory nerve. | Soundwave Professional Suite Software 3.2 | 3 | CDRH | 2020-03-04 |  | 2022-06-27 05:57 | Advanced Bionics, LLC | The manufacturer received complaints that customers were attempting to install the fitting software using the supplied USB drive, an error message was received, and installation failed. The failed installation has not impact to the current version of the software, and users are able to continue programming devices with the current software version.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178651 | Z-0914-2020 | iQ200 Series Urine Microscopy Analyzer, Catalog numbers: 700-3345, 700-3375, 700-3325, 700-3347, 700-3300RFB, 700- 3320, 700-3320RFB, 700-3322, 700-3322RFB, 700-3326, 700- 3340, 700-3370, B48999, C10683, C10684 - Product Usage: device used to automate the complete urinalysis profile, including urine test strip chemistry panel and microscopic sediment analysis. | iQ 200 System | 2 | CDRH | 2020-01-29 |  |  | Beckman Coulter Inc. | Customers have reported incidents where two sets of results report the same Specimen Identifier (Specimen ID) with different results, different Medical Record Number (MRN) and different patient identification (demographics).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178650 | Z-0913-2020 | iChemVELOCITY, Urine Chemistry System, Catalog numbers: 700-7176-001, 700-7177-001, B75502 - Product Usage: used to automate the urine chemistry analysis profile using iChem¿ VELOCITY Urine Chemistry Strips. | iChem¿VELOCITY" Automated Urine Chemistry System | 2 | CDRH | 2020-01-29 |  |  | Beckman Coulter Inc. | Customers have reported incidents where two sets of results report the same Specimen Identifier (Specimen ID) with different results, different Medical Record Number (MRN) and different patient identification (demographics).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178163 | Z-0725-2020 | ApexPro Telemetry System - Product Usage: The ApexPro Telemetry System is intended for use under the direct supervision of a licensed healthcare practitioner. The system is designed to acquire and monitor physiological data for ambulating patients within a defined coverage area. The system processes this physiological data to detect various ECG arrhythmia events and select physiological parameter limit violations. The ApexPro Telemetry System is intended to be installed in the hospital or clinical environment in order to provide clinicians with patient physiological data, while allowing for patient mobility. These systems are typically deployed in sub acute care areas in hospitals or clinical sites where patient mobility can enhance the effectiveness of the medical procedures administered. | ApexPro Telemetry System | 2 | CDRH | 2019-12-20 |  |  | GE Healthcare, LLC | When connected to the Mission Critical (MC) and /or Information Exchange (IX) networks, certain versions of the CARESCAPE Telemetry Server, ApexPro Telemetry Server, CARESCAPE Central Station (CSCS) version 1 and Central Information Center (CIC) systems were identified to have vulnerabilities to a cyber-attack.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178165 | Z-0726-2020 | CIC Pro Clinical Information Center Central Station - Product Usage: The intended use is to provide clinicians with adult, pediatric and neonatal patient data in a centralized location within a hospital or clinical environment. CIC Pro Clinical Information Center central station is intended to collect information from a network and display this data. This data includes physiological, patient demographic and/or other nonmedical information. | CIC Pro Clinical Information Center Central Station | 2 | CDRH | 2019-12-20 |  |  | GE Healthcare, LLC | When connected to the Mission Critical (MC) and /or Information Exchange (IX) networks, certain versions of the CARESCAPE Telemetry Server, ApexPro Telemetry Server, CARESCAPE Central Station (CSCS) version 1 and Central Information Center (CIC) systems were identified to have vulnerabilities to a cyber-attack.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178167 | Z-0727-2020 | CARESCAPE Central Station - Product Usage: The intended use is to provide clinicians with adult, pediatric and neonatal patient data within a hospital or clinical environment. The CARESCAPE Central Station is intended to collect information from a network and display this data. This data includes physiological, patient demographic and/or other nonmedical information. | CARESCAPE Central Station | 2 | CDRH | 2019-12-20 |  |  | GE Healthcare, LLC | When connected to the Mission Critical (MC) and /or Information Exchange (IX) networks, certain versions of the CARESCAPE Telemetry Server, ApexPro Telemetry Server, CARESCAPE Central Station (CSCS) version 1 and Central Information Center (CIC) systems were identified to have vulnerabilities to a cyber-attack.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175186 | Z-0117-2020 | Aestiva 7100, Model Numbers: 1. 1006-9320-000 2. 1006-9321-000 3. 1006-9320-000-001101 4. 1006-9021-000 5. 1006-9012-000 6. 1006-9111-000 7. 1006-9321-000-007639 8. 1006-9321-000-000023 9. 1006-9321-000-015180 10. 1006-9320-000-010369 11. 1006-9321-000-019486 12. 1006-9321-000-007601 13. 1006-9321-000-009266 14. 1006-9321-000-008526 15. 1006-9321-000-116507 16. 1006-9321-000-009244 17. 1006-9321-000-012769 18. 1006-9320-000-004497 19. 1006-9320-000-009697 20. 1006-9320-000-014484 21. 1006-9321-000-008072 22. 1006-9321-000-007618 23. 1006-9321-000-013784 24. 1006-9320-000-010963 25. 1006-9320-000-014485 26. 1006-9320-000-008839 27. 1006-9321-000-017374 28. 1006-9321-000-014553 29. 1006-9031-000 30. 1006-9321-000-000533 31. 1006-9321-000-007785 32. 1006-9320-000-016418 33. 1006-9320-000-013302 34. 1006-9320-000-020339 35. 1006-9320-000-002141 36. 1006-9320-000-009193 37. 1006-9320-000-014443 38. 1006-9320-000-008124 39. 1006-9321-000-009347 40. 1006-9320-000-009310 41. 1006-9320-000-004623 42. 1006-9321-000-020437 43. 1006-9321-000-025830 44. 1006-9321-000-004690 45. 1006-9321-000-011348 46. 1006-9114-000 47. 1006-9322-000 48. 1006-9321-000-006678 49. 1006-9321-000-043401 50. 1006-9322-000-043399 51. 1006-9320-000-018778 52. 1006-9320-000-010161 53. 1006-9320-000-009400 54. 1006-9320-000-007948 55. 1006-9321-000-001309 56. 1006-9321-000-008305 57. 1006-9321-000-027472 58. 1006-9321-000-008643 59. 1006-9320-000-008841 60. 1006-9320-000-010102 61. 1006-9322-000-020407 62. 1006-9321-000-004831 63. 1006-9320-000-010684 64. 1006-9320-000-014654 65. 1006-9320-000-006392 66. 1006-9320-000-009056 67. 1006-9321-000-025850 68. 1006-9320-000-000473 69. 1006-9320-000-017639 70. 1006-9321-000-003110 71. 1006-9321-000-019653 72. 1006-9321-000-008529 73. 1006-9321-000-010029 74. 1006-9321-000-005790 75. 1006-9320-000-010177 76. 1006-9320-000-028131 77. 1006-9321-000-006318 78. 1006-9322-000-106884 79. 1006-9320-000-004596 80. 1006-9321-000-006645 81. 1006-9321-000-010584 82. 1006-9320-000-007779 83. 1006-9320-000-025679 84. 1006-9320-000-011509 85. 1006-9320-000-008650 86. 1006-9321-000-020045 87. 1006-9321-000-014996 88. 1006-9320-000-015449 89. 1006-9320-000-010685 90. 1006-9321-000-000266 91. 1006-9321-000-019680 92. 1006-9321-000-004778 93. 1006-9321-000-004424 94. 1006-9321-000-008747 95. 1006-9320-000-009249 96. 1006-9320-000-010834 97. 1006-9320-000-022596 98. 1006-9024-000 99. 1006-9026-000 100. 1006-9322-000-011934 101. 1006-9322-000-049412 102. 1006-9322-000-110551 103. 1006-9320-000-019424 104. 1006-9321-000-017187 105. 1006-9321-000-005997 106. 1006-9321-000-091883 107. 1006-9321-000-004524 108. 1006-9321-000-013425 109. 1006-9320-000-015413 110. 1006-9321-000-007378 111. 1006-9320-000-010661 112. 1006-9321-000-005998 113. 1006-9321-000-156642 114. 1006-9320-000-015195 115. 1006-9320-000-017086 116. 1006-9320-000-017852 117. 1006-9321-000-001974 118. 1006-9321-000-010236 119. 1006-9321-000-007641 120. 1006-9321-000-014028 121. 1006-9321-000-015552 122. 1006-9321-000-011524 123. 1006-9321-000-009297 124. 1006-9320-000-017427 125. 1006-9320-000-017498 126. 1006-9321-000-022470 127. 1006-9321-000-024751 128. 1006-9321-000-007642 129. 1006-9321-000-007636 130. 1006-9321-000-029968 131. 1006-9321-000-039161 132. 1006-9321-000-035978 133. 1006-9321-000-011702 134. 1006-9321-000-043612 models truncated | Aestiva 7100 | 2 | CDRH | 2019-12-17 13:52 |  |  | GE Healthcare, LLC | Certain Aespire and Aestiva Anesthesia Systems were noted to have a vulnerability to a cyber-attack when connected to the hospital network. An insufficiently secured terminal server may provide an opportunity for a malicious actor that has already penetrated the hospital network to send fraudulent flow sensor correction parameters.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175184 | Z-0115-2020 | Aestiva 7900, Model Numbers: a) 1006-9300-000 b) 1006-9305-000 c) 1006-9118-000 d) 1006-9112-000 e) 1006-9113-000 f) 1006-9020-000 g) 1006-9033-000 h) 1006-9117-000 i) 1006-9119-000 j) 1006-9030-000 k) 1006-9032-000 l) 1006-9022-000 m) To be provided n) 1006-9120-000 o) 1006-9027-000 p) 1006-9003-000 q) 1006-9025-000 | Aestiva 7900 | 2 | CDRH | 2019-12-17 13:52 |  |  | GE Healthcare, LLC | Certain Aespire and Aestiva Anesthesia Systems were noted to have a vulnerability to a cyber-attack when connected to the hospital network. An insufficiently secured terminal server may provide an opportunity for a malicious actor that has already penetrated the hospital network to send fraudulent flow sensor correction parameters.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175183 | Z-0114-2020 | Aestiva MRI, Model Numbers: a) 1006-9310-000 b) 1006-9110-000 c) 1006-9023-000 d) 1006-9028-000 e) 1006-9310-000-305077 f) 1006-9310-000-015243 g) 1006-9310-000-017602 h) 1006-9310-000-103785 I) 1006-9310-000-025109 j) 1006-9310-000-009650 k) 1006-9310-000-015224 l) 1006-9310-000-031881 m) 1006-9310-000-031854 n) 1006-9310-000-026571 | Aestiva MRI | 2 | CDRH | 2019-12-17 13:52 |  |  | GE Healthcare, LLC | Certain Aespire and Aestiva Anesthesia Systems were noted to have a vulnerability to a cyber-attack when connected to the hospital network. An insufficiently secured terminal server may provide an opportunity for a malicious actor that has already penetrated the hospital network to send fraudulent flow sensor correction parameters.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175185 | Z-0116-2020 | Aespire 7100, Model Numbers: 1. 1009-9000-000-009469 2. 1009-9000-000 3. 1009-9000-000-013913 4. 1009-9000-000-001366 5. 1009-9000-000-009857 6. 1009-9003-000 7. 1009-9000-000-008120 8. 1009-9000-000-009600 9. 1009-9000-000-014290 10. 1009-9000-000-010585 11. 1009-9000-000-004993 12. 1009-9011-000-004795 13. 1009-9000-000-013407 14. 1009-9000-000-004604 15. 1009-9000-000-017674 16. 1009-9000-000-008763 17. 1009-9000-000-008306 18. 1009-9000-000-003277 19. 1009-9000-000-004166 20. 1009-9000-000-008560 21. 1009-9000-000-000299 22. 1009-9000-000-011444 23. 1009-9000-000-003529 24. 1009-9000-000-005666 25. 1009-9000-000-019821 26. 1009-9011-000 27. 1009-9000-000-017345 28. 1009-9000-000-009893 29. 1009-9000-000-007258 30. 1009-9000-000-004959 31. 1009-9000-000-010185 32. 1009-9000-000-016672 33. 1009-9011-000-017769 34. 1009-9000-000-014329 35. 1009-9011-000-020685 36. 1009-9000-000-026022 37. 1009-9011-000-002912 38. 1009-9000-000-014723 39. 1009-9000-000-013385 40. 1009-9000-000-006725 41. 1009-9000-000-005163 42. 1009-9000-000-004022 43. 1009-9000-000-017939 44. 1009-9000-000-000300 45. 1009-9000-000-018855 46. 1009-9000-000-016724 47. 1009-9000-000-010268 48. 1009-9000-000-011595 49. 1009-9000-000-019277 50. 1009-9011-000-023222 51. 1009-9011-000-015929 52. 1009-9000-000-017824 53. 1009-9000-000-008270 54. 1009-9000-000-010416 55. 1009-9000-000-019089 56. 1009-9000-000-013395 57. 1009-9011-000-021989 58. 1009-9000-000-000517 59. 1009-9000-000-002407 60. 1009-9000-000-008415 61. 1009-9000-000-009597 62. 1009-9000-000-010178 63. 1009-9011-000-010400 64. 1009-9011-000-011366 65. 1009-9000-000-012360 66. 1009-9101-000 67. 1009-9000-000-013787 68. 1009-9000-000-006266 69. 1009-9000-000-014915 70. 1009-9000-000-000290 71. 1009-9000-000-013507 72. 1009-9000-000-011600 73. 1009-9000-000-010595 74. 1009-9000-000-009543 75. 1009-9000-000-011111 76. 1009-9000-000-013378 77. 1009-9000-000-000526 78. 1009-9000-000-010536 79. 1009-9000-000-010116 80. 1009-9000-000-005939 81. 1009-9000-000-008216 82. 1009-9000-000-006786 83. 1009-9000-000-012155 84. 1009-9000-000-012379 85. 1009-9000-000-014000 86. 1009-9011-000-022466 87. 1009-9000-000-009912 88. 1009-9000-000-006906 89. 1009-9000-000-004380 90. 1009-9000-000-010245 91. 1009-9000-000-010413 92. 1009-9000-000-013616 93. 1009-9000-000-009428 94. 1009-9000-000-010597 95. 1009-9000-000-020606 96. 1009-9000-000-010059 97. 1009-9000-000-014122 98. 1009-9000-000-016795 99. 1009-9000-000-012149 100. 1009-9000-000-005760 101. 1009-9000-000-009836 102. 1009-9000-000-016818 103. 1009-9000-000-017699 104. 1009-9000-000-004020 105. 1009-9000-000-013953 106. 1009-9000-000-009583 107. 1009-9000-000-007938 108. 1009-9000-000-007336 109. 1009-9000-000-010435 110. 1009-9011-000-016719 111. 1009-9000-000-024098 112. 1009-9011-000-004634 113. 1009-9011-000-016794 114. 1009-9011-000-005431 115. 1009-9000-000-000598 116. 1009-9011-000-005442 117. 1009-9000-000-022792 118. 1009-9000-000-003950 119. 1009-9000-000-013921 120. 1009-9000-000-011116 121. 1009-9000-000-017203 122. 1009-9000-000-003335 123. 1009-9000-000-008264 124. 1009-9000-000-009868 125. 1009-9011-000-015245 truncated text | Aespire 7100 | 2 | CDRH | 2019-12-17 13:52 |  |  | GE Healthcare, LLC | Certain Aespire and Aestiva Anesthesia Systems were noted to have a vulnerability to a cyber-attack when connected to the hospital network. An insufficiently secured terminal server may provide an opportunity for a malicious actor that has already penetrated the hospital network to send fraudulent flow sensor correction parameters.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=176041 | Z-0172-2020 | Rosa Brain 3.0, Model No. ROSA 3.0. The device is delivered assembled in a wooden case and protected with plastic bubble and plastic film. Seals are put on the crate. The product is not shipped sterile. The device is intended for the spatial positioning and orientation of instrument holders or tool guides to be used by neurosurgeons. | ROSA Brain 3.0 | 1 | CDRH | 2019-11-01 | 2019-11-01 21:01 | 2021-11-02 08:32 | MEDTECH SAS | The instrument holder may be sent on a trajectory that is not within the intended target. If it is not corrected, the associated device may be placed incorrectly.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175749 | Z-2753-2019 | Accu-Chek Connect Diabetes Management App | Accu-Chek Connect App | 2 | CDRH | 2019-09-26 08:07 |  | 2020-10-30 09:50 | Roche Diabetes Care, Inc. | Users with Android OS 8.0 and above may be unable to transfer values obtained on their meter to the App.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=174125 | Z-2363-2019 | Alinity ci -series System Control¿Module, Part Number 03R70 01 Product Usage: The Alinity ci series System Control Module has a scalable design to provide full integration of clinical chemistry and immunoassay analysis. The Alinity ci series System Control Module includes a clinical chemistry module and an immunoassay module, each performing all sample processing activities, and a system control module to provide a single user friendly interface. Each system can be customized by combining one or more sample processing modules, in multiple configurations, with a system control module to form a single workstation. The Alinity i Processing Module is a fully automated immunoassay analyzer allowing random and continuous access, as well as priority and automated retest processing using chemiluminescent microparticle immunoassay (CMIA) technology. CMIA technology is used to determine the presence of antigens, antibodies, and analytes in samples. | Alinity ciseries System Control Module | 2 | CDRH | 2019-08-23 |  | 2023-06-14 14:24 | Abbott Gmbh & Co. KG | Abbott has identified an issue with all on market versions (v2.6.0 and v2.6.1) of Alinity ci series Software where incorrect results may occur after a system Stop due to the Alinity i re use of reaction vessels (RVs). This issue only occurs if the system is transitioned from Processing to Stopped to Idle.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=168856 | Z-0478-2019 | Insulia Diabetes Management Companion | Insulia Diabetes Management Companion | 2 | CDRH | 2018-11-16 |  | 2019-03-12 | Voluntis | The firm identified an issue with the basal calculator identified on the Android version of the Insulia application. The bug could lead from low-impact to high-impact hypoglycemia depending on the circumstances. A new update of Insulia is now available (version 1.6.42), which resolves this issue.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=166868 | Z-2992-2018 | VidiStar(TM) PACS & DICOM Viewer Software system | VidiStar(TM) PACS & DICOM Viewer Software system | 2 | CDRH | 2018-09-06 |  | 2021-02-11 19:43 | Hitachi Healthcare Americas Corp Informatics Division | The secure filesystem client software used in the interface between the Vidistar PACS and an EHR system may cause intermixed images from multiple patients showing in a single study.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=163850 | Z-1956-2018 | MagNA Pure 24 System, GMMI: 07290519001 Product Usage: The MagNA Pure 24 System is a fully automated clinical sample extractor. It extracts nucleic acids from various sample types, sample tubes, and volume inputs for multiple downstream applications. The main features of the system are : Scalable purification of 1 to 24 samples Primary/secondary tube handling Inventory and sample management via barcodes Efficient order management via host The MagNA Pure 24 System offers 4 run types: Purification (rack) run Purification (cartridge) run Post Elution run Sample transfer run | MagNA Pure 24 System - Pathogen Protocols | 2 | CDRH | 2018-05-30 |  | 2019-04-30 14:34 | Roche Molecular Systems, Inc. | Cross-contamination of samples has been reported when running the existing protocol on the system with the rack-based purification run.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=160858 | Z-0487-2018 | GE HEALTHCARE CARESCAPE Central Station (CSCS) software version 2.0.2 The CARESCAPE Central Station is intended to collect information from a network and display this data. This data includes physiological, patient demographic and/or other non-medical information. Physiological parameters and waveforms from monitors and telemetry systems can be displayed, provide alarm annunciation, and can be printed from the CARESCAPE Central Station. | GE HEALTHCARE CARESCAPE Central Station (CSCS) | 2 | CDRH | 2018-01-29 | 2018-01-25 21:02 | 2021-07-07 12:44 | GE Medical Systems Information Technologies, Inc. | CARESCAPE Central Station (CSCS) software version 2.0.2 units may experience unexpected NO COMM (No Communication) and network communication issues after boot-up or system restart.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=159600 | Z-0231-2018 | ABL800 analyzer with FLEXQ module. Device intended for in vitro testing of samples of whole blood for the parameters pCO2, cK+, cNa+, cCa2+, cCl , cGlu, cLac, cCrea, ctBil, and co-oximetry parameters (ctHb, sO2, and the hemoglobin fractions FO2Hb, FCOHb, FMetHb, FHHb and FHbF) - in vitro testing of samples of expired air for the parameters pO2 and pCO2. It is also used for in vitro testing of pleura samples for the pH parameter | ABL800 FLEX | 2 | CDRH | 2017-12-08 |  | 2023-10-10 09:48 | Radiometer America Inc | Due to misinterpretation of the barcode by the scanner, when the registration receipt barcode is scanned by the analyzer , a result from a different patient will be printed or displayed on the analyzer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=157834 | Z-3110-2017 | MOSAIQ Medical Charged-Particle Radiation Therapy System | MOSAIQ Oncology Information System | 2 | CDRH | 2017-09-06 |  | 2020-10-26 16:41 | Elekta, Inc. | The table shift in the vertical direction was inverted during the use of CMA (Couch Move Assistant). Therapy may be directed to an incorrect area.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=155812 | Z-2746-2017 | Install CD, Spine tools, Plus and S7, Version 25. Model Number 9731958. | StealthStation System Update | 2 | CDRH | 2017-07-19 | 2017-06-26 21:00 | 2021-01-08 17:17 | Medtronic Navigation, Inc. | Software issue related to the StealthStation S7 system and the Synergy Spine application Version 2.1 configured with Spine Tool Install CD version 25. Issue may result in user being unable to navigate the Navigated Elevate Inserter with the StealthAiR Spine Frame during spine surgical procedures.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=154521 | Z-1874-2017 | Newport Medical HT70 and HT70- Plus Ventilators, Rx Only Product Usage: The Newport HT70 family of ventilators is intended to provide continuous or intermittent positive pressure mechanical ventilatory support for the care of individuals who require mechanical ventilation through invasive or non-invasive interfaces. Specifically, the Newport HT70 family of ventilators is applicable for infant, pediatric and adult patients greater than or equal to 5 kg (11 lbs). The Newport HT70 is a restricted medical device intended for use by qualified, trained personnel under the direction of a physician. The Newport HT70 is suitable for use in hospital, sub-acute, emergency room, and home care environments, as well as for transport and emergency response applications. | Newport Medical" HT70 and HT70- Plus Ventilators | 1 | CDRH | 2017-05-01 | 2017-04-10 21:08 | 2020-07-18 10:05 | Newport Medical Instruments Inc | Medtronic is issuing a voluntary field corrective action for all its Newport" HT70 and Newport" HT70 Plus ventilators because they may shutdown spontaneously during normal operation without an accompanying alarm.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=154388 | Z-1828-2017 | Merge Eye Station and Merge Eye Care PACS. | Merge Eye Station and Merge Eye Care PACS | 2 | CDRH | 2017-04-15 |  | 2020-06-22 17:53 | Merge Healthcare, Inc. | During an antivirus program scan of the Eye Station or Eye Care PACS, the antivirus program may detect a newly captured image to be an unwanted file and delete the image permanently.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152590 | Z-1682-2017 | Plum 360 Infusion Pump, Software Version 15.02. The infusion pump is capable of delivering fluids for a variety of therapies such as parenteral, enteral, or epidural infusions | Plum 360 Infusion Pump | 2 | CDRH | 2017-03-24 |  | 2019-07-09 09:08 | Hospira Inc., A Pfizer Company | Depleted Battery alarm shows instead of Replace Battery. On battery power, ongoing therapy stops, alarms show and sound, pump shuts down after 3 minutes. On AC power, the pump reboots (takes about 25 seconds). Therapy can resume. AC power interruption of 3-7 seconds prompts an incorrect E323 alarm, ongoing therapy stop, and pump reboot. Longer power loss results in battery power operation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152992 | Z-1540-2017 | Hitachi Oasis MRI System | Hitachi Oasis MRI System | 2 | CDRH | 2017-03-23 | 2017-03-23 21:00 | 2017-06-29 | Hitachi Medical Systems America Inc | Image data transferred from the MRI system to a workstation showed errors on the slice position reference image.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152993 | Z-1541-2017 | Hitachi Echelon Oval MRI System | Hitachi Echelon Oval MRI System | 2 | CDRH | 2017-03-23 | 2017-03-23 21:00 | 2017-06-29 | Hitachi Medical Systems America Inc | Image data transferred from the MRI system to a workstation showed errors on the slice position reference image.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152994 | Z-1542-2017 | Hitachi Echelon MRI System | Hitachi Echelon MRI System | 2 | CDRH | 2017-03-23 | 2017-03-23 21:00 | 2017-06-29 | Hitachi Medical Systems America Inc | Image data transferred from the MRI system to a workstation showed errors on the slice position reference image.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152632 | Z-1200-2017 | VIDAS 3 software v. 1.1.4 | VIDAS 3 software v. 1.1.4 | 2 | CDRH | 2017-02-14 | 2017-02-14 21:00 | 2017-08-28 | BioMerieux SA | During development of the VIDAS 3 software version 1.2, some anomalies have been identified and observed to be already present in the current software version VIDAS 3 version 1.1.4. available in the field.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151942 | Z-0971-2017 | Artis zee/zeego, Angiographic x-ray system Stand alone system The Artis systems are a family of dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed with the Artis zee/ zeego and Q/ Q.zen include cardiac angiography, neuro angiography, general angiography, rotational angiography, operating room angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. Artis zee/ zeego and Q/ Q.zen can also support the acquisition of position triggered imaging for spatial data synthesis. | Siemens Artis Zee/Zeego | 2 | CDRH | 2017-01-09 |  | 2019-04-09 14:16 | Siemens Medical Solutions USA, Inc | Siemens initiated a corrective action to address two possible, mutually independent causes of a system defect related to the following: - In Artis Systems with A100 Plus or A100G generators, an attempt to resume operation following detection of a fault can result in the failure of a module in the high-voltage generator. - For biplane systems delivered with software version VD11B since April of this year, software problems in conjunction with graphics card may, in rare cases, result in the loss of image display in the examination room.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151943 | Z-0972-2017 | Artis Q/Q.zen, Angiographic x-ray system Stand alone system The Artis systems are a family of dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed with the Artis zee/ zeego and Q/ Q.zen include cardiac angiography, neuro angiography, general angiography, rotational angiography, operating room angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. Artis zee/ zeego and Q/ Q.zen can also support the acquisition of position triggered imaging for spatial data synthesis | Siemens Artis Q/Q.zen | 2 | CDRH | 2017-01-09 |  | 2019-04-09 14:16 | Siemens Medical Solutions USA, Inc | Siemens initiated a corrective action to address two possible, mutually independent causes of a system defect related to the following: - In Artis Systems with A100 Plus or A100G generators, an attempt to resume operation following detection of a fault can result in the failure of a module in the high-voltage generator. - For biplane systems delivered with software version VD11B since April of this year, software problems in conjunction with graphics card may, in rare cases, result in the loss of image display in the examination room.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150897 | Z-0603-2017 | AXIOM Artis, Angiographic x-ray system Product Usage: The Artis systems are a family of dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed with the Artis zee / zeego and Q/ Q.zen include cardiac angiography, neuro angiography, general angiography, rotational angiography, operating room angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. Artis zee / zeego and Q/ Q.zen can also support the acquisition of position triggered imaging for spatial data synthesis. | Artis Qzee/zeego | 2 | CDRH | 2016-11-16 | 2016-11-17 21:00 | 2017-06-21 | Siemens Medical Solutions USA, Inc | The possibility exists that the semi-automatic focus switch may not work as intended when pressing the footswitch multiple times.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150896 | Z-0602-2017 | Artis zee/zeego, Angiographic x-ray system Product Usage: The Artis systems are a family of dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed with the Artis zee / zeego and Q/ Q.zen include cardiac angiography, neuro angiography, general angiography, rotational angiography, operating room angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. Artis zee / zeego and Q/ Q.zen can also support the acquisition of position triggered imaging for spatial data synthesis. | Artis Qzee/zeego | 2 | CDRH | 2016-11-16 | 2016-11-17 21:00 | 2017-06-21 | Siemens Medical Solutions USA, Inc | The possibility exists that the semi-automatic focus switch may not work as intended when pressing the footswitch multiple times.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150895 | Z-0601-2017 | Artis Q, Angiographic x-ray system Product Usage: The Artis systems are a family of dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed with the Artis zee / zeego and Q/ Q.zen include cardiac angiography, neuro angiography, general angiography, rotational angiography, operating room angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. Artis zee / zeego and Q/ Q.zen can also support the acquisition of position triggered imaging for spatial data synthesis. | Artis Q | 2 | CDRH | 2016-11-16 | 2016-11-17 21:00 | 2017-06-21 | Siemens Medical Solutions USA, Inc | The possibility exists that the semi-automatic focus switch may not work as intended when pressing the footswitch multiple times.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=147097 | Z-2867-2016 | Philips Expression Information Portal (IP5) Model 865471. The IP5 is intended to be used as a secondary display control unit for the Expression MRI Patient Monitoring System and the Expression MR200 MRI Patient Monitoring System. | Philips Expression Information Portal (IP5) | 2 | CDRH | 2016-09-21 |  | 2016-12-14 | Invivo Corporation | Frozen Display Numerics and Disabled Menu Keys after extended run time. This customer notification was sent August 22, 2014.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148568 | Z-2624-2016 | SOMATOM Definition AS with software version VA48A-SP2; Model # 10430603, computed tomography x-ray system | SOMATOM Definition AS | 2 | CDRH | 2016-08-23 | 2016-08-23 21:00 | 2017-01-24 | Siemens Medical Solutions USA, Inc | Due to an internal communication error between the firmware and the software of the components, the planned CT scan executes properly, but the injector is not started. Therefore, the contrast agent is not injected and the desired examination result is not achieved. This error only affects the automatic mode or coupled mode and does not affect the manual control of the injector.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148570 | Z-2626-2016 | SOMATOM Definition Edge with software version VA48A-SP2; Model # 8098027 computed tomography x-ray system | SOMATOM Definition Edge | 2 | CDRH | 2016-08-23 | 2016-08-23 21:00 | 2017-01-24 | Siemens Medical Solutions USA, Inc | Due to an internal communication error between the firmware and the software of the components, the planned CT scan executes properly, but the injector is not started. Therefore, the contrast agent is not injected and the desired examination result is not achieved. This error only affects the automatic mode or coupled mode and does not affect the manual control of the injector.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148569 | Z-2625-2016 | SOMATOM Definition Flash with software version VA48A-SP2; Model # 10590000, computed tomography x-ray system | SOMATOM Definition Flash | 2 | CDRH | 2016-08-23 | 2016-08-23 21:00 | 2017-01-24 | Siemens Medical Solutions USA, Inc | Due to an internal communication error between the firmware and the software of the components, the planned CT scan executes properly, but the injector is not started. Therefore, the contrast agent is not injected and the desired examination result is not achieved. This error only affects the automatic mode or coupled mode and does not affect the manual control of the injector.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148011 | Z-2587-2016 | OPTICAL BIOMETER AL-Scan: Software Version: 1.09 and earlier. Ophthalmic: The OPTICAL BIOMETER AL-Scan is a measuring device that can measure as a single unit the values necessary to calculate the power of an IOL for cataract surgery such as axial length, corneal curvature radius, and anterior chamber depth. It measures these necessary values successively through a non-contact optical measurement method. | OPTICAL BIOMETER | 2 | CDRH | 2016-08-17 |  | 2016-12-22 | Nidek Inc | Nidek Inc. received information from our manufacturer NIDEK CO. Japan that the shown axial lengths were much shorter at AL-Scan measurement.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=145824 | Z-1813-2016 | Volcano s5, s5i, CORE and CORE Mobile systems with software version 3.4 and v3.4 software kits. Cardiology: The Volcano Universal (Core /Core Mobile), S5iz Platform Imaging System is used for the qualitative and quantitative evaluation of vascular morphology in the coronary arteries and vessels of the peripheral vasculature. | Volcano s5, s5i, CORE and CORE Mobile systems | 2 | CDRH | 2016-05-25 |  | 2018-01-02 | Volcano Corporation | Volcano has become aware of an incompatibility issue between Impacted Systems and hospital network scans. In specific circumstances, an impacted System will encounter unexpected data from the hospital network and be forced to reboot
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144567 | Z-1282-2016 | Artis One, Interventional, Fluoroscopic, x-ray system Artis one is an angiography system developed for diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed with the Artis One include cardiac angiography, neuro-angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures as well as procedures next to the table for patient extremities. | Artis One | 2 | CDRH | 2016-03-31 |  | 2017-11-17 | Siemens Medical Solutions USA, Inc | After importing segmentation results of the left atrium created on the Artis One system, the possibility exists for the results to be mirrored to an electro-anatomical 3D mapping system, CARTO from Biosense Webster Inc.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144566 | Z-1278-2016 | Cios Alpha, mobile X-ray system | CIOS ALPHA | 2 | CDRH | 2016-03-31 | 2016-03-31 21:00 | 2017-03-20 | Siemens Medical Solutions USA, Inc | Software issues on Cios Alpha mobile C-Arm system
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144032 | Z-1092-2016 | Spacelabs Healthcare Xhibit Central Station, Model 96102 is used to provide clinicians with central monitoring of patient data for those patients connected to networked Spacelabs Healthcare patient monitors and telemetry transmitters. | Spacelabs Healthcare Xhibit Central Station | 2 | CDRH | 2016-03-10 |  | 2016-10-24 | Spacelabs Healthcare Inc | The firm has received one report of values for patient height and weight being switched when input at the Xhibit Central Station, Model 96102, causing a bedside monitor Body Surface Area (BSA) calculation to be in error.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139360 | Z-0598-2016 | Brainlab Cranial Navigation System: An Image Guided Surgery System / Stereotactic. Radiology Departments. The BrainLAB Cranial IGS System is intended to be an intra-operative image guided localization system to enable minimally invasive surgery. | Brainlab Cranial Navigation System | 1 | CDRH | 2016-01-19 | 2016-01-19 21:00 | 2018-04-17 | Brainlab AG | Software Error: The effect of setup on overall navigation accuracy could potentially intensify small inaccuracies arising from individual steps of a complex navigation procedure that may cause an inaccurate display of instruments by the navigation system compared to the actual patient anatomy.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142336 | Z-0550-2016 | Syngo Imaging XS is a Picture Archiving and Communication System (PACS) | Siemens Syngo Imaging XS is a Picture Archiving and Communication System (PACS) | 2 | CDRH | 2015-12-24 |  | 2016-12-05 | Siemens Medical Solutions USA, Inc | For the Syngo Workflow SLR System with Software Ver: VB10C: Printouts may be printed in incorrect anatomical size when using syngo Imaging XS filming application in conjunction with a printer not released for anatomical print usage. This may happen due to an invalid combination of printer and the syngo Imaging XS printer configuration file.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139302 | Z-2440-2015 | ExacTrac 6.0.x Patient Positioning System, Radiation therapy. Intended to be used to place patients at an accurately defined point within the treatment beam of a medical accelerator for stereotactic radiosurgery or radiotherapy procedures. | Brainlab AG ExacTrac 6.0.x | 2 | CDRH | 2015-08-21 | 2015-08-21 21:02 | 2017-01-05 | Brainlab AG | ExacTrac 6.0 Patient Positioning System: Display of potentially incorrect Digitally Reconstructed Radiograph (DRR) for x-ray correction and verification.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=137250 | Z-1719-2015 | Laksell GammaPlan, a computer based dose planning system specifically designed for use with Leksell Gamma Knife, radiation therapy treatment. | Leksell GammaPlan | 2 | CDRH | 2015-06-05 | 2015-06-05 21:00 | 2016-01-12 | Elekta, Inc. | Memory can become corrupted when creating a fused study via drag and drop in Leksell GammaPlan 10.2.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=136871 | Z-1682-2015 | Ingenuity TF Positron Emission Tomography (PET) and Magnetic Resonance Imaging (MRI), 882380, Philips Medical Systems. The device is a diagnostic imaging system for fixed installations that combines Positron Emission Tomography (PET) and Magnetic Resonance Imaging (MRI). | Ingenuity TF PET/MR | 2 | CDRH | 2015-05-28 | 2015-05-28 21:01 | 2018-06-26 | Philips Medical Systems (Cleveland) Inc | It was noticed that the spectroscopy voxel map did not align with the anatomical images of the phantom. This resulted in incorrect orientation of the Chemical Shift Imaging (CSI) data.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=136653 | Z-1641-2015 | ENDOPATH ECHELON" Vascular White Reload for Advanced Placement Tip, Product Usage: The instruments have application in multiple open or minimally invasive general, gynecologic, urologic, thoracic and pediatric surgical procedures. | Echelon Flex Powered Vascular Stapler With Advanced Placement Tip and White Reloads | 2 | CDRH | 2015-05-21 | 2015-05-21 21:01 | 2015-11-17 | Ethicon Endo-Surgery Inc | During an internal inspection of the production process the firm discovered an issue which may cause the cartridge to deploy an incomplete staple line.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135411 | Z-1612-2015 | Fujifilm Synapse Cardiovascular I ProSolv CardioVascular Synapse Cardiovascular v4.0.8 Synapse Cardiovascular v4.0.8 SR1 Intended Use: Picture Archiving and Communication System Computer, Diagnostic, Programmable | Fujifilm Medical Systems | 2 | CDRH | 2015-05-14 | 2015-05-14 21:00 | 2018-09-20 | Fujifilm Medical Systems U.S.A., Inc. | Under certain circumstances the patient Demographics in a report exported into the EMR may not match the demographics shown in the corresponding Synapse CV clinical report.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135415 | Z-1611-2015 | FujiMedical Synapse Cardiovascular I ProSolv CardioVascular. A picture Archiving and Communication System Software versions: Synapse Cardiovascular v4.0.8 Synapse Cardiovascular v4.0.8 SR1 | FujiMedical | 2 | CDRH | 2015-05-14 | 2015-05-14 21:00 | 2018-09-20 | Fujifilm Medical Systems U.S.A., Inc. | Measurement values in the EMR may not accurately reflect the actual measurement value on the Synapse Cardiovascular clinical report.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135814 | Z-1582-2015 | ExacTrac 6.x. is software used to place patients at an accurately defined point within the treatment beam of a medical accelerator for stereotactic radiosurgery or radiotherapy procedures. | Brainlab ExacTrac 6.x Patient Positioning System | 2 | CDRH | 2015-05-06 | 2015-05-06 21:01 | 2017-04-24 | Brainlab AG | ExacTrac 6.x Patient Positioning System: Potentially incorrect patient positioning when using the ExacTrac Cone Beam CT (CBCT) with a CBCT acquired at a couch angle other than 0.0 degrees.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=134489 | Z-1521-2015 | Software Version 2.8 & Below on VITROS 5,1 FS Chemistry Systems, Catalog Number 6801375, Global Trade Item Number 10758750001132, and VITROS 5,1 FS Chemistry System Refurbished, Catalog Number 6801890, Global Trade Item Number 10758750001644; IVD. Intended for use in the in vitro quantitative, semi quantitative, and qualitative measurement of a variety of analytes of clinical interest, using VITROS Chemistry Products MicroSlides and VITROS Chemistry Products MicroTip Reagents. | Ortho Clinical Diagnostics VITROS 5,1 FS Chemistry Systems | 2 | CDRH | 2015-04-24 | 2015-04-25 21:00 | 2018-06-01 | Ortho-Clinical Diagnostics | Software Anomaly; It is possible for the device to process samples with a cartridge other than the intended cartridge, potentially leading to erroneous patient results. If this anomaly occurred previously, an indication would be a series of results that were similar across multiple samples (i.e., results could be believable or outside of the Reportable Range for the intended assay).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=134353 | Z-1316-2015 | ExacTrac versions 6.x patient positioning systems are used to position patients during radiosurgery or radiotherapy procedures.. | Brainlab | 2 | CDRH | 2015-03-25 | 2015-03-25 21:00 | 2017-02-28 | Brainlab AG | The ExacTrac 6.x Patient Positioning System may incorrectly position the patient when using the ExacTrac Cone Beam CT (CBCT) with a TrueBeam-specific optional subvolume-CBCT.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=133318 | Z-1049-2015 | FORUM Archive and Viewer, version 3.1, v 3.1.1, (DVD Format) and v 3.2, v 3.2.1.(DVD and USB Flash Drive Format). Catalog numbers: 000000-20107-750 (DVD with either FORUM v 3.1 or v 3.1.1) 000000-2058-601 (DVD with either FORUM v 3.2 or 3.2.1); 000000-2084-928 (USB drive with FORUM 3.2.1) Ophthalmic image management system. | FORUM Archive and Viewer | 2 | CDRH | 2015-02-04 | 2015-02-04 21:00 | 2016-03-04 | Carl Zeiss Meditec AG | Software defect in the FORUM Viewer versions 3.1 and 3.2 which may lead to misinterpretation of the optical coherence tomography (OCT) data.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132038 | Z-0957-2015 | iPlan RT is a radiation treatment planning system that is intended for use in stereotactic, conformal, computer planned, Linac based radiation treatment of cranial, head and neck, and extracranial lesions. | Brainlab | 2 | CDRH | 2015-01-13 09:26 | 2015-01-15 21:00 | 2018-03-28 | Brainlab AG | iPlan RT Radiation Treatment Planning Software: Potentially incorrect patient positioning when using multiple localized CT image data sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=132036 | Z-0956-2015 | iPlan RT Dose is a stereotactic radiation treatment planning system that is intended for use in stereotactic, conformal, computer planned, Linac based radiation treatment of cranial, head and neck, and extracranial lesions. | Brainlab | 2 | CDRH | 2015-01-13 09:26 | 2015-01-15 21:00 | 2018-03-28 | Brainlab AG | iPlan RT Radiation Treatment Planning Software: Potentially incorrect patient positioning when using multiple localized CT image data sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=131428 | Z-0464-2015 | AlignRT- Intended for prescription use. The system is indicated for use during simulation, setup and stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body where radiation is indicated. | AlignRT | 2 | CDRH | 2014-11-25 | 2014-12-16 12:45 | 2016-05-10 | Vision Rt Inc | Potential failure of AlignRT to assert interlock.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=128845 | Z-2213-2014 | Navigation Spine & Trauma 3D Version 2.0 and 2.1 Is intended as an intraoperative image-guided localization system to enable minimally invasive surgery. It links a freehand probe, tracked by a passive marker sensor system to virtual computer image space on a patient's preoperative or Intraoperative 2D or 3D image data. The system is indicated for any medical condition in which the use of stereotactic surgery may be appropriate and where a reference to a rigid anatomical structure, such as the skull, the pelvis, a long bone or vertebra can be identified relative to the acquired image (CT, MR, 2D fluoroscopic image or 3D fluoroscopic image reconstruction) and/or an image data based model of the anatomy. | Brainlab | 2 | CDRH | 2014-08-13 | 2014-08-13 21:00 | 2017-03-13 | Brainlab AG | Brainlab Navigation Software Spine &amp; Trauma 3D 2.0/2.1 offers automatic registration of intraoperatively acquired CT image data sets. To enable automatic registration, the software requires the gantry position of the scanner. The gantry position can either be entered manually or submitted automatically from the CT scanner. If automatic gantry communication is available, the so-called &quot
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=128684 | Z-2140-2014 | VITROS 4600 Chemistry System (VITROS 5,1 FS System family member) --- The VITROS 4600 Chemistry System is intended for use in the in vitro quantitative measurement of a variety of analytes of clinical interest, using both VITROS Chemistry Products Slides (colorimetric endpoint, rate, ion-selective electrode, and immunorate methods) and VITROS Chemistry Products MicroTip liquid reagents (spectrophotometric and spectrophotometric immunoassay methods.) The VITROS Software is a part of the VITROS 3600, 4600, & 5600 Systems. The VITROS 3600, 4600, & 5600 Systems share the same software base. VITROS Sample Programming consists of Sample Identification (SID), test requests, and patient demographics. Patient demographics include patient name, address, date of birth, gender, physician, and patient identification (PID). | VITROS 4600 Chemistry System | 2 | CDRH | 2014-08-01 | 2014-08-01 21:00 | 2018-05-21 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics has identified an anomaly with VITROS System Software Version 3.0 and below. Internal testing confirmed that under very specific conditions, incorrect patient demographics were associated with a sample identification (ID) number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=128683 | Z-2139-2014 | VITROS 3600 Immunodiagnostic System --- For use in the in vitro quantitative, semi-quantitative and qualitative measurement of a variety of analytes of clinical interest, using VITROS Immunodiagnostic Products Reagents. The VITROS Software is a part of the VITROS 3600, 4600, & 5600 Systems. The VITROS 3600, 4600, & 5600 Systems share the same software base. VITROS Sample Programming consists of Sample Identification (SID), test requests, and patient demographics. Patient demographics include patient name, address, date of birth, gender, physician, and patient identification (PID). | VITROS 3600 Immunodiagnostic System | 2 | CDRH | 2014-08-01 | 2014-08-01 21:00 | 2018-05-21 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics has identified an anomaly with VITROS System Software Version 3.0 and below. Internal testing confirmed that under very specific conditions, incorrect patient demographics were associated with a sample identification (ID) number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=128685 | Z-2141-2014 | VITROS 5600 Integrated System --- For use in the in vitro quantitative, semi-quantitative, and qualitative measurement of a variety of analytes of clinical interest, using VITROS Chemistry Products Slides, VITROS Chemistry Products MicroTip Reagents and VITROS Immunodiagnostic Products Reagents. The VITROS Software is a part of the VITROS 3600, 4600, & 5600 Systems. The VITROS 3600, 4600, & 5600 Systems share the same software base. VITROS Sample Programming consists of Sample Identification (SID), test requests, and patient demographics. Patient demographics include patient name, address, date of birth, gender, physician, and patient identification (PID). | VITROS 5600 Integrated System | 2 | CDRH | 2014-08-01 | 2014-08-01 21:00 | 2018-05-21 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics has identified an anomaly with VITROS System Software Version 3.0 and below. Internal testing confirmed that under very specific conditions, incorrect patient demographics were associated with a sample identification (ID) number.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126256 | Z-1464-2014 | Ingenuity Core 128, Computed Tomography X-ray systems, Philips Medical Systems (Cleveland), Inc. | Ingenuity Core128 | 2 | CDRH | 2014-04-16 | 2014-04-16 21:00 | 2014-11-05 | Philips Medical Systems (Cleveland) Inc | A customer reported that after upgrading to software version 3.5.5 from 3.5.4, reconstructions would intermittently not complete.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126320 | Z-1442-2014 | SpaceLabs Arkon Anesthesia Delivery System with Version 2.0 Software installed. This system includes a ventilator. | Arkon Anesthesia Delivery System | 1 | CDRH | 2014-04-11 | 2014-04-12 21:00 | 2014-08-20 | Del Mar Reynolds Medical, Ltd. | The System may go into a controlled failed state if a series of parameters on the user interface are activated in a specific sequence.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122863 | Z-0334-2014 | Zeiss brand IOLMaster 500, Model 500, Biomicroscope, Slit-Lamp, AC-Powered, Product is manufactured by Carl Zeiss Meditec AG (Jena), Carl Zeiss Promenade 10, Jena, Germany 07745; and distributed by Carl Zeiss Meditec, Inc., Dublin, CA This device is intended for use primarily by physicians and health care workers and may be only used under the supervision of a physician. This device will not be sold to the general public. The IOLMaster is intended for the biometric determination of ocular measurements of axial length, anterior chamber depth, corneal radius, white-to white (WTW), and for the measurement of pupil size and deviation of the visual axis from the center of the pupil. For patients who are candidates for intraocular lens (IOL) implantation, the device also performs calculations to assist physicians in determining the appropriate IOL power and type for implantation | IOLMaster 500 | 2 | CDRH | 2013-11-15 | 2013-11-15 21:00 | 2014-04-17 | Carl Zeiss Meditec, Inc. | The user database delivered for download contained an erroneous value. Specifically the "aO" constant of the lens AMO Sensar AR40E is listed as "0.0" in the internal lens database of the IOLMaster. The correct value is -2.420. This could lead to the implantation of lens with a wrong diopter value.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122862 | Z-0333-2014 | Zeiss brand IOLMaster, Version 1-5, Biomicroscope, Slit-Lamp, AC-Powered, Product is manufactured by Carl Zeiss Meditec AG (Jena), Carl Zeiss Promenade 10, Jena, Germany 07745; and distributed by Carl Zeiss Meditec, Inc., Dublin, CA This device is intended for use primarily by physicians and health care workers and may be only used under the supervision of a physician. This device will not be sold to the general public. The IOLMaster is intended for the biometric determination of ocular measurements of axial length, anterior chamber depth, corneal radius, white-to white (WTW), and for the measurement of pupil size and deviation of the visual axis from the center of the pupil. For patients who are candidates for intraocular lens (IOL) implantation, the device also performs calculations to assist physicians in determining the appropriate IOL power and type for implantation | IOLMaster | 2 | CDRH | 2013-11-15 | 2013-11-15 21:00 | 2014-04-17 | Carl Zeiss Meditec, Inc. | The user database delivered for download contained an erroneous value. Specifically the "aO" constant of the lens AMO Sensar AR40E is listed as "0.0" in the internal lens database of the IOLMaster. The correct value is -2.420. This could lead to the implantation of lens with a wrong diopter value.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122607 | Z-0269-2014 | Philips brand IntelliSpace PACS 4.4, Image Management System; The system is a software package used with general purpose computing hardware to acquire, store, distribute, process and display images and associated data throughout a clinical environment. Product is manufactured and distributed by Philips Healthcare Informatics, Inc. Foster City, CA IntelliSpace PACS are image management systems intended to be used by trained professionals, including but not limited to physicians, nurses and medical technicians. The system is a software package used with general purpose computing hardware to acquire, store, distribute, process and display images and associated data throughout a clinical environment. The software performs digital image processing, measurement, communication and storage. iSite supports receiving, sending, printing, storing and displaying studies received from the following modality types via DICOM: CT, MR, NM, US, XA, PET, DX, DR, RE, RT, MG, SC, VL, as well as hospital/radiology information systems. | IntelliSpace PACS | 2 | CDRH | 2013-11-13 | 2013-11-13 21:00 | 2014-06-11 | Philips Healthcare Informatics, Inc. | Philips Healthcare became aware of a software problem with the IntelliSpace PACS system, in that a timing issue has been identified that may result in a condition where the device does not store additional images received before processing is complete. The missing images may contain pathology which may result in a different diagnosis if they had been available at time of interpretation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122368 | Z-0035-2014 | IMPAX 5.2 DB Server (running Oracle 10.1.0.4.0), medical imaging Picture Archiving and Communication System (PACS). | IMPAX | 2 | CDRH | 2013-10-21 | 2013-10-21 21:00 | 2013-10-24 | AGFA Corp. | System downtime and/or slow performance may result if software in use is an old version of IMPAX and is not at current supported levels.